## Dat_Individ
Data File: | k3r_2239_0_3_i1.akf
Curve Alignment Program: | 1.5.7
------------------------MEASURING INFOS------------------------
Diagram Title #1 | Forces and Moments at Knee Joint
Diagram Title #2 | Measured Relative to Implant
Comment #1 | Knee Joint; Gaitanalysis; Deep Knee Bend;
Comment #2 | K3R, 8 Months PO *
BodyWeight [N]: | 960
Patient(en): | K3R | [7 empty] | Subject | K3R
Implant Type: | Knee Joint
Varus Angle Leg [°] | 3.5 | [7 empty] | Individual loads in N (forces) and Nm (moments)
Axes Directions X, Y, Z: | Lateral, Anterior, Superior
Axes Origin: | Tibial Plateau at Stem Axis | [7 empty] | BW [kg] = | 97.8593
Transformation X Y Z (°): | 10 | 0.5 | 10.5
Displacement X, Y, Z (mm): | 0 | 0 | 0
Activity Code: | KneeActivities | 1P_1_15_3
---------------------------DATA--------------------------------
Number of Data Sets = | 601
Number of analog channels = | 0
Time-Offset(sec) = | 0
Time | Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My | Time
[s] | [N] | [N] | [N] | [N] | [Nm] | [Nm] | [Nm] | [Nm] | [N] | (Nm] | [%Cycle]
0 | -29.6 | -11.51 | -920.39 | 921.09 | 6.91 | -2.38 | -6.17 | 9.92 | 920.39 | 2.38 | 0
0.0018 | -29.86 | -11.63 | -915.72 | 916.42 | 6.94 | -2.35 | -6.22 | 9.96 | 915.72 | 2.35 | 0.16735
0.0036 | -30.33 | -12.52 | -914.68 | 915.38 | 6.92 | -2.33 | -6.25 | 9.96 | 914.68 | 2.33 | 0.3347
0.0054 | -30.87 | -13.44 | -916.3 | 917.03 | 6.9 | -2.31 | -6.26 | 9.95 | 916.3 | 2.31 | 0.50205
0.0072 | -31.92 | -15.22 | -919.17 | 920 | 6.85 | -2.27 | -6.26 | 9.91 | 919.17 | 2.27 | 0.66939
0.009 | -32.69 | -15.88 | -925.15 | 926.02 | 6.81 | -2.27 | -6.2 | 9.86 | 925.15 | 2.27 | 0.83674
0.0108 | -32.63 | -15.28 | -928.49 | 929.38 | 6.82 | -2.29 | -6.16 | 9.85 | 928.49 | 2.29 | 1.0041
0.0125 | -33.1 | -15.46 | -931.57 | 932.5 | 6.8 | -2.28 | -6.13 | 9.83 | 931.57 | 2.28 | 1.1621
0.0143 | -34.34 | -17.1 | -933.19 | 934.24 | 6.74 | -2.23 | -6.13 | 9.79 | 933.19 | 2.23 | 1.3295
0.0161 | -33.93 | -16.55 | -931.96 | 932.98 | 6.8 | -2.24 | -6.14 | 9.84 | 931.96 | 2.24 | 1.4968
0.0179 | -33.9 | -16.96 | -927.29 | 928.27 | 6.83 | -2.21 | -6.19 | 9.86 | 927.29 | 2.21 | 1.6642
0.0197 | -35.71 | -19.66 | -921.16 | 922.18 | 6.76 | -2.11 | -6.27 | 9.82 | 921.16 | 2.11 | 1.8315
0.0215 | -36.73 | -21.49 | -919.78 | 920.84 | 6.73 | -2.06 | -6.3 | 9.78 | 919.78 | 2.06 | 1.9989
0.0233 | -37.03 | -21.67 | -919.46 | 920.54 | 6.77 | -2.04 | -6.31 | 9.81 | 919.46 | 2.04 | 2.1662
0.0251 | -37.73 | -22.18 | -920.14 | 921.26 | 6.81 | -2.01 | -6.3 | 9.83 | 920.14 | 2.01 | 2.3336
0.0269 | -38.57 | -22.88 | -921.05 | 922.22 | 6.86 | -1.98 | -6.3 | 9.86 | 921.05 | 1.98 | 2.5009
0.0287 | -39.08 | -23.29 | -922.71 | 923.9 | 6.92 | -1.96 | -6.29 | 9.9 | 922.71 | 1.96 | 2.6683
0.0305 | -40.06 | -24.19 | -924.94 | 926.21 | 6.95 | -1.92 | -6.27 | 9.91 | 924.94 | 1.92 | 2.8356
0.0323 | -41.01 | -25 | -929.22 | 930.53 | 6.95 | -1.9 | -6.21 | 9.87 | 929.22 | 1.9 | 3.003
0.0341 | -41.32 | -25.15 | -932.81 | 934.14 | 6.99 | -1.89 | -6.16 | 9.87 | 932.81 | 1.89 | 3.1703
0.0359 | -41.63 | -25.3 | -936.4 | 937.75 | 7.03 | -1.88 | -6.1 | 9.87 | 936.4 | 1.88 | 3.3377
0.0376 | -42.59 | -26.61 | -938.2 | 939.61 | 7.07 | -1.82 | -6.09 | 9.88 | 938.2 | 1.82 | 3.4957
0.0394 | -43.16 | -27.22 | -938.02 | 939.47 | 7.13 | -1.78 | -6.1 | 9.92 | 938.02 | 1.78 | 3.6631
0.0412 | -43.08 | -26.8 | -937.91 | 939.34 | 7.25 | -1.76 | -6.1 | 10.01 | 937.91 | 1.76 | 3.8304
0.043 | -43.61 | -27.53 | -938.73 | 940.22 | 7.33 | -1.72 | -6.1 | 10.07 | 938.73 | 1.72 | 3.9978
0.0448 | -44.58 | -28.84 | -940.03 | 941.62 | 7.37 | -1.68 | -6.11 | 10.11 | 940.03 | 1.68 | 4.1651
0.0466 | -45.62 | -30.15 | -941.02 | 942.73 | 7.39 | -1.64 | -6.11 | 10.14 | 941.02 | 1.64 | 4.3325
0.0484 | -46.16 | -31.24 | -941.23 | 943.02 | 7.41 | -1.62 | -6.11 | 10.15 | 941.23 | 1.62 | 4.4998
0.0502 | -46.46 | -32.11 | -940.82 | 942.66 | 7.44 | -1.6 | -6.11 | 10.17 | 940.82 | 1.6 | 4.6672
0.052 | -46.57 | -32.76 | -938.33 | 940.16 | 7.52 | -1.56 | -6.13 | 10.22 | 938.33 | 1.56 | 4.8345
0.0538 | -47.49 | -34.1 | -934.45 | 936.35 | 7.62 | -1.44 | -6.19 | 10.29 | 934.45 | 1.44 | 5.0019
0.0556 | -48.42 | -35.43 | -930.56 | 932.54 | 7.71 | -1.33 | -6.25 | 10.35 | 930.56 | 1.33 | 5.1692
0.0574 | -49.16 | -36.52 | -927.31 | 929.36 | 7.81 | -1.22 | -6.3 | 10.42 | 927.31 | 1.22 | 5.3366
0.0592 | -48.64 | -36.14 | -927.49 | 929.49 | 7.92 | -1.21 | -6.27 | 10.46 | 927.49 | 1.21 | 5.5039
0.061 | -48 | -35.77 | -927.72 | 929.68 | 8.02 | -1.2 | -6.23 | 10.52 | 927.72 | 1.2 | 5.6713
0.0627 | -47.94 | -35.78 | -928.34 | 930.28 | 8.16 | -1.1 | -6.24 | 10.61 | 928.34 | 1.1 | 5.8293
0.0645 | -48.32 | -36.75 | -926.7 | 928.7 | 8.15 | -1.07 | -6.24 | 10.6 | 926.7 | 1.07 | 5.9967
0.0663 | -48.49 | -37.44 | -924.11 | 926.15 | 8.17 | -1.03 | -6.25 | 10.62 | 924.11 | 1.03 | 6.164
0.0681 | -48.01 | -37.28 | -918.09 | 920.11 | 8.27 | -0.96 | -6.3 | 10.71 | 918.09 | 0.96 | 6.3313
0.0699 | -46.69 | -34.7 | -913.93 | 915.78 | 8.52 | -0.86 | -6.31 | 10.91 | 913.93 | 0.86 | 6.4987
0.0717 | -45.64 | -33.03 | -906.58 | 908.33 | 8.65 | -0.82 | -6.33 | 11.03 | 906.58 | 0.82 | 6.666
0.0735 | -44.74 | -31.99 | -898.5 | 900.18 | 8.74 | -0.8 | -6.36 | 11.12 | 898.5 | 0.8 | 6.8334
0.0753 | -44.34 | -32.43 | -890.57 | 892.28 | 8.76 | -0.75 | -6.41 | 11.17 | 890.57 | 0.75 | 7.0007
0.0771 | -42.76 | -31.74 | -883.73 | 885.35 | 8.84 | -0.77 | -6.43 | 11.25 | 883.73 | 0.77 | 7.1681
0.0789 | -40.46 | -30.37 | -877.53 | 879.02 | 8.96 | -0.82 | -6.43 | 11.34 | 877.53 | 0.82 | 7.3354
0.0807 | -38.16 | -28.99 | -871.33 | 872.69 | 9.07 | -0.87 | -6.43 | 11.44 | 871.33 | 0.87 | 7.5028
0.0825 | -35.48 | -27.53 | -868.53 | 869.76 | 9.21 | -0.96 | -6.4 | 11.54 | 868.53 | 0.96 | 7.6701
0.0843 | -34.72 | -26.75 | -869.11 | 870.34 | 9.34 | -0.94 | -6.34 | 11.61 | 869.11 | 0.94 | 7.8375
0.086 | -35.27 | -28.09 | -869.01 | 870.35 | 9.37 | -0.88 | -6.31 | 11.61 | 869.01 | 0.88 | 7.9955
0.0878 | -36.16 | -28.33 | -870.44 | 871.87 | 9.43 | -0.83 | -6.27 | 11.62 | 870.44 | 0.83 | 8.1629
0.0896 | -37.38 | -28.66 | -872.74 | 874.26 | 9.49 | -0.77 | -6.22 | 11.61 | 872.74 | 0.77 | 8.3302
0.0914 | -42.42 | -33.16 | -879.26 | 881.08 | 9.39 | -0.57 | -6.2 | 11.46 | 879.26 | 0.57 | 8.4976
0.0932 | -47.8 | -37.67 | -883.02 | 885.16 | 9.3 | -0.33 | -6.21 | 11.33 | 883.02 | 0.33 | 8.6649
0.095 | -50.67 | -39.99 | -881.62 | 884.02 | 9.29 | -0.19 | -6.23 | 11.3 | 881.62 | 0.19 | 8.8323
0.0968 | -51.43 | -40.34 | -878.9 | 881.38 | 9.31 | -0.16 | -6.22 | 11.29 | 878.9 | 0.16 | 8.9996
0.0986 | -52.51 | -41.39 | -875.75 | 878.35 | 9.3 | -0.11 | -6.22 | 11.26 | 875.75 | 0.11 | 9.167
0.1004 | -53.77 | -43.29 | -872.15 | 874.94 | 9.26 | -0.04 | -6.22 | 11.22 | 872.15 | 0.04 | 9.3343
0.1022 | -52.86 | -43.46 | -868 | 870.75 | 9.26 | -0.07 | -6.19 | 11.21 | 868 | 0.07 | 9.5017
0.104 | -52.33 | -44.81 | -865.1 | 867.89 | 9.2 | -0.1 | -6.15 | 11.14 | 865.1 | 0.1 | 9.669
0.1058 | -52.55 | -46.55 | -861.79 | 864.71 | 9.12 | -0.09 | -6.13 | 11.06 | 861.79 | 0.09 | 9.8364
0.1076 | -51.9 | -46.41 | -854.38 | 857.29 | 9.09 | -0.08 | -6.1 | 11.02 | 854.38 | 0.08 | 10.0037
0.1094 | -50.65 | -45.52 | -850.39 | 853.19 | 9.04 | -0.11 | -6.01 | 10.93 | 850.39 | 0.11 | 10.1711
0.1111 | -49.57 | -44.95 | -848.11 | 850.82 | 8.98 | -0.13 | -5.92 | 10.83 | 848.11 | 0.13 | 10.3291
0.1129 | -49.31 | -44.72 | -840.46 | 843.18 | 8.93 | -0.09 | -5.9 | 10.77 | 840.46 | 0.09 | 10.4965
0.1147 | -48.68 | -43.2 | -829.64 | 832.28 | 8.86 | -0.01 | -5.85 | 10.68 | 829.64 | 0.01 | 10.6638
0.1165 | -48.29 | -41.49 | -819.55 | 822.11 | 8.77 | 0.08 | -5.8 | 10.58 | 819.55 | -0.08 | 10.8312
0.1183 | -47.46 | -40.37 | -817.77 | 820.23 | 8.63 | 0.12 | -5.67 | 10.39 | 817.77 | -0.12 | 10.9985
0.1201 | -45.4 | -38.01 | -813.75 | 816 | 8.59 | 0.13 | -5.57 | 10.31 | 813.75 | -0.13 | 11.1659
0.1219 | -42.5 | -34.6 | -806.28 | 808.26 | 8.63 | 0.12 | -5.5 | 10.3 | 806.28 | -0.12 | 11.3332
0.1237 | -39.16 | -30.52 | -795.48 | 797.17 | 8.7 | 0.12 | -5.46 | 10.34 | 795.48 | -0.12 | 11.5006
0.1255 | -37.33 | -27.38 | -784.45 | 785.99 | 8.69 | 0.18 | -5.42 | 10.32 | 784.45 | -0.18 | 11.6679
0.1273 | -36.13 | -24.94 | -774.51 | 775.93 | 8.7 | 0.27 | -5.42 | 10.34 | 774.51 | -0.27 | 11.8353
0.1291 | -35.88 | -23.2 | -766.22 | 767.56 | 8.78 | 0.36 | -5.48 | 10.43 | 766.22 | -0.36 | 12.0026
0.1309 | -35.64 | -21.07 | -760.16 | 761.41 | 8.85 | 0.44 | -5.5 | 10.49 | 760.16 | -0.44 | 12.17
0.1327 | -34.73 | -18.58 | -755.38 | 756.53 | 8.87 | 0.46 | -5.47 | 10.5 | 755.38 | -0.46 | 12.3373
0.1345 | -34.28 | -16.94 | -750.52 | 751.62 | 8.82 | 0.48 | -5.44 | 10.44 | 750.52 | -0.48 | 12.5046
0.1362 | -35.39 | -16.76 | -744.77 | 745.9 | 8.75 | 0.57 | -5.45 | 10.4 | 744.77 | -0.57 | 12.6627
0.138 | -37.53 | -17.11 | -743.9 | 745.11 | 8.68 | 0.72 | -5.45 | 10.36 | 743.9 | -0.72 | 12.83
0.1398 | -40.3 | -18.28 | -746.57 | 747.94 | 8.6 | 0.88 | -5.45 | 10.31 | 746.57 | -0.88 | 12.9974
0.1416 | -42.74 | -18.81 | -749.53 | 751.01 | 8.52 | 1.02 | -5.43 | 10.26 | 749.53 | -1.02 | 13.1647
0.1434 | -42.55 | -18.85 | -754.73 | 756.17 | 8.49 | 1 | -5.38 | 10.22 | 754.73 | -1 | 13.3321
0.1452 | -43.87 | -19.52 | -762.56 | 764.08 | 8.43 | 1.04 | -5.33 | 10.15 | 762.56 | -1.04 | 13.4994
0.147 | -45.66 | -21.03 | -773.2 | 774.84 | 8.31 | 1.08 | -5.26 | 10.04 | 773.2 | -1.08 | 13.6668
0.1488 | -46.95 | -22.02 | -784.34 | 786.06 | 8.23 | 1.1 | -5.17 | 9.93 | 784.34 | -1.1 | 13.8341
0.1506 | -46.43 | -22.55 | -792.73 | 794.42 | 8.17 | 1.05 | -5.06 | 9.83 | 792.73 | -1.05 | 14.0015
0.1524 | -45.58 | -22.04 | -798.1 | 799.71 | 8.13 | 0.97 | -4.97 | 9.75 | 798.1 | -0.97 | 14.1688
0.1542 | -44.4 | -21.54 | -801.89 | 803.41 | 8.09 | 0.9 | -4.89 | 9.67 | 801.89 | -0.9 | 14.3362
0.156 | -44.03 | -21.27 | -803.74 | 805.24 | 8.05 | 0.84 | -4.84 | 9.62 | 803.74 | -0.84 | 14.5035
0.1578 | -43.7 | -20.94 | -803.59 | 805.06 | 8.02 | 0.8 | -4.8 | 9.57 | 803.59 | -0.8 | 14.6709
0.1595 | -43.55 | -20.84 | -801.99 | 803.46 | 7.99 | 0.8 | -4.78 | 9.53 | 801.99 | -0.8 | 14.8289
0.1613 | -44.41 | -21.9 | -800.26 | 801.81 | 7.91 | 0.83 | -4.79 | 9.48 | 800.26 | -0.83 | 14.9963
0.1631 | -45.67 | -22.87 | -797.91 | 799.58 | 7.86 | 0.89 | -4.82 | 9.44 | 797.91 | -0.89 | 15.1636
0.1649 | -46.89 | -23.71 | -795.24 | 797.03 | 7.81 | 0.95 | -4.84 | 9.41 | 795.24 | -0.95 | 15.331
0.1667 | -46.03 | -22.47 | -788.77 | 790.47 | 7.86 | 0.94 | -4.85 | 9.46 | 788.77 | -0.94 | 15.4983
0.1685 | -43.74 | -20.25 | -780.63 | 782.15 | 7.94 | 0.86 | -4.86 | 9.55 | 780.63 | -0.86 | 15.6657
0.1703 | -40.45 | -17.39 | -771.09 | 772.36 | 8.06 | 0.73 | -4.88 | 9.68 | 771.09 | -0.73 | 15.833
0.1721 | -36.11 | -13.85 | -765.36 | 766.46 | 8.16 | 0.55 | -4.83 | 9.79 | 765.36 | -0.55 | 16.0004
0.1739 | -32.44 | -10.69 | -760.87 | 761.85 | 8.24 | 0.41 | -4.77 | 9.88 | 760.87 | -0.41 | 16.1677
0.1757 | -33.49 | -10.38 | -760.96 | 762.07 | 8.2 | 0.45 | -4.7 | 9.84 | 760.96 | -0.45 | 16.3351
0.1775 | -35.17 | -10.41 | -762.32 | 763.57 | 8.14 | 0.53 | -4.64 | 9.8 | 762.32 | -0.53 | 16.5024
0.1793 | -38.21 | -12.4 | -768.3 | 769.74 | 8 | 0.65 | -4.57 | 9.66 | 768.3 | -0.65 | 16.6698
0.1811 | -41.49 | -15.13 | -778.02 | 779.74 | 7.84 | 0.79 | -4.47 | 9.52 | 778.02 | -0.79 | 16.8371
0.1829 | -43.08 | -16.89 | -789.11 | 790.98 | 7.78 | 0.88 | -4.35 | 9.45 | 789.11 | -0.88 | 17.0045
0.1846 | -43.84 | -18.07 | -800.45 | 802.39 | 7.76 | 0.95 | -4.23 | 9.42 | 800.45 | -0.95 | 17.1625
0.1864 | -43.35 | -18.12 | -807.63 | 809.47 | 7.76 | 0.98 | -4.12 | 9.35 | 807.63 | -0.98 | 17.3299
0.1882 | -42.3 | -17.57 | -810.36 | 811.99 | 7.77 | 1 | -4.03 | 9.28 | 810.36 | -1 | 17.4972
0.19 | -42.6 | -18.29 | -810.24 | 811.79 | 7.72 | 1.11 | -4 | 9.19 | 810.24 | -1.11 | 17.6646
0.1918 | -43.16 | -19.32 | -808.1 | 809.63 | 7.65 | 1.24 | -4 | 9.09 | 808.1 | -1.24 | 17.8319
0.1936 | -43.04 | -19.41 | -803.05 | 804.56 | 7.6 | 1.38 | -4.02 | 9.03 | 803.05 | -1.38 | 17.9993
0.1954 | -41.94 | -18.9 | -798.83 | 800.3 | 7.58 | 1.46 | -4.01 | 9 | 798.83 | -1.46 | 18.1666
0.1972 | -40.8 | -18.94 | -797.31 | 798.71 | 7.59 | 1.51 | -3.98 | 9 | 797.31 | -1.51 | 18.334
0.199 | -37.9 | -17.93 | -796.75 | 798.05 | 7.64 | 1.49 | -3.92 | 9.01 | 796.75 | -1.49 | 18.5013
0.2008 | -35.26 | -16.97 | -796.23 | 797.5 | 7.71 | 1.49 | -3.86 | 9.05 | 796.23 | -1.49 | 18.6687
0.2026 | -32.87 | -16.13 | -796.87 | 798.12 | 7.79 | 1.52 | -3.78 | 9.08 | 796.87 | -1.52 | 18.836
0.2044 | -33.11 | -17.08 | -800.18 | 801.42 | 7.81 | 1.72 | -3.71 | 9.03 | 800.18 | -1.72 | 19.0033
0.2062 | -34.25 | -18.61 | -802.83 | 804.07 | 7.81 | 1.95 | -3.64 | 9 | 802.83 | -1.95 | 19.1707
0.208 | -35.45 | -20.51 | -802.36 | 803.65 | 7.85 | 2.16 | -3.6 | 9.02 | 802.36 | -2.16 | 19.338
0.2097 | -36.98 | -23.09 | -802.03 | 803.39 | 7.89 | 2.41 | -3.55 | 9.05 | 802.03 | -2.41 | 19.4961
0.2115 | -39.53 | -25.41 | -803.21 | 804.72 | 7.92 | 2.66 | -3.5 | 9.11 | 803.21 | -2.66 | 19.6634
0.2133 | -43.31 | -29.27 | -808.82 | 810.62 | 7.93 | 2.99 | -3.48 | 9.21 | 808.82 | -2.99 | 19.8308
0.2151 | -48.04 | -33.35 | -815.28 | 817.44 | 7.96 | 3.35 | -3.44 | 9.33 | 815.28 | -3.35 | 19.9981
0.2169 | -52.74 | -37.77 | -825.2 | 827.82 | 7.91 | 3.65 | -3.43 | 9.4 | 825.2 | -3.65 | 20.1655
0.2187 | -58.09 | -42.98 | -838.89 | 842.08 | 7.84 | 3.96 | -3.42 | 9.46 | 838.89 | -3.96 | 20.3328
0.2205 | -63.24 | -48.02 | -856.44 | 860.24 | 7.83 | 4.29 | -3.4 | 9.61 | 856.44 | -4.29 | 20.5002
0.2223 | -65.29 | -50.26 | -873.62 | 877.64 | 7.91 | 4.47 | -3.35 | 9.75 | 873.62 | -4.47 | 20.6675
0.2241 | -64.54 | -49.6 | -884.16 | 888.04 | 8.11 | 4.55 | -3.33 | 9.94 | 884.16 | -4.55 | 20.8349
0.2259 | -65.47 | -51.34 | -896.21 | 900.21 | 8.19 | 4.65 | -3.32 | 10.05 | 896.21 | -4.65 | 21.0022
0.2277 | -66.18 | -52.79 | -909.7 | 913.79 | 8.28 | 4.74 | -3.31 | 10.16 | 909.7 | -4.74 | 21.1696
0.2295 | -65.36 | -52.5 | -923.86 | 927.81 | 8.42 | 4.77 | -3.3 | 10.29 | 923.86 | -4.77 | 21.3369
0.2313 | -64.09 | -51.93 | -934.5 | 938.28 | 8.56 | 4.8 | -3.29 | 10.41 | 934.5 | -4.8 | 21.5043
0.2331 | -64.22 | -53.04 | -947.59 | 951.4 | 8.62 | 4.87 | -3.27 | 10.48 | 947.59 | -4.87 | 21.6716
0.2348 | -64.83 | -55 | -965.49 | 969.39 | 8.66 | 4.93 | -3.25 | 10.52 | 965.49 | -4.93 | 21.8297
0.2366 | -65.22 | -57.06 | -986.82 | 990.78 | 8.73 | 4.96 | -3.24 | 10.59 | 986.82 | -4.96 | 21.997
0.2384 | -64.61 | -58.27 | -1007.57 | 1011.47 | 8.87 | 4.97 | -3.22 | 10.7 | 1007.57 | -4.97 | 22.1644
0.2402 | -62.94 | -58.69 | -1024.87 | 1028.62 | 9.01 | 4.99 | -3.2 | 10.82 | 1024.87 | -4.99 | 22.3317
0.242 | -61.05 | -58.86 | -1041.22 | 1044.81 | 9.15 | 5.03 | -3.19 | 10.95 | 1041.22 | -5.03 | 22.4991
0.2438 | -58.55 | -57.57 | -1052.1 | 1055.45 | 9.3 | 5.08 | -3.19 | 11.11 | 1052.1 | -5.08 | 22.6664
0.2456 | -55.34 | -56.13 | -1060.36 | 1063.39 | 9.53 | 5.08 | -3.17 | 11.29 | 1060.36 | -5.08 | 22.8338
0.2474 | -52.32 | -53.55 | -1065.95 | 1068.66 | 9.79 | 5.09 | -3.18 | 11.53 | 1065.95 | -5.09 | 23.0011
0.2492 | -49.63 | -51.13 | -1072.96 | 1075.4 | 10.02 | 5.13 | -3.18 | 11.76 | 1072.96 | -5.13 | 23.1685
0.251 | -47.42 | -49.63 | -1084.54 | 1086.79 | 10.16 | 5.2 | -3.17 | 11.91 | 1084.54 | -5.2 | 23.3358
0.2528 | -44.89 | -48.33 | -1096.78 | 1098.84 | 10.31 | 5.21 | -3.16 | 12.04 | 1096.78 | -5.21 | 23.5032
0.2546 | -42.32 | -47.55 | -1107.11 | 1109.01 | 10.47 | 5.21 | -3.17 | 12.19 | 1107.11 | -5.21 | 23.6705
0.2564 | -40.55 | -47.1 | -1118.29 | 1120.11 | 10.59 | 5.22 | -3.18 | 12.3 | 1118.29 | -5.22 | 23.8379
0.2581 | -37.89 | -45.49 | -1130.19 | 1131.89 | 10.7 | 5.19 | -3.18 | 12.38 | 1130.19 | -5.19 | 23.9959
0.2599 | -36 | -44.75 | -1141.25 | 1142.86 | 10.8 | 5.22 | -3.18 | 12.47 | 1141.25 | -5.22 | 24.1633
0.2617 | -35.78 | -45.83 | -1152.1 | 1153.73 | 10.86 | 5.37 | -3.22 | 12.59 | 1152.1 | -5.37 | 24.3306
0.2635 | -36.03 | -47.31 | -1165.18 | 1166.88 | 10.93 | 5.57 | -3.25 | 12.73 | 1165.18 | -5.57 | 24.498
0.2653 | -36.29 | -48.79 | -1178.36 | 1180.13 | 10.98 | 5.77 | -3.28 | 12.86 | 1178.36 | -5.77 | 24.6653
0.2671 | -36.76 | -50.52 | -1194.33 | 1196.18 | 10.97 | 5.97 | -3.3 | 12.96 | 1194.33 | -5.97 | 24.8327
0.2689 | -36.59 | -51.01 | -1209.28 | 1211.15 | 11.05 | 6.12 | -3.34 | 13.11 | 1209.28 | -6.12 | 25
0.2707 | -36.57 | -50.94 | -1225.3 | 1227.16 | 11.12 | 6.29 | -3.36 | 13.27 | 1225.3 | -6.29 | 25.1673
0.2725 | -36.8 | -49.89 | -1243.2 | 1245.02 | 11.17 | 6.5 | -3.37 | 13.43 | 1243.2 | -6.5 | 25.3347
0.2743 | -37.88 | -50.1 | -1264.51 | 1266.31 | 11.19 | 6.7 | -3.38 | 13.56 | 1264.51 | -6.7 | 25.502
0.2761 | -38.68 | -50.99 | -1287.17 | 1288.95 | 11.21 | 6.85 | -3.41 | 13.65 | 1287.17 | -6.85 | 25.6694
0.2779 | -38.24 | -50.99 | -1309.2 | 1310.9 | 11.26 | 6.96 | -3.42 | 13.75 | 1309.2 | -6.96 | 25.8367
0.2797 | -37.12 | -50.98 | -1326.26 | 1327.97 | 11.35 | 7.02 | -3.45 | 13.85 | 1326.26 | -7.02 | 26.0041
0.2815 | -35.77 | -51.31 | -1344.73 | 1346.43 | 11.48 | 7.1 | -3.51 | 14.02 | 1344.73 | -7.1 | 26.1714
0.2832 | -35.44 | -52.6 | -1367.64 | 1369.36 | 11.59 | 7.21 | -3.55 | 14.18 | 1367.64 | -7.21 | 26.3295
0.285 | -34.44 | -53.25 | -1387.86 | 1389.56 | 11.71 | 7.29 | -3.58 | 14.31 | 1387.86 | -7.29 | 26.4968
0.2868 | -32.24 | -52.45 | -1402.54 | 1404.14 | 11.81 | 7.32 | -3.58 | 14.41 | 1402.54 | -7.32 | 26.6642
0.2886 | -30.42 | -51.37 | -1417.19 | 1418.69 | 11.91 | 7.4 | -3.57 | 14.54 | 1417.19 | -7.4 | 26.8315
0.2904 | -28.6 | -50.29 | -1431.84 | 1433.25 | 12 | 7.48 | -3.56 | 14.66 | 1431.84 | -7.48 | 26.9989
0.2922 | -26.65 | -49.77 | -1441.96 | 1443.32 | 12.06 | 7.57 | -3.57 | 14.76 | 1441.96 | -7.57 | 27.1662
0.294 | -24.14 | -47.78 | -1446.44 | 1447.69 | 12.13 | 7.7 | -3.59 | 14.9 | 1446.44 | -7.7 | 27.3336
0.2958 | -21.81 | -43.4 | -1446.8 | 1447.92 | 12.19 | 7.84 | -3.59 | 15.04 | 1446.8 | -7.84 | 27.5009
0.2976 | -20.57 | -38.21 | -1441.13 | 1442.15 | 12.12 | 7.93 | -3.55 | 15.03 | 1441.13 | -7.93 | 27.6683
0.2994 | -20.11 | -34.62 | -1431.57 | 1432.57 | 11.93 | 8.01 | -3.51 | 14.93 | 1431.57 | -8.01 | 27.8356
0.3012 | -20.04 | -31.72 | -1422.88 | 1423.86 | 11.71 | 8.11 | -3.48 | 14.84 | 1422.88 | -8.11 | 28.003
0.303 | -21.09 | -31.96 | -1426.31 | 1427.3 | 11.58 | 8.26 | -3.5 | 14.85 | 1426.31 | -8.26 | 28.1703
0.3048 | -21.3 | -30.96 | -1430.09 | 1431.08 | 11.52 | 8.39 | -3.5 | 14.91 | 1430.09 | -8.39 | 28.3377
0.3066 | -22.06 | -30.56 | -1434.39 | 1435.39 | 11.5 | 8.51 | -3.53 | 14.99 | 1434.39 | -8.51 | 28.505
0.3083 | -23.07 | -30.52 | -1438.99 | 1440 | 11.48 | 8.63 | -3.56 | 15.06 | 1438.99 | -8.63 | 28.6631
0.3101 | -23.66 | -30.86 | -1443.13 | 1444.18 | 11.46 | 8.74 | -3.6 | 15.13 | 1443.13 | -8.74 | 28.8304
0.3119 | -25.26 | -33.04 | -1447.11 | 1448.23 | 11.41 | 8.87 | -3.65 | 15.2 | 1447.11 | -8.87 | 28.9978
0.3137 | -27.64 | -35.78 | -1451.36 | 1452.54 | 11.31 | 9.02 | -3.69 | 15.27 | 1451.36 | -9.02 | 29.1651
0.3155 | -29.17 | -37.18 | -1455.55 | 1456.78 | 11.21 | 9.22 | -3.7 | 15.33 | 1455.55 | -9.22 | 29.3325
0.3173 | -29.34 | -36.35 | -1455.73 | 1456.95 | 11.18 | 9.38 | -3.71 | 15.4 | 1455.73 | -9.38 | 29.4998
0.3191 | -29.39 | -34.62 | -1456.59 | 1457.82 | 11.19 | 9.5 | -3.74 | 15.48 | 1456.59 | -9.5 | 29.6672
0.3209 | -29.43 | -32.85 | -1457.53 | 1458.77 | 11.21 | 9.62 | -3.76 | 15.55 | 1457.53 | -9.62 | 29.8345
0.3227 | -28.13 | -29.37 | -1455.18 | 1456.4 | 11.28 | 9.67 | -3.8 | 15.62 | 1455.18 | -9.67 | 30.0019
0.3245 | -26.94 | -25.33 | -1452.05 | 1453.26 | 11.34 | 9.64 | -3.83 | 15.63 | 1452.05 | -9.64 | 30.1692
0.3263 | -26.61 | -22.98 | -1450.82 | 1452.04 | 11.32 | 9.66 | -3.88 | 15.65 | 1450.82 | -9.66 | 30.3366
0.3281 | -26.49 | -22.01 | -1440.96 | 1442.19 | 11.21 | 9.54 | -3.93 | 15.5 | 1440.96 | -9.54 | 30.5039
0.3299 | -26.54 | -21.95 | -1438.46 | 1439.68 | 11.15 | 9.47 | -3.99 | 15.43 | 1438.46 | -9.47 | 30.6713
0.3316 | -27.36 | -22.46 | -1442.04 | 1443.33 | 11.12 | 9.49 | -4.05 | 15.43 | 1442.04 | -9.49 | 30.8293
0.3334 | -27.92 | -22.53 | -1448.88 | 1450.24 | 11.14 | 9.52 | -4.11 | 15.47 | 1448.88 | -9.52 | 30.9967
0.3352 | -28.38 | -23.2 | -1458.31 | 1459.75 | 11.15 | 9.56 | -4.16 | 15.5 | 1458.31 | -9.56 | 31.164
0.337 | -28.89 | -24.88 | -1469.95 | 1471.46 | 11.17 | 9.59 | -4.23 | 15.53 | 1469.95 | -9.59 | 31.3313
0.3388 | -29.18 | -26.94 | -1483.31 | 1484.85 | 11.22 | 9.6 | -4.3 | 15.57 | 1483.31 | -9.6 | 31.4987
0.3406 | -28.73 | -27.18 | -1500.3 | 1501.76 | 11.29 | 9.7 | -4.3 | 15.68 | 1500.3 | -9.7 | 31.666
0.3424 | -26.96 | -26.52 | -1516.85 | 1518.15 | 11.42 | 9.74 | -4.3 | 15.81 | 1516.85 | -9.74 | 31.8334
0.3442 | -24.28 | -25.16 | -1534.18 | 1535.29 | 11.63 | 9.71 | -4.31 | 15.98 | 1534.18 | -9.71 | 32.0007
0.346 | -21.6 | -23.81 | -1551.51 | 1552.43 | 11.85 | 9.68 | -4.32 | 16.15 | 1551.51 | -9.68 | 32.1681
0.3478 | -21.91 | -23.85 | -1572.68 | 1573.53 | 12.01 | 9.69 | -4.34 | 16.31 | 1572.68 | -9.69 | 32.3354
0.3496 | -21.92 | -23.3 | -1590.65 | 1591.4 | 12.14 | 9.67 | -4.36 | 16.43 | 1590.65 | -9.67 | 32.5028
0.3514 | -21.65 | -23.17 | -1603.69 | 1604.37 | 12.21 | 9.65 | -4.39 | 16.5 | 1603.69 | -9.65 | 32.6701
0.3532 | -21.54 | -23.4 | -1613.97 | 1614.6 | 12.23 | 9.65 | -4.42 | 16.55 | 1613.97 | -9.65 | 32.8375
0.355 | -22.03 | -24.29 | -1622.29 | 1622.9 | 12.23 | 9.68 | -4.45 | 16.57 | 1622.29 | -9.68 | 33.0048
0.3567 | -21.36 | -24.62 | -1628.9 | 1629.47 | 12.26 | 9.66 | -4.49 | 16.59 | 1628.9 | -9.66 | 33.1629
0.3585 | -19.55 | -24.28 | -1633.23 | 1633.75 | 12.31 | 9.6 | -4.51 | 16.61 | 1633.23 | -9.6 | 33.3302
0.3603 | -17.97 | -23.91 | -1636.43 | 1636.91 | 12.37 | 9.53 | -4.5 | 16.6 | 1636.43 | -9.53 | 33.4976
0.3621 | -17.89 | -23.26 | -1641.58 | 1642.05 | 12.42 | 9.5 | -4.46 | 16.6 | 1641.58 | -9.5 | 33.6649
0.3639 | -16.9 | -22.26 | -1643.93 | 1644.39 | 12.46 | 9.45 | -4.46 | 16.59 | 1643.93 | -9.45 | 33.8323
0.3657 | -15.87 | -21.19 | -1646.53 | 1646.97 | 12.5 | 9.38 | -4.46 | 16.57 | 1646.53 | -9.38 | 33.9996
0.3675 | -15.96 | -20.2 | -1655.15 | 1655.59 | 12.51 | 9.34 | -4.45 | 16.54 | 1655.15 | -9.34 | 34.167
0.3693 | -16.03 | -18.81 | -1664.46 | 1664.89 | 12.58 | 9.27 | -4.48 | 16.54 | 1664.46 | -9.27 | 34.3343
0.3711 | -15.96 | -16.79 | -1671.06 | 1671.48 | 12.67 | 9.2 | -4.52 | 16.56 | 1671.06 | -9.2 | 34.5017
0.3729 | -15.48 | -14.35 | -1670.85 | 1671.27 | 12.77 | 9.07 | -4.6 | 16.57 | 1670.85 | -9.07 | 34.669
0.3747 | -14.26 | -12.03 | -1668.61 | 1669.03 | 12.85 | 8.91 | -4.66 | 16.54 | 1668.61 | -8.91 | 34.8364
0.3765 | -14.02 | -12.04 | -1666.75 | 1667.16 | 12.77 | 8.83 | -4.67 | 16.44 | 1666.75 | -8.83 | 35.0037
0.3783 | -14.3 | -12.85 | -1668.21 | 1668.61 | 12.66 | 8.74 | -4.68 | 16.3 | 1668.21 | -8.74 | 35.1711
0.3801 | -14.89 | -13.82 | -1673.75 | 1674.14 | 12.58 | 8.65 | -4.67 | 16.18 | 1673.75 | -8.65 | 35.3384
0.3818 | -15.26 | -14.16 | -1676.85 | 1677.24 | 12.53 | 8.57 | -4.68 | 16.09 | 1676.85 | -8.57 | 35.4965
0.3836 | -15.46 | -14.92 | -1675.04 | 1675.43 | 12.48 | 8.53 | -4.72 | 16.05 | 1675.04 | -8.53 | 35.6638
0.3854 | -16.11 | -16.75 | -1675.97 | 1676.35 | 12.41 | 8.51 | -4.76 | 16 | 1675.97 | -8.51 | 35.8312
0.3872 | -17.71 | -19.4 | -1684.56 | 1684.96 | 12.32 | 8.54 | -4.78 | 15.97 | 1684.56 | -8.54 | 35.9985
0.389 | -19.41 | -21.69 | -1694.73 | 1695.15 | 12.29 | 8.59 | -4.83 | 16 | 1694.73 | -8.59 | 36.1659
0.3908 | -21.33 | -24.56 | -1704.55 | 1705.03 | 12.24 | 8.68 | -4.88 | 16.02 | 1704.55 | -8.68 | 36.3332
0.3926 | -23.34 | -28.35 | -1708.2 | 1708.78 | 12.18 | 8.84 | -4.98 | 16.11 | 1708.2 | -8.84 | 36.5006
0.3944 | -25.61 | -32.59 | -1713.56 | 1714.26 | 12.1 | 8.99 | -5.07 | 16.17 | 1713.56 | -8.99 | 36.6679
0.3962 | -28.66 | -37.42 | -1723.64 | 1724.55 | 11.97 | 9.18 | -5.09 | 16.23 | 1723.64 | -9.18 | 36.8353
0.398 | -31.82 | -41.98 | -1734.07 | 1735.22 | 11.85 | 9.4 | -5.1 | 16.29 | 1734.07 | -9.4 | 37.0026
0.3998 | -33.87 | -44.9 | -1744.02 | 1745.31 | 11.8 | 9.57 | -5.1 | 16.38 | 1744.02 | -9.57 | 37.17
0.4016 | -34.99 | -46.45 | -1758.24 | 1759.59 | 11.85 | 9.68 | -5.12 | 16.49 | 1758.24 | -9.68 | 37.3373
0.4034 | -35.73 | -48.48 | -1775.29 | 1776.65 | 11.89 | 9.74 | -5.17 | 16.56 | 1775.29 | -9.74 | 37.5046
0.4051 | -36.94 | -50.36 | -1795.14 | 1796.52 | 11.93 | 9.73 | -5.21 | 16.6 | 1795.14 | -9.73 | 37.6627
0.4069 | -38.47 | -52.34 | -1815.65 | 1817.07 | 11.97 | 9.71 | -5.26 | 16.63 | 1815.65 | -9.71 | 37.83
0.4087 | -39.37 | -53.8 | -1833.46 | 1834.92 | 12.04 | 9.7 | -5.31 | 16.68 | 1833.46 | -9.7 | 37.9974
0.4105 | -38.79 | -53.32 | -1848.19 | 1849.63 | 12.19 | 9.67 | -5.36 | 16.79 | 1848.19 | -9.67 | 38.1647
0.4123 | -36.82 | -50.76 | -1865.54 | 1866.89 | 12.46 | 9.62 | -5.37 | 16.94 | 1865.54 | -9.62 | 38.3321
0.4141 | -34.67 | -48.92 | -1886.33 | 1887.57 | 12.71 | 9.58 | -5.36 | 17.07 | 1886.33 | -9.58 | 38.4994
0.4159 | -33.55 | -49.16 | -1910.32 | 1911.52 | 12.88 | 9.57 | -5.34 | 17.18 | 1910.32 | -9.57 | 38.6668
0.4177 | -34.76 | -51.98 | -1934.92 | 1936.16 | 12.93 | 9.62 | -5.34 | 17.25 | 1934.92 | -9.62 | 38.8341
0.4195 | -33.27 | -51.69 | -1957.87 | 1959.08 | 13.15 | 9.59 | -5.37 | 17.4 | 1957.87 | -9.59 | 39.0015
0.4213 | -32.01 | -51.44 | -1981.8 | 1982.97 | 13.37 | 9.57 | -5.41 | 17.57 | 1981.8 | -9.57 | 39.1688
0.4231 | -30.37 | -50.6 | -2002.9 | 2004.03 | 13.61 | 9.53 | -5.47 | 17.74 | 2002.9 | -9.53 | 39.3362
0.4249 | -27.91 | -48.47 | -2020.46 | 2021.46 | 13.89 | 9.5 | -5.53 | 17.94 | 2020.46 | -9.5 | 39.5035
0.4267 | -25.53 | -47.01 | -2033.96 | 2034.9 | 14.06 | 9.37 | -5.55 | 18.02 | 2033.96 | -9.37 | 39.6709
0.4285 | -24.96 | -47.23 | -2047.93 | 2048.84 | 14.17 | 9.35 | -5.59 | 18.11 | 2047.93 | -9.35 | 39.8382
0.4302 | -23.73 | -45.72 | -2056.2 | 2057.02 | 14.31 | 9.35 | -5.6 | 18.23 | 2056.2 | -9.35 | 39.9963
0.432 | -21.24 | -43.23 | -2057.98 | 2058.68 | 14.47 | 9.32 | -5.61 | 18.33 | 2057.98 | -9.32 | 40.1636
0.4338 | -19.38 | -41.4 | -2057.79 | 2058.43 | 14.58 | 9.32 | -5.63 | 18.42 | 2057.79 | -9.32 | 40.331
0.4356 | -17.56 | -39.06 | -2053.76 | 2054.34 | 14.69 | 9.32 | -5.64 | 18.52 | 2053.76 | -9.32 | 40.4983
0.4374 | -16.57 | -37.44 | -2047.23 | 2047.79 | 14.72 | 9.29 | -5.68 | 18.54 | 2047.23 | -9.29 | 40.6657
0.4392 | -15.27 | -36.5 | -2043.3 | 2043.85 | 14.72 | 9.29 | -5.68 | 18.54 | 2043.3 | -9.29 | 40.833
0.441 | -13.76 | -35.79 | -2041.16 | 2041.69 | 14.71 | 9.36 | -5.68 | 18.56 | 2041.16 | -9.36 | 41.0004
0.4428 | -11.81 | -33.79 | -2035.22 | 2035.72 | 14.77 | 9.43 | -5.71 | 18.65 | 2035.22 | -9.43 | 41.1677
0.4446 | -10.03 | -31.92 | -2030.46 | 2030.93 | 14.85 | 9.49 | -5.73 | 18.73 | 2030.46 | -9.49 | 41.3351
0.4464 | -10.48 | -32.47 | -2031.31 | 2031.8 | 14.82 | 9.65 | -5.74 | 18.78 | 2031.31 | -9.65 | 41.5024
0.4482 | -13.05 | -35.62 | -2036.59 | 2037.18 | 14.68 | 9.88 | -5.74 | 18.78 | 2036.59 | -9.88 | 41.6698
0.45 | -14.56 | -37.27 | -2045.03 | 2045.66 | 14.62 | 10.04 | -5.71 | 18.79 | 2045.03 | -10.04 | 41.8371
0.4518 | -16.87 | -39.7 | -2056.43 | 2057.15 | 14.51 | 10.23 | -5.68 | 18.79 | 2056.43 | -10.23 | 42.0045
0.4536 | -18.02 | -39.89 | -2067.21 | 2067.93 | 14.46 | 10.38 | -5.65 | 18.81 | 2067.21 | -10.38 | 42.1718
0.4553 | -18.44 | -39.37 | -2076 | 2076.7 | 14.45 | 10.46 | -5.63 | 18.86 | 2076 | -10.46 | 42.3299
0.4571 | -18.93 | -39.05 | -2081.66 | 2082.33 | 14.44 | 10.55 | -5.63 | 18.9 | 2081.66 | -10.55 | 42.4972
0.4589 | -18.55 | -37.68 | -2083.37 | 2083.97 | 14.48 | 10.6 | -5.63 | 18.95 | 2083.37 | -10.6 | 42.6646
0.4607 | -17.9 | -36.23 | -2083.28 | 2083.82 | 14.52 | 10.65 | -5.62 | 19 | 2083.28 | -10.65 | 42.8319
0.4625 | -17.82 | -35.72 | -2082.6 | 2083.11 | 14.53 | 10.72 | -5.62 | 19.03 | 2082.6 | -10.72 | 42.9993
0.4643 | -18.07 | -35.46 | -2083.85 | 2084.34 | 14.52 | 10.79 | -5.63 | 19.07 | 2083.85 | -10.79 | 43.1666
0.4661 | -19.2 | -36 | -2090.29 | 2090.8 | 14.48 | 10.88 | -5.64 | 19.09 | 2090.29 | -10.88 | 43.334
0.4679 | -19 | -35.21 | -2099.52 | 2100.02 | 14.46 | 10.89 | -5.69 | 19.1 | 2099.52 | -10.89 | 43.5013
0.4697 | -18.97 | -33.83 | -2106.38 | 2106.87 | 14.48 | 10.9 | -5.73 | 19.12 | 2106.38 | -10.9 | 43.6687
0.4715 | -18.63 | -33.18 | -2109.43 | 2109.96 | 14.45 | 10.86 | -5.72 | 19.09 | 2109.43 | -10.86 | 43.836
0.4733 | -18.89 | -32.86 | -2110.21 | 2110.75 | 14.4 | 10.86 | -5.72 | 19.06 | 2110.21 | -10.86 | 44.0033
0.4751 | -21.26 | -34.89 | -2110.88 | 2111.46 | 14.26 | 10.96 | -5.71 | 19 | 2110.88 | -10.96 | 44.1707
0.4769 | -23.08 | -35.74 | -2107.3 | 2107.89 | 14.18 | 11.05 | -5.75 | 18.98 | 2107.3 | -11.05 | 44.338
0.4786 | -22.97 | -35.17 | -2097.54 | 2098.12 | 14.15 | 11.09 | -5.77 | 18.99 | 2097.54 | -11.09 | 44.4961
0.4804 | -22.37 | -33.84 | -2086.09 | 2086.64 | 14.14 | 11.13 | -5.79 | 19 | 2086.09 | -11.13 | 44.6634
0.4822 | -21.88 | -32.11 | -2072.44 | 2072.96 | 14.15 | 11.17 | -5.8 | 19.03 | 2072.44 | -11.17 | 44.8308
0.484 | -21.37 | -30.37 | -2058.83 | 2059.32 | 14.15 | 11.22 | -5.81 | 19.06 | 2058.83 | -11.22 | 44.9981
0.4858 | -21.43 | -29.62 | -2050.65 | 2051.15 | 14.13 | 11.28 | -5.82 | 19.08 | 2050.65 | -11.28 | 45.1655
0.4876 | -20.97 | -28.83 | -2045.52 | 2046.01 | 14.12 | 11.31 | -5.81 | 19.09 | 2045.52 | -11.31 | 45.3328
0.4894 | -20.81 | -28.93 | -2043.29 | 2043.79 | 14.09 | 11.36 | -5.79 | 19.09 | 2043.29 | -11.36 | 45.5002
0.4912 | -21.83 | -30.21 | -2044.68 | 2045.24 | 14.02 | 11.45 | -5.77 | 19.09 | 2044.68 | -11.45 | 45.6675
0.493 | -23.17 | -31.96 | -2048.02 | 2048.63 | 13.94 | 11.56 | -5.74 | 19.09 | 2048.02 | -11.56 | 45.8349
0.4948 | -23.51 | -32.66 | -2051.6 | 2052.21 | 13.91 | 11.62 | -5.7 | 19.09 | 2051.6 | -11.62 | 46.0022
0.4966 | -25.28 | -34.76 | -2056.39 | 2057.01 | 13.82 | 11.72 | -5.68 | 19.06 | 2056.39 | -11.72 | 46.1696
0.4984 | -25.3 | -34.96 | -2054.92 | 2055.53 | 13.81 | 11.77 | -5.67 | 19.07 | 2054.92 | -11.77 | 46.3369
0.5002 | -26.95 | -35.66 | -2051 | 2051.62 | 13.76 | 11.91 | -5.7 | 19.12 | 2051 | -11.91 | 46.5043
0.502 | -25.82 | -33.08 | -2045.92 | 2046.5 | 13.83 | 11.91 | -5.69 | 19.17 | 2045.92 | -11.91 | 46.6716
0.5037 | -26.11 | -32.87 | -2043.73 | 2044.31 | 13.79 | 11.94 | -5.68 | 19.16 | 2043.73 | -11.94 | 46.8297
0.5055 | -29.58 | -36.84 | -2045.42 | 2046.11 | 13.56 | 12.09 | -5.7 | 19.09 | 2045.42 | -12.09 | 46.997
0.5073 | -31.71 | -39.57 | -2045.05 | 2045.82 | 13.41 | 12.17 | -5.71 | 19.04 | 2045.05 | -12.17 | 47.1644
0.5091 | -31.48 | -39.58 | -2040.27 | 2041.04 | 13.37 | 12.17 | -5.73 | 19.02 | 2040.27 | -12.17 | 47.3317
0.5109 | -31.64 | -39.35 | -2034.93 | 2035.69 | 13.32 | 12.17 | -5.74 | 18.98 | 2034.93 | -12.17 | 47.4991
0.5127 | -31.95 | -38.97 | -2028.66 | 2029.42 | 13.27 | 12.17 | -5.77 | 18.96 | 2028.66 | -12.17 | 47.6664
0.5145 | -30.16 | -36.08 | -2019.22 | 2019.91 | 13.33 | 12.1 | -5.77 | 18.95 | 2019.22 | -12.1 | 47.8338
0.5163 | -29.2 | -34.81 | -2007.66 | 2008.32 | 13.3 | 12.05 | -5.76 | 18.9 | 2007.66 | -12.05 | 48.0011
0.5181 | -29.51 | -35.22 | -1995.45 | 1996.14 | 13.2 | 12.06 | -5.77 | 18.83 | 1995.45 | -12.06 | 48.1685
0.5199 | -29.35 | -34.7 | -1984 | 1984.69 | 13.14 | 12.06 | -5.77 | 18.79 | 1984 | -12.06 | 48.3358
0.5217 | -28 | -31.33 | -1975.25 | 1975.91 | 13.21 | 12.02 | -5.77 | 18.82 | 1975.25 | -12.02 | 48.5032
0.5235 | -25.99 | -28.88 | -1969.81 | 1970.44 | 13.23 | 11.91 | -5.75 | 18.76 | 1969.81 | -11.91 | 48.6705
0.5253 | -27.08 | -30.36 | -1967.14 | 1967.84 | 13.11 | 11.95 | -5.75 | 18.71 | 1967.14 | -11.95 | 48.8379
0.5271 | -26.91 | -30.32 | -1965.69 | 1966.39 | 13.1 | 11.97 | -5.73 | 18.71 | 1965.69 | -11.97 | 49.0052
0.5288 | -26.43 | -29.83 | -1963.85 | 1964.54 | 13.12 | 11.98 | -5.71 | 18.72 | 1963.85 | -11.98 | 49.1633
0.5306 | -24.62 | -27.32 | -1960.31 | 1960.93 | 13.22 | 11.95 | -5.68 | 18.76 | 1960.31 | -11.95 | 49.3306
0.5324 | -22.81 | -24.81 | -1956.77 | 1957.33 | 13.31 | 11.91 | -5.65 | 18.8 | 1956.77 | -11.91 | 49.498
0.5342 | -21.24 | -23.01 | -1954.34 | 1954.82 | 13.34 | 11.81 | -5.63 | 18.73 | 1954.34 | -11.81 | 49.6653
0.536 | -21.34 | -22.5 | -1955.08 | 1955.56 | 13.33 | 11.81 | -5.6 | 18.72 | 1955.08 | -11.81 | 49.8327
0.5378 | -22.12 | -22.27 | -1955.85 | 1956.33 | 13.33 | 11.89 | -5.57 | 18.75 | 1955.85 | -11.89 | 50
0.5396 | -22.28 | -21.92 | -1956.93 | 1957.4 | 13.33 | 11.91 | -5.54 | 18.76 | 1956.93 | -11.91 | 50.1673
0.5414 | -21.77 | -21.74 | -1959.49 | 1959.96 | 13.32 | 11.87 | -5.52 | 18.72 | 1959.49 | -11.87 | 50.3347
0.5432 | -23.68 | -24.59 | -1966.54 | 1967.08 | 13.19 | 11.91 | -5.52 | 18.66 | 1966.54 | -11.91 | 50.502
0.545 | -24.72 | -25.46 | -1974.08 | 1974.62 | 13.15 | 11.94 | -5.5 | 18.64 | 1974.08 | -11.94 | 50.6694
0.5468 | -25.74 | -26.48 | -1982.44 | 1982.99 | 13.11 | 11.95 | -5.49 | 18.61 | 1982.44 | -11.95 | 50.8367
0.5486 | -29.13 | -31.4 | -1995.96 | 1996.59 | 12.91 | 12.03 | -5.49 | 18.51 | 1995.96 | -12.03 | 51.0041
0.5504 | -32.36 | -36.21 | -2009.84 | 2010.54 | 12.72 | 12.09 | -5.5 | 18.42 | 2009.84 | -12.09 | 51.1714
0.5521 | -34.03 | -39.04 | -2021.2 | 2021.95 | 12.63 | 12.11 | -5.5 | 18.36 | 2021.2 | -12.11 | 51.3295
0.5539 | -35.5 | -41.86 | -2031.87 | 2032.69 | 12.52 | 12.13 | -5.48 | 18.3 | 2031.87 | -12.13 | 51.4968
0.5557 | -36.39 | -44.42 | -2041.72 | 2042.58 | 12.45 | 12.13 | -5.5 | 18.26 | 2041.72 | -12.13 | 51.6642
0.5575 | -37.55 | -46.68 | -2051.85 | 2052.77 | 12.42 | 12.18 | -5.57 | 18.29 | 2051.85 | -12.18 | 51.8315
0.5593 | -38.17 | -48.05 | -2064.17 | 2065.15 | 12.43 | 12.2 | -5.62 | 18.33 | 2064.17 | -12.2 | 51.9989
0.5611 | -37.56 | -47.82 | -2078.1 | 2079.09 | 12.54 | 12.14 | -5.68 | 18.4 | 2078.1 | -12.14 | 52.1662
0.5629 | -38.22 | -49.44 | -2095.77 | 2096.82 | 12.57 | 12.15 | -5.72 | 18.45 | 2095.77 | -12.15 | 52.3336
0.5647 | -39.48 | -51.14 | -2119.31 | 2120.4 | 12.65 | 12.14 | -5.79 | 18.54 | 2119.31 | -12.14 | 52.5009
0.5665 | -39.59 | -51.23 | -2141.69 | 2142.77 | 12.77 | 12.11 | -5.82 | 18.62 | 2141.69 | -12.11 | 52.6683
0.5683 | -38.74 | -50.49 | -2157.19 | 2158.24 | 12.88 | 12.05 | -5.83 | 18.67 | 2157.19 | -12.05 | 52.8356
0.5701 | -37.01 | -49.24 | -2170.66 | 2171.67 | 13 | 11.95 | -5.84 | 18.71 | 2170.66 | -11.95 | 53.003
0.5719 | -35.64 | -48.36 | -2182.49 | 2183.48 | 13.1 | 11.86 | -5.86 | 18.74 | 2182.49 | -11.86 | 53.1703
0.5737 | -33.73 | -46.31 | -2189.23 | 2190.14 | 13.23 | 11.79 | -5.88 | 18.79 | 2189.23 | -11.79 | 53.3377
0.5755 | -31.74 | -44.56 | -2192.34 | 2193.15 | 13.35 | 11.68 | -5.9 | 18.81 | 2192.34 | -11.68 | 53.505
0.5772 | -30.17 | -43.29 | -2196 | 2196.74 | 13.45 | 11.59 | -5.93 | 18.83 | 2196 | -11.59 | 53.6631
0.579 | -29.81 | -42.69 | -2204.27 | 2204.99 | 13.49 | 11.58 | -5.93 | 18.85 | 2204.27 | -11.58 | 53.8304
0.5808 | -31.2 | -44.25 | -2213.58 | 2214.37 | 13.43 | 11.64 | -5.94 | 18.85 | 2213.58 | -11.64 | 53.9978
0.5826 | -30.84 | -44.16 | -2220.08 | 2220.89 | 13.47 | 11.59 | -5.96 | 18.87 | 2220.08 | -11.59 | 54.1651
0.5844 | -30.24 | -43.78 | -2225.96 | 2226.8 | 13.53 | 11.54 | -5.99 | 18.89 | 2225.96 | -11.54 | 54.3325
0.5862 | -28.92 | -42.31 | -2230.08 | 2230.89 | 13.62 | 11.47 | -6.02 | 18.93 | 2230.08 | -11.47 | 54.4998
0.588 | -26.87 | -39.51 | -2237.64 | 2238.34 | 13.73 | 11.41 | -6.01 | 18.97 | 2237.64 | -11.41 | 54.6672
0.5898 | -28.2 | -41.24 | -2245.88 | 2246.65 | 13.65 | 11.43 | -6.02 | 18.94 | 2245.88 | -11.43 | 54.8345
0.5916 | -29.62 | -42.78 | -2250.08 | 2250.94 | 13.57 | 11.45 | -6.03 | 18.91 | 2250.08 | -11.45 | 55.0019
0.5934 | -29.81 | -42.27 | -2246.22 | 2247.09 | 13.53 | 11.42 | -6.04 | 18.86 | 2246.22 | -11.42 | 55.1692
0.5952 | -29.8 | -41.54 | -2241.36 | 2242.2 | 13.47 | 11.39 | -6.03 | 18.79 | 2241.36 | -11.39 | 55.3366
0.597 | -29.79 | -41.14 | -2234.39 | 2235.21 | 13.4 | 11.35 | -6.04 | 18.71 | 2234.39 | -11.35 | 55.5039
0.5988 | -29.79 | -41.33 | -2223.85 | 2224.68 | 13.29 | 11.3 | -6.05 | 18.61 | 2223.85 | -11.3 | 55.6713
0.6006 | -30.01 | -41.75 | -2212.07 | 2212.9 | 13.17 | 11.29 | -6.05 | 18.51 | 2212.07 | -11.29 | 55.8386
0.6023 | -29.51 | -41.29 | -2198.47 | 2199.28 | 13.07 | 11.29 | -6.04 | 18.42 | 2198.47 | -11.29 | 55.9967
0.6041 | -30.75 | -41.88 | -2186.2 | 2187 | 12.94 | 11.36 | -6.03 | 18.34 | 2186.2 | -11.36 | 56.164
0.6059 | -33.14 | -42.3 | -2173.49 | 2174.28 | 12.84 | 11.45 | -6.03 | 18.31 | 2173.49 | -11.45 | 56.3313
0.6077 | -34.34 | -42.07 | -2164.3 | 2165.06 | 12.79 | 11.49 | -6.01 | 18.28 | 2164.3 | -11.49 | 56.4987
0.6095 | -34.02 | -41.83 | -2164.69 | 2165.43 | 12.78 | 11.43 | -5.97 | 18.23 | 2164.69 | -11.43 | 56.666
0.6113 | -36.1 | -44.52 | -2170.93 | 2171.75 | 12.68 | 11.47 | -5.95 | 18.16 | 2170.93 | -11.47 | 56.8334
0.6131 | -38.88 | -46.72 | -2176.41 | 2177.32 | 12.62 | 11.54 | -5.98 | 18.17 | 2176.41 | -11.54 | 57.0007
0.6149 | -39.59 | -46.64 | -2178.04 | 2178.95 | 12.64 | 11.57 | -6.01 | 18.22 | 2178.04 | -11.57 | 57.1681
0.6167 | -38.07 | -43.88 | -2177.57 | 2178.39 | 12.77 | 11.52 | -6.01 | 18.3 | 2177.57 | -11.52 | 57.3354
0.6185 | -35.47 | -39.76 | -2175.5 | 2176.18 | 12.99 | 11.44 | -6.01 | 18.4 | 2175.5 | -11.44 | 57.5028
0.6203 | -32.08 | -35.66 | -2180.74 | 2181.28 | 13.2 | 11.29 | -5.96 | 18.46 | 2180.74 | -11.29 | 57.6701
0.6221 | -32.69 | -36.26 | -2190.44 | 2191 | 13.18 | 11.34 | -5.94 | 18.47 | 2190.44 | -11.34 | 57.8375
0.6239 | -33.73 | -37.84 | -2200.69 | 2201.29 | 13.17 | 11.37 | -5.97 | 18.49 | 2200.69 | -11.37 | 58.0048
0.6256 | -32.29 | -36.99 | -2209.26 | 2209.82 | 13.28 | 11.3 | -6 | 18.54 | 2209.26 | -11.3 | 58.1629
0.6274 | -30.51 | -36.43 | -2220.03 | 2220.55 | 13.37 | 11.18 | -6.01 | 18.55 | 2220.03 | -11.18 | 58.3302
0.6292 | -29.59 | -37.85 | -2234.3 | 2234.83 | 13.36 | 11.05 | -5.98 | 18.45 | 2234.3 | -11.05 | 58.4976
0.631 | -31.34 | -41.78 | -2250.36 | 2250.98 | 13.25 | 11.03 | -5.97 | 18.36 | 2250.36 | -11.03 | 58.6649
0.6328 | -32.27 | -44.09 | -2261.06 | 2261.72 | 13.17 | 11.05 | -5.96 | 18.33 | 2261.06 | -11.05 | 58.8323
0.6346 | -32.4 | -44.89 | -2263.59 | 2264.29 | 13.15 | 11.03 | -5.98 | 18.3 | 2263.59 | -11.03 | 58.9996
0.6364 | -30.75 | -43.57 | -2264.68 | 2265.32 | 13.2 | 10.95 | -5.98 | 18.3 | 2264.68 | -10.95 | 59.167
0.6382 | -28.8 | -41.99 | -2265.66 | 2266.24 | 13.26 | 10.84 | -5.96 | 18.29 | 2265.66 | -10.84 | 59.3343
0.64 | -28.21 | -42.19 | -2272.57 | 2273.15 | 13.21 | 10.79 | -5.91 | 18.21 | 2272.57 | -10.79 | 59.5017
0.6418 | -29.13 | -43.94 | -2279.75 | 2280.38 | 13.09 | 10.79 | -5.85 | 18.1 | 2279.75 | -10.79 | 59.669
0.6436 | -28.94 | -43.32 | -2279.06 | 2279.68 | 13.09 | 10.81 | -5.85 | 18.11 | 2279.06 | -10.81 | 59.8364
0.6454 | -27.35 | -41.01 | -2274.25 | 2274.82 | 13.18 | 10.77 | -5.84 | 18.16 | 2274.25 | -10.77 | 60.0037
0.6472 | -25.47 | -39.95 | -2269.8 | 2270.32 | 13.2 | 10.73 | -5.82 | 18.15 | 2269.8 | -10.73 | 60.1711
0.649 | -24.9 | -40.17 | -2271 | 2271.53 | 13.15 | 10.75 | -5.77 | 18.1 | 2271 | -10.75 | 60.3384
0.6507 | -24.28 | -39.57 | -2268.35 | 2268.88 | 13.13 | 10.78 | -5.74 | 18.08 | 2268.35 | -10.78 | 60.4965
0.6525 | -22.58 | -36.33 | -2255.18 | 2255.65 | 13.22 | 10.82 | -5.72 | 18.17 | 2255.18 | -10.82 | 60.6638
0.6543 | -20.99 | -33.95 | -2244.2 | 2244.59 | 13.25 | 10.89 | -5.67 | 18.21 | 2244.2 | -10.89 | 60.8312
0.6561 | -20.9 | -32.85 | -2235.21 | 2235.57 | 13.23 | 11.01 | -5.64 | 18.25 | 2235.21 | -11.01 | 60.9985
0.6579 | -21.65 | -32.27 | -2226.71 | 2227.07 | 13.19 | 11.16 | -5.62 | 18.29 | 2226.71 | -11.16 | 61.1659
0.6597 | -20.78 | -29.82 | -2216.88 | 2217.19 | 13.17 | 11.21 | -5.55 | 18.28 | 2216.88 | -11.21 | 61.3332
0.6615 | -19.74 | -27.35 | -2207.67 | 2207.93 | 13.15 | 11.24 | -5.48 | 18.25 | 2207.67 | -11.24 | 61.5006
0.6633 | -20.31 | -26.56 | -2192.34 | 2192.6 | 13.03 | 11.32 | -5.43 | 18.19 | 2192.34 | -11.32 | 61.6679
0.6651 | -22.05 | -27.58 | -2178.24 | 2178.54 | 12.84 | 11.44 | -5.41 | 18.12 | 2178.24 | -11.44 | 61.8353
0.6669 | -24.9 | -30.59 | -2173.21 | 2173.58 | 12.63 | 11.6 | -5.39 | 18.05 | 2173.21 | -11.6 | 62.0026
0.6687 | -25.99 | -31.87 | -2172.04 | 2172.45 | 12.52 | 11.69 | -5.36 | 18.01 | 2172.04 | -11.69 | 62.17
0.6705 | -27.61 | -33.81 | -2173.51 | 2173.97 | 12.4 | 11.82 | -5.36 | 18 | 2173.51 | -11.82 | 62.3373
0.6723 | -27.52 | -34.41 | -2174.91 | 2175.39 | 12.35 | 11.88 | -5.34 | 18 | 2174.91 | -11.88 | 62.5046
0.6741 | -26.42 | -32.56 | -2174.18 | 2174.62 | 12.43 | 11.92 | -5.33 | 18.08 | 2174.18 | -11.92 | 62.672
0.6758 | -25.3 | -30.66 | -2173.43 | 2173.83 | 12.52 | 11.96 | -5.32 | 18.17 | 2173.43 | -11.96 | 62.83
0.6776 | -24.26 | -29.12 | -2173.94 | 2174.3 | 12.58 | 11.99 | -5.31 | 18.23 | 2173.94 | -11.99 | 62.9974
0.6794 | -23.35 | -28.39 | -2177.33 | 2177.67 | 12.59 | 11.99 | -5.26 | 18.22 | 2177.33 | -11.99 | 63.1647
0.6812 | -25.21 | -30.81 | -2186.71 | 2187.12 | 12.45 | 12.09 | -5.22 | 18.17 | 2186.71 | -12.09 | 63.3321
0.683 | -27.14 | -33.3 | -2199.54 | 2200.01 | 12.4 | 12.23 | -5.23 | 18.23 | 2199.54 | -12.23 | 63.4994
0.6848 | -28.5 | -35.28 | -2212.71 | 2213.22 | 12.38 | 12.36 | -5.24 | 18.3 | 2212.71 | -12.36 | 63.6668
0.6866 | -28.15 | -35.61 | -2226.15 | 2226.66 | 12.42 | 12.43 | -5.25 | 18.38 | 2226.15 | -12.43 | 63.8341
0.6884 | -26.82 | -34.01 | -2237.48 | 2237.94 | 12.56 | 12.49 | -5.26 | 18.51 | 2237.48 | -12.49 | 64.0015
0.6902 | -24.41 | -31.53 | -2244.5 | 2244.89 | 12.67 | 12.44 | -5.26 | 18.56 | 2244.5 | -12.44 | 64.1688
0.692 | -23.14 | -29.62 | -2252.95 | 2253.3 | 12.73 | 12.4 | -5.24 | 18.58 | 2252.95 | -12.4 | 64.3362
0.6938 | -24.85 | -31.26 | -2261.6 | 2261.98 | 12.63 | 12.47 | -5.24 | 18.56 | 2261.6 | -12.47 | 64.5035
0.6956 | -25.18 | -31.34 | -2262.99 | 2263.39 | 12.58 | 12.49 | -5.25 | 18.54 | 2262.99 | -12.49 | 64.6709
0.6974 | -21.77 | -27.02 | -2256.18 | 2256.47 | 12.68 | 12.41 | -5.2 | 18.54 | 2256.18 | -12.41 | 64.8382
0.6992 | -21.23 | -26.71 | -2248.55 | 2248.83 | 12.53 | 12.4 | -5.13 | 18.41 | 2248.55 | -12.4 | 65.0056
0.7009 | -21.48 | -27.81 | -2240.76 | 2241.05 | 12.29 | 12.41 | -5.07 | 18.23 | 2240.76 | -12.41 | 65.1636
0.7027 | -19.97 | -26.43 | -2232.73 | 2232.99 | 12.21 | 12.37 | -5.03 | 18.14 | 2232.73 | -12.37 | 65.331
0.7045 | -18.12 | -24.55 | -2227.29 | 2227.51 | 12.18 | 12.29 | -5 | 18.07 | 2227.29 | -12.29 | 65.4983
0.7063 | -16.92 | -23.56 | -2226.61 | 2226.81 | 12.13 | 12.25 | -4.98 | 17.99 | 2226.61 | -12.25 | 65.6657
0.7081 | -16.25 | -23.15 | -2226.63 | 2226.83 | 12.04 | 12.26 | -4.96 | 17.93 | 2226.63 | -12.26 | 65.833
0.7099 | -17.24 | -24.14 | -2229.95 | 2230.16 | 11.93 | 12.25 | -4.94 | 17.85 | 2229.95 | -12.25 | 66.0004
0.7117 | -18.29 | -24.8 | -2232.7 | 2232.94 | 11.85 | 12.25 | -4.92 | 17.79 | 2232.7 | -12.25 | 66.1677
0.7135 | -16.67 | -22.83 | -2229.88 | 2230.08 | 11.91 | 12.11 | -4.91 | 17.74 | 2229.88 | -12.11 | 66.3351
0.7153 | -15.63 | -21.97 | -2229.85 | 2230.03 | 11.94 | 11.98 | -4.92 | 17.67 | 2229.85 | -11.98 | 66.5024
0.7171 | -15.16 | -22.09 | -2231.99 | 2232.19 | 11.93 | 11.85 | -4.93 | 17.58 | 2231.99 | -11.85 | 66.6698
0.7189 | -15.08 | -23.63 | -2237.65 | 2237.85 | 11.88 | 11.72 | -4.94 | 17.46 | 2237.65 | -11.72 | 66.8371
0.7207 | -14.25 | -24.63 | -2242.42 | 2242.62 | 11.86 | 11.57 | -4.94 | 17.34 | 2242.42 | -11.57 | 67.0045
0.7225 | -13.71 | -25.52 | -2243.07 | 2243.29 | 11.88 | 11.46 | -4.96 | 17.29 | 2243.07 | -11.46 | 67.1718
0.7242 | -11.16 | -24.2 | -2239.52 | 2239.7 | 12.01 | 11.27 | -4.97 | 17.28 | 2239.52 | -11.27 | 67.3299
0.726 | -11.07 | -25.96 | -2239.14 | 2239.35 | 12 | 11.19 | -4.99 | 17.22 | 2239.14 | -11.19 | 67.4972
0.7278 | -10.3 | -26.68 | -2236.46 | 2236.68 | 11.97 | 11.08 | -4.97 | 17.13 | 2236.46 | -11.08 | 67.6646
0.7296 | -9.34 | -27.01 | -2233.23 | 2233.46 | 11.96 | 10.96 | -4.96 | 17.04 | 2233.23 | -10.96 | 67.8319
0.7314 | -7.18 | -24.69 | -2226.39 | 2226.59 | 12.03 | 10.79 | -4.94 | 16.97 | 2226.39 | -10.79 | 67.9993
0.7332 | -5.97 | -22.67 | -2220.39 | 2220.56 | 12.05 | 10.63 | -4.92 | 16.88 | 2220.39 | -10.63 | 68.1666
0.735 | -4.32 | -19.96 | -2211.97 | 2212.13 | 12.11 | 10.44 | -4.9 | 16.8 | 2211.97 | -10.44 | 68.334
0.7368 | -1.68 | -16.2 | -2199.98 | 2200.14 | 12.22 | 10.21 | -4.9 | 16.73 | 2199.98 | -10.21 | 68.5013
0.7386 | 0.41 | -12.4 | -2188.35 | 2188.52 | 12.29 | 9.97 | -4.87 | 16.61 | 2188.35 | -9.97 | 68.6687
0.7404 | 0.98 | -11.13 | -2177.84 | 2178.02 | 12.29 | 9.78 | -4.84 | 16.49 | 2177.84 | -9.78 | 68.836
0.7422 | 2.23 | -8.98 | -2167.64 | 2167.83 | 12.28 | 9.53 | -4.77 | 16.32 | 2167.64 | -9.53 | 69.0033
0.744 | 3.47 | -6.83 | -2157.44 | 2157.64 | 12.27 | 9.28 | -4.71 | 16.15 | 2157.44 | -9.28 | 69.1707
0.7458 | 3.55 | -6.52 | -2145 | 2145.23 | 12.17 | 9.06 | -4.68 | 15.94 | 2145 | -9.06 | 69.338
0.7476 | 4.1 | -5.1 | -2130.01 | 2130.27 | 12.12 | 8.8 | -4.65 | 15.74 | 2130.01 | -8.8 | 69.5054
0.7493 | 5.78 | -1.47 | -2114.93 | 2115.28 | 12.08 | 8.46 | -4.64 | 15.5 | 2114.93 | -8.46 | 69.6634
0.7511 | 7.69 | 2.6 | -2099.94 | 2100.39 | 12.04 | 8.09 | -4.63 | 15.25 | 2099.94 | -8.09 | 69.8308
0.7529 | 5.52 | 0.26 | -2086.07 | 2086.53 | 11.84 | 7.9 | -4.65 | 14.99 | 2086.07 | -7.9 | 69.9981
0.7547 | 4.79 | -0.55 | -2068.18 | 2068.66 | 11.71 | 7.69 | -4.66 | 14.79 | 2068.18 | -7.69 | 70.1655
0.7565 | 5.87 | 1.04 | -2046.27 | 2046.75 | 11.65 | 7.4 | -4.67 | 14.59 | 2046.27 | -7.4 | 70.3328
0.7583 | 6.64 | 3.27 | -2020.56 | 2021.08 | 11.6 | 7.13 | -4.7 | 14.43 | 2020.56 | -7.13 | 70.5002
0.7601 | 7.45 | 5.31 | -1994.65 | 1995.24 | 11.47 | 6.79 | -4.72 | 14.17 | 1994.65 | -6.79 | 70.6675
0.7619 | 8.31 | 7.38 | -1968.71 | 1969.36 | 11.34 | 6.46 | -4.74 | 13.9 | 1968.71 | -6.46 | 70.8349
0.7637 | 9.26 | 9.37 | -1942.46 | 1943.19 | 11.2 | 6.11 | -4.74 | 13.63 | 1942.46 | -6.11 | 71.0022
0.7655 | 10.88 | 11.42 | -1914.33 | 1915.16 | 11.09 | 5.72 | -4.7 | 13.35 | 1914.33 | -5.72 | 71.1696
0.7673 | 12.99 | 13.68 | -1882.54 | 1883.39 | 10.99 | 5.33 | -4.64 | 13.1 | 1882.54 | -5.33 | 71.3369
0.7691 | 14.38 | 15.35 | -1851.71 | 1852.56 | 10.86 | 5 | -4.59 | 12.84 | 1851.71 | -5 | 71.5043
0.7709 | 14.77 | 15.77 | -1824.25 | 1825.08 | 10.67 | 4.72 | -4.55 | 12.56 | 1824.25 | -4.72 | 71.6716
0.7727 | 14.62 | 15.43 | -1800.23 | 1801.05 | 10.45 | 4.48 | -4.51 | 12.27 | 1800.23 | -4.48 | 71.839
0.7744 | 14.77 | 15.67 | -1778.5 | 1779.33 | 10.28 | 4.32 | -4.49 | 12.05 | 1778.5 | -4.32 | 71.997
0.7762 | 15.38 | 16.52 | -1750.9 | 1751.7 | 10.1 | 4.04 | -4.41 | 11.78 | 1750.9 | -4.04 | 72.1644
0.778 | 14.58 | 15.69 | -1726.13 | 1726.93 | 9.87 | 3.79 | -4.35 | 11.48 | 1726.13 | -3.79 | 72.3317
0.7798 | 13.88 | 14.71 | -1704.46 | 1705.25 | 9.67 | 3.58 | -4.32 | 11.24 | 1704.46 | -3.58 | 72.4991
0.7816 | 12.89 | 13.23 | -1683.81 | 1684.58 | 9.45 | 3.39 | -4.29 | 10.99 | 1683.81 | -3.39 | 72.6664
0.7834 | 9.24 | 8.66 | -1669.29 | 1669.89 | 9.13 | 3.37 | -4.28 | 10.72 | 1669.29 | -3.37 | 72.8338
0.7852 | 4.95 | 4.64 | -1654.31 | 1654.77 | 8.85 | 3.39 | -4.27 | 10.49 | 1654.31 | -3.39 | 73.0011
0.787 | 2.68 | 3.94 | -1634.74 | 1635.25 | 8.67 | 3.33 | -4.23 | 10.32 | 1634.74 | -3.33 | 73.1685
0.7888 | 0.96 | 3.55 | -1614.37 | 1614.94 | 8.5 | 3.24 | -4.17 | 10.14 | 1614.37 | -3.24 | 73.3358
0.7906 | 0.72 | 4.76 | -1586.33 | 1586.95 | 8.35 | 3.11 | -4.12 | 9.94 | 1586.33 | -3.11 | 73.5032
0.7924 | 0.63 | 7.24 | -1556.12 | 1556.77 | 8.23 | 2.98 | -4.07 | 9.79 | 1556.12 | -2.98 | 73.6705
0.7942 | 1.27 | 13.5 | -1523.26 | 1523.93 | 8.2 | 2.83 | -3.99 | 9.69 | 1523.26 | -2.83 | 73.8379
0.796 | 1.33 | 17.69 | -1495.68 | 1496.4 | 8.12 | 2.65 | -3.89 | 9.52 | 1495.68 | -2.65 | 74.0052
0.7977 | 1.55 | 20.84 | -1468.84 | 1469.6 | 8.05 | 2.47 | -3.81 | 9.38 | 1468.84 | -2.47 | 74.1633
0.7995 | 1.46 | 22.66 | -1445.67 | 1446.49 | 7.94 | 2.32 | -3.73 | 9.21 | 1445.67 | -2.32 | 74.3306
0.8013 | 1.91 | 24.31 | -1423.09 | 1423.99 | 7.79 | 2.15 | -3.66 | 9 | 1423.09 | -2.15 | 74.498
0.8031 | 2.41 | 25.38 | -1401.82 | 1402.76 | 7.63 | 1.97 | -3.59 | 8.78 | 1401.82 | -1.97 | 74.6653
0.8049 | 2.81 | 24.99 | -1385 | 1385.93 | 7.43 | 1.78 | -3.52 | 8.53 | 1385 | -1.78 | 74.8327
0.8067 | 2.84 | 24.82 | -1370.8 | 1371.74 | 7.19 | 1.6 | -3.44 | 8.23 | 1370.8 | -1.6 | 75
0.8085 | 5.33 | 27.76 | -1350.74 | 1351.76 | 7.15 | 1.34 | -3.35 | 8.11 | 1350.74 | -1.34 | 75.1673
0.8103 | 8.46 | 32.06 | -1327.95 | 1328.97 | 7.19 | 1.05 | -3.24 | 8.09 | 1327.95 | -1.05 | 75.3347
0.8121 | 11.02 | 36.43 | -1303.76 | 1304.91 | 7.23 | 0.82 | -3.15 | 8.08 | 1303.76 | -0.82 | 75.502
0.8139 | 12.77 | 39.89 | -1278.99 | 1280.29 | 7.22 | 0.65 | -3.05 | 8.04 | 1278.99 | -0.65 | 75.6694
0.8157 | 14.15 | 43.1 | -1254.02 | 1255.36 | 7.2 | 0.53 | -2.96 | 7.99 | 1254.02 | -0.53 | 75.8367
0.8175 | 14.15 | 44.42 | -1234.73 | 1236.16 | 7.1 | 0.45 | -2.89 | 7.85 | 1234.73 | -0.45 | 76.0041
0.8193 | 12.91 | 43.67 | -1218.6 | 1220 | 6.92 | 0.41 | -2.82 | 7.66 | 1218.6 | -0.41 | 76.1714
0.8211 | 11.83 | 43.13 | -1202.57 | 1203.97 | 6.74 | 0.36 | -2.75 | 7.47 | 1202.57 | -0.36 | 76.3388
0.8228 | 10.97 | 42.48 | -1187.53 | 1188.9 | 6.53 | 0.32 | -2.71 | 7.25 | 1187.53 | -0.32 | 76.4968
0.8246 | 10.32 | 42.55 | -1170.06 | 1171.44 | 6.32 | 0.29 | -2.67 | 7.04 | 1170.06 | -0.29 | 76.6642
0.8264 | 9.91 | 44.4 | -1153.07 | 1154.61 | 6.18 | 0.28 | -2.63 | 6.89 | 1153.07 | -0.28 | 76.8315
0.8282 | 8.42 | 43.78 | -1140.57 | 1142.13 | 6 | 0.31 | -2.58 | 6.71 | 1140.57 | -0.31 | 76.9989
0.83 | 7.17 | 43.9 | -1129.37 | 1130.94 | 5.88 | 0.37 | -2.57 | 6.6 | 1129.37 | -0.37 | 77.1662
0.8318 | 6.04 | 44.07 | -1118.71 | 1120.31 | 5.76 | 0.43 | -2.57 | 6.5 | 1118.71 | -0.43 | 77.3336
0.8336 | 5.04 | 43.35 | -1111.32 | 1112.94 | 5.62 | 0.52 | -2.56 | 6.37 | 1111.32 | -0.52 | 77.5009
0.8354 | 4.01 | 43.14 | -1104.01 | 1105.67 | 5.48 | 0.57 | -2.54 | 6.24 | 1104.01 | -0.57 | 77.6683
0.8372 | 2.48 | 43.18 | -1097.53 | 1099.34 | 5.35 | 0.64 | -2.54 | 6.11 | 1097.53 | -0.64 | 77.8356
0.839 | 0.1 | 42.01 | -1092.19 | 1094.18 | 5.19 | 0.76 | -2.54 | 5.97 | 1092.19 | -0.76 | 78.003
0.8408 | -0.15 | 43.28 | -1085.45 | 1087.65 | 5.1 | 0.81 | -2.56 | 5.9 | 1085.45 | -0.81 | 78.1703
0.8426 | -0.13 | 45.34 | -1078.31 | 1080.82 | 5.09 | 0.91 | -2.6 | 5.91 | 1078.31 | -0.91 | 78.3377
0.8444 | -1.47 | 46.14 | -1071.96 | 1074.7 | 4.99 | 1.04 | -2.64 | 5.87 | 1071.96 | -1.04 | 78.505
0.8462 | -3.53 | 46.16 | -1067.17 | 1070.07 | 4.83 | 1.18 | -2.68 | 5.81 | 1067.17 | -1.18 | 78.6724
0.8479 | -6.09 | 46.2 | -1064.28 | 1067.27 | 4.66 | 1.36 | -2.74 | 5.77 | 1064.28 | -1.36 | 78.8304
0.8497 | -7.47 | 47.35 | -1063.11 | 1066.19 | 4.61 | 1.52 | -2.82 | 5.82 | 1063.11 | -1.52 | 78.9978
0.8515 | -11.09 | 45.32 | -1064.15 | 1067.34 | 4.45 | 1.77 | -2.86 | 5.82 | 1064.15 | -1.77 | 79.1651
0.8533 | -13.82 | 43.6 | -1066.68 | 1069.77 | 4.32 | 1.98 | -2.9 | 5.83 | 1066.68 | -1.98 | 79.3325
0.8551 | -16.35 | 41.79 | -1068.72 | 1071.65 | 4.19 | 2.2 | -2.93 | 5.87 | 1068.72 | -2.2 | 79.4998
0.8569 | -19.27 | 39.35 | -1068.74 | 1071.69 | 4.07 | 2.45 | -2.96 | 5.92 | 1068.74 | -2.45 | 79.6672
0.8587 | -21.77 | 37.38 | -1067.87 | 1070.8 | 3.95 | 2.71 | -2.97 | 6 | 1067.87 | -2.71 | 79.8345
0.8605 | -23.73 | 36.11 | -1066.77 | 1069.63 | 3.87 | 2.97 | -2.99 | 6.14 | 1066.77 | -2.97 | 80.0019
0.8623 | -25.36 | 35.19 | -1065.14 | 1067.97 | 3.82 | 3.21 | -3.04 | 6.33 | 1065.14 | -3.21 | 80.1692
0.8641 | -26.99 | 34.19 | -1063.45 | 1066.25 | 3.77 | 3.46 | -3.08 | 6.52 | 1063.45 | -3.46 | 80.3366
0.8659 | -28.54 | 32.05 | -1060.83 | 1063.5 | 3.69 | 3.66 | -3.04 | 6.6 | 1060.83 | -3.66 | 80.5039
0.8677 | -30.52 | 29.65 | -1056.8 | 1059.4 | 3.62 | 3.85 | -3 | 6.66 | 1056.8 | -3.85 | 80.6713
0.8695 | -31.89 | 28.07 | -1051.54 | 1054.13 | 3.61 | 4.01 | -2.98 | 6.73 | 1051.54 | -4.01 | 80.8386
0.8712 | -31.86 | 27.89 | -1047.59 | 1050.18 | 3.67 | 4.07 | -2.97 | 6.79 | 1047.59 | -4.07 | 80.9967
0.873 | -30.46 | 28.75 | -1041.6 | 1044.24 | 3.75 | 4.1 | -2.96 | 6.84 | 1041.6 | -4.1 | 81.164
0.8748 | -30.89 | 27.86 | -1038.26 | 1040.99 | 3.8 | 4.17 | -2.95 | 6.87 | 1038.26 | -4.17 | 81.3313
0.8766 | -32.37 | 24.53 | -1037.47 | 1040.03 | 3.75 | 4.19 | -2.96 | 6.88 | 1037.47 | -4.19 | 81.4987
0.8784 | -34.35 | 20.6 | -1037.84 | 1040.4 | 3.69 | 4.25 | -2.95 | 6.85 | 1037.84 | -4.25 | 81.666
0.8802 | -34.68 | 18.05 | -1035 | 1037.51 | 3.66 | 4.24 | -2.94 | 6.83 | 1035 | -4.24 | 81.8334
0.882 | -34.73 | 15.41 | -1031.91 | 1034.32 | 3.67 | 4.15 | -2.98 | 6.8 | 1031.91 | -4.15 | 82.0007
0.8838 | -35.11 | 12.68 | -1028.72 | 1031.03 | 3.65 | 4.06 | -2.99 | 6.77 | 1028.72 | -4.06 | 82.1681
0.8856 | -36.93 | 9.76 | -1025.45 | 1027.63 | 3.51 | 4.08 | -2.87 | 6.72 | 1025.45 | -4.08 | 82.3354
0.8874 | -37.16 | 7.89 | -1019.5 | 1021.57 | 3.47 | 4.01 | -2.8 | 6.63 | 1019.5 | -4.01 | 82.5028
0.8892 | -36.06 | 6.91 | -1011.27 | 1013.25 | 3.51 | 3.85 | -2.76 | 6.5 | 1011.27 | -3.85 | 82.6701
0.891 | -34.3 | 8.43 | -1000.01 | 1001.8 | 3.59 | 3.76 | -2.73 | 6.5 | 1000.01 | -3.76 | 82.8375
0.8928 | -32.02 | 10.71 | -990.75 | 992.33 | 3.6 | 3.63 | -2.66 | 6.42 | 990.75 | -3.63 | 83.0048
0.8946 | -28.09 | 13.02 | -977.69 | 978.97 | 3.63 | 3.38 | -2.59 | 6.27 | 977.69 | -3.38 | 83.1722
0.8963 | -23.19 | 17.44 | -958.55 | 959.53 | 3.65 | 3.15 | -2.57 | 6.18 | 958.55 | -3.15 | 83.3302
0.8981 | -20.42 | 19.9 | -942.42 | 943.27 | 3.61 | 3 | -2.58 | 6.04 | 942.42 | -3 | 83.4976
0.8999 | -19.48 | 20.82 | -930.3 | 931.08 | 3.55 | 2.86 | -2.59 | 5.81 | 930.3 | -2.86 | 83.6649
0.9017 | -20.18 | 20.47 | -921.66 | 922.4 | 3.47 | 2.75 | -2.58 | 5.56 | 921.66 | -2.75 | 83.8323
0.9035 | -21.74 | 18.1 | -917.13 | 917.79 | 3.48 | 2.57 | -2.61 | 5.38 | 917.13 | -2.57 | 83.9996
0.9053 | -22.47 | 15.93 | -913.21 | 913.83 | 3.55 | 2.34 | -2.66 | 5.28 | 913.21 | -2.34 | 84.167
0.9071 | -23.2 | 13.76 | -909.29 | 909.86 | 3.63 | 2.12 | -2.7 | 5.19 | 909.29 | -2.12 | 84.3343
0.9089 | -25.1 | 12.44 | -907.82 | 908.42 | 3.59 | 2.02 | -2.7 | 5.09 | 907.82 | -2.02 | 84.5017
0.9107 | -26.16 | 12.54 | -906.21 | 906.84 | 3.54 | 1.92 | -2.68 | 5 | 906.21 | -1.92 | 84.669
0.9125 | -27.13 | 11.18 | -907.4 | 908.03 | 3.48 | 1.78 | -2.69 | 4.91 | 907.4 | -1.78 | 84.8364
0.9143 | -26.19 | 11.7 | -906.95 | 907.51 | 3.48 | 1.5 | -2.71 | 4.85 | 906.95 | -1.5 | 85.0037
0.9161 | -25.3 | 14.69 | -906.08 | 906.63 | 3.58 | 1.12 | -2.74 | 4.95 | 906.08 | -1.12 | 85.1711
0.9179 | -24.41 | 17.69 | -905.21 | 905.75 | 3.69 | 0.74 | -2.78 | 5.05 | 905.21 | -0.74 | 85.3384
0.9197 | -23.29 | 19 | -902.79 | 903.3 | 3.75 | 0.36 | -2.87 | 5.19 | 902.79 | -0.36 | 85.5058
0.9214 | -22.32 | 19.75 | -900.89 | 901.39 | 3.82 | -0.04 | -2.97 | 5.4 | 900.89 | 0.04 | 85.6638
0.9232 | -21.53 | 20.32 | -899.79 | 900.28 | 3.88 | -0.42 | -3.07 | 5.6 | 899.79 | 0.42 | 85.8312
0.925 | -21.81 | 20.07 | -901.6 | 902.1 | 3.94 | -0.74 | -3.17 | 5.76 | 901.6 | 0.74 | 85.9985
0.9268 | -23.63 | 19.11 | -905.3 | 905.83 | 3.97 | -1.03 | -3.25 | 5.9 | 905.3 | 1.03 | 86.1659
0.9286 | -25.92 | 15.83 | -911.46 | 912 | 3.87 | -1.3 | -3.36 | 6 | 911.46 | 1.3 | 86.3332
0.9304 | -26.32 | 14.51 | -914.92 | 915.45 | 3.84 | -1.61 | -3.48 | 6.16 | 914.92 | 1.61 | 86.5006
0.9322 | -24.12 | 16.42 | -914.3 | 914.81 | 3.95 | -1.98 | -3.59 | 6.42 | 914.3 | 1.98 | 86.6679
0.934 | -21.85 | 20.35 | -912.17 | 912.71 | 4.12 | -2.32 | -3.72 | 6.7 | 912.17 | 2.32 | 86.8353
0.9358 | -19.83 | 24.77 | -911.33 | 911.94 | 4.31 | -2.63 | -3.85 | 6.99 | 911.33 | 2.63 | 87.0026
0.9376 | -17.54 | 28.54 | -912.24 | 912.91 | 4.46 | -2.96 | -3.98 | 7.31 | 912.24 | 2.96 | 87.17
0.9394 | -15.34 | 31.58 | -914.51 | 915.23 | 4.58 | -3.29 | -4.11 | 7.62 | 914.51 | 3.29 | 87.3373
0.9412 | -13.04 | 33.11 | -919.67 | 920.42 | 4.62 | -3.58 | -4.24 | 7.9 | 919.67 | 3.58 | 87.5046
0.943 | -12.56 | 33.17 | -925.62 | 926.35 | 4.58 | -3.69 | -4.43 | 8.04 | 925.62 | 3.69 | 87.672
0.9447 | -11.6 | 35.53 | -925.48 | 926.31 | 4.62 | -3.72 | -4.65 | 8.17 | 925.48 | 3.72 | 87.83
0.9465 | -10.56 | 38.02 | -930.4 | 931.33 | 4.68 | -3.86 | -4.78 | 8.3 | 930.4 | 3.86 | 87.9974
0.9483 | -10.24 | 39.44 | -937.21 | 938.2 | 4.69 | -4.09 | -4.9 | 8.46 | 937.21 | 4.09 | 88.1647
0.9501 | -10.27 | 40.73 | -944.58 | 945.61 | 4.68 | -4.32 | -5.05 | 8.59 | 944.58 | 4.32 | 88.3321
0.9519 | -9.41 | 42.58 | -952.53 | 953.62 | 4.7 | -4.56 | -5.18 | 8.74 | 952.53 | 4.56 | 88.4994
0.9537 | -9.16 | 43.53 | -961.44 | 962.55 | 4.7 | -4.73 | -5.31 | 8.86 | 961.44 | 4.73 | 88.6668
0.9555 | -9.84 | 43.01 | -971.95 | 973.02 | 4.66 | -4.79 | -5.44 | 8.92 | 971.95 | 4.79 | 88.8341
0.9573 | -7.71 | 43.33 | -984.61 | 985.65 | 4.68 | -5.07 | -5.49 | 9.15 | 984.61 | 5.07 | 89.0015
0.9591 | -6.8 | 42.16 | -1000.53 | 1001.5 | 4.65 | -5.24 | -5.53 | 9.28 | 1000.53 | 5.24 | 89.1688
0.9609 | -5.52 | 42.02 | -1022.14 | 1023.08 | 4.66 | -5.49 | -5.52 | 9.47 | 1022.14 | 5.49 | 89.3362
0.9627 | -6.22 | 41.17 | -1037.57 | 1038.48 | 4.63 | -5.56 | -5.58 | 9.53 | 1037.57 | 5.56 | 89.5035
0.9645 | -7.51 | 39.28 | -1052.48 | 1053.35 | 4.59 | -5.61 | -5.62 | 9.57 | 1052.48 | 5.61 | 89.6709
0.9663 | -9.58 | 34.6 | -1068.3 | 1069.09 | 4.44 | -5.7 | -5.68 | 9.6 | 1068.3 | 5.7 | 89.8382
0.9681 | -11.65 | 29.91 | -1084.12 | 1084.84 | 4.29 | -5.79 | -5.73 | 9.63 | 1084.12 | 5.79 | 90.0056
0.9698 | -12.28 | 27.81 | -1095.38 | 1096.08 | 4.23 | -5.91 | -5.81 | 9.71 | 1095.38 | 5.91 | 90.1636
0.9716 | -11.99 | 27.33 | -1102.96 | 1103.67 | 4.23 | -6.03 | -5.89 | 9.83 | 1102.96 | 6.03 | 90.331
0.9734 | -12.05 | 26.16 | -1109.65 | 1110.32 | 4.21 | -6.17 | -5.97 | 9.94 | 1109.65 | 6.17 | 90.4983
0.9752 | -11.69 | 25.02 | -1115.69 | 1116.32 | 4.2 | -6.31 | -6.03 | 10.07 | 1115.69 | 6.31 | 90.6657
0.977 | -11.18 | 25.04 | -1119.81 | 1120.42 | 4.26 | -6.38 | -6.08 | 10.17 | 1119.81 | 6.38 | 90.833
0.9788 | -10.19 | 25.62 | -1121.05 | 1121.65 | 4.34 | -6.44 | -6.15 | 10.27 | 1121.05 | 6.44 | 91.0004
0.9806 | -8.43 | 27.37 | -1122.07 | 1122.7 | 4.45 | -6.52 | -6.19 | 10.39 | 1122.07 | 6.52 | 91.1677
0.9824 | -7.34 | 27.76 | -1122.66 | 1123.29 | 4.47 | -6.59 | -6.23 | 10.46 | 1122.66 | 6.59 | 91.3351
0.9842 | -7.54 | 26.74 | -1125.89 | 1126.48 | 4.42 | -6.55 | -6.28 | 10.47 | 1125.89 | 6.55 | 91.5024
0.986 | -6.32 | 28.3 | -1126.96 | 1127.55 | 4.51 | -6.57 | -6.33 | 10.56 | 1126.96 | 6.57 | 91.6698
0.9878 | -3.92 | 32.38 | -1123.76 | 1124.41 | 4.72 | -6.62 | -6.39 | 10.74 | 1123.76 | 6.62 | 91.8371
0.9896 | -1.94 | 36.83 | -1118.09 | 1118.86 | 4.96 | -6.63 | -6.46 | 10.91 | 1118.09 | 6.63 | 92.0045
0.9914 | -1.13 | 39.38 | -1111.39 | 1112.25 | 5.16 | -6.63 | -6.55 | 11.05 | 1111.39 | 6.63 | 92.1718
0.9932 | -1.08 | 40.57 | -1103.46 | 1104.37 | 5.35 | -6.62 | -6.67 | 11.18 | 1103.46 | 6.62 | 92.3392
0.9949 | -0.31 | 43.17 | -1093.06 | 1094.04 | 5.63 | -6.66 | -6.76 | 11.39 | 1093.06 | 6.66 | 92.4972
0.9967 | 0.86 | 46.33 | -1083.29 | 1084.39 | 5.93 | -6.72 | -6.84 | 11.62 | 1083.29 | 6.72 | 92.6646
0.9985 | 3.3 | 49.08 | -1076.98 | 1078.2 | 6.16 | -6.86 | -6.86 | 11.87 | 1076.98 | 6.86 | 92.8319
1.0003 | 5.74 | 51.82 | -1070.67 | 1072.01 | 6.4 | -6.99 | -6.89 | 12.11 | 1070.67 | 6.99 | 92.9993
1.0021 | 6.6 | 54.09 | -1066.88 | 1068.37 | 6.58 | -7.02 | -6.92 | 12.27 | 1066.88 | 7.02 | 93.1666
1.0039 | 6.75 | 54.23 | -1067.86 | 1069.37 | 6.65 | -7.03 | -6.94 | 12.34 | 1067.86 | 7.03 | 93.334
1.0057 | 6.59 | 53.66 | -1070.08 | 1071.57 | 6.69 | -7.02 | -6.95 | 12.37 | 1070.08 | 7.02 | 93.5013
1.0075 | 5.06 | 51.23 | -1074.28 | 1075.64 | 6.67 | -6.94 | -7 | 12.35 | 1074.28 | 6.94 | 93.6687
1.0093 | 2.58 | 47.71 | -1078.81 | 1080.03 | 6.61 | -6.78 | -7.09 | 12.3 | 1078.81 | 6.78 | 93.836
1.0111 | 1.69 | 46.63 | -1084.71 | 1085.9 | 6.65 | -6.69 | -7.12 | 12.29 | 1084.71 | 6.69 | 94.0033
1.0129 | 0.38 | 45.59 | -1090.32 | 1091.49 | 6.68 | -6.54 | -7.18 | 12.26 | 1090.32 | 6.54 | 94.1707
1.0147 | 0.39 | 45.57 | -1093.29 | 1094.44 | 6.76 | -6.51 | -7.2 | 12.3 | 1093.29 | 6.51 | 94.338
1.0165 | -0.06 | 44.8 | -1096.61 | 1097.73 | 6.82 | -6.46 | -7.23 | 12.32 | 1096.61 | 6.46 | 94.5054
1.0182 | -0.67 | 44.27 | -1098.25 | 1099.34 | 6.87 | -6.48 | -7.22 | 12.36 | 1098.25 | 6.48 | 94.6634
1.02 | 0.76 | 47 | -1096.82 | 1098 | 7.06 | -6.54 | -7.23 | 12.52 | 1096.82 | 6.54 | 94.8308
1.0218 | 1.37 | 49.11 | -1094.72 | 1096 | 7.2 | -6.59 | -7.26 | 12.67 | 1094.72 | 6.59 | 94.9981
1.0236 | 0.44 | 48.41 | -1098.17 | 1099.4 | 7.21 | -6.6 | -7.28 | 12.69 | 1098.17 | 6.6 | 95.1655
1.0254 | -1.84 | 45.02 | -1104.42 | 1105.49 | 7.08 | -6.58 | -7.29 | 12.63 | 1104.42 | 6.58 | 95.3328
1.0272 | -4.22 | 41.58 | -1111.84 | 1112.76 | 6.94 | -6.56 | -7.29 | 12.54 | 1111.84 | 6.56 | 95.5002
1.029 | -6.75 | 37.46 | -1119.44 | 1120.22 | 6.8 | -6.51 | -7.28 | 12.44 | 1119.44 | 6.51 | 95.6675
1.0308 | -9.33 | 33.21 | -1126.69 | 1127.33 | 6.65 | -6.45 | -7.27 | 12.34 | 1126.69 | 6.45 | 95.8349
1.0326 | -10.54 | 30.95 | -1131.45 | 1132.03 | 6.59 | -6.43 | -7.27 | 12.29 | 1131.45 | 6.43 | 96.0022
1.0344 | -11.38 | 28.89 | -1134.5 | 1135.02 | 6.52 | -6.41 | -7.28 | 12.22 | 1134.5 | 6.41 | 96.1696
1.0362 | -12.06 | 27.26 | -1135.42 | 1135.91 | 6.48 | -6.37 | -7.29 | 12.18 | 1135.42 | 6.37 | 96.3369
1.038 | -12.55 | 26.23 | -1134.11 | 1134.58 | 6.51 | -6.34 | -7.31 | 12.19 | 1134.11 | 6.34 | 96.5043
1.0398 | -12.59 | 25.48 | -1132.04 | 1132.51 | 6.58 | -6.3 | -7.31 | 12.2 | 1132.04 | 6.3 | 96.6716
1.0416 | -13.28 | 24.65 | -1128.65 | 1129.13 | 6.68 | -6.25 | -7.3 | 12.22 | 1128.65 | 6.25 | 96.839
1.0433 | -14.78 | 22.55 | -1128.55 | 1129.04 | 6.73 | -6.2 | -7.27 | 12.21 | 1128.55 | 6.2 | 96.997
1.0451 | -15.6 | 21.32 | -1125.54 | 1126.06 | 6.8 | -6.16 | -7.27 | 12.21 | 1125.54 | 6.16 | 97.1644
1.0469 | -14.72 | 22.32 | -1120.35 | 1120.88 | 6.94 | -6.16 | -7.25 | 12.27 | 1120.35 | 6.16 | 97.3317
1.0487 | -14.9 | 21.48 | -1115.56 | 1116.13 | 7.03 | -6.22 | -7.19 | 12.3 | 1115.56 | 6.22 | 97.4991
1.0505 | -13.33 | 23.47 | -1106.85 | 1107.4 | 7.21 | -6.34 | -7.13 | 12.42 | 1106.85 | 6.34 | 97.6664
1.0523 | -11.17 | 26.78 | -1097.03 | 1097.56 | 7.39 | -6.46 | -7.09 | 12.57 | 1097.03 | 6.46 | 97.8338
1.0541 | -9.84 | 27.84 | -1091.02 | 1091.54 | 7.47 | -6.52 | -7.06 | 12.62 | 1091.02 | 6.52 | 98.0011
1.0559 | -10 | 26.82 | -1087.96 | 1088.46 | 7.46 | -6.54 | -7.04 | 12.61 | 1087.96 | 6.54 | 98.1685
1.0577 | -11.55 | 23.23 | -1085.69 | 1086.2 | 7.3 | -6.51 | -7.01 | 12.47 | 1085.69 | 6.51 | 98.3358
1.0595 | -12.02 | 22.81 | -1083.13 | 1083.64 | 7.28 | -6.48 | -6.98 | 12.42 | 1083.13 | 6.48 | 98.5032
1.0613 | -12.38 | 22.61 | -1080.02 | 1080.53 | 7.27 | -6.44 | -6.96 | 12.38 | 1080.02 | 6.44 | 98.6705
1.0631 | -13.15 | 21.08 | -1079.42 | 1079.92 | 7.24 | -6.4 | -6.94 | 12.31 | 1079.42 | 6.4 | 98.8379
1.0649 | -14.64 | 18.01 | -1082.34 | 1082.8 | 7.15 | -6.34 | -6.9 | 12.18 | 1082.34 | 6.34 | 99.0052
1.0667 | -16.5 | 15.66 | -1090.05 | 1090.48 | 7.05 | -6.2 | -6.84 | 12 | 1090.05 | 6.2 | 99.1726
1.0684 | -17.32 | 15.23 | -1097.31 | 1097.72 | 7.05 | -6.07 | -6.77 | 11.87 | 1097.31 | 6.07 | 99.3306
1.0702 | -17.35 | 16.61 | -1095.7 | 1096.12 | 7.11 | -6.05 | -6.73 | 11.87 | 1095.7 | 6.05 | 99.498
1.072 | -16.6 | 17.32 | -1092.63 | 1093.02 | 7.14 | -6.1 | -6.68 | 11.9 | 1092.63 | 6.1 | 99.6653
1.0738 | -15.84 | 18.03 | -1089.56 | 1089.92 | 7.16 | -6.16 | -6.64 | 11.92 | 1089.56 | 6.16 | 99.8327
1.0756 | -15.86 | 18.52 | -1089.84 | 1090.21 | 7.23 | -6.14 | -6.59 | 11.94 | 1089.84 | 6.14 | 100
| Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My
Minimum | -66.18 | -58.86 | -2279.75 | 745.11 | 3.47 | -7.03 | -7.31 | 4.85 | 743.9 | -12.49
Maximum | 15.38 | 54.23 | -743.9 | 2280.38 | 14.85 | 12.49 | -2.54 | 19.17 | 2279.75 | 7.03

## Diagr_Individ
Individual Loads
Contact forces and moments, measured telemetrically in knee implants of 8 subjects.
Averages of one individual subject.
Fx  lateral, Fy  anterior, -Fz  distal, Fres = Resultant Force
Mx around x, My around y, Mz around z, Mres = Resultant Moment
z-axis = axis of implant stem!
From Bergmann et al, PlosOne (2014) "Standardized Loads Acting in Knee Implants"
Subject K3R
Subject K3R

## Dat_BW
Data File: | k3r_2239_0_3_i1.akf
Curve Alignment Program: | 1.5.7
------------------------MEASURING INFOS------------------------
Diagram Title #1 | Forces and Moments at Knee Joint
Diagram Title #2 | Measured Relative to Implant
Comment #1 | Knee Joint; Gaitanalysis; Deep Knee Bend;
Comment #2 | K3R, 8 Months PO *
BodyWeight [N]: | 960
Patient(en): | K3R | [7 empty] | Subject | K3R
Implant Type: | Knee Joint
Varus Angle Leg [°] | 3.5 | [7 empty] | Normalized Loads in %BW (Forces) and %BWm (moments)
Axes Directions X, Y, Z: | Lateral, Anterior, Superior
Axes Origin: | Tibial Plateau at Stem Axis | [7 empty] | BW [kg] = | 97.8593
Transformation X Y Z (°): | 10 | 0.5 | 10.5
Displacement X, Y, Z (mm): | 0 | 0 | 0
Activity Code: | KneeActivities | 1P_1_15_3
---------------------------DATA--------------------------------
Number of Data Sets = | 601
Number of analog channels = | 0
Time-Offset(sec) = | 0
Time | Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My | Time
[s] | [%BW] | [%BW] | [%BW] | [%BW] | [%BWm] | [%BWm] | [%BWm] | [%BWm] | [%BW] | [%BWm] | [%Cycle]
0 | -3.0833 | -1.199 | -95.874 | 95.9469 | 0.71979 | -0.24792 | -0.64271 | 1.0333 | 95.874 | 0.24792 | 0
0.0018 | -3.1104 | -1.2115 | -95.3875 | 95.4604 | 0.72292 | -0.24479 | -0.64792 | 1.0375 | 95.3875 | 0.24479 | 0.16735
0.0036 | -3.1594 | -1.3042 | -95.2792 | 95.3521 | 0.72083 | -0.24271 | -0.65104 | 1.0375 | 95.2792 | 0.24271 | 0.3347
0.0054 | -3.2156 | -1.4 | -95.4479 | 95.524 | 0.71875 | -0.24063 | -0.65208 | 1.0365 | 95.4479 | 0.24063 | 0.50205
0.0072 | -3.325 | -1.5854 | -95.7469 | 95.8333 | 0.71354 | -0.23646 | -0.65208 | 1.0323 | 95.7469 | 0.23646 | 0.66939
0.009 | -3.4052 | -1.6542 | -96.3698 | 96.4604 | 0.70937 | -0.23646 | -0.64583 | 1.0271 | 96.3698 | 0.23646 | 0.83674
0.0108 | -3.399 | -1.5917 | -96.7177 | 96.8104 | 0.71042 | -0.23854 | -0.64167 | 1.026 | 96.7177 | 0.23854 | 1.0041
0.0125 | -3.4479 | -1.6104 | -97.0385 | 97.1354 | 0.70833 | -0.2375 | -0.63854 | 1.024 | 97.0385 | 0.2375 | 1.1621
0.0143 | -3.5771 | -1.7813 | -97.2073 | 97.3167 | 0.70208 | -0.23229 | -0.63854 | 1.0198 | 97.2073 | 0.23229 | 1.3295
0.0161 | -3.5344 | -1.724 | -97.0792 | 97.1854 | 0.70833 | -0.23333 | -0.63958 | 1.025 | 97.0792 | 0.23333 | 1.4968
0.0179 | -3.5312 | -1.7667 | -96.5927 | 96.6948 | 0.71146 | -0.23021 | -0.64479 | 1.0271 | 96.5927 | 0.23021 | 1.6642
0.0197 | -3.7198 | -2.0479 | -95.9542 | 96.0604 | 0.70417 | -0.21979 | -0.65312 | 1.0229 | 95.9542 | 0.21979 | 1.8315
0.0215 | -3.826 | -2.2385 | -95.8104 | 95.9208 | 0.70104 | -0.21458 | -0.65625 | 1.0187 | 95.8104 | 0.21458 | 1.9989
0.0233 | -3.8573 | -2.2573 | -95.7771 | 95.8896 | 0.70521 | -0.2125 | -0.65729 | 1.0219 | 95.7771 | 0.2125 | 2.1662
0.0251 | -3.9302 | -2.3104 | -95.8479 | 95.9646 | 0.70937 | -0.20937 | -0.65625 | 1.024 | 95.8479 | 0.20937 | 2.3336
0.0269 | -4.0177 | -2.3833 | -95.9427 | 96.0646 | 0.71458 | -0.20625 | -0.65625 | 1.0271 | 95.9427 | 0.20625 | 2.5009
0.0287 | -4.0708 | -2.426 | -96.1156 | 96.2396 | 0.72083 | -0.20417 | -0.65521 | 1.0312 | 96.1156 | 0.20417 | 2.6683
0.0305 | -4.1729 | -2.5198 | -96.3479 | 96.4802 | 0.72396 | -0.2 | -0.65312 | 1.0323 | 96.3479 | 0.2 | 2.8356
0.0323 | -4.2719 | -2.6042 | -96.7938 | 96.9302 | 0.72396 | -0.19792 | -0.64687 | 1.0281 | 96.7938 | 0.19792 | 3.003
0.0341 | -4.3042 | -2.6198 | -97.1677 | 97.3063 | 0.72813 | -0.19687 | -0.64167 | 1.0281 | 97.1677 | 0.19687 | 3.1703
0.0359 | -4.3365 | -2.6354 | -97.5417 | 97.6823 | 0.73229 | -0.19583 | -0.63542 | 1.0281 | 97.5417 | 0.19583 | 3.3377
0.0376 | -4.4365 | -2.7719 | -97.7292 | 97.876 | 0.73646 | -0.18958 | -0.63438 | 1.0292 | 97.7292 | 0.18958 | 3.4957
0.0394 | -4.4958 | -2.8354 | -97.7104 | 97.8615 | 0.74271 | -0.18542 | -0.63542 | 1.0333 | 97.7104 | 0.18542 | 3.6631
0.0412 | -4.4875 | -2.7917 | -97.699 | 97.8479 | 0.75521 | -0.18333 | -0.63542 | 1.0427 | 97.699 | 0.18333 | 3.8304
0.043 | -4.5427 | -2.8677 | -97.7844 | 97.9396 | 0.76354 | -0.17917 | -0.63542 | 1.049 | 97.7844 | 0.17917 | 3.9978
0.0448 | -4.6437 | -3.0042 | -97.9198 | 98.0854 | 0.76771 | -0.175 | -0.63646 | 1.0531 | 97.9198 | 0.175 | 4.1651
0.0466 | -4.7521 | -3.1406 | -98.0229 | 98.201 | 0.76979 | -0.17083 | -0.63646 | 1.0562 | 98.0229 | 0.17083 | 4.3325
0.0484 | -4.8083 | -3.2542 | -98.0448 | 98.2313 | 0.77187 | -0.16875 | -0.63646 | 1.0573 | 98.0448 | 0.16875 | 4.4998
0.0502 | -4.8396 | -3.3448 | -98.0021 | 98.1937 | 0.775 | -0.16667 | -0.63646 | 1.0594 | 98.0021 | 0.16667 | 4.6672
0.052 | -4.851 | -3.4125 | -97.7427 | 97.9333 | 0.78333 | -0.1625 | -0.63854 | 1.0646 | 97.7427 | 0.1625 | 4.8345
0.0538 | -4.9469 | -3.5521 | -97.3385 | 97.5365 | 0.79375 | -0.15 | -0.64479 | 1.0719 | 97.3385 | 0.15 | 5.0019
0.0556 | -5.0438 | -3.6906 | -96.9333 | 97.1396 | 0.80312 | -0.13854 | -0.65104 | 1.0781 | 96.9333 | 0.13854 | 5.1692
0.0574 | -5.1208 | -3.8042 | -96.5948 | 96.8083 | 0.81354 | -0.12708 | -0.65625 | 1.0854 | 96.5948 | 0.12708 | 5.3366
0.0592 | -5.0667 | -3.7646 | -96.6135 | 96.8219 | 0.825 | -0.12604 | -0.65312 | 1.0896 | 96.6135 | 0.12604 | 5.5039
0.061 | -5 | -3.726 | -96.6375 | 96.8417 | 0.83542 | -0.125 | -0.64896 | 1.0958 | 96.6375 | 0.125 | 5.6713
0.0627 | -4.9938 | -3.7271 | -96.7021 | 96.9042 | 0.85 | -0.11458 | -0.65 | 1.1052 | 96.7021 | 0.11458 | 5.8293
0.0645 | -5.0333 | -3.8281 | -96.5312 | 96.7396 | 0.84896 | -0.11146 | -0.65 | 1.1042 | 96.5312 | 0.11146 | 5.9967
0.0663 | -5.051 | -3.9 | -96.2615 | 96.474 | 0.85104 | -0.10729 | -0.65104 | 1.1062 | 96.2615 | 0.10729 | 6.164
0.0681 | -5.001 | -3.8833 | -95.6344 | 95.8448 | 0.86146 | -0.1 | -0.65625 | 1.1156 | 95.6344 | 0.1 | 6.3313
0.0699 | -4.8635 | -3.6146 | -95.201 | 95.3937 | 0.8875 | -0.089583 | -0.65729 | 1.1365 | 95.201 | 0.089583 | 6.4987
0.0717 | -4.7542 | -3.4406 | -94.4354 | 94.6177 | 0.90104 | -0.085417 | -0.65938 | 1.149 | 94.4354 | 0.085417 | 6.666
0.0735 | -4.6604 | -3.3323 | -93.5938 | 93.7687 | 0.91042 | -0.083333 | -0.6625 | 1.1583 | 93.5938 | 0.083333 | 6.8334
0.0753 | -4.6188 | -3.3781 | -92.7677 | 92.9458 | 0.9125 | -0.078125 | -0.66771 | 1.1635 | 92.7677 | 0.078125 | 7.0007
0.0771 | -4.4542 | -3.3062 | -92.0552 | 92.224 | 0.92083 | -0.080208 | -0.66979 | 1.1719 | 92.0552 | 0.080208 | 7.1681
0.0789 | -4.2146 | -3.1635 | -91.4094 | 91.5646 | 0.93333 | -0.085417 | -0.66979 | 1.1812 | 91.4094 | 0.085417 | 7.3354
0.0807 | -3.975 | -3.0198 | -90.7635 | 90.9052 | 0.94479 | -0.090625 | -0.66979 | 1.1917 | 90.7635 | 0.090625 | 7.5028
0.0825 | -3.6958 | -2.8677 | -90.4719 | 90.6 | 0.95938 | -0.1 | -0.66667 | 1.2021 | 90.4719 | 0.1 | 7.6701
0.0843 | -3.6167 | -2.7865 | -90.5323 | 90.6604 | 0.97292 | -0.097917 | -0.66042 | 1.2094 | 90.5323 | 0.097917 | 7.8375
0.086 | -3.674 | -2.926 | -90.5219 | 90.6615 | 0.97604 | -0.091667 | -0.65729 | 1.2094 | 90.5219 | 0.091667 | 7.9955
0.0878 | -3.7667 | -2.951 | -90.6708 | 90.8198 | 0.98229 | -0.086458 | -0.65312 | 1.2104 | 90.6708 | 0.086458 | 8.1629
0.0896 | -3.8938 | -2.9854 | -90.9104 | 91.0687 | 0.98854 | -0.080208 | -0.64792 | 1.2094 | 90.9104 | 0.080208 | 8.3302
0.0914 | -4.4188 | -3.4542 | -91.5896 | 91.7792 | 0.97813 | -0.059375 | -0.64583 | 1.1938 | 91.5896 | 0.059375 | 8.4976
0.0932 | -4.9792 | -3.924 | -91.9813 | 92.2042 | 0.96875 | -0.034375 | -0.64687 | 1.1802 | 91.9813 | 0.034375 | 8.6649
0.095 | -5.2781 | -4.1656 | -91.8354 | 92.0854 | 0.96771 | -0.019792 | -0.64896 | 1.1771 | 91.8354 | 0.019792 | 8.8323
0.0968 | -5.3573 | -4.2021 | -91.5521 | 91.8104 | 0.96979 | -0.016667 | -0.64792 | 1.176 | 91.5521 | 0.016667 | 8.9996
0.0986 | -5.4698 | -4.3115 | -91.224 | 91.4948 | 0.96875 | -0.011458 | -0.64792 | 1.1729 | 91.224 | 0.011458 | 9.167
0.1004 | -5.601 | -4.5094 | -90.849 | 91.1396 | 0.96458 | -0.0041667 | -0.64792 | 1.1687 | 90.849 | 0.0041667 | 9.3343
0.1022 | -5.5062 | -4.5271 | -90.4167 | 90.7031 | 0.96458 | -0.0072917 | -0.64479 | 1.1677 | 90.4167 | 0.0072917 | 9.5017
0.104 | -5.451 | -4.6677 | -90.1146 | 90.4052 | 0.95833 | -0.010417 | -0.64062 | 1.1604 | 90.1146 | 0.010417 | 9.669
0.1058 | -5.474 | -4.849 | -89.7698 | 90.074 | 0.95 | -0.009375 | -0.63854 | 1.1521 | 89.7698 | 0.009375 | 9.8364
0.1076 | -5.4062 | -4.8344 | -88.9979 | 89.301 | 0.94688 | -0.0083333 | -0.63542 | 1.1479 | 88.9979 | 0.0083333 | 10.0037
0.1094 | -5.276 | -4.7417 | -88.5823 | 88.874 | 0.94167 | -0.011458 | -0.62604 | 1.1385 | 88.5823 | 0.011458 | 10.1711
0.1111 | -5.1635 | -4.6823 | -88.3448 | 88.6271 | 0.93542 | -0.013542 | -0.61667 | 1.1281 | 88.3448 | 0.013542 | 10.3291
0.1129 | -5.1365 | -4.6583 | -87.5479 | 87.8312 | 0.93021 | -0.009375 | -0.61458 | 1.1219 | 87.5479 | 0.009375 | 10.4965
0.1147 | -5.0708 | -4.5 | -86.4208 | 86.6958 | 0.92292 | -0.0010417 | -0.60938 | 1.1125 | 86.4208 | 0.0010417 | 10.6638
0.1165 | -5.0302 | -4.3219 | -85.3698 | 85.6365 | 0.91354 | 0.0083333 | -0.60417 | 1.1021 | 85.3698 | -0.0083333 | 10.8312
0.1183 | -4.9437 | -4.2052 | -85.1844 | 85.4406 | 0.89896 | 0.0125 | -0.59062 | 1.0823 | 85.1844 | -0.0125 | 10.9985
0.1201 | -4.7292 | -3.9594 | -84.7656 | 85 | 0.89479 | 0.013542 | -0.58021 | 1.074 | 84.7656 | -0.013542 | 11.1659
0.1219 | -4.4271 | -3.6042 | -83.9875 | 84.1937 | 0.89896 | 0.0125 | -0.57292 | 1.0729 | 83.9875 | -0.0125 | 11.3332
0.1237 | -4.0792 | -3.1792 | -82.8625 | 83.0385 | 0.90625 | 0.0125 | -0.56875 | 1.0771 | 82.8625 | -0.0125 | 11.5006
0.1255 | -3.8885 | -2.8521 | -81.7135 | 81.874 | 0.90521 | 0.01875 | -0.56458 | 1.075 | 81.7135 | -0.01875 | 11.6679
0.1273 | -3.7635 | -2.5979 | -80.6781 | 80.826 | 0.90625 | 0.028125 | -0.56458 | 1.0771 | 80.6781 | -0.028125 | 11.8353
0.1291 | -3.7375 | -2.4167 | -79.8146 | 79.9542 | 0.91458 | 0.0375 | -0.57083 | 1.0865 | 79.8146 | -0.0375 | 12.0026
0.1309 | -3.7125 | -2.1948 | -79.1833 | 79.3135 | 0.92188 | 0.045833 | -0.57292 | 1.0927 | 79.1833 | -0.045833 | 12.17
0.1327 | -3.6177 | -1.9354 | -78.6854 | 78.8052 | 0.92396 | 0.047917 | -0.56979 | 1.0938 | 78.6854 | -0.047917 | 12.3373
0.1345 | -3.5708 | -1.7646 | -78.1792 | 78.2938 | 0.91875 | 0.05 | -0.56667 | 1.0875 | 78.1792 | -0.05 | 12.5046
0.1362 | -3.6865 | -1.7458 | -77.5802 | 77.6979 | 0.91146 | 0.059375 | -0.56771 | 1.0833 | 77.5802 | -0.059375 | 12.6627
0.138 | -3.9094 | -1.7823 | -77.4896 | 77.6156 | 0.90417 | 0.075 | -0.56771 | 1.0792 | 77.4896 | -0.075 | 12.83
0.1398 | -4.1979 | -1.9042 | -77.7677 | 77.9104 | 0.89583 | 0.091667 | -0.56771 | 1.074 | 77.7677 | -0.091667 | 12.9974
0.1416 | -4.4521 | -1.9594 | -78.076 | 78.2302 | 0.8875 | 0.10625 | -0.56563 | 1.0688 | 78.076 | -0.10625 | 13.1647
0.1434 | -4.4323 | -1.9635 | -78.6177 | 78.7677 | 0.88438 | 0.10417 | -0.56042 | 1.0646 | 78.6177 | -0.10417 | 13.3321
0.1452 | -4.5698 | -2.0333 | -79.4333 | 79.5917 | 0.87813 | 0.10833 | -0.55521 | 1.0573 | 79.4333 | -0.10833 | 13.4994
0.147 | -4.7562 | -2.1906 | -80.5417 | 80.7125 | 0.86562 | 0.1125 | -0.54792 | 1.0458 | 80.5417 | -0.1125 | 13.6668
0.1488 | -4.8906 | -2.2938 | -81.7021 | 81.8812 | 0.85729 | 0.11458 | -0.53854 | 1.0344 | 81.7021 | -0.11458 | 13.8341
0.1506 | -4.8365 | -2.349 | -82.576 | 82.7521 | 0.85104 | 0.10938 | -0.52708 | 1.024 | 82.576 | -0.10938 | 14.0015
0.1524 | -4.7479 | -2.2958 | -83.1354 | 83.3031 | 0.84688 | 0.10104 | -0.51771 | 1.0156 | 83.1354 | -0.10104 | 14.1688
0.1542 | -4.625 | -2.2437 | -83.5302 | 83.6885 | 0.84271 | 0.09375 | -0.50937 | 1.0073 | 83.5302 | -0.09375 | 14.3362
0.156 | -4.5865 | -2.2156 | -83.7229 | 83.8792 | 0.83854 | 0.0875 | -0.50417 | 1.0021 | 83.7229 | -0.0875 | 14.5035
0.1578 | -4.5521 | -2.1812 | -83.7073 | 83.8604 | 0.83542 | 0.083333 | -0.5 | 0.99687 | 83.7073 | -0.083333 | 14.6709
0.1595 | -4.5365 | -2.1708 | -83.5406 | 83.6937 | 0.83229 | 0.083333 | -0.49792 | 0.99271 | 83.5406 | -0.083333 | 14.8289
0.1613 | -4.626 | -2.2812 | -83.3604 | 83.5219 | 0.82396 | 0.086458 | -0.49896 | 0.9875 | 83.3604 | -0.086458 | 14.9963
0.1631 | -4.7573 | -2.3823 | -83.1156 | 83.2896 | 0.81875 | 0.092708 | -0.50208 | 0.98333 | 83.1156 | -0.092708 | 15.1636
0.1649 | -4.8844 | -2.4698 | -82.8375 | 83.024 | 0.81354 | 0.098958 | -0.50417 | 0.98021 | 82.8375 | -0.098958 | 15.331
0.1667 | -4.7948 | -2.3406 | -82.1635 | 82.3406 | 0.81875 | 0.097917 | -0.50521 | 0.98542 | 82.1635 | -0.097917 | 15.4983
0.1685 | -4.5563 | -2.1094 | -81.3156 | 81.474 | 0.82708 | 0.089583 | -0.50625 | 0.99479 | 81.3156 | -0.089583 | 15.6657
0.1703 | -4.2135 | -1.8115 | -80.3219 | 80.4542 | 0.83958 | 0.076042 | -0.50833 | 1.0083 | 80.3219 | -0.076042 | 15.833
0.1721 | -3.7615 | -1.4427 | -79.725 | 79.8396 | 0.85 | 0.057292 | -0.50313 | 1.0198 | 79.725 | -0.057292 | 16.0004
0.1739 | -3.3792 | -1.1135 | -79.2573 | 79.3594 | 0.85833 | 0.042708 | -0.49687 | 1.0292 | 79.2573 | -0.042708 | 16.1677
0.1757 | -3.4885 | -1.0813 | -79.2667 | 79.3823 | 0.85417 | 0.046875 | -0.48958 | 1.025 | 79.2667 | -0.046875 | 16.3351
0.1775 | -3.6635 | -1.0844 | -79.4083 | 79.5385 | 0.84792 | 0.055208 | -0.48333 | 1.0208 | 79.4083 | -0.055208 | 16.5024
0.1793 | -3.9802 | -1.2917 | -80.0312 | 80.1813 | 0.83333 | 0.067708 | -0.47604 | 1.0063 | 80.0312 | -0.067708 | 16.6698
0.1811 | -4.3219 | -1.576 | -81.0438 | 81.2229 | 0.81667 | 0.082292 | -0.46563 | 0.99167 | 81.0438 | -0.082292 | 16.8371
0.1829 | -4.4875 | -1.7594 | -82.199 | 82.3937 | 0.81042 | 0.091667 | -0.45312 | 0.98437 | 82.199 | -0.091667 | 17.0045
0.1846 | -4.5667 | -1.8823 | -83.3802 | 83.5823 | 0.80833 | 0.098958 | -0.44063 | 0.98125 | 83.3802 | -0.098958 | 17.1625
0.1864 | -4.5156 | -1.8875 | -84.1281 | 84.3198 | 0.80833 | 0.10208 | -0.42917 | 0.97396 | 84.1281 | -0.10208 | 17.3299
0.1882 | -4.4062 | -1.8302 | -84.4125 | 84.5823 | 0.80937 | 0.10417 | -0.41979 | 0.96667 | 84.4125 | -0.10417 | 17.4972
0.19 | -4.4375 | -1.9052 | -84.4 | 84.5615 | 0.80417 | 0.11563 | -0.41667 | 0.95729 | 84.4 | -0.11563 | 17.6646
0.1918 | -4.4958 | -2.0125 | -84.1771 | 84.3365 | 0.79688 | 0.12917 | -0.41667 | 0.94688 | 84.1771 | -0.12917 | 17.8319
0.1936 | -4.4833 | -2.0219 | -83.651 | 83.8083 | 0.79167 | 0.14375 | -0.41875 | 0.94062 | 83.651 | -0.14375 | 17.9993
0.1954 | -4.3688 | -1.9687 | -83.2115 | 83.3646 | 0.78958 | 0.15208 | -0.41771 | 0.9375 | 83.2115 | -0.15208 | 18.1666
0.1972 | -4.25 | -1.9729 | -83.0531 | 83.199 | 0.79063 | 0.15729 | -0.41458 | 0.9375 | 83.0531 | -0.15729 | 18.334
0.199 | -3.9479 | -1.8677 | -82.9948 | 83.1302 | 0.79583 | 0.15521 | -0.40833 | 0.93854 | 82.9948 | -0.15521 | 18.5013
0.2008 | -3.6729 | -1.7677 | -82.9406 | 83.0729 | 0.80312 | 0.15521 | -0.40208 | 0.94271 | 82.9406 | -0.15521 | 18.6687
0.2026 | -3.424 | -1.6802 | -83.0073 | 83.1375 | 0.81146 | 0.15833 | -0.39375 | 0.94583 | 83.0073 | -0.15833 | 18.836
0.2044 | -3.449 | -1.7792 | -83.3521 | 83.4813 | 0.81354 | 0.17917 | -0.38646 | 0.94062 | 83.3521 | -0.17917 | 19.0033
0.2062 | -3.5677 | -1.9385 | -83.6281 | 83.7573 | 0.81354 | 0.20312 | -0.37917 | 0.9375 | 83.6281 | -0.20312 | 19.1707
0.208 | -3.6927 | -2.1365 | -83.5792 | 83.7135 | 0.81771 | 0.225 | -0.375 | 0.93958 | 83.5792 | -0.225 | 19.338
0.2097 | -3.8521 | -2.4052 | -83.5448 | 83.6865 | 0.82188 | 0.25104 | -0.36979 | 0.94271 | 83.5448 | -0.25104 | 19.4961
0.2115 | -4.1177 | -2.6469 | -83.6677 | 83.825 | 0.825 | 0.27708 | -0.36458 | 0.94896 | 83.6677 | -0.27708 | 19.6634
0.2133 | -4.5115 | -3.049 | -84.2521 | 84.4396 | 0.82604 | 0.31146 | -0.3625 | 0.95938 | 84.2521 | -0.31146 | 19.8308
0.2151 | -5.0042 | -3.474 | -84.925 | 85.15 | 0.82917 | 0.34896 | -0.35833 | 0.97188 | 84.925 | -0.34896 | 19.9981
0.2169 | -5.4938 | -3.9344 | -85.9583 | 86.2313 | 0.82396 | 0.38021 | -0.35729 | 0.97917 | 85.9583 | -0.38021 | 20.1655
0.2187 | -6.051 | -4.4771 | -87.3844 | 87.7167 | 0.81667 | 0.4125 | -0.35625 | 0.98542 | 87.3844 | -0.4125 | 20.3328
0.2205 | -6.5875 | -5.0021 | -89.2125 | 89.6083 | 0.81563 | 0.44688 | -0.35417 | 1.001 | 89.2125 | -0.44688 | 20.5002
0.2223 | -6.801 | -5.2354 | -91.0021 | 91.4208 | 0.82396 | 0.46563 | -0.34896 | 1.0156 | 91.0021 | -0.46563 | 20.6675
0.2241 | -6.7229 | -5.1667 | -92.1 | 92.5042 | 0.84479 | 0.47396 | -0.34687 | 1.0354 | 92.1 | -0.47396 | 20.8349
0.2259 | -6.8198 | -5.3479 | -93.3552 | 93.7719 | 0.85313 | 0.48438 | -0.34583 | 1.0469 | 93.3552 | -0.48438 | 21.0022
0.2277 | -6.8938 | -5.499 | -94.7604 | 95.1865 | 0.8625 | 0.49375 | -0.34479 | 1.0583 | 94.7604 | -0.49375 | 21.1696
0.2295 | -6.8083 | -5.4688 | -96.2354 | 96.6469 | 0.87708 | 0.49687 | -0.34375 | 1.0719 | 96.2354 | -0.49687 | 21.3369
0.2313 | -6.676 | -5.4094 | -97.3438 | 97.7375 | 0.89167 | 0.5 | -0.34271 | 1.0844 | 97.3438 | -0.5 | 21.5043
0.2331 | -6.6896 | -5.525 | -98.7073 | 99.1042 | 0.89792 | 0.50729 | -0.34063 | 1.0917 | 98.7073 | -0.50729 | 21.6716
0.2348 | -6.7531 | -5.7292 | -100.5719 | 100.9781 | 0.90208 | 0.51354 | -0.33854 | 1.0958 | 100.5719 | -0.51354 | 21.8297
0.2366 | -6.7938 | -5.9437 | -102.7938 | 103.2062 | 0.90938 | 0.51667 | -0.3375 | 1.1031 | 102.7938 | -0.51667 | 21.997
0.2384 | -6.7302 | -6.0698 | -104.9552 | 105.3615 | 0.92396 | 0.51771 | -0.33542 | 1.1146 | 104.9552 | -0.51771 | 22.1644
0.2402 | -6.5563 | -6.1135 | -106.7573 | 107.1479 | 0.93854 | 0.51979 | -0.33333 | 1.1271 | 106.7573 | -0.51979 | 22.3317
0.242 | -6.3594 | -6.1312 | -108.4604 | 108.8344 | 0.95312 | 0.52396 | -0.33229 | 1.1406 | 108.4604 | -0.52396 | 22.4991
0.2438 | -6.099 | -5.9969 | -109.5937 | 109.9427 | 0.96875 | 0.52917 | -0.33229 | 1.1573 | 109.5937 | -0.52917 | 22.6664
0.2456 | -5.7646 | -5.8469 | -110.4542 | 110.7698 | 0.99271 | 0.52917 | -0.33021 | 1.176 | 110.4542 | -0.52917 | 22.8338
0.2474 | -5.45 | -5.5781 | -111.0365 | 111.3188 | 1.0198 | 0.53021 | -0.33125 | 1.201 | 111.0365 | -0.53021 | 23.0011
0.2492 | -5.1698 | -5.326 | -111.7667 | 112.0208 | 1.0437 | 0.53438 | -0.33125 | 1.225 | 111.7667 | -0.53438 | 23.1685
0.251 | -4.9396 | -5.1698 | -112.9729 | 113.2073 | 1.0583 | 0.54167 | -0.33021 | 1.2406 | 112.9729 | -0.54167 | 23.3358
0.2528 | -4.676 | -5.0344 | -114.2479 | 114.4625 | 1.074 | 0.54271 | -0.32917 | 1.2542 | 114.2479 | -0.54271 | 23.5032
0.2546 | -4.4083 | -4.9531 | -115.324 | 115.5219 | 1.0906 | 0.54271 | -0.33021 | 1.2698 | 115.324 | -0.54271 | 23.6705
0.2564 | -4.224 | -4.9062 | -116.4885 | 116.6781 | 1.1031 | 0.54375 | -0.33125 | 1.2812 | 116.4885 | -0.54375 | 23.8379
0.2581 | -3.9469 | -4.7385 | -117.7281 | 117.9052 | 1.1146 | 0.54063 | -0.33125 | 1.2896 | 117.7281 | -0.54063 | 23.9959
0.2599 | -3.75 | -4.6615 | -118.8802 | 119.0479 | 1.125 | 0.54375 | -0.33125 | 1.299 | 118.8802 | -0.54375 | 24.1633
0.2617 | -3.7271 | -4.774 | -120.0104 | 120.1802 | 1.1313 | 0.55937 | -0.33542 | 1.3115 | 120.0104 | -0.55937 | 24.3306
0.2635 | -3.7531 | -4.9281 | -121.3729 | 121.55 | 1.1385 | 0.58021 | -0.33854 | 1.326 | 121.3729 | -0.58021 | 24.498
0.2653 | -3.7802 | -5.0823 | -122.7458 | 122.9302 | 1.1438 | 0.60104 | -0.34167 | 1.3396 | 122.7458 | -0.60104 | 24.6653
0.2671 | -3.8292 | -5.2625 | -124.4094 | 124.6021 | 1.1427 | 0.62187 | -0.34375 | 1.35 | 124.4094 | -0.62187 | 24.8327
0.2689 | -3.8115 | -5.3135 | -125.9667 | 126.1615 | 1.151 | 0.6375 | -0.34792 | 1.3656 | 125.9667 | -0.6375 | 25
0.2707 | -3.8094 | -5.3063 | -127.6354 | 127.8292 | 1.1583 | 0.65521 | -0.35 | 1.3823 | 127.6354 | -0.65521 | 25.1673
0.2725 | -3.8333 | -5.1969 | -129.5 | 129.6896 | 1.1635 | 0.67708 | -0.35104 | 1.399 | 129.5 | -0.67708 | 25.3347
0.2743 | -3.9458 | -5.2188 | -131.7198 | 131.9073 | 1.1656 | 0.69792 | -0.35208 | 1.4125 | 131.7198 | -0.69792 | 25.502
0.2761 | -4.0292 | -5.3115 | -134.0802 | 134.2656 | 1.1677 | 0.71354 | -0.35521 | 1.4219 | 134.0802 | -0.71354 | 25.6694
0.2779 | -3.9833 | -5.3115 | -136.375 | 136.5521 | 1.1729 | 0.725 | -0.35625 | 1.4323 | 136.375 | -0.725 | 25.8367
0.2797 | -3.8667 | -5.3104 | -138.1521 | 138.3302 | 1.1823 | 0.73125 | -0.35938 | 1.4427 | 138.1521 | -0.73125 | 26.0041
0.2815 | -3.726 | -5.3448 | -140.076 | 140.2531 | 1.1958 | 0.73958 | -0.36562 | 1.4604 | 140.076 | -0.73958 | 26.1714
0.2832 | -3.6917 | -5.4792 | -142.4625 | 142.6417 | 1.2073 | 0.75104 | -0.36979 | 1.4771 | 142.4625 | -0.75104 | 26.3295
0.285 | -3.5875 | -5.5469 | -144.5687 | 144.7458 | 1.2198 | 0.75938 | -0.37292 | 1.4906 | 144.5687 | -0.75938 | 26.4968
0.2868 | -3.3583 | -5.4635 | -146.0979 | 146.2646 | 1.2302 | 0.7625 | -0.37292 | 1.501 | 146.0979 | -0.7625 | 26.6642
0.2886 | -3.1688 | -5.351 | -147.624 | 147.7802 | 1.2406 | 0.77083 | -0.37188 | 1.5146 | 147.624 | -0.77083 | 26.8315
0.2904 | -2.9792 | -5.2385 | -149.15 | 149.2969 | 1.25 | 0.77917 | -0.37083 | 1.5271 | 149.15 | -0.77917 | 26.9989
0.2922 | -2.776 | -5.1844 | -150.2042 | 150.3458 | 1.2563 | 0.78854 | -0.37188 | 1.5375 | 150.2042 | -0.78854 | 27.1662
0.294 | -2.5146 | -4.9771 | -150.6708 | 150.801 | 1.2635 | 0.80208 | -0.37396 | 1.5521 | 150.6708 | -0.80208 | 27.3336
0.2958 | -2.2719 | -4.5208 | -150.7083 | 150.825 | 1.2698 | 0.81667 | -0.37396 | 1.5667 | 150.7083 | -0.81667 | 27.5009
0.2976 | -2.1427 | -3.9802 | -150.1177 | 150.224 | 1.2625 | 0.82604 | -0.36979 | 1.5656 | 150.1177 | -0.82604 | 27.6683
0.2994 | -2.0948 | -3.6062 | -149.1219 | 149.226 | 1.2427 | 0.83437 | -0.36562 | 1.5552 | 149.1219 | -0.83437 | 27.8356
0.3012 | -2.0875 | -3.3042 | -148.2167 | 148.3187 | 1.2198 | 0.84479 | -0.3625 | 1.5458 | 148.2167 | -0.84479 | 28.003
0.303 | -2.1969 | -3.3292 | -148.574 | 148.6771 | 1.2063 | 0.86042 | -0.36458 | 1.5469 | 148.574 | -0.86042 | 28.1703
0.3048 | -2.2188 | -3.225 | -148.9677 | 149.0708 | 1.2 | 0.87396 | -0.36458 | 1.5531 | 148.9677 | -0.87396 | 28.3377
0.3066 | -2.2979 | -3.1833 | -149.4156 | 149.5198 | 1.1979 | 0.88646 | -0.36771 | 1.5615 | 149.4156 | -0.88646 | 28.505
0.3083 | -2.4031 | -3.1792 | -149.8948 | 150 | 1.1958 | 0.89896 | -0.37083 | 1.5688 | 149.8948 | -0.89896 | 28.6631
0.3101 | -2.4646 | -3.2146 | -150.326 | 150.4354 | 1.1938 | 0.91042 | -0.375 | 1.576 | 150.326 | -0.91042 | 28.8304
0.3119 | -2.6313 | -3.4417 | -150.7406 | 150.8573 | 1.1885 | 0.92396 | -0.38021 | 1.5833 | 150.7406 | -0.92396 | 28.9978
0.3137 | -2.8792 | -3.7271 | -151.1833 | 151.3063 | 1.1781 | 0.93958 | -0.38438 | 1.5906 | 151.1833 | -0.93958 | 29.1651
0.3155 | -3.0385 | -3.8729 | -151.6198 | 151.7479 | 1.1677 | 0.96042 | -0.38542 | 1.5969 | 151.6198 | -0.96042 | 29.3325
0.3173 | -3.0562 | -3.7865 | -151.6385 | 151.7656 | 1.1646 | 0.97708 | -0.38646 | 1.6042 | 151.6385 | -0.97708 | 29.4998
0.3191 | -3.0615 | -3.6062 | -151.7281 | 151.8562 | 1.1656 | 0.98958 | -0.38958 | 1.6125 | 151.7281 | -0.98958 | 29.6672
0.3209 | -3.0656 | -3.4219 | -151.826 | 151.9552 | 1.1677 | 1.0021 | -0.39167 | 1.6198 | 151.826 | -1.0021 | 29.8345
0.3227 | -2.9302 | -3.0594 | -151.5813 | 151.7083 | 1.175 | 1.0073 | -0.39583 | 1.6271 | 151.5813 | -1.0073 | 30.0019
0.3245 | -2.8062 | -2.6385 | -151.2552 | 151.3812 | 1.1812 | 1.0042 | -0.39896 | 1.6281 | 151.2552 | -1.0042 | 30.1692
0.3263 | -2.7719 | -2.3937 | -151.1271 | 151.2542 | 1.1792 | 1.0063 | -0.40417 | 1.6302 | 151.1271 | -1.0063 | 30.3366
0.3281 | -2.7594 | -2.2927 | -150.1 | 150.2281 | 1.1677 | 0.99375 | -0.40937 | 1.6146 | 150.1 | -0.99375 | 30.5039
0.3299 | -2.7646 | -2.2865 | -149.8396 | 149.9667 | 1.1615 | 0.98646 | -0.41563 | 1.6073 | 149.8396 | -0.98646 | 30.6713
0.3316 | -2.85 | -2.3396 | -150.2125 | 150.3469 | 1.1583 | 0.98854 | -0.42188 | 1.6073 | 150.2125 | -0.98854 | 30.8293
0.3334 | -2.9083 | -2.3469 | -150.925 | 151.0667 | 1.1604 | 0.99167 | -0.42813 | 1.6115 | 150.925 | -0.99167 | 30.9967
0.3352 | -2.9562 | -2.4167 | -151.9073 | 152.0573 | 1.1615 | 0.99583 | -0.43333 | 1.6146 | 151.9073 | -0.99583 | 31.164
0.337 | -3.0094 | -2.5917 | -153.1198 | 153.2771 | 1.1635 | 0.99896 | -0.44063 | 1.6177 | 153.1198 | -0.99896 | 31.3313
0.3388 | -3.0396 | -2.8062 | -154.5115 | 154.6719 | 1.1687 | 1 | -0.44792 | 1.6219 | 154.5115 | -1 | 31.4987
0.3406 | -2.9927 | -2.8312 | -156.2812 | 156.4333 | 1.176 | 1.0104 | -0.44792 | 1.6333 | 156.2812 | -1.0104 | 31.666
0.3424 | -2.8083 | -2.7625 | -158.0052 | 158.1406 | 1.1896 | 1.0146 | -0.44792 | 1.6469 | 158.0052 | -1.0146 | 31.8334
0.3442 | -2.5292 | -2.6208 | -159.8104 | 159.926 | 1.2115 | 1.0115 | -0.44896 | 1.6646 | 159.8104 | -1.0115 | 32.0007
0.346 | -2.25 | -2.4802 | -161.6156 | 161.7115 | 1.2344 | 1.0083 | -0.45 | 1.6823 | 161.6156 | -1.0083 | 32.1681
0.3478 | -2.2823 | -2.4844 | -163.8208 | 163.9094 | 1.251 | 1.0094 | -0.45208 | 1.699 | 163.8208 | -1.0094 | 32.3354
0.3496 | -2.2833 | -2.4271 | -165.6927 | 165.7708 | 1.2646 | 1.0073 | -0.45417 | 1.7115 | 165.6927 | -1.0073 | 32.5028
0.3514 | -2.2552 | -2.4135 | -167.051 | 167.1219 | 1.2719 | 1.0052 | -0.45729 | 1.7188 | 167.051 | -1.0052 | 32.6701
0.3532 | -2.2437 | -2.4375 | -168.1219 | 168.1875 | 1.274 | 1.0052 | -0.46042 | 1.724 | 168.1219 | -1.0052 | 32.8375
0.355 | -2.2948 | -2.5302 | -168.9885 | 169.0521 | 1.274 | 1.0083 | -0.46354 | 1.726 | 168.9885 | -1.0083 | 33.0048
0.3567 | -2.225 | -2.5646 | -169.6771 | 169.7365 | 1.2771 | 1.0063 | -0.46771 | 1.7281 | 169.6771 | -1.0063 | 33.1629
0.3585 | -2.0365 | -2.5292 | -170.1281 | 170.1823 | 1.2823 | 1 | -0.46979 | 1.7302 | 170.1281 | -1 | 33.3302
0.3603 | -1.8719 | -2.4906 | -170.4615 | 170.5115 | 1.2885 | 0.99271 | -0.46875 | 1.7292 | 170.4615 | -0.99271 | 33.4976
0.3621 | -1.8635 | -2.4229 | -170.9979 | 171.0469 | 1.2937 | 0.98958 | -0.46458 | 1.7292 | 170.9979 | -0.98958 | 33.6649
0.3639 | -1.7604 | -2.3188 | -171.2427 | 171.2906 | 1.2979 | 0.98437 | -0.46458 | 1.7281 | 171.2427 | -0.98437 | 33.8323
0.3657 | -1.6531 | -2.2073 | -171.5135 | 171.5594 | 1.3021 | 0.97708 | -0.46458 | 1.726 | 171.5135 | -0.97708 | 33.9996
0.3675 | -1.6625 | -2.1042 | -172.4115 | 172.4573 | 1.3031 | 0.97292 | -0.46354 | 1.7229 | 172.4115 | -0.97292 | 34.167
0.3693 | -1.6698 | -1.9594 | -173.3812 | 173.426 | 1.3104 | 0.96562 | -0.46667 | 1.7229 | 173.3812 | -0.96562 | 34.3343
0.3711 | -1.6625 | -1.749 | -174.0687 | 174.1125 | 1.3198 | 0.95833 | -0.47083 | 1.725 | 174.0687 | -0.95833 | 34.5017
0.3729 | -1.6125 | -1.4948 | -174.0469 | 174.0906 | 1.3302 | 0.94479 | -0.47917 | 1.726 | 174.0469 | -0.94479 | 34.669
0.3747 | -1.4854 | -1.2531 | -173.8135 | 173.8573 | 1.3385 | 0.92812 | -0.48542 | 1.7229 | 173.8135 | -0.92812 | 34.8364
0.3765 | -1.4604 | -1.2542 | -173.6198 | 173.6625 | 1.3302 | 0.91979 | -0.48646 | 1.7125 | 173.6198 | -0.91979 | 35.0037
0.3783 | -1.4896 | -1.3385 | -173.7719 | 173.8135 | 1.3188 | 0.91042 | -0.4875 | 1.6979 | 173.7719 | -0.91042 | 35.1711
0.3801 | -1.551 | -1.4396 | -174.349 | 174.3896 | 1.3104 | 0.90104 | -0.48646 | 1.6854 | 174.349 | -0.90104 | 35.3384
0.3818 | -1.5896 | -1.475 | -174.6719 | 174.7125 | 1.3052 | 0.89271 | -0.4875 | 1.676 | 174.6719 | -0.89271 | 35.4965
0.3836 | -1.6104 | -1.5542 | -174.4833 | 174.524 | 1.3 | 0.88854 | -0.49167 | 1.6719 | 174.4833 | -0.88854 | 35.6638
0.3854 | -1.6781 | -1.7448 | -174.5802 | 174.6198 | 1.2927 | 0.88646 | -0.49583 | 1.6667 | 174.5802 | -0.88646 | 35.8312
0.3872 | -1.8448 | -2.0208 | -175.475 | 175.5167 | 1.2833 | 0.88958 | -0.49792 | 1.6635 | 175.475 | -0.88958 | 35.9985
0.389 | -2.0219 | -2.2594 | -176.5344 | 176.5781 | 1.2802 | 0.89479 | -0.50313 | 1.6667 | 176.5344 | -0.89479 | 36.1659
0.3908 | -2.2219 | -2.5583 | -177.5573 | 177.6073 | 1.275 | 0.90417 | -0.50833 | 1.6687 | 177.5573 | -0.90417 | 36.3332
0.3926 | -2.4312 | -2.9531 | -177.9375 | 177.9979 | 1.2688 | 0.92083 | -0.51875 | 1.6781 | 177.9375 | -0.92083 | 36.5006
0.3944 | -2.6677 | -3.3948 | -178.4958 | 178.5687 | 1.2604 | 0.93646 | -0.52812 | 1.6844 | 178.4958 | -0.93646 | 36.6679
0.3962 | -2.9854 | -3.8979 | -179.5458 | 179.6406 | 1.2469 | 0.95625 | -0.53021 | 1.6906 | 179.5458 | -0.95625 | 36.8353
0.398 | -3.3146 | -4.3729 | -180.6323 | 180.7521 | 1.2344 | 0.97917 | -0.53125 | 1.6969 | 180.6323 | -0.97917 | 37.0026
0.3998 | -3.5281 | -4.6771 | -181.6687 | 181.8031 | 1.2292 | 0.99687 | -0.53125 | 1.7063 | 181.6687 | -0.99687 | 37.17
0.4016 | -3.6448 | -4.8385 | -183.15 | 183.2906 | 1.2344 | 1.0083 | -0.53333 | 1.7177 | 183.15 | -1.0083 | 37.3373
0.4034 | -3.7219 | -5.05 | -184.926 | 185.0677 | 1.2385 | 1.0146 | -0.53854 | 1.725 | 184.926 | -1.0146 | 37.5046
0.4051 | -3.8479 | -5.2458 | -186.9938 | 187.1375 | 1.2427 | 1.0135 | -0.54271 | 1.7292 | 186.9938 | -1.0135 | 37.6627
0.4069 | -4.0073 | -5.4521 | -189.1302 | 189.2781 | 1.2469 | 1.0115 | -0.54792 | 1.7323 | 189.1302 | -1.0115 | 37.83
0.4087 | -4.101 | -5.6042 | -190.9854 | 191.1375 | 1.2542 | 1.0104 | -0.55312 | 1.7375 | 190.9854 | -1.0104 | 37.9974
0.4105 | -4.0406 | -5.5542 | -192.5198 | 192.6698 | 1.2698 | 1.0073 | -0.55833 | 1.749 | 192.5198 | -1.0073 | 38.1647
0.4123 | -3.8354 | -5.2875 | -194.3271 | 194.4677 | 1.2979 | 1.0021 | -0.55937 | 1.7646 | 194.3271 | -1.0021 | 38.3321
0.4141 | -3.6115 | -5.0958 | -196.4927 | 196.6219 | 1.324 | 0.99792 | -0.55833 | 1.7781 | 196.4927 | -0.99792 | 38.4994
0.4159 | -3.4948 | -5.1208 | -198.9917 | 199.1167 | 1.3417 | 0.99687 | -0.55625 | 1.7896 | 198.9917 | -0.99687 | 38.6668
0.4177 | -3.6208 | -5.4146 | -201.5542 | 201.6833 | 1.3469 | 1.0021 | -0.55625 | 1.7969 | 201.5542 | -1.0021 | 38.8341
0.4195 | -3.4656 | -5.3844 | -203.9448 | 204.0708 | 1.3698 | 0.99896 | -0.55937 | 1.8125 | 203.9448 | -0.99896 | 39.0015
0.4213 | -3.3344 | -5.3583 | -206.4375 | 206.5594 | 1.3927 | 0.99687 | -0.56354 | 1.8302 | 206.4375 | -0.99687 | 39.1688
0.4231 | -3.1635 | -5.2708 | -208.6354 | 208.7531 | 1.4177 | 0.99271 | -0.56979 | 1.8479 | 208.6354 | -0.99271 | 39.3362
0.4249 | -2.9073 | -5.049 | -210.4646 | 210.5687 | 1.4469 | 0.98958 | -0.57604 | 1.8688 | 210.4646 | -0.98958 | 39.5035
0.4267 | -2.6594 | -4.8969 | -211.8708 | 211.9688 | 1.4646 | 0.97604 | -0.57812 | 1.8771 | 211.8708 | -0.97604 | 39.6709
0.4285 | -2.6 | -4.9198 | -213.326 | 213.4208 | 1.476 | 0.97396 | -0.58229 | 1.8865 | 213.326 | -0.97396 | 39.8382
0.4302 | -2.4719 | -4.7625 | -214.1875 | 214.2729 | 1.4906 | 0.97396 | -0.58333 | 1.899 | 214.1875 | -0.97396 | 39.9963
0.432 | -2.2125 | -4.5031 | -214.3729 | 214.4458 | 1.5073 | 0.97083 | -0.58437 | 1.9094 | 214.3729 | -0.97083 | 40.1636
0.4338 | -2.0187 | -4.3125 | -214.3531 | 214.4198 | 1.5188 | 0.97083 | -0.58646 | 1.9188 | 214.3531 | -0.97083 | 40.331
0.4356 | -1.8292 | -4.0687 | -213.9333 | 213.9938 | 1.5302 | 0.97083 | -0.5875 | 1.9292 | 213.9333 | -0.97083 | 40.4983
0.4374 | -1.726 | -3.9 | -213.2531 | 213.3115 | 1.5333 | 0.96771 | -0.59167 | 1.9312 | 213.2531 | -0.96771 | 40.6657
0.4392 | -1.5906 | -3.8021 | -212.8438 | 212.901 | 1.5333 | 0.96771 | -0.59167 | 1.9312 | 212.8438 | -0.96771 | 40.833
0.441 | -1.4333 | -3.7281 | -212.6208 | 212.676 | 1.5323 | 0.975 | -0.59167 | 1.9333 | 212.6208 | -0.975 | 41.0004
0.4428 | -1.2302 | -3.5198 | -212.0021 | 212.0542 | 1.5385 | 0.98229 | -0.59479 | 1.9427 | 212.0021 | -0.98229 | 41.1677
0.4446 | -1.0448 | -3.325 | -211.5062 | 211.5552 | 1.5469 | 0.98854 | -0.59688 | 1.951 | 211.5062 | -0.98854 | 41.3351
0.4464 | -1.0917 | -3.3823 | -211.5948 | 211.6458 | 1.5437 | 1.0052 | -0.59792 | 1.9563 | 211.5948 | -1.0052 | 41.5024
0.4482 | -1.3594 | -3.7104 | -212.1448 | 212.2063 | 1.5292 | 1.0292 | -0.59792 | 1.9563 | 212.1448 | -1.0292 | 41.6698
0.45 | -1.5167 | -3.8823 | -213.024 | 213.0896 | 1.5229 | 1.0458 | -0.59479 | 1.9573 | 213.024 | -1.0458 | 41.8371
0.4518 | -1.7573 | -4.1354 | -214.2115 | 214.2865 | 1.5115 | 1.0656 | -0.59167 | 1.9573 | 214.2115 | -1.0656 | 42.0045
0.4536 | -1.8771 | -4.1552 | -215.3344 | 215.4094 | 1.5063 | 1.0813 | -0.58854 | 1.9594 | 215.3344 | -1.0813 | 42.1718
0.4553 | -1.9208 | -4.101 | -216.25 | 216.3229 | 1.5052 | 1.0896 | -0.58646 | 1.9646 | 216.25 | -1.0896 | 42.3299
0.4571 | -1.9719 | -4.0677 | -216.8396 | 216.9094 | 1.5042 | 1.099 | -0.58646 | 1.9687 | 216.8396 | -1.099 | 42.4972
0.4589 | -1.9323 | -3.925 | -217.0177 | 217.0802 | 1.5083 | 1.1042 | -0.58646 | 1.974 | 217.0177 | -1.1042 | 42.6646
0.4607 | -1.8646 | -3.774 | -217.0083 | 217.0646 | 1.5125 | 1.1094 | -0.58542 | 1.9792 | 217.0083 | -1.1094 | 42.8319
0.4625 | -1.8562 | -3.7208 | -216.9375 | 216.9906 | 1.5135 | 1.1167 | -0.58542 | 1.9823 | 216.9375 | -1.1167 | 42.9993
0.4643 | -1.8823 | -3.6938 | -217.0677 | 217.1188 | 1.5125 | 1.124 | -0.58646 | 1.9865 | 217.0677 | -1.124 | 43.1666
0.4661 | -2 | -3.75 | -217.7385 | 217.7917 | 1.5083 | 1.1333 | -0.5875 | 1.9885 | 217.7385 | -1.1333 | 43.334
0.4679 | -1.9792 | -3.6677 | -218.7 | 218.7521 | 1.5063 | 1.1344 | -0.59271 | 1.9896 | 218.7 | -1.1344 | 43.5013
0.4697 | -1.976 | -3.524 | -219.4146 | 219.4656 | 1.5083 | 1.1354 | -0.59688 | 1.9917 | 219.4146 | -1.1354 | 43.6687
0.4715 | -1.9406 | -3.4562 | -219.7323 | 219.7875 | 1.5052 | 1.1313 | -0.59583 | 1.9885 | 219.7323 | -1.1313 | 43.836
0.4733 | -1.9677 | -3.4229 | -219.8135 | 219.8698 | 1.5 | 1.1313 | -0.59583 | 1.9854 | 219.8135 | -1.1313 | 44.0033
0.4751 | -2.2146 | -3.6344 | -219.8833 | 219.9437 | 1.4854 | 1.1417 | -0.59479 | 1.9792 | 219.8833 | -1.1417 | 44.1707
0.4769 | -2.4042 | -3.7229 | -219.5104 | 219.5719 | 1.4771 | 1.151 | -0.59896 | 1.9771 | 219.5104 | -1.151 | 44.338
0.4786 | -2.3927 | -3.6635 | -218.4938 | 218.5542 | 1.474 | 1.1552 | -0.60104 | 1.9781 | 218.4938 | -1.1552 | 44.4961
0.4804 | -2.3302 | -3.525 | -217.301 | 217.3583 | 1.4729 | 1.1594 | -0.60313 | 1.9792 | 217.301 | -1.1594 | 44.6634
0.4822 | -2.2792 | -3.3448 | -215.8792 | 215.9333 | 1.474 | 1.1635 | -0.60417 | 1.9823 | 215.8792 | -1.1635 | 44.8308
0.484 | -2.226 | -3.1635 | -214.4615 | 214.5125 | 1.474 | 1.1687 | -0.60521 | 1.9854 | 214.4615 | -1.1687 | 44.9981
0.4858 | -2.2323 | -3.0854 | -213.6094 | 213.6615 | 1.4719 | 1.175 | -0.60625 | 1.9875 | 213.6094 | -1.175 | 45.1655
0.4876 | -2.1844 | -3.0031 | -213.075 | 213.126 | 1.4708 | 1.1781 | -0.60521 | 1.9885 | 213.075 | -1.1781 | 45.3328
0.4894 | -2.1677 | -3.0135 | -212.8427 | 212.8948 | 1.4677 | 1.1833 | -0.60313 | 1.9885 | 212.8427 | -1.1833 | 45.5002
0.4912 | -2.274 | -3.1469 | -212.9875 | 213.0458 | 1.4604 | 1.1927 | -0.60104 | 1.9885 | 212.9875 | -1.1927 | 45.6675
0.493 | -2.4135 | -3.3292 | -213.3354 | 213.399 | 1.4521 | 1.2042 | -0.59792 | 1.9885 | 213.3354 | -1.2042 | 45.8349
0.4948 | -2.449 | -3.4021 | -213.7083 | 213.7719 | 1.449 | 1.2104 | -0.59375 | 1.9885 | 213.7083 | -1.2104 | 46.0022
0.4966 | -2.6333 | -3.6208 | -214.2073 | 214.2719 | 1.4396 | 1.2208 | -0.59167 | 1.9854 | 214.2073 | -1.2208 | 46.1696
0.4984 | -2.6354 | -3.6417 | -214.0542 | 214.1177 | 1.4385 | 1.226 | -0.59062 | 1.9865 | 214.0542 | -1.226 | 46.3369
0.5002 | -2.8073 | -3.7146 | -213.6458 | 213.7104 | 1.4333 | 1.2406 | -0.59375 | 1.9917 | 213.6458 | -1.2406 | 46.5043
0.502 | -2.6896 | -3.4458 | -213.1167 | 213.1771 | 1.4406 | 1.2406 | -0.59271 | 1.9969 | 213.1167 | -1.2406 | 46.6716
0.5037 | -2.7198 | -3.424 | -212.8885 | 212.949 | 1.4365 | 1.2437 | -0.59167 | 1.9958 | 212.8885 | -1.2437 | 46.8297
0.5055 | -3.0812 | -3.8375 | -213.0646 | 213.1365 | 1.4125 | 1.2594 | -0.59375 | 1.9885 | 213.0646 | -1.2594 | 46.997
0.5073 | -3.3031 | -4.1219 | -213.026 | 213.1062 | 1.3969 | 1.2677 | -0.59479 | 1.9833 | 213.026 | -1.2677 | 47.1644
0.5091 | -3.2792 | -4.1229 | -212.5281 | 212.6083 | 1.3927 | 1.2677 | -0.59688 | 1.9812 | 212.5281 | -1.2677 | 47.3317
0.5109 | -3.2958 | -4.099 | -211.9719 | 212.051 | 1.3875 | 1.2677 | -0.59792 | 1.9771 | 211.9719 | -1.2677 | 47.4991
0.5127 | -3.3281 | -4.0594 | -211.3187 | 211.3979 | 1.3823 | 1.2677 | -0.60104 | 1.975 | 211.3187 | -1.2677 | 47.6664
0.5145 | -3.1417 | -3.7583 | -210.3354 | 210.4073 | 1.3885 | 1.2604 | -0.60104 | 1.974 | 210.3354 | -1.2604 | 47.8338
0.5163 | -3.0417 | -3.626 | -209.1312 | 209.2 | 1.3854 | 1.2552 | -0.6 | 1.9687 | 209.1312 | -1.2552 | 48.0011
0.5181 | -3.074 | -3.6688 | -207.8594 | 207.9313 | 1.375 | 1.2563 | -0.60104 | 1.9615 | 207.8594 | -1.2563 | 48.1685
0.5199 | -3.0573 | -3.6146 | -206.6667 | 206.7385 | 1.3687 | 1.2563 | -0.60104 | 1.9573 | 206.6667 | -1.2563 | 48.3358
0.5217 | -2.9167 | -3.2635 | -205.7552 | 205.824 | 1.376 | 1.2521 | -0.60104 | 1.9604 | 205.7552 | -1.2521 | 48.5032
0.5235 | -2.7073 | -3.0083 | -205.1885 | 205.2542 | 1.3781 | 1.2406 | -0.59896 | 1.9542 | 205.1885 | -1.2406 | 48.6705
0.5253 | -2.8208 | -3.1625 | -204.9104 | 204.9833 | 1.3656 | 1.2448 | -0.59896 | 1.949 | 204.9104 | -1.2448 | 48.8379
0.5271 | -2.8031 | -3.1583 | -204.7594 | 204.8323 | 1.3646 | 1.2469 | -0.59688 | 1.949 | 204.7594 | -1.2469 | 49.0052
0.5288 | -2.7531 | -3.1073 | -204.5677 | 204.6396 | 1.3667 | 1.2479 | -0.59479 | 1.95 | 204.5677 | -1.2479 | 49.1633
0.5306 | -2.5646 | -2.8458 | -204.199 | 204.2635 | 1.3771 | 1.2448 | -0.59167 | 1.9542 | 204.199 | -1.2448 | 49.3306
0.5324 | -2.376 | -2.5844 | -203.8302 | 203.8885 | 1.3865 | 1.2406 | -0.58854 | 1.9583 | 203.8302 | -1.2406 | 49.498
0.5342 | -2.2125 | -2.3969 | -203.5771 | 203.6271 | 1.3896 | 1.2302 | -0.58646 | 1.951 | 203.5771 | -1.2302 | 49.6653
0.536 | -2.2229 | -2.3438 | -203.6542 | 203.7042 | 1.3885 | 1.2302 | -0.58333 | 1.95 | 203.6542 | -1.2302 | 49.8327
0.5378 | -2.3042 | -2.3198 | -203.7344 | 203.7844 | 1.3885 | 1.2385 | -0.58021 | 1.9531 | 203.7344 | -1.2385 | 50
0.5396 | -2.3208 | -2.2833 | -203.8469 | 203.8958 | 1.3885 | 1.2406 | -0.57708 | 1.9542 | 203.8469 | -1.2406 | 50.1673
0.5414 | -2.2677 | -2.2646 | -204.1135 | 204.1625 | 1.3875 | 1.2365 | -0.575 | 1.95 | 204.1135 | -1.2365 | 50.3347
0.5432 | -2.4667 | -2.5615 | -204.8479 | 204.9042 | 1.374 | 1.2406 | -0.575 | 1.9438 | 204.8479 | -1.2406 | 50.502
0.545 | -2.575 | -2.6521 | -205.6333 | 205.6896 | 1.3698 | 1.2437 | -0.57292 | 1.9417 | 205.6333 | -1.2437 | 50.6694
0.5468 | -2.6812 | -2.7583 | -206.5042 | 206.5615 | 1.3656 | 1.2448 | -0.57188 | 1.9385 | 206.5042 | -1.2448 | 50.8367
0.5486 | -3.0344 | -3.2708 | -207.9125 | 207.9781 | 1.3448 | 1.2531 | -0.57188 | 1.9281 | 207.9125 | -1.2531 | 51.0041
0.5504 | -3.3708 | -3.7719 | -209.3583 | 209.4313 | 1.325 | 1.2594 | -0.57292 | 1.9188 | 209.3583 | -1.2594 | 51.1714
0.5521 | -3.5448 | -4.0667 | -210.5417 | 210.6198 | 1.3156 | 1.2615 | -0.57292 | 1.9125 | 210.5417 | -1.2615 | 51.3295
0.5539 | -3.6979 | -4.3604 | -211.6531 | 211.7385 | 1.3042 | 1.2635 | -0.57083 | 1.9062 | 211.6531 | -1.2635 | 51.4968
0.5557 | -3.7906 | -4.6271 | -212.6792 | 212.7688 | 1.2969 | 1.2635 | -0.57292 | 1.9021 | 212.6792 | -1.2635 | 51.6642
0.5575 | -3.9115 | -4.8625 | -213.7344 | 213.8302 | 1.2937 | 1.2688 | -0.58021 | 1.9052 | 213.7344 | -1.2688 | 51.8315
0.5593 | -3.976 | -5.0052 | -215.0177 | 215.1198 | 1.2948 | 1.2708 | -0.58542 | 1.9094 | 215.0177 | -1.2708 | 51.9989
0.5611 | -3.9125 | -4.9813 | -216.4688 | 216.5719 | 1.3062 | 1.2646 | -0.59167 | 1.9167 | 216.4688 | -1.2646 | 52.1662
0.5629 | -3.9813 | -5.15 | -218.3094 | 218.4188 | 1.3094 | 1.2656 | -0.59583 | 1.9219 | 218.3094 | -1.2656 | 52.3336
0.5647 | -4.1125 | -5.3271 | -220.7615 | 220.875 | 1.3177 | 1.2646 | -0.60313 | 1.9312 | 220.7615 | -1.2646 | 52.5009
0.5665 | -4.124 | -5.3365 | -223.0927 | 223.2052 | 1.3302 | 1.2615 | -0.60625 | 1.9396 | 223.0927 | -1.2615 | 52.6683
0.5683 | -4.0354 | -5.2594 | -224.7073 | 224.8167 | 1.3417 | 1.2552 | -0.60729 | 1.9448 | 224.7073 | -1.2552 | 52.8356
0.5701 | -3.8552 | -5.1292 | -226.1104 | 226.2156 | 1.3542 | 1.2448 | -0.60833 | 1.949 | 226.1104 | -1.2448 | 53.003
0.5719 | -3.7125 | -5.0375 | -227.3427 | 227.4458 | 1.3646 | 1.2354 | -0.61042 | 1.9521 | 227.3427 | -1.2354 | 53.1703
0.5737 | -3.5135 | -4.824 | -228.0448 | 228.1396 | 1.3781 | 1.2281 | -0.6125 | 1.9573 | 228.0448 | -1.2281 | 53.3377
0.5755 | -3.3062 | -4.6417 | -228.3688 | 228.4531 | 1.3906 | 1.2167 | -0.61458 | 1.9594 | 228.3688 | -1.2167 | 53.505
0.5772 | -3.1427 | -4.5094 | -228.75 | 228.8271 | 1.401 | 1.2073 | -0.61771 | 1.9615 | 228.75 | -1.2073 | 53.6631
0.579 | -3.1052 | -4.4469 | -229.6115 | 229.6865 | 1.4052 | 1.2063 | -0.61771 | 1.9635 | 229.6115 | -1.2063 | 53.8304
0.5808 | -3.25 | -4.6094 | -230.5813 | 230.6635 | 1.399 | 1.2125 | -0.61875 | 1.9635 | 230.5813 | -1.2125 | 53.9978
0.5826 | -3.2125 | -4.6 | -231.2583 | 231.3427 | 1.4031 | 1.2073 | -0.62083 | 1.9656 | 231.2583 | -1.2073 | 54.1651
0.5844 | -3.15 | -4.5604 | -231.8708 | 231.9583 | 1.4094 | 1.2021 | -0.62396 | 1.9677 | 231.8708 | -1.2021 | 54.3325
0.5862 | -3.0125 | -4.4073 | -232.3 | 232.3844 | 1.4187 | 1.1948 | -0.62708 | 1.9719 | 232.3 | -1.1948 | 54.4998
0.588 | -2.799 | -4.1156 | -233.0875 | 233.1604 | 1.4302 | 1.1885 | -0.62604 | 1.976 | 233.0875 | -1.1885 | 54.6672
0.5898 | -2.9375 | -4.2958 | -233.9458 | 234.026 | 1.4219 | 1.1906 | -0.62708 | 1.9729 | 233.9458 | -1.1906 | 54.8345
0.5916 | -3.0854 | -4.4562 | -234.3833 | 234.4729 | 1.4135 | 1.1927 | -0.62813 | 1.9698 | 234.3833 | -1.1927 | 55.0019
0.5934 | -3.1052 | -4.4031 | -233.9812 | 234.0719 | 1.4094 | 1.1896 | -0.62917 | 1.9646 | 233.9812 | -1.1896 | 55.1692
0.5952 | -3.1042 | -4.3271 | -233.475 | 233.5625 | 1.4031 | 1.1865 | -0.62813 | 1.9573 | 233.475 | -1.1865 | 55.3366
0.597 | -3.1031 | -4.2854 | -232.749 | 232.8344 | 1.3958 | 1.1823 | -0.62917 | 1.949 | 232.749 | -1.1823 | 55.5039
0.5988 | -3.1031 | -4.3052 | -231.651 | 231.7375 | 1.3844 | 1.1771 | -0.63021 | 1.9385 | 231.651 | -1.1771 | 55.6713
0.6006 | -3.126 | -4.349 | -230.424 | 230.5104 | 1.3719 | 1.176 | -0.63021 | 1.9281 | 230.424 | -1.176 | 55.8386
0.6023 | -3.074 | -4.301 | -229.0073 | 229.0917 | 1.3615 | 1.176 | -0.62917 | 1.9188 | 229.0073 | -1.176 | 55.9967
0.6041 | -3.2031 | -4.3625 | -227.7292 | 227.8125 | 1.3479 | 1.1833 | -0.62813 | 1.9104 | 227.7292 | -1.1833 | 56.164
0.6059 | -3.4521 | -4.4062 | -226.4052 | 226.4875 | 1.3375 | 1.1927 | -0.62813 | 1.9073 | 226.4052 | -1.1927 | 56.3313
0.6077 | -3.5771 | -4.3823 | -225.4479 | 225.5271 | 1.3323 | 1.1969 | -0.62604 | 1.9042 | 225.4479 | -1.1969 | 56.4987
0.6095 | -3.5438 | -4.3573 | -225.4885 | 225.5656 | 1.3313 | 1.1906 | -0.62187 | 1.899 | 225.4885 | -1.1906 | 56.666
0.6113 | -3.7604 | -4.6375 | -226.1385 | 226.224 | 1.3208 | 1.1948 | -0.61979 | 1.8917 | 226.1385 | -1.1948 | 56.8334
0.6131 | -4.05 | -4.8667 | -226.7094 | 226.8042 | 1.3146 | 1.2021 | -0.62292 | 1.8927 | 226.7094 | -1.2021 | 57.0007
0.6149 | -4.124 | -4.8583 | -226.8792 | 226.974 | 1.3167 | 1.2052 | -0.62604 | 1.8979 | 226.8792 | -1.2052 | 57.1681
0.6167 | -3.9656 | -4.5708 | -226.8302 | 226.9156 | 1.3302 | 1.2 | -0.62604 | 1.9062 | 226.8302 | -1.2 | 57.3354
0.6185 | -3.6948 | -4.1417 | -226.6146 | 226.6854 | 1.3531 | 1.1917 | -0.62604 | 1.9167 | 226.6146 | -1.1917 | 57.5028
0.6203 | -3.3417 | -3.7146 | -227.1604 | 227.2167 | 1.375 | 1.176 | -0.62083 | 1.9229 | 227.1604 | -1.176 | 57.6701
0.6221 | -3.4052 | -3.7771 | -228.1708 | 228.2292 | 1.3729 | 1.1812 | -0.61875 | 1.924 | 228.1708 | -1.1812 | 57.8375
0.6239 | -3.5135 | -3.9417 | -229.2385 | 229.301 | 1.3719 | 1.1844 | -0.62187 | 1.926 | 229.2385 | -1.1844 | 58.0048
0.6256 | -3.3635 | -3.8531 | -230.1313 | 230.1896 | 1.3833 | 1.1771 | -0.625 | 1.9312 | 230.1313 | -1.1771 | 58.1629
0.6274 | -3.1781 | -3.7948 | -231.2531 | 231.3073 | 1.3927 | 1.1646 | -0.62604 | 1.9323 | 231.2531 | -1.1646 | 58.3302
0.6292 | -3.0823 | -3.9427 | -232.7396 | 232.7948 | 1.3917 | 1.151 | -0.62292 | 1.9219 | 232.7396 | -1.151 | 58.4976
0.631 | -3.2646 | -4.3521 | -234.4125 | 234.4771 | 1.3802 | 1.149 | -0.62187 | 1.9125 | 234.4125 | -1.149 | 58.6649
0.6328 | -3.3615 | -4.5927 | -235.5271 | 235.5958 | 1.3719 | 1.151 | -0.62083 | 1.9094 | 235.5271 | -1.151 | 58.8323
0.6346 | -3.375 | -4.676 | -235.7906 | 235.8635 | 1.3698 | 1.149 | -0.62292 | 1.9062 | 235.7906 | -1.149 | 58.9996
0.6364 | -3.2031 | -4.5385 | -235.9042 | 235.9708 | 1.375 | 1.1406 | -0.62292 | 1.9062 | 235.9042 | -1.1406 | 59.167
0.6382 | -3 | -4.374 | -236.0062 | 236.0667 | 1.3813 | 1.1292 | -0.62083 | 1.9052 | 236.0062 | -1.1292 | 59.3343
0.64 | -2.9385 | -4.3948 | -236.726 | 236.7865 | 1.376 | 1.124 | -0.61562 | 1.8969 | 236.726 | -1.124 | 59.5017
0.6418 | -3.0344 | -4.5771 | -237.474 | 237.5396 | 1.3635 | 1.124 | -0.60938 | 1.8854 | 237.474 | -1.124 | 59.669
0.6436 | -3.0146 | -4.5125 | -237.4021 | 237.4667 | 1.3635 | 1.126 | -0.60938 | 1.8865 | 237.4021 | -1.126 | 59.8364
0.6454 | -2.849 | -4.2719 | -236.901 | 236.9604 | 1.3729 | 1.1219 | -0.60833 | 1.8917 | 236.901 | -1.1219 | 60.0037
0.6472 | -2.6531 | -4.1615 | -236.4375 | 236.4917 | 1.375 | 1.1177 | -0.60625 | 1.8906 | 236.4375 | -1.1177 | 60.1711
0.649 | -2.5938 | -4.1844 | -236.5625 | 236.6177 | 1.3698 | 1.1198 | -0.60104 | 1.8854 | 236.5625 | -1.1198 | 60.3384
0.6507 | -2.5292 | -4.1219 | -236.2865 | 236.3417 | 1.3677 | 1.1229 | -0.59792 | 1.8833 | 236.2865 | -1.1229 | 60.4965
0.6525 | -2.3521 | -3.7844 | -234.9146 | 234.9635 | 1.3771 | 1.1271 | -0.59583 | 1.8927 | 234.9146 | -1.1271 | 60.6638
0.6543 | -2.1865 | -3.5365 | -233.7708 | 233.8115 | 1.3802 | 1.1344 | -0.59062 | 1.8969 | 233.7708 | -1.1344 | 60.8312
0.6561 | -2.1771 | -3.4219 | -232.8344 | 232.8719 | 1.3781 | 1.1469 | -0.5875 | 1.901 | 232.8344 | -1.1469 | 60.9985
0.6579 | -2.2552 | -3.3615 | -231.949 | 231.9865 | 1.374 | 1.1625 | -0.58542 | 1.9052 | 231.949 | -1.1625 | 61.1659
0.6597 | -2.1646 | -3.1063 | -230.925 | 230.9573 | 1.3719 | 1.1677 | -0.57812 | 1.9042 | 230.925 | -1.1677 | 61.3332
0.6615 | -2.0562 | -2.849 | -229.9656 | 229.9927 | 1.3698 | 1.1708 | -0.57083 | 1.901 | 229.9656 | -1.1708 | 61.5006
0.6633 | -2.1156 | -2.7667 | -228.3688 | 228.3958 | 1.3573 | 1.1792 | -0.56563 | 1.8948 | 228.3688 | -1.1792 | 61.6679
0.6651 | -2.2969 | -2.8729 | -226.9 | 226.9313 | 1.3375 | 1.1917 | -0.56354 | 1.8875 | 226.9 | -1.1917 | 61.8353
0.6669 | -2.5938 | -3.1865 | -226.376 | 226.4146 | 1.3156 | 1.2083 | -0.56146 | 1.8802 | 226.376 | -1.2083 | 62.0026
0.6687 | -2.7073 | -3.3198 | -226.2542 | 226.2969 | 1.3042 | 1.2177 | -0.55833 | 1.876 | 226.2542 | -1.2177 | 62.17
0.6705 | -2.876 | -3.5219 | -226.4073 | 226.4552 | 1.2917 | 1.2312 | -0.55833 | 1.875 | 226.4073 | -1.2312 | 62.3373
0.6723 | -2.8667 | -3.5844 | -226.5531 | 226.6031 | 1.2865 | 1.2375 | -0.55625 | 1.875 | 226.5531 | -1.2375 | 62.5046
0.6741 | -2.7521 | -3.3917 | -226.4771 | 226.5229 | 1.2948 | 1.2417 | -0.55521 | 1.8833 | 226.4771 | -1.2417 | 62.672
0.6758 | -2.6354 | -3.1938 | -226.399 | 226.4406 | 1.3042 | 1.2458 | -0.55417 | 1.8927 | 226.399 | -1.2458 | 62.83
0.6776 | -2.5271 | -3.0333 | -226.4521 | 226.4896 | 1.3104 | 1.249 | -0.55312 | 1.899 | 226.4521 | -1.249 | 62.9974
0.6794 | -2.4323 | -2.9573 | -226.8052 | 226.8406 | 1.3115 | 1.249 | -0.54792 | 1.8979 | 226.8052 | -1.249 | 63.1647
0.6812 | -2.626 | -3.2094 | -227.7823 | 227.825 | 1.2969 | 1.2594 | -0.54375 | 1.8927 | 227.7823 | -1.2594 | 63.3321
0.683 | -2.8271 | -3.4687 | -229.1188 | 229.1677 | 1.2917 | 1.274 | -0.54479 | 1.899 | 229.1188 | -1.274 | 63.4994
0.6848 | -2.9688 | -3.675 | -230.4906 | 230.5437 | 1.2896 | 1.2875 | -0.54583 | 1.9062 | 230.4906 | -1.2875 | 63.6668
0.6866 | -2.9323 | -3.7094 | -231.8906 | 231.9437 | 1.2937 | 1.2948 | -0.54688 | 1.9146 | 231.8906 | -1.2948 | 63.8341
0.6884 | -2.7938 | -3.5427 | -233.0708 | 233.1188 | 1.3083 | 1.301 | -0.54792 | 1.9281 | 233.0708 | -1.301 | 64.0015
0.6902 | -2.5427 | -3.2844 | -233.8021 | 233.8427 | 1.3198 | 1.2958 | -0.54792 | 1.9333 | 233.8021 | -1.2958 | 64.1688
0.692 | -2.4104 | -3.0854 | -234.6823 | 234.7188 | 1.326 | 1.2917 | -0.54583 | 1.9354 | 234.6823 | -1.2917 | 64.3362
0.6938 | -2.5885 | -3.2563 | -235.5833 | 235.6229 | 1.3156 | 1.299 | -0.54583 | 1.9333 | 235.5833 | -1.299 | 64.5035
0.6956 | -2.6229 | -3.2646 | -235.7281 | 235.7698 | 1.3104 | 1.301 | -0.54688 | 1.9312 | 235.7281 | -1.301 | 64.6709
0.6974 | -2.2677 | -2.8146 | -235.0187 | 235.049 | 1.3208 | 1.2927 | -0.54167 | 1.9312 | 235.0187 | -1.2927 | 64.8382
0.6992 | -2.2115 | -2.7823 | -234.224 | 234.2531 | 1.3052 | 1.2917 | -0.53438 | 1.9177 | 234.224 | -1.2917 | 65.0056
0.7009 | -2.2375 | -2.8969 | -233.4125 | 233.4427 | 1.2802 | 1.2927 | -0.52812 | 1.899 | 233.4125 | -1.2927 | 65.1636
0.7027 | -2.0802 | -2.7531 | -232.576 | 232.6031 | 1.2719 | 1.2885 | -0.52396 | 1.8896 | 232.576 | -1.2885 | 65.331
0.7045 | -1.8875 | -2.5573 | -232.0094 | 232.0323 | 1.2688 | 1.2802 | -0.52083 | 1.8823 | 232.0094 | -1.2802 | 65.4983
0.7063 | -1.7625 | -2.4542 | -231.9385 | 231.9594 | 1.2635 | 1.276 | -0.51875 | 1.874 | 231.9385 | -1.276 | 65.6657
0.7081 | -1.6927 | -2.4115 | -231.9406 | 231.9615 | 1.2542 | 1.2771 | -0.51667 | 1.8677 | 231.9406 | -1.2771 | 65.833
0.7099 | -1.7958 | -2.5146 | -232.2865 | 232.3083 | 1.2427 | 1.276 | -0.51458 | 1.8594 | 232.2865 | -1.276 | 66.0004
0.7117 | -1.9052 | -2.5833 | -232.5729 | 232.5979 | 1.2344 | 1.276 | -0.5125 | 1.8531 | 232.5729 | -1.276 | 66.1677
0.7135 | -1.7365 | -2.3781 | -232.2792 | 232.3 | 1.2406 | 1.2615 | -0.51146 | 1.8479 | 232.2792 | -1.2615 | 66.3351
0.7153 | -1.6281 | -2.2885 | -232.276 | 232.2948 | 1.2437 | 1.2479 | -0.5125 | 1.8406 | 232.276 | -1.2479 | 66.5024
0.7171 | -1.5792 | -2.301 | -232.499 | 232.5198 | 1.2427 | 1.2344 | -0.51354 | 1.8312 | 232.499 | -1.2344 | 66.6698
0.7189 | -1.5708 | -2.4615 | -233.0885 | 233.1094 | 1.2375 | 1.2208 | -0.51458 | 1.8188 | 233.0885 | -1.2208 | 66.8371
0.7207 | -1.4844 | -2.5656 | -233.5854 | 233.6062 | 1.2354 | 1.2052 | -0.51458 | 1.8062 | 233.5854 | -1.2052 | 67.0045
0.7225 | -1.4281 | -2.6583 | -233.6531 | 233.676 | 1.2375 | 1.1938 | -0.51667 | 1.801 | 233.6531 | -1.1938 | 67.1718
0.7242 | -1.1625 | -2.5208 | -233.2833 | 233.3021 | 1.251 | 1.174 | -0.51771 | 1.8 | 233.2833 | -1.174 | 67.3299
0.726 | -1.1531 | -2.7042 | -233.2438 | 233.2656 | 1.25 | 1.1656 | -0.51979 | 1.7937 | 233.2438 | -1.1656 | 67.4972
0.7278 | -1.0729 | -2.7792 | -232.9646 | 232.9875 | 1.2469 | 1.1542 | -0.51771 | 1.7844 | 232.9646 | -1.1542 | 67.6646
0.7296 | -0.97292 | -2.8135 | -232.6281 | 232.6521 | 1.2458 | 1.1417 | -0.51667 | 1.775 | 232.6281 | -1.1417 | 67.8319
0.7314 | -0.74792 | -2.5719 | -231.9156 | 231.9365 | 1.2531 | 1.124 | -0.51458 | 1.7677 | 231.9156 | -1.124 | 67.9993
0.7332 | -0.62187 | -2.3615 | -231.2906 | 231.3083 | 1.2552 | 1.1073 | -0.5125 | 1.7583 | 231.2906 | -1.1073 | 68.1666
0.735 | -0.45 | -2.0792 | -230.4135 | 230.4302 | 1.2615 | 1.0875 | -0.51042 | 1.75 | 230.4135 | -1.0875 | 68.334
0.7368 | -0.175 | -1.6875 | -229.1646 | 229.1813 | 1.2729 | 1.0635 | -0.51042 | 1.7427 | 229.1646 | -1.0635 | 68.5013
0.7386 | 0.042708 | -1.2917 | -227.9531 | 227.9708 | 1.2802 | 1.0385 | -0.50729 | 1.7302 | 227.9531 | -1.0385 | 68.6687
0.7404 | 0.10208 | -1.1594 | -226.8583 | 226.8771 | 1.2802 | 1.0187 | -0.50417 | 1.7177 | 226.8583 | -1.0187 | 68.836
0.7422 | 0.23229 | -0.93542 | -225.7958 | 225.8156 | 1.2792 | 0.99271 | -0.49687 | 1.7 | 225.7958 | -0.99271 | 69.0033
0.744 | 0.36146 | -0.71146 | -224.7333 | 224.7542 | 1.2781 | 0.96667 | -0.49062 | 1.6823 | 224.7333 | -0.96667 | 69.1707
0.7458 | 0.36979 | -0.67917 | -223.4375 | 223.4615 | 1.2677 | 0.94375 | -0.4875 | 1.6604 | 223.4375 | -0.94375 | 69.338
0.7476 | 0.42708 | -0.53125 | -221.876 | 221.9031 | 1.2625 | 0.91667 | -0.48438 | 1.6396 | 221.876 | -0.91667 | 69.5054
0.7493 | 0.60208 | -0.15313 | -220.3052 | 220.3417 | 1.2583 | 0.88125 | -0.48333 | 1.6146 | 220.3052 | -0.88125 | 69.6634
0.7511 | 0.80104 | 0.27083 | -218.7438 | 218.7906 | 1.2542 | 0.84271 | -0.48229 | 1.5885 | 218.7438 | -0.84271 | 69.8308
0.7529 | 0.575 | 0.027083 | -217.299 | 217.3469 | 1.2333 | 0.82292 | -0.48438 | 1.5615 | 217.299 | -0.82292 | 69.9981
0.7547 | 0.49896 | -0.057292 | -215.4354 | 215.4854 | 1.2198 | 0.80104 | -0.48542 | 1.5406 | 215.4354 | -0.80104 | 70.1655
0.7565 | 0.61146 | 0.10833 | -213.1531 | 213.2031 | 1.2135 | 0.77083 | -0.48646 | 1.5198 | 213.1531 | -0.77083 | 70.3328
0.7583 | 0.69167 | 0.34063 | -210.475 | 210.5292 | 1.2083 | 0.74271 | -0.48958 | 1.5031 | 210.475 | -0.74271 | 70.5002
0.7601 | 0.77604 | 0.55312 | -207.776 | 207.8375 | 1.1948 | 0.70729 | -0.49167 | 1.476 | 207.776 | -0.70729 | 70.6675
0.7619 | 0.86562 | 0.76875 | -205.074 | 205.1417 | 1.1812 | 0.67292 | -0.49375 | 1.4479 | 205.074 | -0.67292 | 70.8349
0.7637 | 0.96458 | 0.97604 | -202.3396 | 202.4156 | 1.1667 | 0.63646 | -0.49375 | 1.4198 | 202.3396 | -0.63646 | 71.0022
0.7655 | 1.1333 | 1.1896 | -199.4094 | 199.4958 | 1.1552 | 0.59583 | -0.48958 | 1.3906 | 199.4094 | -0.59583 | 71.1696
0.7673 | 1.3531 | 1.425 | -196.0979 | 196.1865 | 1.1448 | 0.55521 | -0.48333 | 1.3646 | 196.0979 | -0.55521 | 71.3369
0.7691 | 1.4979 | 1.599 | -192.8865 | 192.975 | 1.1313 | 0.52083 | -0.47813 | 1.3375 | 192.8865 | -0.52083 | 71.5043
0.7709 | 1.5385 | 1.6427 | -190.026 | 190.1125 | 1.1115 | 0.49167 | -0.47396 | 1.3083 | 190.026 | -0.49167 | 71.6716
0.7727 | 1.5229 | 1.6073 | -187.524 | 187.6094 | 1.0885 | 0.46667 | -0.46979 | 1.2781 | 187.524 | -0.46667 | 71.839
0.7744 | 1.5385 | 1.6323 | -185.2604 | 185.3469 | 1.0708 | 0.45 | -0.46771 | 1.2552 | 185.2604 | -0.45 | 71.997
0.7762 | 1.6021 | 1.7208 | -182.3854 | 182.4688 | 1.0521 | 0.42083 | -0.45937 | 1.2271 | 182.3854 | -0.42083 | 72.1644
0.778 | 1.5188 | 1.6344 | -179.8052 | 179.8885 | 1.0281 | 0.39479 | -0.45312 | 1.1958 | 179.8052 | -0.39479 | 72.3317
0.7798 | 1.4458 | 1.5323 | -177.5479 | 177.6302 | 1.0073 | 0.37292 | -0.45 | 1.1708 | 177.5479 | -0.37292 | 72.4991
0.7816 | 1.3427 | 1.3781 | -175.3969 | 175.4771 | 0.98437 | 0.35313 | -0.44688 | 1.1448 | 175.3969 | -0.35313 | 72.6664
0.7834 | 0.9625 | 0.90208 | -173.8844 | 173.9469 | 0.95104 | 0.35104 | -0.44583 | 1.1167 | 173.8844 | -0.35104 | 72.8338
0.7852 | 0.51562 | 0.48333 | -172.324 | 172.3719 | 0.92188 | 0.35313 | -0.44479 | 1.0927 | 172.324 | -0.35313 | 73.0011
0.787 | 0.27917 | 0.41042 | -170.2854 | 170.3385 | 0.90312 | 0.34687 | -0.44063 | 1.075 | 170.2854 | -0.34687 | 73.1685
0.7888 | 0.1 | 0.36979 | -168.1635 | 168.2229 | 0.88542 | 0.3375 | -0.43438 | 1.0562 | 168.1635 | -0.3375 | 73.3358
0.7906 | 0.075 | 0.49583 | -165.2427 | 165.3073 | 0.86979 | 0.32396 | -0.42917 | 1.0354 | 165.2427 | -0.32396 | 73.5032
0.7924 | 0.065625 | 0.75417 | -162.0958 | 162.1635 | 0.85729 | 0.31042 | -0.42396 | 1.0198 | 162.0958 | -0.31042 | 73.6705
0.7942 | 0.13229 | 1.4062 | -158.6729 | 158.7427 | 0.85417 | 0.29479 | -0.41563 | 1.0094 | 158.6729 | -0.29479 | 73.8379
0.796 | 0.13854 | 1.8427 | -155.8 | 155.875 | 0.84583 | 0.27604 | -0.40521 | 0.99167 | 155.8 | -0.27604 | 74.0052
0.7977 | 0.16146 | 2.1708 | -153.0042 | 153.0833 | 0.83854 | 0.25729 | -0.39687 | 0.97708 | 153.0042 | -0.25729 | 74.1633
0.7995 | 0.15208 | 2.3604 | -150.5906 | 150.676 | 0.82708 | 0.24167 | -0.38854 | 0.95938 | 150.5906 | -0.24167 | 74.3306
0.8013 | 0.19896 | 2.5323 | -148.2385 | 148.3323 | 0.81146 | 0.22396 | -0.38125 | 0.9375 | 148.2385 | -0.22396 | 74.498
0.8031 | 0.25104 | 2.6437 | -146.0229 | 146.1208 | 0.79479 | 0.20521 | -0.37396 | 0.91458 | 146.0229 | -0.20521 | 74.6653
0.8049 | 0.29271 | 2.6031 | -144.2708 | 144.3677 | 0.77396 | 0.18542 | -0.36667 | 0.88854 | 144.2708 | -0.18542 | 74.8327
0.8067 | 0.29583 | 2.5854 | -142.7917 | 142.8896 | 0.74896 | 0.16667 | -0.35833 | 0.85729 | 142.7917 | -0.16667 | 75
0.8085 | 0.55521 | 2.8917 | -140.7021 | 140.8083 | 0.74479 | 0.13958 | -0.34896 | 0.84479 | 140.7021 | -0.13958 | 75.1673
0.8103 | 0.88125 | 3.3396 | -138.3281 | 138.4344 | 0.74896 | 0.10938 | -0.3375 | 0.84271 | 138.3281 | -0.10938 | 75.3347
0.8121 | 1.1479 | 3.7948 | -135.8083 | 135.9281 | 0.75313 | 0.085417 | -0.32812 | 0.84167 | 135.8083 | -0.085417 | 75.502
0.8139 | 1.3302 | 4.1552 | -133.2281 | 133.3635 | 0.75208 | 0.067708 | -0.31771 | 0.8375 | 133.2281 | -0.067708 | 75.6694
0.8157 | 1.474 | 4.4896 | -130.6271 | 130.7667 | 0.75 | 0.055208 | -0.30833 | 0.83229 | 130.6271 | -0.055208 | 75.8367
0.8175 | 1.474 | 4.6271 | -128.6177 | 128.7667 | 0.73958 | 0.046875 | -0.30104 | 0.81771 | 128.6177 | -0.046875 | 76.0041
0.8193 | 1.3448 | 4.549 | -126.9375 | 127.0833 | 0.72083 | 0.042708 | -0.29375 | 0.79792 | 126.9375 | -0.042708 | 76.1714
0.8211 | 1.2323 | 4.4927 | -125.2677 | 125.4135 | 0.70208 | 0.0375 | -0.28646 | 0.77812 | 125.2677 | -0.0375 | 76.3388
0.8228 | 1.1427 | 4.425 | -123.701 | 123.8438 | 0.68021 | 0.033333 | -0.28229 | 0.75521 | 123.701 | -0.033333 | 76.4968
0.8246 | 1.075 | 4.4323 | -121.8812 | 122.025 | 0.65833 | 0.030208 | -0.27813 | 0.73333 | 121.8812 | -0.030208 | 76.6642
0.8264 | 1.0323 | 4.625 | -120.1115 | 120.2719 | 0.64375 | 0.029167 | -0.27396 | 0.71771 | 120.1115 | -0.029167 | 76.8315
0.8282 | 0.87708 | 4.5604 | -118.8094 | 118.9719 | 0.625 | 0.032292 | -0.26875 | 0.69896 | 118.8094 | -0.032292 | 76.9989
0.83 | 0.74687 | 4.5729 | -117.6427 | 117.8063 | 0.6125 | 0.038542 | -0.26771 | 0.6875 | 117.6427 | -0.038542 | 77.1662
0.8318 | 0.62917 | 4.5906 | -116.5323 | 116.699 | 0.6 | 0.044792 | -0.26771 | 0.67708 | 116.5323 | -0.044792 | 77.3336
0.8336 | 0.525 | 4.5156 | -115.7625 | 115.9313 | 0.58542 | 0.054167 | -0.26667 | 0.66354 | 115.7625 | -0.054167 | 77.5009
0.8354 | 0.41771 | 4.4938 | -115.001 | 115.174 | 0.57083 | 0.059375 | -0.26458 | 0.65 | 115.001 | -0.059375 | 77.6683
0.8372 | 0.25833 | 4.4979 | -114.326 | 114.5146 | 0.55729 | 0.066667 | -0.26458 | 0.63646 | 114.326 | -0.066667 | 77.8356
0.839 | 0.010417 | 4.376 | -113.7698 | 113.9771 | 0.54063 | 0.079167 | -0.26458 | 0.62187 | 113.7698 | -0.079167 | 78.003
0.8408 | -0.015625 | 4.5083 | -113.0677 | 113.2969 | 0.53125 | 0.084375 | -0.26667 | 0.61458 | 113.0677 | -0.084375 | 78.1703
0.8426 | -0.013542 | 4.7229 | -112.324 | 112.5854 | 0.53021 | 0.094792 | -0.27083 | 0.61562 | 112.324 | -0.094792 | 78.3377
0.8444 | -0.15313 | 4.8063 | -111.6625 | 111.9479 | 0.51979 | 0.10833 | -0.275 | 0.61146 | 111.6625 | -0.10833 | 78.505
0.8462 | -0.36771 | 4.8083 | -111.1635 | 111.4656 | 0.50313 | 0.12292 | -0.27917 | 0.60521 | 111.1635 | -0.12292 | 78.6724
0.8479 | -0.63438 | 4.8125 | -110.8625 | 111.174 | 0.48542 | 0.14167 | -0.28542 | 0.60104 | 110.8625 | -0.14167 | 78.8304
0.8497 | -0.77812 | 4.9323 | -110.7406 | 111.0615 | 0.48021 | 0.15833 | -0.29375 | 0.60625 | 110.7406 | -0.15833 | 78.9978
0.8515 | -1.1552 | 4.7208 | -110.849 | 111.1812 | 0.46354 | 0.18438 | -0.29792 | 0.60625 | 110.849 | -0.18438 | 79.1651
0.8533 | -1.4396 | 4.5417 | -111.1125 | 111.4344 | 0.45 | 0.20625 | -0.30208 | 0.60729 | 111.1125 | -0.20625 | 79.3325
0.8551 | -1.7031 | 4.3531 | -111.325 | 111.6302 | 0.43646 | 0.22917 | -0.30521 | 0.61146 | 111.325 | -0.22917 | 79.4998
0.8569 | -2.0073 | 4.099 | -111.3271 | 111.6344 | 0.42396 | 0.25521 | -0.30833 | 0.61667 | 111.3271 | -0.25521 | 79.6672
0.8587 | -2.2677 | 3.8938 | -111.2365 | 111.5417 | 0.41146 | 0.28229 | -0.30938 | 0.625 | 111.2365 | -0.28229 | 79.8345
0.8605 | -2.4719 | 3.7615 | -111.1219 | 111.4198 | 0.40313 | 0.30938 | -0.31146 | 0.63958 | 111.1219 | -0.30938 | 80.0019
0.8623 | -2.6417 | 3.6656 | -110.9521 | 111.2469 | 0.39792 | 0.33437 | -0.31667 | 0.65938 | 110.9521 | -0.33437 | 80.1692
0.8641 | -2.8115 | 3.5615 | -110.776 | 111.0677 | 0.39271 | 0.36042 | -0.32083 | 0.67917 | 110.776 | -0.36042 | 80.3366
0.8659 | -2.9729 | 3.3385 | -110.5031 | 110.7812 | 0.38438 | 0.38125 | -0.31667 | 0.6875 | 110.5031 | -0.38125 | 80.5039
0.8677 | -3.1792 | 3.0885 | -110.0833 | 110.3542 | 0.37708 | 0.40104 | -0.3125 | 0.69375 | 110.0833 | -0.40104 | 80.6713
0.8695 | -3.3219 | 2.924 | -109.5354 | 109.8052 | 0.37604 | 0.41771 | -0.31042 | 0.70104 | 109.5354 | -0.41771 | 80.8386
0.8712 | -3.3188 | 2.9052 | -109.124 | 109.3937 | 0.38229 | 0.42396 | -0.30938 | 0.70729 | 109.124 | -0.42396 | 80.9967
0.873 | -3.1729 | 2.9948 | -108.5 | 108.775 | 0.39062 | 0.42708 | -0.30833 | 0.7125 | 108.5 | -0.42708 | 81.164
0.8748 | -3.2177 | 2.9021 | -108.1521 | 108.4365 | 0.39583 | 0.43438 | -0.30729 | 0.71562 | 108.1521 | -0.43438 | 81.3313
0.8766 | -3.3719 | 2.5552 | -108.0698 | 108.3365 | 0.39062 | 0.43646 | -0.30833 | 0.71667 | 108.0698 | -0.43646 | 81.4987
0.8784 | -3.5781 | 2.1458 | -108.1083 | 108.375 | 0.38438 | 0.44271 | -0.30729 | 0.71354 | 108.1083 | -0.44271 | 81.666
0.8802 | -3.6125 | 1.8802 | -107.8125 | 108.074 | 0.38125 | 0.44167 | -0.30625 | 0.71146 | 107.8125 | -0.44167 | 81.8334
0.882 | -3.6177 | 1.6052 | -107.4906 | 107.7417 | 0.38229 | 0.43229 | -0.31042 | 0.70833 | 107.4906 | -0.43229 | 82.0007
0.8838 | -3.6573 | 1.3208 | -107.1583 | 107.399 | 0.38021 | 0.42292 | -0.31146 | 0.70521 | 107.1583 | -0.42292 | 82.1681
0.8856 | -3.8469 | 1.0167 | -106.8177 | 107.0448 | 0.36562 | 0.425 | -0.29896 | 0.7 | 106.8177 | -0.425 | 82.3354
0.8874 | -3.8708 | 0.82188 | -106.1979 | 106.4135 | 0.36146 | 0.41771 | -0.29167 | 0.69063 | 106.1979 | -0.41771 | 82.5028
0.8892 | -3.7563 | 0.71979 | -105.3406 | 105.5469 | 0.36562 | 0.40104 | -0.2875 | 0.67708 | 105.3406 | -0.40104 | 82.6701
0.891 | -3.5729 | 0.87813 | -104.1677 | 104.3542 | 0.37396 | 0.39167 | -0.28437 | 0.67708 | 104.1677 | -0.39167 | 82.8375
0.8928 | -3.3354 | 1.1156 | -103.2031 | 103.3677 | 0.375 | 0.37812 | -0.27708 | 0.66875 | 103.2031 | -0.37812 | 83.0048
0.8946 | -2.926 | 1.3562 | -101.8427 | 101.976 | 0.37812 | 0.35208 | -0.26979 | 0.65312 | 101.8427 | -0.35208 | 83.1722
0.8963 | -2.4156 | 1.8167 | -99.849 | 99.951 | 0.38021 | 0.32812 | -0.26771 | 0.64375 | 99.849 | -0.32812 | 83.3302
0.8981 | -2.1271 | 2.0729 | -98.1688 | 98.2573 | 0.37604 | 0.3125 | -0.26875 | 0.62917 | 98.1688 | -0.3125 | 83.4976
0.8999 | -2.0292 | 2.1688 | -96.9062 | 96.9875 | 0.36979 | 0.29792 | -0.26979 | 0.60521 | 96.9062 | -0.29792 | 83.6649
0.9017 | -2.1021 | 2.1323 | -96.0062 | 96.0833 | 0.36146 | 0.28646 | -0.26875 | 0.57917 | 96.0062 | -0.28646 | 83.8323
0.9035 | -2.2646 | 1.8854 | -95.5344 | 95.6031 | 0.3625 | 0.26771 | -0.27187 | 0.56042 | 95.5344 | -0.26771 | 83.9996
0.9053 | -2.3406 | 1.6594 | -95.126 | 95.1906 | 0.36979 | 0.24375 | -0.27708 | 0.55 | 95.126 | -0.24375 | 84.167
0.9071 | -2.4167 | 1.4333 | -94.7177 | 94.7771 | 0.37812 | 0.22083 | -0.28125 | 0.54063 | 94.7177 | -0.22083 | 84.3343
0.9089 | -2.6146 | 1.2958 | -94.5646 | 94.6271 | 0.37396 | 0.21042 | -0.28125 | 0.53021 | 94.5646 | -0.21042 | 84.5017
0.9107 | -2.725 | 1.3062 | -94.3969 | 94.4625 | 0.36875 | 0.2 | -0.27917 | 0.52083 | 94.3969 | -0.2 | 84.669
0.9125 | -2.826 | 1.1646 | -94.5208 | 94.5865 | 0.3625 | 0.18542 | -0.28021 | 0.51146 | 94.5208 | -0.18542 | 84.8364
0.9143 | -2.7281 | 1.2188 | -94.474 | 94.5323 | 0.3625 | 0.15625 | -0.28229 | 0.50521 | 94.474 | -0.15625 | 85.0037
0.9161 | -2.6354 | 1.5302 | -94.3833 | 94.4406 | 0.37292 | 0.11667 | -0.28542 | 0.51562 | 94.3833 | -0.11667 | 85.1711
0.9179 | -2.5427 | 1.8427 | -94.2927 | 94.349 | 0.38438 | 0.077083 | -0.28958 | 0.52604 | 94.2927 | -0.077083 | 85.3384
0.9197 | -2.426 | 1.9792 | -94.0406 | 94.0938 | 0.39062 | 0.0375 | -0.29896 | 0.54063 | 94.0406 | -0.0375 | 85.5058
0.9214 | -2.325 | 2.0573 | -93.8427 | 93.8948 | 0.39792 | -0.0041667 | -0.30938 | 0.5625 | 93.8427 | 0.0041667 | 85.6638
0.9232 | -2.2427 | 2.1167 | -93.7281 | 93.7792 | 0.40417 | -0.04375 | -0.31979 | 0.58333 | 93.7281 | 0.04375 | 85.8312
0.925 | -2.2719 | 2.0906 | -93.9167 | 93.9688 | 0.41042 | -0.077083 | -0.33021 | 0.6 | 93.9167 | 0.077083 | 85.9985
0.9268 | -2.4615 | 1.9906 | -94.3021 | 94.3573 | 0.41354 | -0.10729 | -0.33854 | 0.61458 | 94.3021 | 0.10729 | 86.1659
0.9286 | -2.7 | 1.649 | -94.9437 | 95 | 0.40313 | -0.13542 | -0.35 | 0.625 | 94.9437 | 0.13542 | 86.3332
0.9304 | -2.7417 | 1.5115 | -95.3042 | 95.3594 | 0.4 | -0.16771 | -0.3625 | 0.64167 | 95.3042 | 0.16771 | 86.5006
0.9322 | -2.5125 | 1.7104 | -95.2396 | 95.2927 | 0.41146 | -0.20625 | -0.37396 | 0.66875 | 95.2396 | 0.20625 | 86.6679
0.934 | -2.276 | 2.1198 | -95.0177 | 95.074 | 0.42917 | -0.24167 | -0.3875 | 0.69792 | 95.0177 | 0.24167 | 86.8353
0.9358 | -2.0656 | 2.5802 | -94.9302 | 94.9938 | 0.44896 | -0.27396 | -0.40104 | 0.72813 | 94.9302 | 0.27396 | 87.0026
0.9376 | -1.8271 | 2.9729 | -95.025 | 95.0948 | 0.46458 | -0.30833 | -0.41458 | 0.76146 | 95.025 | 0.30833 | 87.17
0.9394 | -1.5979 | 3.2896 | -95.2615 | 95.3365 | 0.47708 | -0.34271 | -0.42813 | 0.79375 | 95.2615 | 0.34271 | 87.3373
0.9412 | -1.3583 | 3.449 | -95.799 | 95.8771 | 0.48125 | -0.37292 | -0.44167 | 0.82292 | 95.799 | 0.37292 | 87.5046
0.943 | -1.3083 | 3.4552 | -96.4188 | 96.4948 | 0.47708 | -0.38438 | -0.46146 | 0.8375 | 96.4188 | 0.38438 | 87.672
0.9447 | -1.2083 | 3.701 | -96.4042 | 96.4906 | 0.48125 | -0.3875 | -0.48438 | 0.85104 | 96.4042 | 0.3875 | 87.83
0.9465 | -1.1 | 3.9604 | -96.9167 | 97.0135 | 0.4875 | -0.40208 | -0.49792 | 0.86458 | 96.9167 | 0.40208 | 87.9974
0.9483 | -1.0667 | 4.1083 | -97.626 | 97.7292 | 0.48854 | -0.42604 | -0.51042 | 0.88125 | 97.626 | 0.42604 | 88.1647
0.9501 | -1.0698 | 4.2427 | -98.3937 | 98.501 | 0.4875 | -0.45 | -0.52604 | 0.89479 | 98.3937 | 0.45 | 88.3321
0.9519 | -0.98021 | 4.4354 | -99.2219 | 99.3354 | 0.48958 | -0.475 | -0.53958 | 0.91042 | 99.2219 | 0.475 | 88.4994
0.9537 | -0.95417 | 4.5344 | -100.15 | 100.2656 | 0.48958 | -0.49271 | -0.55312 | 0.92292 | 100.15 | 0.49271 | 88.6668
0.9555 | -1.025 | 4.4802 | -101.2448 | 101.3563 | 0.48542 | -0.49896 | -0.56667 | 0.92917 | 101.2448 | 0.49896 | 88.8341
0.9573 | -0.80312 | 4.5135 | -102.5635 | 102.6719 | 0.4875 | -0.52812 | -0.57188 | 0.95312 | 102.5635 | 0.52812 | 89.0015
0.9591 | -0.70833 | 4.3917 | -104.2219 | 104.3229 | 0.48438 | -0.54583 | -0.57604 | 0.96667 | 104.2219 | 0.54583 | 89.1688
0.9609 | -0.575 | 4.3771 | -106.4729 | 106.5708 | 0.48542 | -0.57188 | -0.575 | 0.98646 | 106.4729 | 0.57188 | 89.3362
0.9627 | -0.64792 | 4.2885 | -108.0802 | 108.175 | 0.48229 | -0.57917 | -0.58125 | 0.99271 | 108.0802 | 0.57917 | 89.5035
0.9645 | -0.78229 | 4.0917 | -109.6333 | 109.724 | 0.47813 | -0.58437 | -0.58542 | 0.99687 | 109.6333 | 0.58437 | 89.6709
0.9663 | -0.99792 | 3.6042 | -111.2812 | 111.3635 | 0.4625 | -0.59375 | -0.59167 | 1 | 111.2812 | 0.59375 | 89.8382
0.9681 | -1.2135 | 3.1156 | -112.9292 | 113.0042 | 0.44688 | -0.60313 | -0.59688 | 1.0031 | 112.9292 | 0.60313 | 90.0056
0.9698 | -1.2792 | 2.8969 | -114.1021 | 114.175 | 0.44063 | -0.61562 | -0.60521 | 1.0115 | 114.1021 | 0.61562 | 90.1636
0.9716 | -1.249 | 2.8469 | -114.8917 | 114.9656 | 0.44063 | -0.62813 | -0.61354 | 1.024 | 114.8917 | 0.62813 | 90.331
0.9734 | -1.2552 | 2.725 | -115.5885 | 115.6583 | 0.43854 | -0.64271 | -0.62187 | 1.0354 | 115.5885 | 0.64271 | 90.4983
0.9752 | -1.2177 | 2.6063 | -116.2177 | 116.2833 | 0.4375 | -0.65729 | -0.62813 | 1.049 | 116.2177 | 0.65729 | 90.6657
0.977 | -1.1646 | 2.6083 | -116.6469 | 116.7104 | 0.44375 | -0.66458 | -0.63333 | 1.0594 | 116.6469 | 0.66458 | 90.833
0.9788 | -1.0615 | 2.6688 | -116.776 | 116.8385 | 0.45208 | -0.67083 | -0.64062 | 1.0698 | 116.776 | 0.67083 | 91.0004
0.9806 | -0.87813 | 2.851 | -116.8823 | 116.9479 | 0.46354 | -0.67917 | -0.64479 | 1.0823 | 116.8823 | 0.67917 | 91.1677
0.9824 | -0.76458 | 2.8917 | -116.9438 | 117.0094 | 0.46563 | -0.68646 | -0.64896 | 1.0896 | 116.9438 | 0.68646 | 91.3351
0.9842 | -0.78542 | 2.7854 | -117.2802 | 117.3417 | 0.46042 | -0.68229 | -0.65417 | 1.0906 | 117.2802 | 0.68229 | 91.5024
0.986 | -0.65833 | 2.9479 | -117.3917 | 117.4531 | 0.46979 | -0.68437 | -0.65938 | 1.1 | 117.3917 | 0.68437 | 91.6698
0.9878 | -0.40833 | 3.3729 | -117.0583 | 117.126 | 0.49167 | -0.68958 | -0.66563 | 1.1187 | 117.0583 | 0.68958 | 91.8371
0.9896 | -0.20208 | 3.8365 | -116.4677 | 116.5479 | 0.51667 | -0.69063 | -0.67292 | 1.1365 | 116.4677 | 0.69063 | 92.0045
0.9914 | -0.11771 | 4.1021 | -115.7698 | 115.8594 | 0.5375 | -0.69063 | -0.68229 | 1.151 | 115.7698 | 0.69063 | 92.1718
0.9932 | -0.1125 | 4.226 | -114.9437 | 115.0385 | 0.55729 | -0.68958 | -0.69479 | 1.1646 | 114.9437 | 0.68958 | 92.3392
0.9949 | -0.032292 | 4.4969 | -113.8604 | 113.9625 | 0.58646 | -0.69375 | -0.70417 | 1.1865 | 113.8604 | 0.69375 | 92.4972
0.9967 | 0.089583 | 4.826 | -112.8427 | 112.9573 | 0.61771 | -0.7 | -0.7125 | 1.2104 | 112.8427 | 0.7 | 92.6646
0.9985 | 0.34375 | 5.1125 | -112.1854 | 112.3125 | 0.64167 | -0.71458 | -0.71458 | 1.2365 | 112.1854 | 0.71458 | 92.8319
1.0003 | 0.59792 | 5.3979 | -111.5281 | 111.6677 | 0.66667 | -0.72813 | -0.71771 | 1.2615 | 111.5281 | 0.72813 | 92.9993
1.0021 | 0.6875 | 5.6344 | -111.1333 | 111.2885 | 0.68542 | -0.73125 | -0.72083 | 1.2781 | 111.1333 | 0.73125 | 93.1666
1.0039 | 0.70312 | 5.649 | -111.2354 | 111.3927 | 0.69271 | -0.73229 | -0.72292 | 1.2854 | 111.2354 | 0.73229 | 93.334
1.0057 | 0.68646 | 5.5896 | -111.4667 | 111.6219 | 0.69688 | -0.73125 | -0.72396 | 1.2885 | 111.4667 | 0.73125 | 93.5013
1.0075 | 0.52708 | 5.3365 | -111.9042 | 112.0458 | 0.69479 | -0.72292 | -0.72917 | 1.2865 | 111.9042 | 0.72292 | 93.6687
1.0093 | 0.26875 | 4.9698 | -112.376 | 112.5031 | 0.68854 | -0.70625 | -0.73854 | 1.2812 | 112.376 | 0.70625 | 93.836
1.0111 | 0.17604 | 4.8573 | -112.9906 | 113.1146 | 0.69271 | -0.69688 | -0.74167 | 1.2802 | 112.9906 | 0.69688 | 94.0033
1.0129 | 0.039583 | 4.749 | -113.575 | 113.6969 | 0.69583 | -0.68125 | -0.74792 | 1.2771 | 113.575 | 0.68125 | 94.1707
1.0147 | 0.040625 | 4.7469 | -113.8844 | 114.0042 | 0.70417 | -0.67812 | -0.75 | 1.2812 | 113.8844 | 0.67812 | 94.338
1.0165 | -0.00625 | 4.6667 | -114.2302 | 114.3469 | 0.71042 | -0.67292 | -0.75313 | 1.2833 | 114.2302 | 0.67292 | 94.5054
1.0182 | -0.069792 | 4.6115 | -114.401 | 114.5146 | 0.71562 | -0.675 | -0.75208 | 1.2875 | 114.401 | 0.675 | 94.6634
1.02 | 0.079167 | 4.8958 | -114.2521 | 114.375 | 0.73542 | -0.68125 | -0.75313 | 1.3042 | 114.2521 | 0.68125 | 94.8308
1.0218 | 0.14271 | 5.1156 | -114.0333 | 114.1667 | 0.75 | -0.68646 | -0.75625 | 1.3198 | 114.0333 | 0.68646 | 94.9981
1.0236 | 0.045833 | 5.0427 | -114.3927 | 114.5208 | 0.75104 | -0.6875 | -0.75833 | 1.3219 | 114.3927 | 0.6875 | 95.1655
1.0254 | -0.19167 | 4.6896 | -115.0438 | 115.1552 | 0.7375 | -0.68542 | -0.75938 | 1.3156 | 115.0438 | 0.68542 | 95.3328
1.0272 | -0.43958 | 4.3312 | -115.8167 | 115.9125 | 0.72292 | -0.68333 | -0.75938 | 1.3062 | 115.8167 | 0.68333 | 95.5002
1.029 | -0.70312 | 3.9021 | -116.6083 | 116.6896 | 0.70833 | -0.67812 | -0.75833 | 1.2958 | 116.6083 | 0.67812 | 95.6675
1.0308 | -0.97188 | 3.4594 | -117.3635 | 117.4302 | 0.69271 | -0.67188 | -0.75729 | 1.2854 | 117.3635 | 0.67188 | 95.8349
1.0326 | -1.0979 | 3.224 | -117.8594 | 117.9198 | 0.68646 | -0.66979 | -0.75729 | 1.2802 | 117.8594 | 0.66979 | 96.0022
1.0344 | -1.1854 | 3.0094 | -118.1771 | 118.2313 | 0.67917 | -0.66771 | -0.75833 | 1.2729 | 118.1771 | 0.66771 | 96.1696
1.0362 | -1.2563 | 2.8396 | -118.2729 | 118.324 | 0.675 | -0.66354 | -0.75938 | 1.2688 | 118.2729 | 0.66354 | 96.3369
1.038 | -1.3073 | 2.7323 | -118.1365 | 118.1854 | 0.67812 | -0.66042 | -0.76146 | 1.2698 | 118.1365 | 0.66042 | 96.5043
1.0398 | -1.3115 | 2.6542 | -117.9208 | 117.9698 | 0.68542 | -0.65625 | -0.76146 | 1.2708 | 117.9208 | 0.65625 | 96.6716
1.0416 | -1.3833 | 2.5677 | -117.5677 | 117.6177 | 0.69583 | -0.65104 | -0.76042 | 1.2729 | 117.5677 | 0.65104 | 96.839
1.0433 | -1.5396 | 2.349 | -117.5573 | 117.6083 | 0.70104 | -0.64583 | -0.75729 | 1.2719 | 117.5573 | 0.64583 | 96.997
1.0451 | -1.625 | 2.2208 | -117.2438 | 117.2979 | 0.70833 | -0.64167 | -0.75729 | 1.2719 | 117.2438 | 0.64167 | 97.1644
1.0469 | -1.5333 | 2.325 | -116.7031 | 116.7583 | 0.72292 | -0.64167 | -0.75521 | 1.2781 | 116.7031 | 0.64167 | 97.3317
1.0487 | -1.5521 | 2.2375 | -116.2042 | 116.2635 | 0.73229 | -0.64792 | -0.74896 | 1.2812 | 116.2042 | 0.64792 | 97.4991
1.0505 | -1.3885 | 2.4448 | -115.2969 | 115.3542 | 0.75104 | -0.66042 | -0.74271 | 1.2937 | 115.2969 | 0.66042 | 97.6664
1.0523 | -1.1635 | 2.7896 | -114.274 | 114.3292 | 0.76979 | -0.67292 | -0.73854 | 1.3094 | 114.274 | 0.67292 | 97.8338
1.0541 | -1.025 | 2.9 | -113.6479 | 113.7021 | 0.77812 | -0.67917 | -0.73542 | 1.3146 | 113.6479 | 0.67917 | 98.0011
1.0559 | -1.0417 | 2.7938 | -113.3292 | 113.3812 | 0.77708 | -0.68125 | -0.73333 | 1.3135 | 113.3292 | 0.68125 | 98.1685
1.0577 | -1.2031 | 2.4198 | -113.0927 | 113.1458 | 0.76042 | -0.67812 | -0.73021 | 1.299 | 113.0927 | 0.67812 | 98.3358
1.0595 | -1.2521 | 2.376 | -112.826 | 112.8792 | 0.75833 | -0.675 | -0.72708 | 1.2937 | 112.826 | 0.675 | 98.5032
1.0613 | -1.2896 | 2.3552 | -112.5021 | 112.5552 | 0.75729 | -0.67083 | -0.725 | 1.2896 | 112.5021 | 0.67083 | 98.6705
1.0631 | -1.3698 | 2.1958 | -112.4396 | 112.4917 | 0.75417 | -0.66667 | -0.72292 | 1.2823 | 112.4396 | 0.66667 | 98.8379
1.0649 | -1.525 | 1.876 | -112.7437 | 112.7917 | 0.74479 | -0.66042 | -0.71875 | 1.2688 | 112.7437 | 0.66042 | 99.0052
1.0667 | -1.7188 | 1.6313 | -113.5469 | 113.5917 | 0.73438 | -0.64583 | -0.7125 | 1.25 | 113.5469 | 0.64583 | 99.1726
1.0684 | -1.8042 | 1.5865 | -114.3031 | 114.3458 | 0.73438 | -0.63229 | -0.70521 | 1.2365 | 114.3031 | 0.63229 | 99.3306
1.0702 | -1.8073 | 1.7302 | -114.1354 | 114.1792 | 0.74062 | -0.63021 | -0.70104 | 1.2365 | 114.1354 | 0.63021 | 99.498
1.072 | -1.7292 | 1.8042 | -113.8156 | 113.8563 | 0.74375 | -0.63542 | -0.69583 | 1.2396 | 113.8156 | 0.63542 | 99.6653
1.0738 | -1.65 | 1.8781 | -113.4958 | 113.5333 | 0.74583 | -0.64167 | -0.69167 | 1.2417 | 113.4958 | 0.64167 | 99.8327
1.0756 | -1.6521 | 1.9292 | -113.525 | 113.5635 | 0.75313 | -0.63958 | -0.68646 | 1.2437 | 113.525 | 0.63958 | 100
| Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My
Minimum | -6.8938 | -6.1312 | -237.474 | 77.6156 | 0.36146 | -0.73229 | -0.76146 | 0.50521 | 77.4896 | -1.301
Maximum | 1.6021 | 5.649 | -77.4896 | 237.5396 | 1.5469 | 1.301 | -0.26458 | 1.9969 | 237.474 | 0.73229

## Diagr_BW
Loads in %BW / %BWm
Contact forces and moments, measured telemetrically in knee implants of 8 subjects.
Averages of one individual subject.
Fx  lateral, Fy  anterior, -Fz  distal, Fres = Resultant Force
Mx around x, My around y, Mz around z, Mres = Resultant Moment
z-axis = axis of implant stem!
From Bergmann et al, PlosOne (2014) "Standardized Loads Acting in Knee Implants"
Subject K3R
Subject K3R

## Dat_AVER75
Data File: | k3r_2239_0_3_i1.akf
Curve Alignment Program: | 1.5.7
------------------------MEASURING INFOS------------------------
Diagram Title #1 | Forces and Moments at Knee Joint
Diagram Title #2 | Measured Relative to Implant
Comment #1 | Knee Joint; Gaitanalysis; Deep Knee Bend;
Comment #2 | K3R, 8 Months PO *
BodyWeight [N]: | 960
Patient(en): | K3R | [7 empty] | AVER75 loads in N (forces) and Nm (moments)
Implant Type: | Knee Joint
Varus Angle Leg [°] | 3.5 | [7 empty] | Subject | K3R
Axes Directions X, Y, Z: | Lateral, Anterior, Superior
Axes Origin: | Tibial Plateau at Stem Axis | [7 empty] | The individual peak value of Fres was adapted to the average peak of Fres from all subjects!
Transformation X Y Z (°): | 10 | 0.5 | 10.5 | [5 empty] | The AVER75 forces Fres of all subjects therefore have the same peak values.
Displacement X, Y, Z (mm): | 0 | 0 | 0 | [5 empty] | All 6 load components were then adapted by the same Multiplication Factor.
Activity Code: | KneeActivities | 1P_1_15_3
[9 empty] | Multiplication Factor = |  | Max(Fres, AVER75 of 8 Subjects) / |  |  | Max(Fres, Individual Subject)
[9 empty] | Max(Fres, AVER75 of 8 Subjects) = |  |  | 1935.87 | N
[9 empty] | Max(Fres, Individual Subject) = |  |  | 2280.38 | N
---------------------------DATA--------------------------------
Number of Data Sets = | 601
Number of analog channels = | 0
Time-Offset(sec) = | 0
Time | Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My | Time
[s] | [N] | [N] | [N] | [N] | [Nm] | [Nm] | [Nm] | [Nm] | [N] | (Nm] | [%Cycle]
0 | -25.1282 | -9.7711 | -781.3414 | 781.9357 | 5.8661 | -2.0204 | -5.2379 | 8.4213 | 781.3414 | 2.0204 | 0
0.0018 | -25.3489 | -9.873 | -777.377 | 777.9712 | 5.8915 | -1.995 | -5.2803 | 8.4553 | 777.377 | 1.995 | 0.16735
0.0036 | -25.7479 | -10.6285 | -776.4941 | 777.0883 | 5.8746 | -1.978 | -5.3058 | 8.4553 | 776.4941 | 1.978 | 0.3347
0.0054 | -26.2063 | -11.4095 | -777.8693 | 778.4891 | 5.8576 | -1.961 | -5.3143 | 8.4468 | 777.8693 | 1.961 | 0.50205
0.0072 | -27.0977 | -12.9206 | -780.3058 | 781.0104 | 5.8151 | -1.9271 | -5.3143 | 8.4128 | 780.3058 | 1.9271 | 0.66939
0.009 | -27.7513 | -13.4809 | -785.3823 | 786.1209 | 5.7812 | -1.9271 | -5.2633 | 8.3704 | 785.3823 | 1.9271 | 0.83674
0.0108 | -27.7004 | -12.9716 | -788.2177 | 788.9733 | 5.7897 | -1.944 | -5.2294 | 8.3619 | 788.2177 | 1.944 | 1.0041
0.0125 | -28.0994 | -13.1244 | -790.8324 | 791.6219 | 5.7727 | -1.9355 | -5.2039 | 8.3449 | 790.8324 | 1.9355 | 1.1621
0.0143 | -29.1521 | -14.5166 | -792.2077 | 793.099 | 5.7217 | -1.8931 | -5.2039 | 8.311 | 792.2077 | 1.8931 | 1.3295
0.0161 | -28.804 | -14.0497 | -791.1635 | 792.0294 | 5.7727 | -1.9016 | -5.2124 | 8.3534 | 791.1635 | 1.9016 | 1.4968
0.0179 | -28.7785 | -14.3978 | -787.199 | 788.031 | 5.7982 | -1.8761 | -5.2548 | 8.3704 | 787.199 | 1.8761 | 1.6642
0.0197 | -30.3151 | -16.6899 | -781.9951 | 782.861 | 5.7387 | -1.7912 | -5.3228 | 8.3364 | 781.9951 | 1.7912 | 1.8315
0.0215 | -31.181 | -18.2434 | -780.8236 | 781.7235 | 5.7133 | -1.7488 | -5.3482 | 8.3025 | 780.8236 | 1.7488 | 1.9989
0.0233 | -31.4357 | -18.3962 | -780.5519 | 781.4688 | 5.7472 | -1.7318 | -5.3567 | 8.3279 | 780.5519 | 1.7318 | 2.1662
0.0251 | -32.0299 | -18.8291 | -781.1292 | 782.08 | 5.7812 | -1.7063 | -5.3482 | 8.3449 | 781.1292 | 1.7063 | 2.3336
0.0269 | -32.743 | -19.4234 | -781.9017 | 782.895 | 5.8236 | -1.6809 | -5.3482 | 8.3704 | 781.9017 | 1.6809 | 2.5009
0.0287 | -33.176 | -19.7714 | -783.3109 | 784.3212 | 5.8746 | -1.6639 | -5.3397 | 8.4044 | 783.3109 | 1.6639 | 2.6683
0.0305 | -34.0079 | -20.5355 | -785.204 | 786.2822 | 5.9 | -1.6299 | -5.3228 | 8.4128 | 785.204 | 1.6299 | 2.8356
0.0323 | -34.8144 | -21.2231 | -788.8374 | 789.9495 | 5.9 | -1.613 | -5.2718 | 8.3789 | 788.8374 | 1.613 | 3.003
0.0341 | -35.0776 | -21.3504 | -791.8851 | 793.0141 | 5.934 | -1.6045 | -5.2294 | 8.3789 | 791.8851 | 1.6045 | 3.1703
0.0359 | -35.3407 | -21.4778 | -794.9327 | 796.0788 | 5.9679 | -1.596 | -5.1784 | 8.3789 | 794.9327 | 1.596 | 3.3377
0.0376 | -36.1557 | -22.5899 | -796.4608 | 797.6578 | 6.0019 | -1.545 | -5.1699 | 8.3874 | 796.4608 | 1.545 | 3.4957
0.0394 | -36.6396 | -23.1077 | -796.308 | 797.5389 | 6.0528 | -1.5111 | -5.1784 | 8.4213 | 796.308 | 1.5111 | 3.6631
0.0412 | -36.5717 | -22.7512 | -796.2146 | 797.4286 | 6.1547 | -1.4941 | -5.1784 | 8.4977 | 796.2146 | 1.4941 | 3.8304
0.043 | -37.0216 | -23.3709 | -796.9107 | 798.1756 | 6.2226 | -1.4601 | -5.1784 | 8.5487 | 796.9107 | 1.4601 | 3.9978
0.0448 | -37.845 | -24.483 | -798.0143 | 799.3641 | 6.2566 | -1.4262 | -5.1869 | 8.5826 | 798.0143 | 1.4262 | 4.1651
0.0466 | -38.7279 | -25.5951 | -798.8547 | 800.3064 | 6.2736 | -1.3922 | -5.1869 | 8.6081 | 798.8547 | 1.3922 | 4.3325
0.0484 | -39.1863 | -26.5204 | -799.033 | 800.5526 | 6.2905 | -1.3753 | -5.1869 | 8.6166 | 799.033 | 1.3753 | 4.4998
0.0502 | -39.441 | -27.259 | -798.685 | 800.247 | 6.316 | -1.3583 | -5.1869 | 8.6336 | 798.685 | 1.3583 | 4.6672
0.052 | -39.5344 | -27.8108 | -796.5711 | 798.1247 | 6.3839 | -1.3243 | -5.2039 | 8.676 | 796.5711 | 1.3243 | 4.8345
0.0538 | -40.3154 | -28.9483 | -793.2773 | 794.8903 | 6.4688 | -1.2225 | -5.2548 | 8.7354 | 793.2773 | 1.2225 | 5.0019
0.0556 | -41.1049 | -30.0774 | -789.975 | 791.6559 | 6.5452 | -1.1291 | -5.3058 | 8.7864 | 789.975 | 1.1291 | 5.1692
0.0574 | -41.7331 | -31.0027 | -787.216 | 788.9563 | 6.6301 | -1.0357 | -5.3482 | 8.8458 | 787.216 | 1.0357 | 5.3366
0.0592 | -41.2917 | -30.6801 | -787.3688 | 789.0666 | 6.7235 | -1.0272 | -5.3228 | 8.8797 | 787.3688 | 1.0272 | 5.5039
0.061 | -40.7484 | -30.366 | -787.5641 | 789.2279 | 6.8084 | -1.0187 | -5.2888 | 8.9307 | 787.5641 | 1.0187 | 5.6713
0.0627 | -40.6974 | -30.3745 | -788.0904 | 789.7373 | 6.9272 | -0.93382 | -5.2973 | 9.0071 | 788.0904 | 0.93382 | 5.8293
0.0645 | -41.02 | -31.198 | -786.6982 | 788.396 | 6.9187 | -0.90835 | -5.2973 | 8.9986 | 786.6982 | 0.90835 | 5.9967
0.0663 | -41.1643 | -31.7837 | -784.4994 | 786.2312 | 6.9357 | -0.87439 | -5.3058 | 9.0156 | 784.4994 | 0.87439 | 6.164
0.0681 | -40.7569 | -31.6479 | -779.3889 | 781.1037 | 7.0206 | -0.81497 | -5.3482 | 9.092 | 779.3889 | 0.81497 | 6.3313
0.0699 | -39.6363 | -29.4577 | -775.8574 | 777.4279 | 7.2328 | -0.73007 | -5.3567 | 9.2618 | 775.8574 | 0.73007 | 6.4987
0.0717 | -38.7449 | -28.04 | -769.6178 | 771.1034 | 7.3432 | -0.69612 | -5.3737 | 9.3636 | 769.6178 | 0.69612 | 6.666
0.0735 | -37.9809 | -27.1571 | -762.7585 | 764.1847 | 7.4196 | -0.67914 | -5.3992 | 9.44 | 762.7585 | 0.67914 | 6.8334
0.0753 | -37.6413 | -27.5306 | -756.0265 | 757.4782 | 7.4366 | -0.63669 | -5.4416 | 9.4825 | 756.0265 | 0.63669 | 7.0007
0.0771 | -36.3 | -26.9449 | -750.2199 | 751.5951 | 7.5045 | -0.65367 | -5.4586 | 9.5504 | 750.2199 | 0.65367 | 7.1681
0.0789 | -34.3475 | -25.7818 | -744.9565 | 746.2214 | 7.6064 | -0.69612 | -5.4586 | 9.6268 | 744.9565 | 0.69612 | 7.3354
0.0807 | -32.395 | -24.6103 | -739.6932 | 740.8477 | 7.6997 | -0.73856 | -5.4586 | 9.7117 | 739.6932 | 0.73856 | 7.5028
0.0825 | -30.1198 | -23.3709 | -737.3162 | 738.3604 | 7.8186 | -0.81497 | -5.4331 | 9.7966 | 737.3162 | 0.81497 | 7.6701
0.0843 | -29.4747 | -22.7087 | -737.8086 | 738.8528 | 7.929 | -0.79799 | -5.3822 | 9.856 | 737.8086 | 0.79799 | 7.8375
0.086 | -29.9416 | -23.8463 | -737.7237 | 738.8613 | 7.9544 | -0.74705 | -5.3567 | 9.856 | 737.7237 | 0.74705 | 7.9955
0.0878 | -30.6971 | -24.05 | -738.9377 | 740.1516 | 8.0054 | -0.70461 | -5.3228 | 9.8645 | 738.9377 | 0.70461 | 8.1629
0.0896 | -31.7328 | -24.3302 | -740.8902 | 742.1806 | 8.0563 | -0.65367 | -5.2803 | 9.856 | 740.8902 | 0.65367 | 8.3302
0.0914 | -36.0114 | -28.1503 | -746.4252 | 747.9702 | 7.9714 | -0.48389 | -5.2633 | 9.7287 | 746.4252 | 0.48389 | 8.4976
0.0932 | -40.5786 | -31.979 | -749.6171 | 751.4338 | 7.895 | -0.28015 | -5.2718 | 9.6183 | 749.6171 | 0.28015 | 8.6649
0.095 | -43.015 | -33.9485 | -748.4286 | 750.4661 | 7.8865 | -0.1613 | -5.2888 | 9.5928 | 748.4286 | 0.1613 | 8.8323
0.0968 | -43.6602 | -34.2456 | -746.1196 | 748.2249 | 7.9035 | -0.13583 | -5.2803 | 9.5844 | 746.1196 | 0.13583 | 8.9996
0.0986 | -44.577 | -35.137 | -743.4455 | 745.6527 | 7.895 | -0.093382 | -5.2803 | 9.5589 | 743.4455 | 0.093382 | 9.167
0.1004 | -45.6467 | -36.7499 | -740.3893 | 742.7578 | 7.861 | -0.033957 | -5.2803 | 9.5249 | 740.3893 | 0.033957 | 9.3343
0.1022 | -44.8741 | -36.8943 | -736.8663 | 739.2008 | 7.861 | -0.059425 | -5.2548 | 9.5164 | 736.8663 | 0.059425 | 9.5017
0.104 | -44.4242 | -38.0403 | -734.4044 | 736.7729 | 7.8101 | -0.084892 | -5.2209 | 9.457 | 734.4044 | 0.084892 | 9.669
0.1058 | -44.611 | -39.5174 | -731.5945 | 734.0733 | 7.7422 | -0.076403 | -5.2039 | 9.3891 | 731.5945 | 0.076403 | 9.8364
0.1076 | -44.0592 | -39.3986 | -725.3039 | 727.7743 | 7.7167 | -0.067914 | -5.1784 | 9.3551 | 725.3039 | 0.067914 | 10.0037
0.1094 | -42.998 | -38.643 | -721.9167 | 724.2937 | 7.6743 | -0.093382 | -5.102 | 9.2787 | 721.9167 | 0.093382 | 10.1711
0.1111 | -42.0812 | -38.1591 | -719.9812 | 722.2818 | 7.6233 | -0.11036 | -5.0256 | 9.1939 | 719.9812 | 0.11036 | 10.3291
0.1129 | -41.8605 | -37.9639 | -713.4869 | 715.796 | 7.5809 | -0.076403 | -5.0087 | 9.1429 | 713.4869 | 0.076403 | 10.4965
0.1147 | -41.3256 | -36.6735 | -704.3016 | 706.5427 | 7.5215 | -0.0084892 | -4.9662 | 9.0665 | 704.3016 | 0.0084892 | 10.6638
0.1165 | -40.9946 | -35.2219 | -695.7359 | 697.9092 | 7.4451 | 0.067914 | -4.9238 | 8.9816 | 695.7359 | -0.067914 | 10.8312
0.1183 | -40.2899 | -34.2711 | -694.2248 | 696.3132 | 7.3262 | 0.10187 | -4.8134 | 8.8203 | 694.2248 | -0.10187 | 10.9985
0.1201 | -38.5412 | -32.2676 | -690.8122 | 692.7222 | 7.2923 | 0.11036 | -4.7285 | 8.7524 | 690.8122 | -0.11036 | 11.1659
0.1219 | -36.0793 | -29.3728 | -684.4707 | 686.1516 | 7.3262 | 0.10187 | -4.6691 | 8.7439 | 684.4707 | -0.10187 | 11.3332
0.1237 | -33.2439 | -25.9092 | -675.3023 | 676.737 | 7.3856 | 0.10187 | -4.6351 | 8.7779 | 675.3023 | -0.10187 | 11.5006
0.1255 | -31.6903 | -23.2435 | -665.9387 | 667.246 | 7.3772 | 0.15281 | -4.6012 | 8.7609 | 665.9387 | -0.15281 | 11.6679
0.1273 | -30.6716 | -21.1722 | -657.5004 | 658.7058 | 7.3856 | 0.22921 | -4.6012 | 8.7779 | 657.5004 | -0.22921 | 11.8353
0.1291 | -30.4594 | -19.695 | -650.4628 | 651.6003 | 7.4536 | 0.30561 | -4.6521 | 8.8543 | 650.4628 | -0.30561 | 12.0026
0.1309 | -30.2557 | -17.8868 | -645.3183 | 646.3795 | 7.513 | 0.37353 | -4.6691 | 8.9052 | 645.3183 | -0.37353 | 12.17
0.1327 | -29.4831 | -15.773 | -641.2604 | 642.2367 | 7.53 | 0.39051 | -4.6436 | 8.9137 | 641.2604 | -0.39051 | 12.3373
0.1345 | -29.1011 | -14.3808 | -637.1347 | 638.0685 | 7.4875 | 0.40748 | -4.6181 | 8.8628 | 637.1347 | -0.40748 | 12.5046
0.1362 | -30.0434 | -14.228 | -632.2534 | 633.2126 | 7.4281 | 0.48389 | -4.6266 | 8.8288 | 632.2534 | -0.48389 | 12.6627
0.138 | -31.8601 | -14.5251 | -631.5148 | 632.542 | 7.3687 | 0.61123 | -4.6266 | 8.7949 | 631.5148 | -0.61123 | 12.83
0.1398 | -34.2116 | -15.5183 | -633.7814 | 634.9444 | 7.3007 | 0.74705 | -4.6266 | 8.7524 | 633.7814 | -0.74705 | 12.9974
0.1416 | -36.283 | -15.9683 | -636.2942 | 637.5506 | 7.2328 | 0.8659 | -4.6097 | 8.71 | 636.2942 | -0.8659 | 13.1647
0.1434 | -36.1217 | -16.0022 | -640.7086 | 641.9311 | 7.2074 | 0.84892 | -4.5672 | 8.676 | 640.7086 | -0.84892 | 13.3321
0.1452 | -37.2423 | -16.571 | -647.3557 | 648.6461 | 7.1564 | 0.88288 | -4.5248 | 8.6166 | 647.3557 | -0.88288 | 13.4994
0.147 | -38.7619 | -17.8529 | -656.3883 | 657.7805 | 7.0546 | 0.91684 | -4.4653 | 8.5232 | 656.3883 | -0.91684 | 13.6668
0.1488 | -39.857 | -18.6933 | -665.8453 | 667.3054 | 6.9866 | 0.93382 | -4.3889 | 8.4298 | 665.8453 | -0.93382 | 13.8341
0.1506 | -39.4156 | -19.1432 | -672.9678 | 674.4024 | 6.9357 | 0.89137 | -4.2956 | 8.3449 | 672.9678 | -0.89137 | 14.0015
0.1524 | -38.694 | -18.7103 | -677.5265 | 678.8933 | 6.9018 | 0.82346 | -4.2192 | 8.277 | 677.5265 | -0.82346 | 14.1688
0.1542 | -37.6922 | -18.2858 | -680.7439 | 682.0343 | 6.8678 | 0.76403 | -4.1512 | 8.2091 | 680.7439 | -0.76403 | 14.3362
0.156 | -37.3781 | -18.0566 | -682.3144 | 683.5878 | 6.8338 | 0.7131 | -4.1088 | 8.1667 | 682.3144 | -0.7131 | 14.5035
0.1578 | -37.098 | -17.7765 | -682.1871 | 683.435 | 6.8084 | 0.67914 | -4.0748 | 8.1242 | 682.1871 | -0.67914 | 14.6709
0.1595 | -36.9707 | -17.6916 | -680.8288 | 682.0767 | 6.7829 | 0.67914 | -4.0579 | 8.0902 | 680.8288 | -0.67914 | 14.8289
0.1613 | -37.7007 | -18.5914 | -679.3602 | 680.676 | 6.715 | 0.70461 | -4.0663 | 8.0478 | 679.3602 | -0.70461 | 14.9963
0.1631 | -38.7704 | -19.4149 | -677.3652 | 678.7829 | 6.6725 | 0.75554 | -4.0918 | 8.0138 | 677.3652 | -0.75554 | 15.1636
0.1649 | -39.8061 | -20.128 | -675.0986 | 676.6181 | 6.6301 | 0.80648 | -4.1088 | 7.9884 | 675.0986 | -0.80648 | 15.331
0.1667 | -39.076 | -19.0753 | -669.606 | 671.0492 | 6.6725 | 0.79799 | -4.1173 | 8.0308 | 669.606 | -0.79799 | 15.4983
0.1685 | -37.1319 | -17.1907 | -662.6958 | 663.9861 | 6.7405 | 0.73007 | -4.1258 | 8.1072 | 662.6958 | -0.73007 | 15.6657
0.1703 | -34.339 | -14.7628 | -654.597 | 655.6752 | 6.8423 | 0.61971 | -4.1428 | 8.2176 | 654.597 | -0.61971 | 15.833
0.1721 | -30.6547 | -11.7576 | -649.7327 | 650.6665 | 6.9272 | 0.46691 | -4.1003 | 8.311 | 649.7327 | -0.46691 | 16.0004
0.1739 | -27.5391 | -9.075 | -645.921 | 646.753 | 6.9951 | 0.34806 | -4.0494 | 8.3874 | 645.921 | -0.34806 | 16.1677
0.1757 | -28.4305 | -8.8118 | -645.9974 | 646.9397 | 6.9612 | 0.38202 | -3.9899 | 8.3534 | 645.9974 | -0.38202 | 16.3351
0.1775 | -29.8567 | -8.8373 | -647.152 | 648.2131 | 6.9102 | 0.44993 | -3.939 | 8.3195 | 647.152 | -0.44993 | 16.5024
0.1793 | -32.4374 | -10.5267 | -652.2285 | 653.451 | 6.7914 | 0.5518 | -3.8796 | 8.2006 | 652.2285 | -0.5518 | 16.6698
0.1811 | -35.2219 | -12.8442 | -660.4801 | 661.9402 | 6.6556 | 0.67065 | -3.7947 | 8.0818 | 660.4801 | -0.67065 | 16.8371
0.1829 | -36.5717 | -14.3383 | -669.8947 | 671.4821 | 6.6046 | 0.74705 | -3.6928 | 8.0223 | 669.8947 | -0.74705 | 17.0045
0.1846 | -37.2168 | -15.3401 | -679.5215 | 681.1684 | 6.5877 | 0.80648 | -3.5909 | 7.9969 | 679.5215 | -0.80648 | 17.1625
0.1864 | -36.8009 | -15.3825 | -685.6167 | 687.1788 | 6.5877 | 0.83195 | -3.4976 | 7.9374 | 685.6167 | -0.83195 | 17.3299
0.1882 | -35.9095 | -14.9156 | -687.9343 | 689.318 | 6.5961 | 0.84892 | -3.4212 | 7.878 | 687.9343 | -0.84892 | 17.4972
0.19 | -36.1642 | -15.5268 | -687.8324 | 689.1483 | 6.5537 | 0.94231 | -3.3957 | 7.8016 | 687.8324 | -0.94231 | 17.6646
0.1918 | -36.6396 | -16.4012 | -686.0157 | 687.3146 | 6.4943 | 1.0527 | -3.3957 | 7.7167 | 686.0157 | -1.0527 | 17.8319
0.1936 | -36.5377 | -16.4776 | -681.7287 | 683.0105 | 6.4518 | 1.1715 | -3.4127 | 7.6658 | 681.7287 | -1.1715 | 17.9993
0.1954 | -35.6039 | -16.0447 | -678.1462 | 679.3941 | 6.4348 | 1.2394 | -3.4042 | 7.6403 | 678.1462 | -1.2394 | 18.1666
0.1972 | -34.6361 | -16.0786 | -676.8558 | 678.0443 | 6.4433 | 1.2819 | -3.3787 | 7.6403 | 676.8558 | -1.2819 | 18.334
0.199 | -32.1742 | -15.2212 | -676.3804 | 677.484 | 6.4858 | 1.2649 | -3.3278 | 7.6488 | 676.3804 | -1.2649 | 18.5013
0.2008 | -29.9331 | -14.4062 | -675.939 | 677.0171 | 6.5452 | 1.2649 | -3.2768 | 7.6828 | 675.939 | -1.2649 | 18.6687
0.2026 | -27.9041 | -13.6931 | -676.4823 | 677.5435 | 6.6131 | 1.2904 | -3.2089 | 7.7082 | 676.4823 | -1.2904 | 18.836
0.2044 | -28.1079 | -14.4996 | -679.2922 | 680.3449 | 6.6301 | 1.4601 | -3.1495 | 7.6658 | 679.2922 | -1.4601 | 19.0033
0.2062 | -29.0757 | -15.7985 | -681.5419 | 682.5946 | 6.6301 | 1.6554 | -3.0901 | 7.6403 | 681.5419 | -1.6554 | 19.1707
0.208 | -30.0944 | -17.4114 | -681.1429 | 682.238 | 6.6641 | 1.8337 | -3.0561 | 7.6573 | 681.1429 | -1.8337 | 19.338
0.2097 | -31.3932 | -19.6017 | -680.8628 | 682.0173 | 6.698 | 2.0459 | -3.0137 | 7.6828 | 680.8628 | -2.0459 | 19.4961
0.2115 | -33.558 | -21.5712 | -681.8645 | 683.1464 | 6.7235 | 2.2581 | -2.9712 | 7.7337 | 681.8645 | -2.2581 | 19.6634
0.2133 | -36.7669 | -24.848 | -686.627 | 688.155 | 6.732 | 2.5383 | -2.9543 | 7.8186 | 686.627 | -2.5383 | 19.8308
0.2151 | -40.7823 | -28.3116 | -692.111 | 693.9447 | 6.7574 | 2.8439 | -2.9203 | 7.9205 | 692.111 | -2.8439 | 19.9981
0.2169 | -44.7723 | -32.0639 | -700.5323 | 702.7565 | 6.715 | 3.0986 | -2.9118 | 7.9799 | 700.5323 | -3.0986 | 20.1655
0.2187 | -49.314 | -36.4868 | -712.1541 | 714.8622 | 6.6556 | 3.3617 | -2.9033 | 8.0308 | 712.1541 | -3.3617 | 20.3328
0.2205 | -53.686 | -40.7653 | -727.0527 | 730.2786 | 6.6471 | 3.6419 | -2.8863 | 8.1582 | 727.0527 | -3.6419 | 20.5002
0.2223 | -55.4263 | -42.6669 | -741.6372 | 745.0499 | 6.715 | 3.7947 | -2.8439 | 8.277 | 741.6372 | -3.7947 | 20.6675
0.2241 | -54.7896 | -42.1066 | -750.5849 | 753.8787 | 6.8848 | 3.8626 | -2.8269 | 8.4383 | 750.5849 | -3.8626 | 20.8349
0.2259 | -55.5791 | -43.5838 | -760.8144 | 764.2101 | 6.9527 | 3.9475 | -2.8184 | 8.5317 | 760.8144 | -3.9475 | 21.0022
0.2277 | -56.1818 | -44.8147 | -772.2664 | 775.7385 | 7.0291 | 4.0239 | -2.8099 | 8.6251 | 772.2664 | -4.0239 | 21.1696
0.2295 | -55.4857 | -44.5685 | -784.2872 | 787.6405 | 7.1479 | 4.0494 | -2.8015 | 8.7354 | 784.2872 | -4.0494 | 21.3369
0.2313 | -54.4076 | -44.0846 | -793.3198 | 796.5287 | 7.2668 | 4.0748 | -2.793 | 8.8373 | 793.3198 | -4.0748 | 21.5043
0.2331 | -54.5179 | -45.0269 | -804.4322 | 807.6666 | 7.3177 | 4.1343 | -2.776 | 8.8967 | 804.4322 | -4.1343 | 21.6716
0.2348 | -55.0358 | -46.6908 | -819.6279 | 822.9387 | 7.3517 | 4.1852 | -2.759 | 8.9307 | 819.6279 | -4.1852 | 21.8297
0.2366 | -55.3668 | -48.4396 | -837.7355 | 841.0972 | 7.4111 | 4.2107 | -2.7505 | 8.9901 | 837.7355 | -4.2107 | 21.997
0.2384 | -54.849 | -49.4668 | -855.3507 | 858.6615 | 7.53 | 4.2192 | -2.7335 | 9.0835 | 855.3507 | -4.2192 | 22.1644
0.2402 | -53.4313 | -49.8234 | -870.037 | 873.2205 | 7.6488 | 4.2361 | -2.7166 | 9.1854 | 870.037 | -4.2361 | 22.3317
0.242 | -51.8268 | -49.9677 | -883.917 | 886.9646 | 7.7677 | 4.2701 | -2.7081 | 9.2957 | 883.917 | -4.2701 | 22.4991
0.2438 | -49.7045 | -48.8726 | -893.1533 | 895.9972 | 7.895 | 4.3125 | -2.7081 | 9.4315 | 893.1533 | -4.3125 | 22.6664
0.2456 | -46.9795 | -47.6501 | -900.1654 | 902.7376 | 8.0902 | 4.3125 | -2.6911 | 9.5844 | 900.1654 | -4.3125 | 22.8338
0.2474 | -44.4157 | -45.4599 | -904.9109 | 907.2114 | 8.311 | 4.321 | -2.6996 | 9.7881 | 904.9109 | -4.321 | 23.0011
0.2492 | -42.1321 | -43.4055 | -910.8618 | 912.9332 | 8.5062 | 4.355 | -2.6996 | 9.9833 | 910.8618 | -4.355 | 23.1685
0.251 | -40.256 | -42.1321 | -920.6924 | 922.6024 | 8.6251 | 4.4144 | -2.6911 | 10.1107 | 920.6924 | -4.4144 | 23.3358
0.2528 | -38.1082 | -41.0285 | -931.0832 | 932.832 | 8.7524 | 4.4229 | -2.6826 | 10.221 | 931.0832 | -4.4229 | 23.5032
0.2546 | -35.9265 | -40.3664 | -939.8526 | 941.4655 | 8.8882 | 4.4229 | -2.6911 | 10.3484 | 939.8526 | -4.4229 | 23.6705
0.2564 | -34.4239 | -39.9843 | -949.3436 | 950.8886 | 8.9901 | 4.4314 | -2.6996 | 10.4418 | 949.3436 | -4.4314 | 23.8379
0.2581 | -32.1657 | -38.6176 | -959.4458 | 960.8889 | 9.0835 | 4.4059 | -2.6996 | 10.5097 | 959.4458 | -4.4059 | 23.9959
0.2599 | -30.5613 | -37.9894 | -968.8349 | 970.2016 | 9.1684 | 4.4314 | -2.6996 | 10.5861 | 968.8349 | -4.4314 | 24.1633
0.2617 | -30.3745 | -38.9062 | -978.0457 | 979.4294 | 9.2193 | 4.5587 | -2.7335 | 10.688 | 978.0457 | -4.5587 | 24.3306
0.2635 | -30.5867 | -40.1626 | -989.1496 | 990.5928 | 9.2787 | 4.7285 | -2.759 | 10.8068 | 989.1496 | -4.7285 | 24.498
0.2653 | -30.8075 | -41.419 | -1000.3384 | 1001.841 | 9.3212 | 4.8983 | -2.7845 | 10.9172 | 1000.3384 | -4.8983 | 24.6653
0.2671 | -31.2065 | -42.8877 | -1013.8958 | 1015.4663 | 9.3127 | 5.0681 | -2.8015 | 11.0021 | 1013.8958 | -5.0681 | 24.8327
0.2689 | -31.0621 | -43.3036 | -1026.5872 | 1028.1747 | 9.3806 | 5.1954 | -2.8354 | 11.1294 | 1026.5872 | -5.1954 | 25
0.2707 | -31.0452 | -43.2442 | -1040.1869 | 1041.7659 | 9.44 | 5.3397 | -2.8524 | 11.2652 | 1040.1869 | -5.3397 | 25.1673
0.2725 | -31.2404 | -42.3528 | -1055.3827 | 1056.9277 | 9.4825 | 5.518 | -2.8609 | 11.4011 | 1055.3827 | -5.518 | 25.3347
0.2743 | -32.1573 | -42.5311 | -1073.4733 | 1075.0013 | 9.4995 | 5.6878 | -2.8694 | 11.5114 | 1073.4733 | -5.6878 | 25.502
0.2761 | -32.8364 | -43.2867 | -1092.7099 | 1094.221 | 9.5164 | 5.8151 | -2.8948 | 11.5878 | 1092.7099 | -5.8151 | 25.6694
0.2779 | -32.4629 | -43.2867 | -1111.4117 | 1112.8549 | 9.5589 | 5.9085 | -2.9033 | 11.6727 | 1111.4117 | -5.9085 | 25.8367
0.2797 | -31.5121 | -43.2782 | -1125.8943 | 1127.346 | 9.6353 | 5.9594 | -2.9288 | 11.7576 | 1125.8943 | -5.9594 | 26.0041
0.2815 | -30.366 | -43.5583 | -1141.574 | 1143.0171 | 9.7457 | 6.0274 | -2.9797 | 11.9019 | 1141.574 | -6.0274 | 26.1714
0.2832 | -30.0859 | -44.6534 | -1161.0228 | 1162.483 | 9.839 | 6.1207 | -3.0137 | 12.0377 | 1161.0228 | -6.1207 | 26.3295
0.285 | -29.237 | -45.2052 | -1178.1881 | 1179.6313 | 9.9409 | 6.1887 | -3.0391 | 12.1481 | 1178.1881 | -6.1887 | 26.4968
0.2868 | -27.3693 | -44.5261 | -1190.6503 | 1192.0086 | 10.0258 | 6.2141 | -3.0391 | 12.233 | 1190.6503 | -6.2141 | 26.6642
0.2886 | -25.8243 | -43.6092 | -1203.087 | 1204.3604 | 10.1107 | 6.282 | -3.0307 | 12.3434 | 1203.087 | -6.282 | 26.8315
0.2904 | -24.2792 | -42.6924 | -1215.5238 | 1216.7208 | 10.1871 | 6.35 | -3.0222 | 12.4452 | 1215.5238 | -6.35 | 26.9989
0.2922 | -22.6238 | -42.251 | -1224.1149 | 1225.2694 | 10.238 | 6.4264 | -3.0307 | 12.5301 | 1224.1149 | -6.4264 | 27.1662
0.294 | -20.493 | -40.5616 | -1227.9181 | 1228.9792 | 10.2975 | 6.5367 | -3.0476 | 12.649 | 1227.9181 | -6.5367 | 27.3336
0.2958 | -18.515 | -36.8433 | -1228.2237 | 1229.1745 | 10.3484 | 6.6556 | -3.0476 | 12.7678 | 1228.2237 | -6.6556 | 27.5009
0.2976 | -17.4624 | -32.4374 | -1223.4103 | 1224.2762 | 10.289 | 6.732 | -3.0137 | 12.7593 | 1223.4103 | -6.732 | 27.6683
0.2994 | -17.0719 | -29.3898 | -1215.2946 | 1216.1435 | 10.1277 | 6.7999 | -2.9797 | 12.6744 | 1215.2946 | -6.7999 | 27.8356
0.3012 | -17.0124 | -26.9279 | -1207.9174 | 1208.7494 | 9.9409 | 6.8848 | -2.9543 | 12.598 | 1207.9174 | -6.8848 | 28.003
0.303 | -17.9038 | -27.1316 | -1210.8292 | 1211.6697 | 9.8305 | 7.0121 | -2.9712 | 12.6065 | 1210.8292 | -7.0121 | 28.1703
0.3048 | -18.0821 | -26.2827 | -1214.0382 | 1214.8786 | 9.7796 | 7.1225 | -2.9712 | 12.6575 | 1214.0382 | -7.1225 | 28.3377
0.3066 | -18.7273 | -25.9431 | -1217.6885 | 1218.5375 | 9.7626 | 7.2243 | -2.9967 | 12.7254 | 1217.6885 | -7.2243 | 28.505
0.3083 | -19.5847 | -25.9092 | -1221.5936 | 1222.451 | 9.7457 | 7.3262 | -3.0222 | 12.7848 | 1221.5936 | -7.3262 | 28.6631
0.3101 | -20.0855 | -26.1978 | -1225.1081 | 1225.9995 | 9.7287 | 7.4196 | -3.0561 | 12.8442 | 1225.1081 | -7.4196 | 28.8304
0.3119 | -21.4438 | -28.0485 | -1228.4868 | 1229.4376 | 9.6862 | 7.53 | -3.0986 | 12.9036 | 1228.4868 | -7.53 | 28.9978
0.3137 | -23.4643 | -30.3745 | -1232.0948 | 1233.0965 | 9.6013 | 7.6573 | -3.1325 | 12.9631 | 1232.0948 | -7.6573 | 29.1651
0.3155 | -24.7631 | -31.563 | -1235.6518 | 1236.6959 | 9.5164 | 7.8271 | -3.141 | 13.014 | 1235.6518 | -7.8271 | 29.3325
0.3173 | -24.9074 | -30.8584 | -1235.8046 | 1236.8403 | 9.491 | 7.9629 | -3.1495 | 13.0734 | 1235.8046 | -7.9629 | 29.4998
0.3191 | -24.9499 | -29.3898 | -1236.5346 | 1237.5788 | 9.4995 | 8.0648 | -3.175 | 13.1413 | 1236.5346 | -8.0648 | 29.6672
0.3209 | -24.9838 | -27.8872 | -1237.3326 | 1238.3853 | 9.5164 | 8.1667 | -3.192 | 13.2008 | 1237.3326 | -8.1667 | 29.8345
0.3227 | -23.8802 | -24.9329 | -1235.3377 | 1236.3734 | 9.5759 | 8.2091 | -3.2259 | 13.2602 | 1235.3377 | -8.2091 | 30.0019
0.3245 | -22.87 | -21.5033 | -1232.6805 | 1233.7077 | 9.6268 | 8.1836 | -3.2514 | 13.2687 | 1232.6805 | -8.1836 | 30.1692
0.3263 | -22.5899 | -19.5083 | -1231.6364 | 1232.672 | 9.6098 | 8.2006 | -3.2938 | 13.2857 | 1231.6364 | -8.2006 | 30.3366
0.3281 | -22.488 | -18.6848 | -1223.266 | 1224.3101 | 9.5164 | 8.0987 | -3.3363 | 13.1583 | 1223.266 | -8.0987 | 30.5039
0.3299 | -22.5305 | -18.6339 | -1221.1437 | 1222.1793 | 9.4655 | 8.0393 | -3.3872 | 13.0989 | 1221.1437 | -8.0393 | 30.6713
0.3316 | -23.2266 | -19.0668 | -1224.1828 | 1225.2779 | 9.44 | 8.0563 | -3.4381 | 13.0989 | 1224.1828 | -8.0563 | 30.8293
0.3334 | -23.702 | -19.1263 | -1229.9894 | 1231.144 | 9.457 | 8.0818 | -3.4891 | 13.1329 | 1229.9894 | -8.0818 | 30.9967
0.3352 | -24.0925 | -19.695 | -1237.9948 | 1239.2172 | 9.4655 | 8.1157 | -3.5315 | 13.1583 | 1237.9948 | -8.1157 | 31.164
0.337 | -24.5254 | -21.1212 | -1247.8763 | 1249.1582 | 9.4825 | 8.1412 | -3.5909 | 13.1838 | 1247.8763 | -8.1412 | 31.3313
0.3388 | -24.7716 | -22.87 | -1259.2179 | 1260.5252 | 9.5249 | 8.1497 | -3.6504 | 13.2178 | 1259.2179 | -8.1497 | 31.4987
0.3406 | -24.3896 | -23.0738 | -1273.6411 | 1274.8806 | 9.5844 | 8.2346 | -3.6504 | 13.3111 | 1273.6411 | -8.2346 | 31.666
0.3424 | -22.887 | -22.5135 | -1287.6908 | 1288.7944 | 9.6947 | 8.2685 | -3.6504 | 13.4215 | 1287.6908 | -8.2685 | 31.8334
0.3442 | -20.6119 | -21.3589 | -1302.4027 | 1303.345 | 9.873 | 8.2431 | -3.6589 | 13.5658 | 1302.4027 | -8.2431 | 32.0007
0.346 | -18.3368 | -20.2129 | -1317.1145 | 1317.8956 | 10.0598 | 8.2176 | -3.6674 | 13.7101 | 1317.1145 | -8.2176 | 32.1681
0.3478 | -18.5999 | -20.2468 | -1335.0863 | 1335.8079 | 10.1956 | 8.2261 | -3.6843 | 13.846 | 1335.0863 | -8.2261 | 32.3354
0.3496 | -18.6084 | -19.7799 | -1350.3414 | 1350.9781 | 10.3059 | 8.2091 | -3.7013 | 13.9478 | 1350.3414 | -8.2091 | 32.5028
0.3514 | -18.3792 | -19.6696 | -1361.4114 | 1361.9887 | 10.3654 | 8.1921 | -3.7268 | 14.0073 | 1361.4114 | -8.1921 | 32.6701
0.3532 | -18.2858 | -19.8648 | -1370.1384 | 1370.6732 | 10.3823 | 8.1921 | -3.7522 | 14.0497 | 1370.1384 | -8.1921 | 32.8375
0.355 | -18.7018 | -20.6204 | -1377.2014 | 1377.7192 | 10.3823 | 8.2176 | -3.7777 | 14.0667 | 1377.2014 | -8.2176 | 33.0048
0.3567 | -18.133 | -20.9005 | -1382.8128 | 1383.2967 | 10.4078 | 8.2006 | -3.8117 | 14.0837 | 1382.8128 | -8.2006 | 33.1629
0.3585 | -16.5965 | -20.6119 | -1386.4886 | 1386.9301 | 10.4503 | 8.1497 | -3.8286 | 14.1006 | 1386.4886 | -8.1497 | 33.3302
0.3603 | -15.2552 | -20.2978 | -1389.2052 | 1389.6127 | 10.5012 | 8.0902 | -3.8202 | 14.0921 | 1389.2052 | -8.0902 | 33.4976
0.3621 | -15.1873 | -19.746 | -1393.5772 | 1393.9762 | 10.5436 | 8.0648 | -3.7862 | 14.0921 | 1393.5772 | -8.0648 | 33.6649
0.3639 | -14.3468 | -18.8971 | -1395.5721 | 1395.9626 | 10.5776 | 8.0223 | -3.7862 | 14.0837 | 1395.5721 | -8.0223 | 33.8323
0.3657 | -13.4724 | -17.9887 | -1397.7793 | 1398.1529 | 10.6116 | 7.9629 | -3.7862 | 14.0667 | 1397.7793 | -7.9629 | 33.9996
0.3675 | -13.5488 | -17.1483 | -1405.0971 | 1405.4706 | 10.62 | 7.929 | -3.7777 | 14.0412 | 1405.0971 | -7.929 | 34.167
0.3693 | -13.6083 | -15.9683 | -1413.0005 | 1413.3656 | 10.6795 | 7.8695 | -3.8032 | 14.0412 | 1413.0005 | -7.8695 | 34.3343
0.3711 | -13.5488 | -14.2534 | -1418.6034 | 1418.96 | 10.7559 | 7.8101 | -3.8371 | 14.0582 | 1418.6034 | -7.8101 | 34.5017
0.3729 | -13.1413 | -12.1821 | -1418.4252 | 1418.7817 | 10.8408 | 7.6997 | -3.9051 | 14.0667 | 1418.4252 | -7.6997 | 34.669
0.3747 | -12.1057 | -10.2126 | -1416.5236 | 1416.8801 | 10.9087 | 7.5639 | -3.956 | 14.0412 | 1416.5236 | -7.5639 | 34.8364
0.3765 | -11.9019 | -10.221 | -1414.9446 | 1415.2926 | 10.8408 | 7.496 | -3.9645 | 13.9563 | 1414.9446 | -7.496 | 35.0037
0.3783 | -12.1396 | -10.9087 | -1416.184 | 1416.5236 | 10.7474 | 7.4196 | -3.973 | 13.8375 | 1416.184 | -7.4196 | 35.1711
0.3801 | -12.6405 | -11.7321 | -1420.8871 | 1421.2181 | 10.6795 | 7.3432 | -3.9645 | 13.7356 | 1420.8871 | -7.3432 | 35.3384
0.3818 | -12.9546 | -12.0208 | -1423.5187 | 1423.8498 | 10.637 | 7.2753 | -3.973 | 13.6592 | 1423.5187 | -7.2753 | 35.4965
0.3836 | -13.1244 | -12.666 | -1421.9822 | 1422.3132 | 10.5946 | 7.2413 | -4.0069 | 13.6252 | 1421.9822 | -7.2413 | 35.6638
0.3854 | -13.6762 | -14.2195 | -1422.7717 | 1423.0943 | 10.5352 | 7.2243 | -4.0409 | 13.5828 | 1422.7717 | -7.2243 | 35.8312
0.3872 | -15.0344 | -16.4691 | -1430.0639 | 1430.4035 | 10.4587 | 7.2498 | -4.0579 | 13.5573 | 1430.0639 | -7.2498 | 35.9985
0.389 | -16.4776 | -18.4132 | -1438.6975 | 1439.054 | 10.4333 | 7.2923 | -4.1003 | 13.5828 | 1438.6975 | -7.2923 | 36.1659
0.3908 | -18.1076 | -20.8496 | -1447.0339 | 1447.4414 | 10.3908 | 7.3687 | -4.1428 | 13.5998 | 1447.0339 | -7.3687 | 36.3332
0.3926 | -19.8139 | -24.067 | -1450.1325 | 1450.6249 | 10.3399 | 7.5045 | -4.2276 | 13.6762 | 1450.1325 | -7.5045 | 36.5006
0.3944 | -21.741 | -27.6664 | -1454.6827 | 1455.277 | 10.272 | 7.6318 | -4.304 | 13.7271 | 1454.6827 | -7.6318 | 36.6679
0.3962 | -24.3302 | -31.7667 | -1463.2399 | 1464.0124 | 10.1616 | 7.7931 | -4.321 | 13.778 | 1463.2399 | -7.7931 | 36.8353
0.398 | -27.0128 | -35.6378 | -1472.0942 | 1473.0704 | 10.0598 | 7.9799 | -4.3295 | 13.829 | 1472.0942 | -7.9799 | 37.0026
0.3998 | -28.7531 | -38.1167 | -1480.541 | 1481.6361 | 10.0173 | 8.1242 | -4.3295 | 13.9054 | 1480.541 | -8.1242 | 37.17
0.4016 | -29.7039 | -39.4325 | -1492.6127 | 1493.7587 | 10.0598 | 8.2176 | -4.3465 | 13.9988 | 1492.6127 | -8.2176 | 37.3373
0.4034 | -30.3321 | -41.1559 | -1507.0868 | 1508.2414 | 10.0937 | 8.2685 | -4.3889 | 14.0582 | 1507.0868 | -8.2685 | 37.5046
0.4051 | -31.3593 | -42.7518 | -1523.938 | 1525.1095 | 10.1277 | 8.26 | -4.4229 | 14.0921 | 1523.938 | -8.26 | 37.6627
0.4069 | -32.6581 | -44.4327 | -1541.3494 | 1542.5549 | 10.1616 | 8.2431 | -4.4653 | 14.1176 | 1541.3494 | -8.2431 | 37.83
0.4087 | -33.4221 | -45.6721 | -1556.4688 | 1557.7082 | 10.221 | 8.2346 | -4.5078 | 14.1601 | 1556.4688 | -8.2346 | 37.9974
0.4105 | -32.9298 | -45.2646 | -1568.9734 | 1570.1959 | 10.3484 | 8.2091 | -4.5502 | 14.2534 | 1568.9734 | -8.2091 | 38.1647
0.4123 | -31.2574 | -43.0914 | -1583.7022 | 1584.8483 | 10.5776 | 8.1667 | -4.5587 | 14.3808 | 1583.7022 | -8.1667 | 38.3321
0.4141 | -29.4322 | -41.5294 | -1601.3514 | 1602.404 | 10.7898 | 8.1327 | -4.5502 | 14.4911 | 1601.3514 | -8.1327 | 38.4994
0.4159 | -28.4814 | -41.7331 | -1621.7171 | 1622.7358 | 10.9341 | 8.1242 | -4.5333 | 14.5845 | 1621.7171 | -8.1242 | 38.6668
0.4177 | -29.5086 | -44.1271 | -1642.6006 | 1643.6533 | 10.9766 | 8.1667 | -4.5333 | 14.6439 | 1642.6006 | -8.1667 | 38.8341
0.4195 | -28.2437 | -43.8809 | -1662.0834 | 1663.1106 | 11.1634 | 8.1412 | -4.5587 | 14.7713 | 1662.0834 | -8.1412 | 39.0015
0.4213 | -27.1741 | -43.6687 | -1682.3982 | 1683.3914 | 11.3501 | 8.1242 | -4.5927 | 14.9156 | 1682.3982 | -8.1242 | 39.1688
0.4231 | -25.7818 | -42.9556 | -1700.3105 | 1701.2698 | 11.5539 | 8.0902 | -4.6436 | 15.0599 | 1700.3105 | -8.0902 | 39.3362
0.4249 | -23.6935 | -41.1474 | -1715.2176 | 1716.0665 | 11.7916 | 8.0648 | -4.6946 | 15.2297 | 1715.2176 | -8.0648 | 39.5035
0.4267 | -21.673 | -39.9079 | -1726.6781 | 1727.4761 | 11.9359 | 7.9544 | -4.7115 | 15.2976 | 1726.6781 | -7.9544 | 39.6709
0.4285 | -21.1892 | -40.0947 | -1738.5375 | 1739.3101 | 12.0293 | 7.9374 | -4.7455 | 15.374 | 1738.5375 | -7.9374 | 39.8382
0.4302 | -20.145 | -38.8128 | -1745.5581 | 1746.2543 | 12.1481 | 7.9374 | -4.754 | 15.4759 | 1745.5581 | -7.9374 | 39.9963
0.432 | -18.0312 | -36.699 | -1747.0692 | 1747.6635 | 12.2839 | 7.912 | -4.7625 | 15.5608 | 1747.0692 | -7.912 | 40.1636
0.4338 | -16.4522 | -35.1455 | -1746.9079 | 1747.4513 | 12.3773 | 7.912 | -4.7794 | 15.6372 | 1746.9079 | -7.912 | 40.331
0.4356 | -14.9071 | -33.159 | -1743.4868 | 1743.9792 | 12.4707 | 7.912 | -4.7879 | 15.7221 | 1743.4868 | -7.912 | 40.4983
0.4374 | -14.0667 | -31.7837 | -1737.9433 | 1738.4187 | 12.4962 | 7.8865 | -4.8219 | 15.7391 | 1737.9433 | -7.8865 | 40.6657
0.4392 | -12.9631 | -30.9857 | -1734.607 | 1735.0739 | 12.4962 | 7.8865 | -4.8219 | 15.7391 | 1734.607 | -7.8865 | 40.833
0.441 | -11.6812 | -30.383 | -1732.7903 | 1733.2403 | 12.4877 | 7.9459 | -4.8219 | 15.756 | 1732.7903 | -7.9459 | 41.0004
0.4428 | -10.0258 | -28.6852 | -1727.7477 | 1728.1722 | 12.5386 | 8.0054 | -4.8474 | 15.8324 | 1727.7477 | -8.0054 | 41.1677
0.4446 | -8.5147 | -27.0977 | -1723.7068 | 1724.1058 | 12.6065 | 8.0563 | -4.8643 | 15.9004 | 1723.7068 | -8.0563 | 41.3351
0.4464 | -8.8967 | -27.5646 | -1724.4284 | 1724.8444 | 12.5811 | 8.1921 | -4.8728 | 15.9428 | 1724.4284 | -8.1921 | 41.5024
0.4482 | -11.0785 | -30.2387 | -1728.9107 | 1729.4116 | 12.4622 | 8.3874 | -4.8728 | 15.9428 | 1728.9107 | -8.3874 | 41.6698
0.45 | -12.3603 | -31.6394 | -1736.0757 | 1736.6105 | 12.4113 | 8.5232 | -4.8474 | 15.9513 | 1736.0757 | -8.5232 | 41.8371
0.4518 | -14.3214 | -33.7023 | -1745.7534 | 1746.3646 | 12.3179 | 8.6845 | -4.8219 | 15.9513 | 1745.7534 | -8.6845 | 42.0045
0.4536 | -15.2976 | -33.8636 | -1754.9048 | 1755.516 | 12.2754 | 8.8118 | -4.7964 | 15.9683 | 1754.9048 | -8.8118 | 42.1718
0.4553 | -15.6542 | -33.4221 | -1762.3669 | 1762.9611 | 12.267 | 8.8797 | -4.7794 | 16.0107 | 1762.3669 | -8.8797 | 42.3299
0.4571 | -16.0701 | -33.1505 | -1767.1718 | 1767.7405 | 12.2585 | 8.9562 | -4.7794 | 16.0447 | 1767.1718 | -8.9562 | 42.4972
0.4589 | -15.7475 | -31.9875 | -1768.6234 | 1769.1328 | 12.2924 | 8.9986 | -4.7794 | 16.0871 | 1768.6234 | -8.9986 | 42.6646
0.4607 | -15.1957 | -30.7565 | -1768.547 | 1769.0054 | 12.3264 | 9.041 | -4.771 | 16.1296 | 1768.547 | -9.041 | 42.8319
0.4625 | -15.1278 | -30.3236 | -1767.9698 | 1768.4027 | 12.3349 | 9.1005 | -4.771 | 16.155 | 1767.9698 | -9.1005 | 42.9993
0.4643 | -15.3401 | -30.1029 | -1769.0309 | 1769.4469 | 12.3264 | 9.1599 | -4.7794 | 16.189 | 1769.0309 | -9.1599 | 43.1666
0.4661 | -16.2993 | -30.5613 | -1774.498 | 1774.9309 | 12.2924 | 9.2363 | -4.7879 | 16.206 | 1774.498 | -9.2363 | 43.334
0.4679 | -16.1296 | -29.8906 | -1782.3336 | 1782.758 | 12.2754 | 9.2448 | -4.8304 | 16.2145 | 1782.3336 | -9.2448 | 43.5013
0.4697 | -16.1041 | -28.7191 | -1788.1572 | 1788.5731 | 12.2924 | 9.2533 | -4.8643 | 16.2314 | 1788.1572 | -9.2533 | 43.6687
0.4715 | -15.8155 | -28.1673 | -1790.7464 | 1791.1963 | 12.267 | 9.2193 | -4.8558 | 16.206 | 1790.7464 | -9.2193 | 43.836
0.4733 | -16.0362 | -27.8957 | -1791.4086 | 1791.867 | 12.2245 | 9.2193 | -4.8558 | 16.1805 | 1791.4086 | -9.2193 | 44.0033
0.4751 | -18.0481 | -29.619 | -1791.9773 | 1792.4697 | 12.1057 | 9.3042 | -4.8474 | 16.1296 | 1791.9773 | -9.3042 | 44.1707
0.4769 | -19.5932 | -30.3406 | -1788.9382 | 1789.439 | 12.0377 | 9.3806 | -4.8813 | 16.1126 | 1788.9382 | -9.3806 | 44.338
0.4786 | -19.4998 | -29.8567 | -1780.6527 | 1781.1451 | 12.0123 | 9.4146 | -4.8983 | 16.1211 | 1780.6527 | -9.4146 | 44.4961
0.4804 | -18.9904 | -28.7276 | -1770.9325 | 1771.3994 | 12.0038 | 9.4485 | -4.9153 | 16.1296 | 1770.9325 | -9.4485 | 44.6634
0.4822 | -18.5745 | -27.259 | -1759.3447 | 1759.7861 | 12.0123 | 9.4825 | -4.9238 | 16.155 | 1759.3447 | -9.4825 | 44.8308
0.484 | -18.1415 | -25.7818 | -1747.7908 | 1748.2068 | 12.0123 | 9.5249 | -4.9323 | 16.1805 | 1747.7908 | -9.5249 | 44.9981
0.4858 | -18.1924 | -25.1451 | -1740.8466 | 1741.2711 | 11.9953 | 9.5759 | -4.9407 | 16.1975 | 1740.8466 | -9.5759 | 45.1655
0.4876 | -17.8019 | -24.4745 | -1736.4916 | 1736.9076 | 11.9868 | 9.6013 | -4.9323 | 16.206 | 1736.4916 | -9.6013 | 45.3328
0.4894 | -17.6661 | -24.5594 | -1734.5985 | 1735.023 | 11.9613 | 9.6438 | -4.9153 | 16.206 | 1734.5985 | -9.6438 | 45.5002
0.4912 | -18.532 | -25.646 | -1735.7785 | 1736.2539 | 11.9019 | 9.7202 | -4.8983 | 16.206 | 1735.7785 | -9.7202 | 45.6675
0.493 | -19.6696 | -27.1316 | -1738.6139 | 1739.1318 | 11.834 | 9.8136 | -4.8728 | 16.206 | 1738.6139 | -9.8136 | 45.8349
0.4948 | -19.9582 | -27.7259 | -1741.6531 | 1742.1709 | 11.8085 | 9.8645 | -4.8389 | 16.206 | 1741.6531 | -9.8645 | 46.0022
0.4966 | -21.4608 | -29.5086 | -1745.7194 | 1746.2458 | 11.7321 | 9.9494 | -4.8219 | 16.1805 | 1745.7194 | -9.9494 | 46.1696
0.4984 | -21.4778 | -29.6784 | -1744.4715 | 1744.9894 | 11.7236 | 9.9918 | -4.8134 | 16.189 | 1744.4715 | -9.9918 | 46.3369
0.5002 | -22.8785 | -30.2726 | -1741.1437 | 1741.6701 | 11.6812 | 10.1107 | -4.8389 | 16.2314 | 1741.1437 | -10.1107 | 46.5043
0.502 | -21.9192 | -28.0824 | -1736.8312 | 1737.3236 | 11.7406 | 10.1107 | -4.8304 | 16.2739 | 1736.8312 | -10.1107 | 46.6716
0.5037 | -22.1654 | -27.9041 | -1734.9721 | 1735.4644 | 11.7067 | 10.1362 | -4.8219 | 16.2654 | 1734.9721 | -10.1362 | 46.8297
0.5055 | -25.1112 | -31.2744 | -1736.4067 | 1736.9925 | 11.5114 | 10.2635 | -4.8389 | 16.206 | 1736.4067 | -10.2635 | 46.997
0.5073 | -26.9194 | -33.5919 | -1736.0926 | 1736.7463 | 11.3841 | 10.3314 | -4.8474 | 16.1635 | 1736.0926 | -10.3314 | 47.1644
0.5091 | -26.7241 | -33.6004 | -1732.0348 | 1732.6885 | 11.3501 | 10.3314 | -4.8643 | 16.1465 | 1732.0348 | -10.3314 | 47.3317
0.5109 | -26.86 | -33.4052 | -1727.5015 | 1728.1467 | 11.3077 | 10.3314 | -4.8728 | 16.1126 | 1727.5015 | -10.3314 | 47.4991
0.5127 | -27.1231 | -33.0826 | -1722.1788 | 1722.824 | 11.2652 | 10.3314 | -4.8983 | 16.0956 | 1722.1788 | -10.3314 | 47.6664
0.5145 | -25.6036 | -30.6292 | -1714.1649 | 1714.7507 | 11.3162 | 10.272 | -4.8983 | 16.0871 | 1714.1649 | -10.272 | 47.8338
0.5163 | -24.7886 | -29.5511 | -1704.3514 | 1704.9117 | 11.2907 | 10.2295 | -4.8898 | 16.0447 | 1704.3514 | -10.2295 | 48.0011
0.5181 | -25.0518 | -29.8991 | -1693.986 | 1694.5718 | 11.2058 | 10.238 | -4.8983 | 15.9852 | 1693.986 | -10.238 | 48.1685
0.5199 | -24.9159 | -29.4577 | -1684.2658 | 1684.8516 | 11.1549 | 10.238 | -4.8983 | 15.9513 | 1684.2658 | -10.238 | 48.3358
0.5217 | -23.7699 | -26.5968 | -1676.8377 | 1677.398 | 11.2143 | 10.2041 | -4.8983 | 15.9768 | 1676.8377 | -10.2041 | 48.5032
0.5235 | -22.0635 | -24.5169 | -1672.2196 | 1672.7544 | 11.2313 | 10.1107 | -4.8813 | 15.9258 | 1672.2196 | -10.1107 | 48.6705
0.5253 | -22.9889 | -25.7733 | -1669.953 | 1670.5472 | 11.1294 | 10.1446 | -4.8813 | 15.8834 | 1669.953 | -10.1446 | 48.8379
0.5271 | -22.8446 | -25.7394 | -1668.722 | 1669.3163 | 11.1209 | 10.1616 | -4.8643 | 15.8834 | 1668.722 | -10.1616 | 49.0052
0.5288 | -22.4371 | -25.3234 | -1667.16 | 1667.7457 | 11.1379 | 10.1701 | -4.8474 | 15.8919 | 1667.16 | -10.1701 | 49.1633
0.5306 | -20.9005 | -23.1926 | -1664.1548 | 1664.6811 | 11.2228 | 10.1446 | -4.8219 | 15.9258 | 1664.1548 | -10.1446 | 49.3306
0.5324 | -19.364 | -21.0618 | -1661.1496 | 1661.625 | 11.2992 | 10.1107 | -4.7964 | 15.9598 | 1661.1496 | -10.1107 | 49.498
0.5342 | -18.0312 | -19.5337 | -1659.0867 | 1659.4942 | 11.3247 | 10.0258 | -4.7794 | 15.9004 | 1659.0867 | -10.0258 | 49.6653
0.536 | -18.116 | -19.1008 | -1659.7149 | 1660.1224 | 11.3162 | 10.0258 | -4.754 | 15.8919 | 1659.7149 | -10.0258 | 49.8327
0.5378 | -18.7782 | -18.9055 | -1660.3686 | 1660.7761 | 11.3162 | 10.0937 | -4.7285 | 15.9173 | 1660.3686 | -10.0937 | 50
0.5396 | -18.914 | -18.6084 | -1661.2854 | 1661.6844 | 11.3162 | 10.1107 | -4.703 | 15.9258 | 1661.2854 | -10.1107 | 50.1673
0.5414 | -18.4811 | -18.4556 | -1663.4587 | 1663.8577 | 11.3077 | 10.0767 | -4.6861 | 15.8919 | 1663.4587 | -10.0767 | 50.3347
0.5432 | -20.1025 | -20.875 | -1669.4436 | 1669.902 | 11.1973 | 10.1107 | -4.6861 | 15.8409 | 1669.4436 | -10.1107 | 50.502
0.545 | -20.9854 | -21.6136 | -1675.8445 | 1676.3029 | 11.1634 | 10.1362 | -4.6691 | 15.8239 | 1675.8445 | -10.1362 | 50.6694
0.5468 | -21.8513 | -22.4795 | -1682.9415 | 1683.4084 | 11.1294 | 10.1446 | -4.6606 | 15.7985 | 1682.9415 | -10.1446 | 50.8367
0.5486 | -24.7292 | -26.6562 | -1694.4189 | 1694.9538 | 10.9596 | 10.2126 | -4.6606 | 15.7136 | 1694.4189 | -10.2126 | 51.0041
0.5504 | -27.4712 | -30.7395 | -1706.202 | 1706.7963 | 10.7983 | 10.2635 | -4.6691 | 15.6372 | 1706.202 | -10.2635 | 51.1714
0.5521 | -28.8889 | -33.142 | -1715.8458 | 1716.4825 | 10.7219 | 10.2805 | -4.6691 | 15.5863 | 1715.8458 | -10.2805 | 51.3295
0.5539 | -30.1368 | -35.536 | -1724.9038 | 1725.5999 | 10.6285 | 10.2975 | -4.6521 | 15.5353 | 1724.9038 | -10.2975 | 51.4968
0.5557 | -30.8924 | -37.7092 | -1733.2657 | 1733.9958 | 10.5691 | 10.2975 | -4.6691 | 15.5014 | 1733.2657 | -10.2975 | 51.6642
0.5575 | -31.8771 | -39.6278 | -1741.8653 | 1742.6463 | 10.5436 | 10.3399 | -4.7285 | 15.5268 | 1741.8653 | -10.3399 | 51.8315
0.5593 | -32.4034 | -40.7908 | -1752.3241 | 1753.156 | 10.5521 | 10.3569 | -4.771 | 15.5608 | 1752.3241 | -10.3569 | 51.9989
0.5611 | -31.8856 | -40.5956 | -1764.1496 | 1764.99 | 10.6455 | 10.3059 | -4.8219 | 15.6202 | 1764.1496 | -10.3059 | 52.1662
0.5629 | -32.4459 | -41.9708 | -1779.1501 | 1780.0415 | 10.671 | 10.3144 | -4.8558 | 15.6627 | 1779.1501 | -10.3144 | 52.3336
0.5647 | -33.5155 | -43.414 | -1799.1338 | 1800.0591 | 10.7389 | 10.3059 | -4.9153 | 15.7391 | 1799.1338 | -10.3059 | 52.5009
0.5665 | -33.6089 | -43.4904 | -1818.1327 | 1819.0495 | 10.8408 | 10.2805 | -4.9407 | 15.807 | 1818.1327 | -10.2805 | 52.6683
0.5683 | -32.8873 | -42.8622 | -1831.291 | 1832.1824 | 10.9341 | 10.2295 | -4.9492 | 15.8494 | 1831.291 | -10.2295 | 52.8356
0.5701 | -31.4187 | -41.801 | -1842.726 | 1843.5834 | 11.036 | 10.1446 | -4.9577 | 15.8834 | 1842.726 | -10.1446 | 53.003
0.5719 | -30.2557 | -41.054 | -1852.7688 | 1853.6092 | 11.1209 | 10.0682 | -4.9747 | 15.9088 | 1852.7688 | -10.0682 | 53.1703
0.5737 | -28.6342 | -39.3137 | -1858.4905 | 1859.2631 | 11.2313 | 10.0088 | -4.9917 | 15.9513 | 1858.4905 | -10.0088 | 53.3377
0.5755 | -26.9449 | -37.8281 | -1861.1307 | 1861.8183 | 11.3331 | 9.9154 | -5.0087 | 15.9683 | 1861.1307 | -9.9154 | 53.505
0.5772 | -25.612 | -36.7499 | -1864.2378 | 1864.866 | 11.418 | 9.839 | -5.0341 | 15.9852 | 1864.2378 | -9.839 | 53.6631
0.579 | -25.3064 | -36.2406 | -1871.2584 | 1871.8696 | 11.452 | 9.8305 | -5.0341 | 16.0022 | 1871.2584 | -9.8305 | 53.8304
0.5808 | -26.4864 | -37.5649 | -1879.1619 | 1879.8325 | 11.4011 | 9.8815 | -5.0426 | 16.0022 | 1879.1619 | -9.8815 | 53.9978
0.5826 | -26.1808 | -37.4885 | -1884.6799 | 1885.3675 | 11.435 | 9.839 | -5.0596 | 16.0192 | 1884.6799 | -9.839 | 54.1651
0.5844 | -25.6715 | -37.1659 | -1889.6715 | 1890.3846 | 11.4859 | 9.7966 | -5.0851 | 16.0362 | 1889.6715 | -9.7966 | 54.3325
0.5862 | -24.5509 | -35.918 | -1893.1691 | 1893.8567 | 11.5623 | 9.7372 | -5.1105 | 16.0701 | 1893.1691 | -9.7372 | 54.4998
0.588 | -22.8106 | -33.541 | -1899.587 | 1900.1812 | 11.6557 | 9.6862 | -5.102 | 16.1041 | 1899.587 | -9.6862 | 54.6672
0.5898 | -23.9397 | -35.0096 | -1906.5821 | 1907.2358 | 11.5878 | 9.7032 | -5.1105 | 16.0786 | 1906.5821 | -9.7032 | 54.8345
0.5916 | -25.1451 | -36.317 | -1910.1476 | 1910.8777 | 11.5199 | 9.7202 | -5.119 | 16.0532 | 1910.1476 | -9.7202 | 55.0019
0.5934 | -25.3064 | -35.884 | -1906.8707 | 1907.6093 | 11.4859 | 9.6947 | -5.1275 | 16.0107 | 1906.8707 | -9.6947 | 55.1692
0.5952 | -25.2979 | -35.2643 | -1902.745 | 1903.4581 | 11.435 | 9.6692 | -5.119 | 15.9513 | 1902.745 | -9.6692 | 55.3366
0.597 | -25.2895 | -34.9247 | -1896.828 | 1897.5241 | 11.3756 | 9.6353 | -5.1275 | 15.8834 | 1896.828 | -9.6353 | 55.5039
0.5988 | -25.2895 | -35.086 | -1887.8803 | 1888.5849 | 11.2822 | 9.5928 | -5.136 | 15.7985 | 1887.8803 | -9.5928 | 55.6713
0.6006 | -25.4762 | -35.4426 | -1877.88 | 1878.5846 | 11.1803 | 9.5844 | -5.136 | 15.7136 | 1877.88 | -9.5844 | 55.8386
0.6023 | -25.0518 | -35.0521 | -1866.3346 | 1867.0222 | 11.0954 | 9.5844 | -5.1275 | 15.6372 | 1866.3346 | -9.5844 | 55.9967
0.6041 | -26.1044 | -35.5529 | -1855.9183 | 1856.5974 | 10.9851 | 9.6438 | -5.119 | 15.5693 | 1855.9183 | -9.6438 | 56.164
0.6059 | -28.1334 | -35.9095 | -1845.1285 | 1845.7991 | 10.9002 | 9.7202 | -5.119 | 15.5438 | 1845.1285 | -9.7202 | 56.3313
0.6077 | -29.1521 | -35.7142 | -1837.3269 | 1837.972 | 10.8577 | 9.7541 | -5.102 | 15.5183 | 1837.3269 | -9.7541 | 56.4987
0.6095 | -28.8804 | -35.5105 | -1837.6579 | 1838.2862 | 10.8493 | 9.7032 | -5.0681 | 15.4759 | 1837.6579 | -9.7032 | 56.666
0.6113 | -30.6462 | -37.7941 | -1842.9552 | 1843.6514 | 10.7644 | 9.7372 | -5.0511 | 15.4165 | 1842.9552 | -9.7372 | 56.8334
0.6131 | -33.0062 | -39.6617 | -1847.6073 | 1848.3799 | 10.7134 | 9.7966 | -5.0766 | 15.425 | 1847.6073 | -9.7966 | 57.0007
0.6149 | -33.6089 | -39.5938 | -1848.9911 | 1849.7636 | 10.7304 | 9.8221 | -5.102 | 15.4674 | 1848.9911 | -9.8221 | 57.1681
0.6167 | -32.3185 | -37.2508 | -1848.5921 | 1849.2882 | 10.8408 | 9.7796 | -5.102 | 15.5353 | 1848.5921 | -9.7796 | 57.3354
0.6185 | -30.1113 | -33.7532 | -1846.8348 | 1847.4121 | 11.0275 | 9.7117 | -5.102 | 15.6202 | 1846.8348 | -9.7117 | 57.5028
0.6203 | -27.2335 | -30.2726 | -1851.2832 | 1851.7416 | 11.2058 | 9.5844 | -5.0596 | 15.6711 | 1851.2832 | -9.5844 | 57.6701
0.6221 | -27.7513 | -30.782 | -1859.5177 | 1859.9931 | 11.1888 | 9.6268 | -5.0426 | 15.6796 | 1859.5177 | -9.6268 | 57.8375
0.6239 | -28.6342 | -32.1233 | -1868.2192 | 1868.7286 | 11.1803 | 9.6523 | -5.0681 | 15.6966 | 1868.2192 | -9.6523 | 58.0048
0.6256 | -27.4118 | -31.4017 | -1875.4945 | 1875.9699 | 11.2737 | 9.5928 | -5.0935 | 15.7391 | 1875.4945 | -9.5928 | 58.1629
0.6274 | -25.9007 | -30.9263 | -1884.6374 | 1885.0789 | 11.3501 | 9.491 | -5.102 | 15.7475 | 1884.6374 | -9.491 | 58.3302
0.6292 | -25.1197 | -32.1318 | -1896.7516 | 1897.2015 | 11.3416 | 9.3806 | -5.0766 | 15.6627 | 1896.7516 | -9.3806 | 58.4976
0.631 | -26.6053 | -35.4681 | -1910.3853 | 1910.9116 | 11.2482 | 9.3636 | -5.0681 | 15.5863 | 1910.3853 | -9.3636 | 58.6649
0.6328 | -27.3948 | -37.4291 | -1919.4688 | 1920.0291 | 11.1803 | 9.3806 | -5.0596 | 15.5608 | 1919.4688 | -9.3806 | 58.8323
0.6346 | -27.5051 | -38.1082 | -1921.6166 | 1922.2108 | 11.1634 | 9.3636 | -5.0766 | 15.5353 | 1921.6166 | -9.3636 | 58.9996
0.6364 | -26.1044 | -36.9876 | -1922.5419 | 1923.0852 | 11.2058 | 9.2957 | -5.0766 | 15.5353 | 1922.5419 | -9.2957 | 59.167
0.6382 | -24.449 | -35.6463 | -1923.3738 | 1923.8662 | 11.2567 | 9.2023 | -5.0596 | 15.5268 | 1923.3738 | -9.2023 | 59.3343
0.64 | -23.9482 | -35.8161 | -1929.2399 | 1929.7323 | 11.2143 | 9.1599 | -5.0171 | 15.4589 | 1929.2399 | -9.1599 | 59.5017
0.6418 | -24.7292 | -37.3017 | -1935.3352 | 1935.87 | 11.1124 | 9.1599 | -4.9662 | 15.3655 | 1935.3352 | -9.1599 | 59.669
0.6436 | -24.5679 | -36.7754 | -1934.7494 | 1935.2758 | 11.1124 | 9.1769 | -4.9662 | 15.374 | 1934.7494 | -9.1769 | 59.8364
0.6454 | -23.2181 | -34.8144 | -1930.6661 | 1931.15 | 11.1888 | 9.1429 | -4.9577 | 15.4165 | 1930.6661 | -9.1429 | 60.0037
0.6472 | -21.6221 | -33.9145 | -1926.8884 | 1927.3298 | 11.2058 | 9.109 | -4.9407 | 15.408 | 1926.8884 | -9.109 | 60.1711
0.649 | -21.1382 | -34.1013 | -1927.9071 | 1928.357 | 11.1634 | 9.1259 | -4.8983 | 15.3655 | 1927.9071 | -9.1259 | 60.3384
0.6507 | -20.6119 | -33.5919 | -1925.6574 | 1926.1074 | 11.1464 | 9.1514 | -4.8728 | 15.3486 | 1925.6574 | -9.1514 | 60.4965
0.6525 | -19.1687 | -30.8414 | -1914.4771 | 1914.8761 | 11.2228 | 9.1854 | -4.8558 | 15.425 | 1914.4771 | -9.1854 | 60.6638
0.6543 | -17.8189 | -28.821 | -1905.1559 | 1905.487 | 11.2482 | 9.2448 | -4.8134 | 15.4589 | 1905.1559 | -9.2448 | 60.8312
0.6561 | -17.7425 | -27.8872 | -1897.5241 | 1897.8297 | 11.2313 | 9.3467 | -4.7879 | 15.4929 | 1897.5241 | -9.3467 | 60.9985
0.6579 | -18.3792 | -27.3948 | -1890.3082 | 1890.6138 | 11.1973 | 9.474 | -4.771 | 15.5268 | 1890.3082 | -9.474 | 61.1659
0.6597 | -17.6406 | -25.3149 | -1881.9633 | 1882.2265 | 11.1803 | 9.5164 | -4.7115 | 15.5183 | 1881.9633 | -9.5164 | 61.3332
0.6615 | -16.7578 | -23.2181 | -1874.1447 | 1874.3654 | 11.1634 | 9.5419 | -4.6521 | 15.4929 | 1874.1447 | -9.5419 | 61.5006
0.6633 | -17.2417 | -22.5474 | -1861.1307 | 1861.3514 | 11.0615 | 9.6098 | -4.6097 | 15.4419 | 1861.1307 | -9.6098 | 61.6679
0.6651 | -18.7188 | -23.4133 | -1849.1609 | 1849.4155 | 10.9002 | 9.7117 | -4.5927 | 15.3825 | 1849.1609 | -9.7117 | 61.8353
0.6669 | -21.1382 | -25.9686 | -1844.8908 | 1845.2049 | 10.7219 | 9.8475 | -4.5757 | 15.3231 | 1844.8908 | -9.8475 | 62.0026
0.6687 | -22.0635 | -27.0552 | -1843.8975 | 1844.2456 | 10.6285 | 9.9239 | -4.5502 | 15.2891 | 1843.8975 | -9.9239 | 62.17
0.6705 | -23.4388 | -28.7021 | -1845.1455 | 1845.536 | 10.5267 | 10.0343 | -4.5502 | 15.2806 | 1845.1455 | -10.0343 | 62.3373
0.6723 | -23.3624 | -29.2115 | -1846.334 | 1846.7414 | 10.4842 | 10.0852 | -4.5333 | 15.2806 | 1846.334 | -10.0852 | 62.5046
0.6741 | -22.4286 | -27.641 | -1845.7142 | 1846.0878 | 10.5521 | 10.1192 | -4.5248 | 15.3486 | 1845.7142 | -10.1192 | 62.672
0.6758 | -21.4778 | -26.028 | -1845.0775 | 1845.4171 | 10.6285 | 10.1531 | -4.5163 | 15.425 | 1845.0775 | -10.1531 | 62.83
0.6776 | -20.5949 | -24.7207 | -1845.5105 | 1845.8161 | 10.6795 | 10.1786 | -4.5078 | 15.4759 | 1845.5105 | -10.1786 | 62.9974
0.6794 | -19.8224 | -24.101 | -1848.3884 | 1848.677 | 10.688 | 10.1786 | -4.4653 | 15.4674 | 1848.3884 | -10.1786 | 63.1647
0.6812 | -21.4014 | -26.1554 | -1856.3513 | 1856.6993 | 10.5691 | 10.2635 | -4.4314 | 15.425 | 1856.3513 | -10.2635 | 63.3321
0.683 | -23.0398 | -28.2692 | -1867.243 | 1867.642 | 10.5267 | 10.3823 | -4.4399 | 15.4759 | 1867.243 | -10.3823 | 63.4994
0.6848 | -24.1943 | -29.95 | -1878.4233 | 1878.8562 | 10.5097 | 10.4927 | -4.4484 | 15.5353 | 1878.4233 | -10.4927 | 63.6668
0.6866 | -23.8972 | -30.2302 | -1889.8328 | 1890.2658 | 10.5436 | 10.5521 | -4.4569 | 15.6032 | 1889.8328 | -10.5521 | 63.8341
0.6884 | -22.7681 | -28.8719 | -1899.4511 | 1899.8417 | 10.6625 | 10.6031 | -4.4653 | 15.7136 | 1899.4511 | -10.6031 | 64.0015
0.6902 | -20.7222 | -26.7666 | -1905.4106 | 1905.7417 | 10.7559 | 10.5606 | -4.4653 | 15.756 | 1905.4106 | -10.5606 | 64.1688
0.692 | -19.6441 | -25.1451 | -1912.584 | 1912.8811 | 10.8068 | 10.5267 | -4.4484 | 15.773 | 1912.584 | -10.5267 | 64.3362
0.6938 | -21.0958 | -26.5374 | -1919.9272 | 1920.2498 | 10.7219 | 10.5861 | -4.4484 | 15.756 | 1919.9272 | -10.5861 | 64.5035
0.6956 | -21.3759 | -26.6053 | -1921.1072 | 1921.4468 | 10.6795 | 10.6031 | -4.4569 | 15.7391 | 1921.1072 | -10.6031 | 64.6709
0.6974 | -18.4811 | -22.9379 | -1915.326 | 1915.5722 | 10.7644 | 10.5352 | -4.4144 | 15.7391 | 1915.326 | -10.5352 | 64.8382
0.6992 | -18.0227 | -22.6748 | -1908.8487 | 1909.0864 | 10.637 | 10.5267 | -4.355 | 15.6287 | 1908.8487 | -10.5267 | 65.0056
0.7009 | -18.2349 | -23.6086 | -1902.2356 | 1902.4818 | 10.4333 | 10.5352 | -4.304 | 15.4759 | 1902.2356 | -10.5352 | 65.1636
0.7027 | -16.953 | -22.4371 | -1895.4188 | 1895.6395 | 10.3654 | 10.5012 | -4.2701 | 15.3995 | 1895.4188 | -10.5012 | 65.331
0.7045 | -15.3825 | -20.8411 | -1890.8006 | 1890.9874 | 10.3399 | 10.4333 | -4.2446 | 15.3401 | 1890.8006 | -10.4333 | 65.4983
0.7063 | -14.3638 | -20.0007 | -1890.2233 | 1890.3931 | 10.2975 | 10.3993 | -4.2276 | 15.2721 | 1890.2233 | -10.3993 | 65.6657
0.7081 | -13.795 | -19.6526 | -1890.2403 | 1890.4101 | 10.221 | 10.4078 | -4.2107 | 15.2212 | 1890.2403 | -10.4078 | 65.833
0.7099 | -14.6355 | -20.493 | -1893.0587 | 1893.237 | 10.1277 | 10.3993 | -4.1937 | 15.1533 | 1893.0587 | -10.3993 | 66.0004
0.7117 | -15.5268 | -21.0533 | -1895.3933 | 1895.597 | 10.0598 | 10.3993 | -4.1767 | 15.1024 | 1895.3933 | -10.3993 | 66.1677
0.7135 | -14.1516 | -19.3809 | -1892.9993 | 1893.1691 | 10.1107 | 10.2805 | -4.1682 | 15.0599 | 1892.9993 | -10.2805 | 66.3351
0.7153 | -13.2687 | -18.6509 | -1892.9739 | 1893.1267 | 10.1362 | 10.1701 | -4.1767 | 15.0005 | 1892.9739 | -10.1701 | 66.5024
0.7171 | -12.8697 | -18.7527 | -1894.7906 | 1894.9603 | 10.1277 | 10.0598 | -4.1852 | 14.9241 | 1894.7906 | -10.0598 | 66.6698
0.7189 | -12.8018 | -20.0601 | -1899.5955 | 1899.7652 | 10.0852 | 9.9494 | -4.1937 | 14.8222 | 1899.5955 | -9.9494 | 66.8371
0.7207 | -12.0972 | -20.909 | -1903.6448 | 1903.8146 | 10.0682 | 9.8221 | -4.1937 | 14.7203 | 1903.6448 | -9.8221 | 67.0045
0.7225 | -11.6388 | -21.6645 | -1904.1966 | 1904.3834 | 10.0852 | 9.7287 | -4.2107 | 14.6779 | 1904.1966 | -9.7287 | 67.1718
0.7242 | -9.474 | -20.544 | -1901.183 | 1901.3358 | 10.1956 | 9.5674 | -4.2192 | 14.6694 | 1901.183 | -9.5674 | 67.3299
0.726 | -9.3976 | -22.0381 | -1900.8604 | 1901.0386 | 10.1871 | 9.4995 | -4.2361 | 14.6185 | 1900.8604 | -9.4995 | 67.4972
0.7278 | -8.7439 | -22.6493 | -1898.5852 | 1898.772 | 10.1616 | 9.4061 | -4.2192 | 14.5421 | 1898.5852 | -9.4061 | 67.6646
0.7296 | -7.929 | -22.9294 | -1895.8432 | 1896.0385 | 10.1531 | 9.3042 | -4.2107 | 14.4657 | 1895.8432 | -9.3042 | 67.8319
0.7314 | -6.0953 | -20.9599 | -1890.0366 | 1890.2064 | 10.2126 | 9.1599 | -4.1937 | 14.4062 | 1890.0366 | -9.1599 | 67.9993
0.7332 | -5.0681 | -19.2451 | -1884.943 | 1885.0873 | 10.2295 | 9.0241 | -4.1767 | 14.3298 | 1884.943 | -9.0241 | 68.1666
0.735 | -3.6674 | -16.9445 | -1877.7951 | 1877.9309 | 10.2805 | 8.8628 | -4.1597 | 14.2619 | 1877.7951 | -8.8628 | 68.334
0.7368 | -1.4262 | -13.7526 | -1867.6165 | 1867.7523 | 10.3739 | 8.6675 | -4.1597 | 14.2025 | 1867.6165 | -8.6675 | 68.5013
0.7386 | 0.34806 | -10.5267 | -1857.7435 | 1857.8878 | 10.4333 | 8.4638 | -4.1343 | 14.1006 | 1857.7435 | -8.4638 | 68.6687
0.7404 | 0.83195 | -9.4485 | -1848.8213 | 1848.9741 | 10.4333 | 8.3025 | -4.1088 | 13.9988 | 1848.8213 | -8.3025 | 68.836
0.7422 | 1.8931 | -7.6233 | -1840.1623 | 1840.3236 | 10.4248 | 8.0902 | -4.0494 | 13.8544 | 1840.1623 | -8.0902 | 69.0033
0.744 | 2.9458 | -5.7982 | -1831.5032 | 1831.673 | 10.4163 | 7.878 | -3.9984 | 13.7101 | 1831.5032 | -7.878 | 69.1707
0.7458 | 3.0137 | -5.535 | -1820.9426 | 1821.1379 | 10.3314 | 7.6913 | -3.973 | 13.5319 | 1820.9426 | -7.6913 | 69.338
0.7476 | 3.4806 | -4.3295 | -1808.2173 | 1808.438 | 10.289 | 7.4705 | -3.9475 | 13.3621 | 1808.2173 | -7.4705 | 69.5054
0.7493 | 4.9068 | -1.2479 | -1795.4155 | 1795.7126 | 10.255 | 7.1819 | -3.939 | 13.1583 | 1795.4155 | -7.1819 | 69.6634
0.7511 | 6.5282 | 2.2072 | -1782.6901 | 1783.0721 | 10.221 | 6.8678 | -3.9305 | 12.9461 | 1782.6901 | -6.8678 | 69.8308
0.7529 | 4.6861 | 0.22072 | -1770.9155 | 1771.306 | 10.0513 | 6.7065 | -3.9475 | 12.7254 | 1770.9155 | -6.7065 | 69.9981
0.7547 | 4.0663 | -0.46691 | -1755.7283 | 1756.1357 | 9.9409 | 6.5282 | -3.956 | 12.5556 | 1755.7283 | -6.5282 | 70.1655
0.7565 | 4.9832 | 0.88288 | -1737.1283 | 1737.5358 | 9.89 | 6.282 | -3.9645 | 12.3858 | 1737.1283 | -6.282 | 70.3328
0.7583 | 5.6369 | 2.776 | -1715.3025 | 1715.7439 | 9.8475 | 6.0528 | -3.9899 | 12.25 | 1715.3025 | -6.0528 | 70.5002
0.7601 | 6.3245 | 4.5078 | -1693.3069 | 1693.8077 | 9.7372 | 5.7642 | -4.0069 | 12.0293 | 1693.3069 | -5.7642 | 70.6675
0.7619 | 7.0546 | 6.2651 | -1671.2858 | 1671.8376 | 9.6268 | 5.4841 | -4.0239 | 11.8 | 1671.2858 | -5.4841 | 70.8349
0.7637 | 7.861 | 7.9544 | -1649.0015 | 1649.6212 | 9.508 | 5.1869 | -4.0239 | 11.5708 | 1649.0015 | -5.1869 | 71.0022
0.7655 | 9.2363 | 9.6947 | -1625.1213 | 1625.8259 | 9.4146 | 4.8558 | -3.9899 | 11.3331 | 1625.1213 | -4.8558 | 71.1696
0.7673 | 11.0275 | 11.6133 | -1598.134 | 1598.8555 | 9.3297 | 4.5248 | -3.939 | 11.1209 | 1598.134 | -4.5248 | 71.3369
0.7691 | 12.2075 | 13.031 | -1571.9616 | 1572.6832 | 9.2193 | 4.2446 | -3.8966 | 10.9002 | 1571.9616 | -4.2446 | 71.5043
0.7709 | 12.5386 | 13.3875 | -1548.6502 | 1549.3548 | 9.058 | 4.0069 | -3.8626 | 10.6625 | 1548.6502 | -4.0069 | 71.6716
0.7727 | 12.4113 | 13.0989 | -1528.259 | 1528.9551 | 8.8713 | 3.8032 | -3.8286 | 10.4163 | 1528.259 | -3.8032 | 71.839
0.7744 | 12.5386 | 13.3026 | -1509.8119 | 1510.5165 | 8.7269 | 3.6674 | -3.8117 | 10.2295 | 1509.8119 | -3.6674 | 71.997
0.7762 | 13.0565 | 14.0242 | -1486.3816 | 1487.0607 | 8.5741 | 3.4297 | -3.7438 | 10.0003 | 1486.3816 | -3.4297 | 72.1644
0.778 | 12.3773 | 13.3196 | -1465.3537 | 1466.0328 | 8.3789 | 3.2174 | -3.6928 | 9.7457 | 1465.3537 | -3.2174 | 72.3317
0.7798 | 11.7831 | 12.4877 | -1446.9575 | 1447.6282 | 8.2091 | 3.0391 | -3.6674 | 9.5419 | 1446.9575 | -3.0391 | 72.4991
0.7816 | 10.9426 | 11.2313 | -1429.4272 | 1430.0809 | 8.0223 | 2.8779 | -3.6419 | 9.3297 | 1429.4272 | -2.8779 | 72.6664
0.7834 | 7.8441 | 7.3517 | -1417.1008 | 1417.6102 | 7.7507 | 2.8609 | -3.6334 | 9.1005 | 1417.1008 | -2.8609 | 72.8338
0.7852 | 4.2022 | 3.939 | -1404.384 | 1404.7745 | 7.513 | 2.8779 | -3.6249 | 8.9052 | 1404.384 | -2.8779 | 73.0011
0.787 | 2.2751 | 3.3448 | -1387.7705 | 1388.2035 | 7.3602 | 2.8269 | -3.5909 | 8.7609 | 1387.7705 | -2.8269 | 73.1685
0.7888 | 0.81497 | 3.0137 | -1370.4779 | 1370.9618 | 7.2159 | 2.7505 | -3.54 | 8.6081 | 1370.4779 | -2.7505 | 73.3358
0.7906 | 0.61123 | 4.0409 | -1346.6741 | 1347.2004 | 7.0885 | 2.6402 | -3.4976 | 8.4383 | 1346.6741 | -2.6402 | 73.5032
0.7924 | 0.53482 | 6.1462 | -1321.0281 | 1321.5799 | 6.9866 | 2.5298 | -3.4551 | 8.311 | 1321.0281 | -2.5298 | 73.6705
0.7942 | 1.0781 | 11.4605 | -1293.1324 | 1293.7012 | 6.9612 | 2.4025 | -3.3872 | 8.2261 | 1293.1324 | -2.4025 | 73.8379
0.796 | 1.1291 | 15.0175 | -1269.7191 | 1270.3303 | 6.8933 | 2.2496 | -3.3023 | 8.0818 | 1269.7191 | -2.2496 | 74.0052
0.7977 | 1.3158 | 17.6916 | -1246.934 | 1247.5792 | 6.8338 | 2.0968 | -3.2344 | 7.9629 | 1246.934 | -2.0968 | 74.1633
0.7995 | 1.2394 | 19.2366 | -1227.2644 | 1227.9605 | 6.7405 | 1.9695 | -3.1665 | 7.8186 | 1227.2644 | -1.9695 | 74.3306
0.8013 | 1.6214 | 20.6373 | -1208.0957 | 1208.8597 | 6.6131 | 1.8252 | -3.1071 | 7.6403 | 1208.0957 | -1.8252 | 74.498
0.8031 | 2.0459 | 21.5457 | -1190.0391 | 1190.8371 | 6.4773 | 1.6724 | -3.0476 | 7.4536 | 1190.0391 | -1.6724 | 74.6653
0.8049 | 2.3855 | 21.2146 | -1175.7602 | 1176.5497 | 6.3075 | 1.5111 | -2.9882 | 7.2413 | 1175.7602 | -1.5111 | 74.8327
0.8067 | 2.4109 | 21.0703 | -1163.7054 | 1164.5034 | 6.1038 | 1.3583 | -2.9203 | 6.9866 | 1163.7054 | -1.3583 | 75
0.8085 | 4.5248 | 23.5661 | -1146.676 | 1147.5419 | 6.0698 | 1.1376 | -2.8439 | 6.8848 | 1146.676 | -1.1376 | 75.1673
0.8103 | 7.1819 | 27.2165 | -1127.329 | 1128.1949 | 6.1038 | 0.89137 | -2.7505 | 6.8678 | 1127.329 | -0.89137 | 75.3347
0.8121 | 9.3551 | 30.9263 | -1106.7935 | 1107.7698 | 6.1377 | 0.69612 | -2.6741 | 6.8593 | 1106.7935 | -0.69612 | 75.502
0.8139 | 10.8408 | 33.8636 | -1085.7657 | 1086.8693 | 6.1292 | 0.5518 | -2.5892 | 6.8254 | 1085.7657 | -0.5518 | 75.6694
0.8157 | 12.0123 | 36.5886 | -1064.5681 | 1065.7056 | 6.1123 | 0.44993 | -2.5128 | 6.7829 | 1064.5681 | -0.44993 | 75.8367
0.8175 | 12.0123 | 37.7092 | -1048.1923 | 1049.4063 | 6.0274 | 0.38202 | -2.4534 | 6.6641 | 1048.1923 | -0.38202 | 76.0041
0.8193 | 10.9596 | 37.0725 | -1034.4992 | 1035.6876 | 5.8746 | 0.34806 | -2.394 | 6.5028 | 1034.4992 | -0.34806 | 76.1714
0.8211 | 10.0428 | 36.6141 | -1020.8909 | 1022.0794 | 5.7217 | 0.30561 | -2.3345 | 6.3415 | 1020.8909 | -0.30561 | 76.3388
0.8228 | 9.3127 | 36.0623 | -1008.1231 | 1009.2861 | 5.5435 | 0.27166 | -2.3006 | 6.1547 | 1008.1231 | -0.27166 | 76.4968
0.8246 | 8.7609 | 36.1217 | -993.2924 | 994.4639 | 5.3652 | 0.24619 | -2.2666 | 5.9764 | 993.2924 | -0.24619 | 76.6642
0.8264 | 8.4128 | 37.6922 | -978.8691 | 980.1765 | 5.2464 | 0.2377 | -2.2327 | 5.8491 | 978.8691 | -0.2377 | 76.8315
0.8282 | 7.1479 | 37.1659 | -968.2576 | 969.5819 | 5.0935 | 0.26317 | -2.1902 | 5.6963 | 968.2576 | -0.26317 | 76.9989
0.83 | 6.0868 | 37.2678 | -958.7496 | 960.0825 | 4.9917 | 0.3141 | -2.1817 | 5.6029 | 958.7496 | -0.3141 | 77.1662
0.8318 | 5.1275 | 37.4121 | -949.7001 | 951.0584 | 4.8898 | 0.36504 | -2.1817 | 5.518 | 949.7001 | -0.36504 | 77.3336
0.8336 | 4.2786 | 36.8009 | -943.4266 | 944.8018 | 4.771 | 0.44144 | -2.1732 | 5.4076 | 943.4266 | -0.44144 | 77.5009
0.8354 | 3.4042 | 36.6226 | -937.2209 | 938.6301 | 4.6521 | 0.48389 | -2.1563 | 5.2973 | 937.2209 | -0.48389 | 77.6683
0.8372 | 2.1053 | 36.6566 | -931.7199 | 933.2564 | 4.5417 | 0.54331 | -2.1563 | 5.1869 | 931.7199 | -0.54331 | 77.8356
0.839 | 0.084892 | 35.6633 | -927.1866 | 928.876 | 4.4059 | 0.64518 | -2.1563 | 5.0681 | 927.1866 | -0.64518 | 78.003
0.8408 | -0.12734 | 36.7414 | -921.4649 | 923.3325 | 4.3295 | 0.68763 | -2.1732 | 5.0087 | 921.4649 | -0.68763 | 78.1703
0.8426 | -0.11036 | 38.4902 | -915.4036 | 917.5344 | 4.321 | 0.77252 | -2.2072 | 5.0171 | 915.4036 | -0.77252 | 78.3377
0.8444 | -1.2479 | 39.1694 | -910.0129 | 912.3389 | 4.2361 | 0.88288 | -2.2412 | 4.9832 | 910.0129 | -0.88288 | 78.505
0.8462 | -2.9967 | 39.1863 | -905.9465 | 908.4084 | 4.1003 | 1.0017 | -2.2751 | 4.9323 | 905.9465 | -1.0017 | 78.6724
0.8479 | -5.1699 | 39.2203 | -903.4932 | 906.0314 | 3.956 | 1.1545 | -2.3261 | 4.8983 | 903.4932 | -1.1545 | 78.8304
0.8497 | -6.3415 | 40.1966 | -902.4999 | 905.1146 | 3.9135 | 1.2904 | -2.394 | 4.9407 | 902.4999 | -1.2904 | 78.9978
0.8515 | -9.4146 | 38.4732 | -903.3828 | 906.0909 | 3.7777 | 1.5026 | -2.4279 | 4.9407 | 903.3828 | -1.5026 | 79.1651
0.8533 | -11.7321 | 37.0131 | -905.5306 | 908.1538 | 3.6674 | 1.6809 | -2.4619 | 4.9492 | 905.5306 | -1.6809 | 79.3325
0.8551 | -13.8799 | 35.4765 | -907.2624 | 909.7497 | 3.557 | 1.8676 | -2.4873 | 4.9832 | 907.2624 | -1.8676 | 79.4998
0.8569 | -16.3588 | 33.4052 | -907.2794 | 909.7837 | 3.4551 | 2.0799 | -2.5128 | 5.0256 | 907.2794 | -2.0799 | 79.6672
0.8587 | -18.4811 | 31.7328 | -906.5408 | 909.0281 | 3.3533 | 2.3006 | -2.5213 | 5.0935 | 906.5408 | -2.3006 | 79.8345
0.8605 | -20.145 | 30.6547 | -905.607 | 908.0349 | 3.2853 | 2.5213 | -2.5383 | 5.2124 | 905.607 | -2.5213 | 80.0019
0.8623 | -21.5287 | 29.8736 | -904.2232 | 906.6257 | 3.2429 | 2.725 | -2.5807 | 5.3737 | 904.2232 | -2.725 | 80.1692
0.8641 | -22.9125 | 29.0247 | -902.7885 | 905.1655 | 3.2004 | 2.9373 | -2.6147 | 5.535 | 902.7885 | -2.9373 | 80.3366
0.8659 | -24.2283 | 27.208 | -900.5644 | 902.831 | 3.1325 | 3.1071 | -2.5807 | 5.6029 | 900.5644 | -3.1071 | 80.5039
0.8677 | -25.9092 | 25.1706 | -897.1432 | 899.3504 | 3.0731 | 3.2684 | -2.5468 | 5.6538 | 897.1432 | -3.2684 | 80.6713
0.8695 | -27.0722 | 23.8293 | -892.6779 | 894.8766 | 3.0646 | 3.4042 | -2.5298 | 5.7133 | 892.6779 | -3.4042 | 80.8386
0.8712 | -27.0467 | 23.6765 | -889.3246 | 891.5233 | 3.1156 | 3.4551 | -2.5213 | 5.7642 | 889.3246 | -3.4551 | 80.9967
0.873 | -25.8582 | 24.4066 | -884.2396 | 886.4807 | 3.1835 | 3.4806 | -2.5128 | 5.8066 | 884.2396 | -3.4806 | 81.164
0.8748 | -26.2233 | 23.651 | -881.4041 | 883.7217 | 3.2259 | 3.54 | -2.5043 | 5.8321 | 881.4041 | -3.54 | 81.3313
0.8766 | -27.4797 | 20.8241 | -880.7335 | 882.9067 | 3.1835 | 3.557 | -2.5128 | 5.8406 | 880.7335 | -3.557 | 81.4987
0.8784 | -29.1605 | 17.4878 | -881.0476 | 883.2208 | 3.1325 | 3.6079 | -2.5043 | 5.8151 | 881.0476 | -3.6079 | 81.666
0.8802 | -29.4407 | 15.3231 | -878.6367 | 880.7675 | 3.1071 | 3.5994 | -2.4958 | 5.7982 | 878.6367 | -3.5994 | 81.8334
0.882 | -29.4831 | 13.0819 | -876.0135 | 878.0594 | 3.1156 | 3.523 | -2.5298 | 5.7727 | 876.0135 | -3.523 | 82.0007
0.8838 | -29.8057 | 10.7644 | -873.3054 | 875.2664 | 3.0986 | 3.4466 | -2.5383 | 5.7472 | 873.3054 | -3.4466 | 82.1681
0.8856 | -31.3508 | 8.2855 | -870.5294 | 872.3801 | 2.9797 | 3.4636 | -2.4364 | 5.7048 | 870.5294 | -3.4636 | 82.3354
0.8874 | -31.546 | 6.698 | -865.4783 | 867.2356 | 2.9458 | 3.4042 | -2.377 | 5.6284 | 865.4783 | -3.4042 | 82.5028
0.8892 | -30.6122 | 5.8661 | -858.4917 | 860.1725 | 2.9797 | 3.2684 | -2.343 | 5.518 | 858.4917 | -3.2684 | 82.6701
0.891 | -29.1181 | 7.1564 | -848.9328 | 850.4524 | 3.0476 | 3.192 | -2.3176 | 5.518 | 848.9328 | -3.192 | 82.8375
0.8928 | -27.1826 | 9.092 | -841.0718 | 842.4131 | 3.0561 | 3.0816 | -2.2581 | 5.4501 | 841.0718 | -3.0816 | 83.0048
0.8946 | -23.8463 | 11.053 | -829.9848 | 831.0714 | 3.0816 | 2.8694 | -2.1987 | 5.3228 | 829.9848 | -2.8694 | 83.1722
0.8963 | -19.6866 | 14.8052 | -813.7364 | 814.5683 | 3.0986 | 2.6741 | -2.1817 | 5.2464 | 813.7364 | -2.6741 | 83.3302
0.8981 | -17.335 | 16.8936 | -800.0432 | 800.7648 | 3.0646 | 2.5468 | -2.1902 | 5.1275 | 800.0432 | -2.5468 | 83.4976
0.8999 | -16.537 | 17.6746 | -789.7543 | 790.4164 | 3.0137 | 2.4279 | -2.1987 | 4.9323 | 789.7543 | -2.4279 | 83.6649
0.9017 | -17.1313 | 17.3775 | -782.4196 | 783.0478 | 2.9458 | 2.3345 | -2.1902 | 4.72 | 782.4196 | -2.3345 | 83.8323
0.9035 | -18.4556 | 15.3655 | -778.5739 | 779.1342 | 2.9543 | 2.1817 | -2.2157 | 4.5672 | 778.5739 | -2.1817 | 83.9996
0.9053 | -19.0753 | 13.5234 | -775.2462 | 775.7725 | 3.0137 | 1.9865 | -2.2581 | 4.4823 | 775.2462 | -1.9865 | 84.167
0.9071 | -19.695 | 11.6812 | -771.9184 | 772.4023 | 3.0816 | 1.7997 | -2.2921 | 4.4059 | 771.9184 | -1.7997 | 84.3343
0.9089 | -21.308 | 10.5606 | -770.6705 | 771.1798 | 3.0476 | 1.7148 | -2.2921 | 4.321 | 770.6705 | -1.7148 | 84.5017
0.9107 | -22.2079 | 10.6455 | -769.3037 | 769.8385 | 3.0052 | 1.6299 | -2.2751 | 4.2446 | 769.3037 | -1.6299 | 84.669
0.9125 | -23.0313 | 9.491 | -770.3139 | 770.8487 | 2.9543 | 1.5111 | -2.2836 | 4.1682 | 770.3139 | -1.5111 | 84.8364
0.9143 | -22.2333 | 9.9324 | -769.9319 | 770.4073 | 2.9543 | 1.2734 | -2.3006 | 4.1173 | 769.9319 | -1.2734 | 85.0037
0.9161 | -21.4778 | 12.4707 | -769.1933 | 769.6602 | 3.0391 | 0.9508 | -2.3261 | 4.2022 | 769.1933 | -0.9508 | 85.1711
0.9179 | -20.7222 | 15.0175 | -768.4548 | 768.9132 | 3.1325 | 0.6282 | -2.36 | 4.2871 | 768.4548 | -0.6282 | 85.3384
0.9197 | -19.7714 | 16.1296 | -766.4004 | 766.8333 | 3.1835 | 0.30561 | -2.4364 | 4.4059 | 766.4004 | -0.30561 | 85.5058
0.9214 | -18.948 | 16.7663 | -764.7874 | 765.2119 | 3.2429 | -0.033957 | -2.5213 | 4.5842 | 764.7874 | 0.033957 | 85.6638
0.9232 | -18.2773 | 17.2501 | -763.8536 | 764.2696 | 3.2938 | -0.35655 | -2.6062 | 4.754 | 763.8536 | 0.35655 | 85.8312
0.925 | -18.515 | 17.0379 | -765.3902 | 765.8146 | 3.3448 | -0.6282 | -2.6911 | 4.8898 | 765.3902 | 0.6282 | 85.9985
0.9268 | -20.0601 | 16.2229 | -768.5312 | 768.9811 | 3.3702 | -0.87439 | -2.759 | 5.0087 | 768.5312 | 0.87439 | 86.1659
0.9286 | -22.0041 | 13.4385 | -773.7605 | 774.219 | 3.2853 | -1.1036 | -2.8524 | 5.0935 | 773.7605 | 1.1036 | 86.3332
0.9304 | -22.3437 | 12.3179 | -776.6978 | 777.1478 | 3.2599 | -1.3668 | -2.9543 | 5.2294 | 776.6978 | 1.3668 | 86.5006
0.9322 | -20.4761 | 13.9393 | -776.1715 | 776.6044 | 3.3533 | -1.6809 | -3.0476 | 5.4501 | 776.1715 | 1.6809 | 86.6679
0.934 | -18.549 | 17.2756 | -774.3633 | 774.8217 | 3.4976 | -1.9695 | -3.158 | 5.6878 | 774.3633 | 1.9695 | 86.8353
0.9358 | -16.8342 | 21.0279 | -773.6502 | 774.168 | 3.6589 | -2.2327 | -3.2684 | 5.934 | 773.6502 | 2.2327 | 87.0026
0.9376 | -14.8901 | 24.2283 | -774.4227 | 774.9915 | 3.7862 | -2.5128 | -3.3787 | 6.2056 | 774.4227 | 2.5128 | 87.17
0.9394 | -13.0225 | 26.809 | -776.3498 | 776.961 | 3.8881 | -2.793 | -3.4891 | 6.4688 | 776.3498 | 2.793 | 87.3373
0.9412 | -11.07 | 28.1079 | -780.7302 | 781.3669 | 3.922 | -3.0391 | -3.5994 | 6.7065 | 780.7302 | 3.0391 | 87.5046
0.943 | -10.6625 | 28.1588 | -785.7813 | 786.401 | 3.8881 | -3.1325 | -3.7607 | 6.8254 | 785.7813 | 3.1325 | 87.672
0.9447 | -9.8475 | 30.1623 | -785.6625 | 786.3671 | 3.922 | -3.158 | -3.9475 | 6.9357 | 785.6625 | 3.158 | 87.83
0.9465 | -8.9646 | 32.2761 | -789.8392 | 790.6287 | 3.973 | -3.2768 | -4.0579 | 7.0461 | 789.8392 | 3.2768 | 87.9974
0.9483 | -8.693 | 33.4816 | -795.6203 | 796.4608 | 3.9815 | -3.4721 | -4.1597 | 7.1819 | 795.6203 | 3.4721 | 88.1647
0.9501 | -8.7185 | 34.5767 | -801.8769 | 802.7513 | 3.973 | -3.6674 | -4.2871 | 7.2923 | 801.8769 | 3.6674 | 88.3321
0.9519 | -7.9884 | 36.1472 | -808.6259 | 809.5512 | 3.9899 | -3.8711 | -4.3974 | 7.4196 | 808.6259 | 3.8711 | 88.4994
0.9537 | -7.7761 | 36.9537 | -816.1898 | 817.1321 | 3.9899 | -4.0154 | -4.5078 | 7.5215 | 816.1898 | 4.0154 | 88.6668
0.9555 | -8.3534 | 36.5122 | -825.112 | 826.0203 | 3.956 | -4.0663 | -4.6181 | 7.5724 | 825.112 | 4.0663 | 88.8341
0.9573 | -6.5452 | 36.7839 | -835.8594 | 836.7422 | 3.973 | -4.304 | -4.6606 | 7.7677 | 835.8594 | 4.304 | 89.0015
0.9591 | -5.7727 | 35.7906 | -849.3742 | 850.1977 | 3.9475 | -4.4484 | -4.6946 | 7.878 | 849.3742 | 4.4484 | 89.1688
0.9609 | -4.6861 | 35.6718 | -867.7195 | 868.5175 | 3.956 | -4.6606 | -4.6861 | 8.0393 | 867.7195 | 4.6606 | 89.3362
0.9627 | -5.2803 | 34.9502 | -880.8184 | 881.5909 | 3.9305 | -4.72 | -4.737 | 8.0902 | 880.8184 | 4.72 | 89.5035
0.9645 | -6.3754 | 33.3457 | -893.4758 | 894.2144 | 3.8966 | -4.7625 | -4.771 | 8.1242 | 893.4758 | 4.7625 | 89.6709
0.9663 | -8.1327 | 29.3728 | -906.9058 | 907.5765 | 3.7692 | -4.8389 | -4.8219 | 8.1497 | 906.9058 | 4.8389 | 89.8382
0.9681 | -9.89 | 25.3913 | -920.3358 | 920.947 | 3.6419 | -4.9153 | -4.8643 | 8.1751 | 920.3358 | 4.9153 | 90.0056
0.9698 | -10.4248 | 23.6086 | -929.8947 | 930.4889 | 3.5909 | -5.0171 | -4.9323 | 8.2431 | 929.8947 | 5.0171 | 90.1636
0.9716 | -10.1786 | 23.2011 | -936.3295 | 936.9323 | 3.5909 | -5.119 | -5.0002 | 8.3449 | 936.3295 | 5.119 | 90.331
0.9734 | -10.2295 | 22.2079 | -942.0089 | 942.5776 | 3.574 | -5.2379 | -5.0681 | 8.4383 | 942.0089 | 5.2379 | 90.4983
0.9752 | -9.9239 | 21.2401 | -947.1364 | 947.6712 | 3.5655 | -5.3567 | -5.119 | 8.5487 | 947.1364 | 5.3567 | 90.6657
0.977 | -9.491 | 21.2571 | -950.6339 | 951.1518 | 3.6164 | -5.4161 | -5.1615 | 8.6336 | 950.6339 | 5.4161 | 90.833
0.9788 | -8.6505 | 21.7494 | -951.6866 | 952.1959 | 3.6843 | -5.4671 | -5.2209 | 8.7185 | 951.6866 | 5.4671 | 91.0004
0.9806 | -7.1564 | 23.2351 | -952.5525 | 953.0873 | 3.7777 | -5.535 | -5.2548 | 8.8203 | 952.5525 | 5.535 | 91.1677
0.9824 | -6.2311 | 23.5661 | -953.0534 | 953.5882 | 3.7947 | -5.5944 | -5.2888 | 8.8797 | 953.0534 | 5.5944 | 91.3351
0.9842 | -6.4009 | 22.7002 | -955.7954 | 956.2962 | 3.7522 | -5.5605 | -5.3312 | 8.8882 | 955.7954 | 5.5605 | 91.5024
0.986 | -5.3652 | 24.0246 | -956.7037 | 957.2046 | 3.8286 | -5.5774 | -5.3737 | 8.9646 | 956.7037 | 5.5774 | 91.6698
0.9878 | -3.3278 | 27.4882 | -953.9872 | 954.539 | 4.0069 | -5.6199 | -5.4246 | 9.1174 | 953.9872 | 5.6199 | 91.8371
0.9896 | -1.6469 | 31.2659 | -949.1738 | 949.8274 | 4.2107 | -5.6284 | -5.4841 | 9.2618 | 949.1738 | 5.6284 | 92.0045
0.9914 | -0.95928 | 33.4306 | -943.486 | 944.2161 | 4.3804 | -5.6284 | -5.5605 | 9.3806 | 943.486 | 5.6284 | 92.1718
0.9932 | -0.91684 | 34.4409 | -936.754 | 937.5265 | 4.5417 | -5.6199 | -5.6623 | 9.491 | 936.754 | 5.6199 | 92.3392
0.9949 | -0.26317 | 36.6481 | -927.9252 | 928.7571 | 4.7794 | -5.6538 | -5.7387 | 9.6692 | 927.9252 | 5.6538 | 92.4972
0.9967 | 0.73007 | 39.3307 | -919.6312 | 920.565 | 5.0341 | -5.7048 | -5.8066 | 9.8645 | 919.6312 | 5.7048 | 92.6646
0.9985 | 2.8015 | 41.6652 | -914.2745 | 915.3102 | 5.2294 | -5.8236 | -5.8236 | 10.0767 | 914.2745 | 5.8236 | 92.8319
1.0003 | 4.8728 | 43.9913 | -908.9178 | 910.0553 | 5.4331 | -5.934 | -5.8491 | 10.2805 | 908.9178 | 5.934 | 92.9993
1.0021 | 5.6029 | 45.9183 | -905.7004 | 906.9653 | 5.5859 | -5.9594 | -5.8746 | 10.4163 | 905.7004 | 5.9594 | 93.1666
1.0039 | 5.7302 | 46.0372 | -906.5323 | 907.8142 | 5.6453 | -5.9679 | -5.8915 | 10.4757 | 906.5323 | 5.9679 | 93.334
1.0057 | 5.5944 | 45.5533 | -908.4169 | 909.6818 | 5.6793 | -5.9594 | -5.9 | 10.5012 | 908.4169 | 5.9594 | 93.5013
1.0075 | 4.2956 | 43.4904 | -911.9824 | 913.1369 | 5.6623 | -5.8915 | -5.9425 | 10.4842 | 911.9824 | 5.8915 | 93.6687
1.0093 | 2.1902 | 40.5022 | -915.828 | 916.8637 | 5.6114 | -5.7557 | -6.0189 | 10.4418 | 915.828 | 5.7557 | 93.836
1.0111 | 1.4347 | 39.5853 | -920.8367 | 921.8469 | 5.6453 | -5.6793 | -6.0443 | 10.4333 | 920.8367 | 5.6793 | 94.0033
1.0129 | 0.32259 | 38.7025 | -925.5991 | 926.5924 | 5.6708 | -5.552 | -6.0953 | 10.4078 | 925.5991 | 5.552 | 94.1707
1.0147 | 0.33108 | 38.6855 | -928.1205 | 929.0967 | 5.7387 | -5.5265 | -6.1123 | 10.4418 | 928.1205 | 5.5265 | 94.338
1.0165 | -0.050935 | 38.0318 | -930.9389 | 931.8897 | 5.7897 | -5.4841 | -6.1377 | 10.4587 | 930.9389 | 5.4841 | 94.5054
1.0182 | -0.56878 | 37.5819 | -932.3311 | 933.2564 | 5.8321 | -5.501 | -6.1292 | 10.4927 | 932.3311 | 5.501 | 94.6634
1.02 | 0.64518 | 39.8994 | -931.1172 | 932.1189 | 5.9934 | -5.552 | -6.1377 | 10.6285 | 931.1172 | 5.552 | 94.8308
1.0218 | 1.163 | 41.6907 | -929.3344 | 930.421 | 6.1123 | -5.5944 | -6.1632 | 10.7559 | 929.3344 | 5.5944 | 94.9981
1.0236 | 0.37353 | 41.0964 | -932.2632 | 933.3074 | 6.1207 | -5.6029 | -6.1802 | 10.7728 | 932.2632 | 5.6029 | 95.1655
1.0254 | -1.562 | 38.2186 | -937.569 | 938.4773 | 6.0104 | -5.5859 | -6.1887 | 10.7219 | 937.569 | 5.5859 | 95.3328
1.0272 | -3.5825 | 35.2983 | -943.868 | 944.649 | 5.8915 | -5.5689 | -6.1887 | 10.6455 | 943.868 | 5.5689 | 95.5002
1.029 | -5.7302 | 31.8007 | -950.3198 | 950.982 | 5.7727 | -5.5265 | -6.1802 | 10.5606 | 950.3198 | 5.5265 | 95.6675
1.0308 | -7.9205 | 28.1928 | -956.4745 | 957.0178 | 5.6453 | -5.4756 | -6.1717 | 10.4757 | 956.4745 | 5.4756 | 95.8349
1.0326 | -8.9477 | 26.2742 | -960.5154 | 961.0078 | 5.5944 | -5.4586 | -6.1717 | 10.4333 | 960.5154 | 5.4586 | 96.0022
1.0344 | -9.6608 | 24.5254 | -963.1046 | 963.5461 | 5.535 | -5.4416 | -6.1802 | 10.3739 | 963.1046 | 5.4416 | 96.1696
1.0362 | -10.238 | 23.1417 | -963.8856 | 964.3016 | 5.501 | -5.4076 | -6.1887 | 10.3399 | 963.8856 | 5.4076 | 96.3369
1.038 | -10.654 | 22.2673 | -962.7735 | 963.1725 | 5.5265 | -5.3822 | -6.2056 | 10.3484 | 962.7735 | 5.3822 | 96.5043
1.0398 | -10.688 | 21.6306 | -961.0163 | 961.4153 | 5.5859 | -5.3482 | -6.2056 | 10.3569 | 961.0163 | 5.3482 | 96.6716
1.0416 | -11.2737 | 20.926 | -958.1384 | 958.5459 | 5.6708 | -5.3058 | -6.1971 | 10.3739 | 958.1384 | 5.3058 | 96.839
1.0433 | -12.5471 | 19.1432 | -958.0535 | 958.4695 | 5.7133 | -5.2633 | -6.1717 | 10.3654 | 958.0535 | 5.2633 | 96.997
1.0451 | -13.2432 | 18.0991 | -955.4983 | 955.9397 | 5.7727 | -5.2294 | -6.1717 | 10.3654 | 955.4983 | 5.2294 | 97.1644
1.0469 | -12.4962 | 18.948 | -951.0923 | 951.5423 | 5.8915 | -5.2294 | -6.1547 | 10.4163 | 951.0923 | 5.2294 | 97.3317
1.0487 | -12.649 | 18.2349 | -947.026 | 947.5099 | 5.9679 | -5.2803 | -6.1038 | 10.4418 | 947.026 | 5.2803 | 97.4991
1.0505 | -11.3162 | 19.9243 | -939.6319 | 940.0988 | 6.1207 | -5.3822 | -6.0528 | 10.5436 | 939.6319 | 5.3822 | 97.6664
1.0523 | -9.4825 | 22.7342 | -931.2954 | 931.7454 | 6.2736 | -5.4841 | -6.0189 | 10.671 | 931.2954 | 5.4841 | 97.8338
1.0541 | -8.3534 | 23.6341 | -926.1934 | 926.6348 | 6.3415 | -5.535 | -5.9934 | 10.7134 | 926.1934 | 5.535 | 98.0011
1.0559 | -8.4892 | 22.7681 | -923.5957 | 924.0201 | 6.333 | -5.552 | -5.9764 | 10.7049 | 923.5957 | 5.552 | 98.1685
1.0577 | -9.8051 | 19.7205 | -921.6686 | 922.1016 | 6.1971 | -5.5265 | -5.951 | 10.5861 | 921.6686 | 5.5265 | 98.3358
1.0595 | -10.2041 | 19.364 | -919.4954 | 919.9283 | 6.1802 | -5.501 | -5.9255 | 10.5436 | 919.4954 | 5.501 | 98.5032
1.0613 | -10.5097 | 19.1942 | -916.8552 | 917.2882 | 6.1717 | -5.4671 | -5.9085 | 10.5097 | 916.8552 | 5.4671 | 98.6705
1.0631 | -11.1634 | 17.8953 | -916.3459 | 916.7703 | 6.1462 | -5.4331 | -5.8915 | 10.4503 | 916.3459 | 5.4331 | 98.8379
1.0649 | -12.4283 | 15.2891 | -918.8247 | 919.2152 | 6.0698 | -5.3822 | -5.8576 | 10.3399 | 918.8247 | 5.3822 | 99.0052
1.0667 | -14.0073 | 13.2942 | -925.3699 | 925.735 | 5.9849 | -5.2633 | -5.8066 | 10.1871 | 925.3699 | 5.2633 | 99.1726
1.0684 | -14.7034 | 12.9291 | -931.5331 | 931.8812 | 5.9849 | -5.153 | -5.7472 | 10.0767 | 931.5331 | 5.153 | 99.3306
1.0702 | -14.7288 | 14.1006 | -930.1664 | 930.5229 | 6.0359 | -5.136 | -5.7133 | 10.0767 | 930.1664 | 5.136 | 99.498
1.072 | -14.0921 | 14.7034 | -927.5602 | 927.8912 | 6.0613 | -5.1784 | -5.6708 | 10.1022 | 927.5602 | 5.1784 | 99.6653
1.0738 | -13.447 | 15.3061 | -924.954 | 925.2596 | 6.0783 | -5.2294 | -5.6369 | 10.1192 | 924.954 | 5.2294 | 99.8327
1.0756 | -13.4639 | 15.7221 | -925.1917 | 925.5058 | 6.1377 | -5.2124 | -5.5944 | 10.1362 | 925.1917 | 5.2124 | 100
| Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My
Minimum | -56.1818 | -49.9677 | -1935.3352 | 632.542 | 2.9458 | -5.9679 | -6.2056 | 4.1173 | 631.5148 | -10.6031
Maximum | 13.0565 | 46.0372 | -631.5148 | 1935.87 | 12.6065 | 10.6031 | -2.1563 | 16.2739 | 1935.3352 | 5.9679

## Diagr_AVER75
AVER75 = Loads at Average Level in Subject with 75 kg Body Weight.
Contact forces and moments, measured telemetrically in knee implants of 8 subjects.
Averages of one individual subject.
Fx  lateral, Fy  anterior, -Fz  distal, Fres = Resultant Force
Mx around x, My around y, Mz around z, Mres = Resultant Moment
z-axis = axis of implant stem!
From Bergmann et al, PlosOne (2014) "Standardized Loads Acting in Knee Implants"
Subject K3R
Subject K3R
G.Bergmann et al., PlosOne, 2014

## Dat_HIGH100
Data File: | k3r_2239_0_3_i1.akf
Curve Alignment Program: | 1.5.7
------------------------MEASURING INFOS------------------------
Diagram Title #1 | Forces and Moments at Knee Joint
Diagram Title #2 | Measured Relative to Implant
Comment #1 | Knee Joint; Gaitanalysis; Deep Knee Bend;
Comment #2 | K3R, 8 Months PO *
BodyWeight [N]: | 960
Patient(en): | K3R | [7 empty] | Subject | K3R
Implant Type: | Knee Joint
Varus Angle Leg [°] | 3.5 | [7 empty] | HIGH100 Loads in N (forces) and Nm (moments)
Axes Directions X, Y, Z: | Lateral, Anterior, Superior
Axes Origin: | Tibial Plateau at Stem Axis | [7 empty] | = AVER75 Loads * 1/Caver
Transformation X Y Z (°): | 10 | 0.5 | 10.5
Displacement X, Y, Z (mm): | 0 | 0 | 0 | [5 empty] | Caver = | 0.57
Activity Code: | KneeActivities | 1P_1_15_3
---------------------------DATA--------------------------------
Number of Data Sets = | 601
Number of analog channels = | 0
Time-Offset(sec) = | 0
Time | Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My | Time
[s] | [N] | [N] | [N] | [N] | [Nm] | [Nm] | [Nm] | [Nm] | [N] | (Nm] | [%Cycle]
0 | -44.0845 | -17.1423 | -1370.7745 | 1371.817 | 10.2913 | -3.5446 | -9.1892 | 14.7743 | 1370.7745 | 3.5446 | 0
0.0018 | -44.4717 | -17.321 | -1363.8192 | 1364.8618 | 10.336 | -3.5 | -9.2637 | 14.8338 | 1363.8192 | 3.5 | 0.16735
0.0036 | -45.1717 | -18.6465 | -1362.2703 | 1363.3129 | 10.3062 | -3.4702 | -9.3084 | 14.8338 | 1362.2703 | 3.4702 | 0.3347
0.0054 | -45.976 | -20.0167 | -1364.683 | 1365.7703 | 10.2765 | -3.4404 | -9.3233 | 14.8189 | 1364.683 | 3.4404 | 0.50205
0.0072 | -47.5398 | -22.6678 | -1368.9575 | 1370.1936 | 10.202 | -3.3808 | -9.3233 | 14.7594 | 1368.9575 | 3.3808 | 0.66939
0.009 | -48.6866 | -23.6507 | -1377.8637 | 1379.1594 | 10.1424 | -3.3808 | -9.2339 | 14.6849 | 1377.8637 | 3.3808 | 0.83674
0.0108 | -48.5972 | -22.7571 | -1382.8381 | 1384.1636 | 10.1573 | -3.4106 | -9.1743 | 14.67 | 1382.8381 | 3.4106 | 1.0041
0.0125 | -49.2972 | -23.0252 | -1387.4253 | 1388.8104 | 10.1275 | -3.3957 | -9.1297 | 14.6402 | 1387.4253 | 3.3957 | 1.1621
0.0143 | -51.144 | -25.4677 | -1389.838 | 1391.4018 | 10.0382 | -3.3212 | -9.1297 | 14.5806 | 1389.838 | 3.3212 | 1.3295
0.0161 | -50.5333 | -24.6486 | -1388.0061 | 1389.5253 | 10.1275 | -3.3361 | -9.1446 | 14.6551 | 1388.0061 | 3.3361 | 1.4968
0.0179 | -50.4887 | -25.2592 | -1381.0509 | 1382.5105 | 10.1722 | -3.2914 | -9.219 | 14.6849 | 1381.0509 | 3.2914 | 1.6642
0.0197 | -53.1844 | -29.2804 | -1371.9212 | 1373.4404 | 10.0679 | -3.1425 | -9.3382 | 14.6253 | 1371.9212 | 3.1425 | 1.8315
0.0215 | -54.7035 | -32.0059 | -1369.866 | 1371.4447 | 10.0233 | -3.068 | -9.3828 | 14.5658 | 1369.866 | 3.068 | 1.9989
0.0233 | -55.1503 | -32.274 | -1369.3894 | 1370.9979 | 10.0828 | -3.0383 | -9.3977 | 14.6104 | 1369.3894 | 3.0383 | 2.1662
0.0251 | -56.1928 | -33.0336 | -1370.4021 | 1372.0702 | 10.1424 | -2.9936 | -9.3828 | 14.6402 | 1370.4021 | 2.9936 | 2.3336
0.0269 | -57.4439 | -34.0761 | -1371.7574 | 1373.4999 | 10.2169 | -2.9489 | -9.3828 | 14.6849 | 1371.7574 | 2.9489 | 2.5009
0.0287 | -58.2034 | -34.6867 | -1374.2297 | 1376.002 | 10.3062 | -2.9191 | -9.368 | 14.7445 | 1374.2297 | 2.9191 | 2.6683
0.0305 | -59.663 | -36.0272 | -1377.551 | 1379.4424 | 10.3509 | -2.8595 | -9.3382 | 14.7594 | 1377.551 | 2.8595 | 2.8356
0.0323 | -61.0779 | -37.2335 | -1383.9253 | 1385.8764 | 10.3509 | -2.8297 | -9.2488 | 14.6998 | 1383.9253 | 2.8297 | 3.003
0.0341 | -61.5396 | -37.4569 | -1389.2721 | 1391.2529 | 10.4105 | -2.8149 | -9.1743 | 14.6998 | 1389.2721 | 2.8149 | 3.1703
0.0359 | -62.0013 | -37.6803 | -1394.6188 | 1396.6294 | 10.4701 | -2.8 | -9.085 | 14.6998 | 1394.6188 | 2.8 | 3.3377
0.0376 | -63.431 | -39.6314 | -1397.2996 | 1399.3996 | 10.5296 | -2.7106 | -9.0701 | 14.7147 | 1397.2996 | 2.7106 | 3.4957
0.0394 | -64.28 | -40.5399 | -1397.0315 | 1399.1911 | 10.619 | -2.651 | -9.085 | 14.7743 | 1397.0315 | 2.651 | 3.6631
0.0412 | -64.1608 | -39.9143 | -1396.8677 | 1398.9975 | 10.7977 | -2.6212 | -9.085 | 14.9083 | 1396.8677 | 2.6212 | 3.8304
0.043 | -64.9502 | -41.0016 | -1398.089 | 1400.3081 | 10.9169 | -2.5617 | -9.085 | 14.9977 | 1398.089 | 2.5617 | 3.9978
0.0448 | -66.3948 | -42.9526 | -1400.0251 | 1402.3932 | 10.9764 | -2.5021 | -9.0999 | 15.0572 | 1400.0251 | 2.5021 | 4.1651
0.0466 | -67.9437 | -44.9036 | -1401.4996 | 1404.0463 | 11.0062 | -2.4425 | -9.0999 | 15.1019 | 1401.4996 | 2.4425 | 4.3325
0.0484 | -68.748 | -46.527 | -1401.8123 | 1404.4782 | 11.036 | -2.4127 | -9.0999 | 15.1168 | 1401.8123 | 2.4127 | 4.4998
0.0502 | -69.1948 | -47.8227 | -1401.2017 | 1403.9421 | 11.0807 | -2.3829 | -9.0999 | 15.1466 | 1401.2017 | 2.3829 | 4.6672
0.052 | -69.3586 | -48.7908 | -1397.4932 | 1400.2187 | 11.1998 | -2.3234 | -9.1297 | 15.2211 | 1397.4932 | 2.3234 | 4.8345
0.0538 | -70.7288 | -50.7865 | -1391.7146 | 1394.5443 | 11.3488 | -2.1447 | -9.219 | 15.3253 | 1391.7146 | 2.1447 | 5.0019
0.0556 | -72.1139 | -52.7673 | -1385.9211 | 1388.8699 | 11.4828 | -1.9808 | -9.3084 | 15.4147 | 1385.9211 | 1.9808 | 5.1692
0.0574 | -73.216 | -54.3907 | -1381.0807 | 1384.1338 | 11.6318 | -1.817 | -9.3828 | 15.5189 | 1381.0807 | 1.817 | 5.3366
0.0592 | -72.4415 | -53.8248 | -1381.3488 | 1384.3275 | 11.7956 | -1.8021 | -9.3382 | 15.5785 | 1381.3488 | 1.8021 | 5.5039
0.061 | -71.4884 | -53.2737 | -1381.6913 | 1384.6104 | 11.9445 | -1.7872 | -9.2786 | 15.6679 | 1381.6913 | 1.7872 | 5.6713
0.0627 | -71.399 | -53.2886 | -1382.6147 | 1385.504 | 12.153 | -1.6383 | -9.2935 | 15.8019 | 1382.6147 | 1.6383 | 5.8293
0.0645 | -71.965 | -54.7333 | -1380.1722 | 1383.1509 | 12.1381 | -1.5936 | -9.2935 | 15.787 | 1380.1722 | 1.5936 | 5.9967
0.0663 | -72.2181 | -55.7609 | -1376.3148 | 1379.3531 | 12.1679 | -1.534 | -9.3084 | 15.8168 | 1376.3148 | 1.534 | 6.164
0.0681 | -71.5033 | -55.5226 | -1367.349 | 1370.3574 | 12.3168 | -1.4298 | -9.3828 | 15.9508 | 1367.349 | 1.4298 | 6.3313
0.0699 | -69.5373 | -51.6801 | -1361.1533 | 1363.9086 | 12.6892 | -1.2808 | -9.3977 | 16.2487 | 1361.1533 | 1.2808 | 6.4987
0.0717 | -67.9735 | -49.1929 | -1350.2067 | 1352.813 | 12.8828 | -1.2213 | -9.4275 | 16.4274 | 1350.2067 | 1.2213 | 6.666
0.0735 | -66.6331 | -47.644 | -1338.1728 | 1340.6749 | 13.0168 | -1.1915 | -9.4722 | 16.5615 | 1338.1728 | 1.1915 | 6.8334
0.0753 | -66.0374 | -48.2993 | -1326.3623 | 1328.9091 | 13.0466 | -1.117 | -9.5467 | 16.6359 | 1326.3623 | 1.117 | 7.0007
0.0771 | -63.6842 | -47.2717 | -1316.1752 | 1318.5879 | 13.1658 | -1.1468 | -9.5765 | 16.7551 | 1316.1752 | 1.1468 | 7.1681
0.0789 | -60.2587 | -45.2313 | -1306.9413 | 1309.1604 | 13.3445 | -1.2213 | -9.5765 | 16.8891 | 1306.9413 | 1.2213 | 7.3354
0.0807 | -56.8332 | -43.176 | -1297.7074 | 1299.7329 | 13.5083 | -1.2957 | -9.5765 | 17.0381 | 1297.7074 | 1.2957 | 7.5028
0.0825 | -52.8418 | -41.0016 | -1293.5372 | 1295.3691 | 13.7168 | -1.4298 | -9.5318 | 17.187 | 1293.5372 | 1.4298 | 7.6701
0.0843 | -51.7099 | -39.8399 | -1294.4011 | 1296.2329 | 13.9104 | -1.4 | -9.4424 | 17.2912 | 1294.4011 | 1.4 | 7.8375
0.086 | -52.5291 | -41.8356 | -1294.2521 | 1296.2478 | 13.9551 | -1.3106 | -9.3977 | 17.2912 | 1294.2521 | 1.3106 | 7.9955
0.0878 | -53.8546 | -42.193 | -1296.3819 | 1298.5116 | 14.0445 | -1.2362 | -9.3382 | 17.3061 | 1296.3819 | 1.2362 | 8.1629
0.0896 | -55.6716 | -42.6845 | -1299.8074 | 1302.0712 | 14.1338 | -1.1468 | -9.2637 | 17.2912 | 1299.8074 | 1.1468 | 8.3302
0.0914 | -63.1778 | -49.3865 | -1309.5179 | 1312.2285 | 13.9849 | -0.84892 | -9.2339 | 17.0678 | 1309.5179 | 0.84892 | 8.4976
0.0932 | -71.1905 | -56.1035 | -1315.1178 | 1318.305 | 13.8509 | -0.49148 | -9.2488 | 16.8742 | 1315.1178 | 0.49148 | 8.6649
0.095 | -75.4649 | -59.5587 | -1313.0327 | 1316.6071 | 13.836 | -0.28297 | -9.2786 | 16.8296 | 1313.0327 | 0.28297 | 8.8323
0.0968 | -76.5968 | -60.08 | -1308.9817 | 1312.6753 | 13.8658 | -0.23829 | -9.2637 | 16.8147 | 1308.9817 | 0.23829 | 8.9996
0.0986 | -78.2053 | -61.6438 | -1304.2903 | 1308.1626 | 13.8509 | -0.16383 | -9.2637 | 16.77 | 1304.2903 | 0.16383 | 9.167
0.1004 | -80.0819 | -64.4736 | -1298.9286 | 1303.0839 | 13.7913 | -0.059574 | -9.2637 | 16.7104 | 1298.9286 | 0.059574 | 9.3343
0.1022 | -78.7266 | -64.7268 | -1292.7479 | 1296.8436 | 13.7913 | -0.10425 | -9.219 | 16.6955 | 1292.7479 | 0.10425 | 9.5017
0.104 | -77.9372 | -66.7374 | -1288.4288 | 1292.5841 | 13.7019 | -0.14893 | -9.1594 | 16.5913 | 1288.4288 | 0.14893 | 9.669
0.1058 | -78.2649 | -69.3288 | -1283.4991 | 1287.848 | 13.5828 | -0.13404 | -9.1297 | 16.4721 | 1283.4991 | 0.13404 | 9.8364
0.1076 | -77.2968 | -69.1203 | -1272.4631 | 1276.797 | 13.5381 | -0.11915 | -9.085 | 16.4125 | 1272.4631 | 0.11915 | 10.0037
0.1094 | -75.4351 | -67.7948 | -1266.5206 | 1270.6907 | 13.4636 | -0.16383 | -8.9509 | 16.2785 | 1266.5206 | 0.16383 | 10.1711
0.1111 | -73.8266 | -66.9459 | -1263.1249 | 1267.161 | 13.3743 | -0.19361 | -8.8169 | 16.1296 | 1263.1249 | 0.19361 | 10.3291
0.1129 | -73.4394 | -66.6033 | -1251.7314 | 1255.7824 | 13.2998 | -0.13404 | -8.7871 | 16.0402 | 1251.7314 | 0.13404 | 10.4965
0.1147 | -72.5011 | -64.3395 | -1235.6168 | 1239.5486 | 13.1956 | -0.014893 | -8.7126 | 15.9062 | 1235.6168 | 0.014893 | 10.6638
0.1165 | -71.9203 | -61.7928 | -1220.5893 | 1224.402 | 13.0615 | 0.11915 | -8.6382 | 15.7572 | 1220.5893 | -0.11915 | 10.8312
0.1183 | -70.6841 | -60.1247 | -1217.9383 | 1221.6021 | 12.853 | 0.17872 | -8.4446 | 15.4743 | 1217.9383 | -0.17872 | 10.9985
0.1201 | -67.6161 | -56.6098 | -1211.9511 | 1215.3022 | 12.7934 | 0.19361 | -8.2956 | 15.3551 | 1211.9511 | -0.19361 | 11.1659
0.1219 | -63.297 | -51.5312 | -1200.8258 | 1203.7747 | 12.853 | 0.17872 | -8.1914 | 15.3402 | 1200.8258 | -0.17872 | 11.3332
0.1237 | -58.3226 | -45.4547 | -1184.7409 | 1187.2579 | 12.9573 | 0.17872 | -8.1318 | 15.3998 | 1184.7409 | -0.17872 | 11.5006
0.1255 | -55.5971 | -40.7782 | -1168.3135 | 1170.607 | 12.9424 | 0.26808 | -8.0722 | 15.37 | 1168.3135 | -0.26808 | 11.6679
0.1273 | -53.8099 | -37.1442 | -1153.5094 | 1155.6243 | 12.9573 | 0.40212 | -8.0722 | 15.3998 | 1153.5094 | -0.40212 | 11.8353
0.1291 | -53.4376 | -34.5527 | -1141.1628 | 1143.1585 | 13.0764 | 0.53616 | -8.1616 | 15.5338 | 1141.1628 | -0.53616 | 12.0026
0.1309 | -53.0801 | -31.3804 | -1132.1374 | 1133.999 | 13.1807 | 0.65531 | -8.1914 | 15.6232 | 1132.1374 | -0.65531 | 12.17
0.1327 | -51.7248 | -27.672 | -1125.0183 | 1126.7311 | 13.2105 | 0.6851 | -8.1467 | 15.6381 | 1125.0183 | -0.6851 | 12.3373
0.1345 | -51.0546 | -25.2294 | -1117.7801 | 1119.4184 | 13.136 | 0.71488 | -8.102 | 15.5487 | 1117.7801 | -0.71488 | 12.5046
0.1362 | -52.7078 | -24.9614 | -1109.2164 | 1110.8994 | 13.0317 | 0.84892 | -8.1169 | 15.4891 | 1109.2164 | -0.84892 | 12.6627
0.138 | -55.895 | -25.4826 | -1107.9207 | 1109.7228 | 12.9275 | 1.0723 | -8.1169 | 15.4296 | 1107.9207 | -1.0723 | 12.83
0.1398 | -60.0204 | -27.2252 | -1111.8972 | 1113.9376 | 12.8083 | 1.3106 | -8.1169 | 15.3551 | 1111.8972 | -1.3106 | 12.9974
0.1416 | -63.6544 | -28.0145 | -1116.3057 | 1118.5099 | 12.6892 | 1.5191 | -8.0871 | 15.2806 | 1116.3057 | -1.5191 | 13.1647
0.1434 | -63.3715 | -28.0741 | -1124.0502 | 1126.1949 | 12.6445 | 1.4893 | -8.0127 | 15.2211 | 1124.0502 | -1.4893 | 13.3321
0.1452 | -65.3374 | -29.0719 | -1135.7118 | 1137.9756 | 12.5551 | 1.5489 | -7.9382 | 15.1168 | 1135.7118 | -1.5489 | 13.4994
0.147 | -68.0033 | -31.3208 | -1151.5584 | 1154.0009 | 12.3764 | 1.6085 | -7.8339 | 14.953 | 1151.5584 | -1.6085 | 13.6668
0.1488 | -69.9246 | -32.7953 | -1168.1496 | 1170.7113 | 12.2573 | 1.6383 | -7.6999 | 14.7892 | 1168.1496 | -1.6383 | 13.8341
0.1506 | -69.1501 | -33.5846 | -1180.6452 | 1183.1622 | 12.1679 | 1.5638 | -7.5361 | 14.6402 | 1180.6452 | -1.5638 | 14.0015
0.1524 | -67.8842 | -32.8251 | -1188.643 | 1191.0408 | 12.1083 | 1.4447 | -7.402 | 14.5211 | 1188.643 | -1.4447 | 14.1688
0.1542 | -66.1267 | -32.0804 | -1194.2876 | 1196.5514 | 12.0488 | 1.3404 | -7.2829 | 14.4019 | 1194.2876 | -1.3404 | 14.3362
0.156 | -65.5757 | -31.6783 | -1197.0428 | 1199.2769 | 11.9892 | 1.251 | -7.2084 | 14.3275 | 1197.0428 | -1.251 | 14.5035
0.1578 | -65.0842 | -31.1868 | -1196.8194 | 1199.0088 | 11.9445 | 1.1915 | -7.1488 | 14.253 | 1196.8194 | -1.1915 | 14.6709
0.1595 | -64.8608 | -31.0379 | -1194.4365 | 1196.6258 | 11.8998 | 1.1915 | -7.119 | 14.1934 | 1194.4365 | -1.1915 | 14.8289
0.1613 | -66.1416 | -32.6166 | -1191.8599 | 1194.1684 | 11.7807 | 1.2362 | -7.1339 | 14.119 | 1191.8599 | -1.2362 | 14.9963
0.1631 | -68.0182 | -34.0612 | -1188.36 | 1190.8472 | 11.7062 | 1.3255 | -7.1786 | 14.0594 | 1188.36 | -1.3255 | 15.1636
0.1649 | -69.8352 | -35.3123 | -1184.3834 | 1187.0494 | 11.6318 | 1.4149 | -7.2084 | 14.0147 | 1184.3834 | -1.4149 | 15.331
0.1667 | -68.5544 | -33.4655 | -1174.7474 | 1177.2793 | 11.7062 | 1.4 | -7.2233 | 14.0892 | 1174.7474 | -1.4 | 15.4983
0.1685 | -65.1438 | -30.1592 | -1162.6242 | 1164.888 | 11.8254 | 1.2808 | -7.2382 | 14.2232 | 1162.6242 | -1.2808 | 15.6657
0.1703 | -60.2438 | -25.8996 | -1148.4159 | 1150.3073 | 12.0041 | 1.0872 | -7.268 | 14.4168 | 1148.4159 | -1.0872 | 15.833
0.1721 | -53.7801 | -20.6274 | -1139.8819 | 1141.5202 | 12.153 | 0.81914 | -7.1935 | 14.5806 | 1139.8819 | -0.81914 | 16.0004
0.1739 | -48.3142 | -15.9211 | -1133.1948 | 1134.6544 | 12.2722 | 0.61063 | -7.1042 | 14.7147 | 1133.1948 | -0.61063 | 16.1677
0.1757 | -49.878 | -15.4594 | -1133.3288 | 1134.982 | 12.2126 | 0.6702 | -6.9999 | 14.6551 | 1133.3288 | -0.6702 | 16.3351
0.1775 | -52.3801 | -15.504 | -1135.3543 | 1137.216 | 12.1232 | 0.78935 | -6.9105 | 14.5955 | 1135.3543 | -0.78935 | 16.5024
0.1793 | -56.9077 | -18.4678 | -1144.2606 | 1146.4052 | 11.9147 | 0.96807 | -6.8063 | 14.387 | 1144.2606 | -0.96807 | 16.6698
0.1811 | -61.7928 | -22.5337 | -1158.737 | 1161.2987 | 11.6764 | 1.1766 | -6.6574 | 14.1785 | 1158.737 | -1.1766 | 16.8371
0.1829 | -64.1608 | -25.155 | -1175.2538 | 1178.0388 | 11.5871 | 1.3106 | -6.4786 | 14.0743 | 1175.2538 | -1.3106 | 17.0045
0.1846 | -65.2927 | -26.9124 | -1192.1429 | 1195.0322 | 11.5573 | 1.4149 | -6.2999 | 14.0296 | 1192.1429 | -1.4149 | 17.1625
0.1864 | -64.5629 | -26.9869 | -1202.8364 | 1205.5768 | 11.5573 | 1.4596 | -6.1361 | 13.9253 | 1202.8364 | -1.4596 | 17.3299
0.1882 | -62.9991 | -26.1677 | -1206.9023 | 1209.3299 | 11.5722 | 1.4893 | -6.002 | 13.8211 | 1206.9023 | -1.4893 | 17.4972
0.19 | -63.4459 | -27.24 | -1206.7236 | 1209.032 | 11.4977 | 1.6532 | -5.9574 | 13.687 | 1206.7236 | -1.6532 | 17.6646
0.1918 | -64.28 | -28.7741 | -1203.5364 | 1205.8151 | 11.3935 | 1.8468 | -5.9574 | 13.5381 | 1203.5364 | -1.8468 | 17.8319
0.1936 | -64.1012 | -28.9081 | -1196.0152 | 1198.2641 | 11.319 | 2.0553 | -5.9872 | 13.4487 | 1196.0152 | -2.0553 | 17.9993
0.1954 | -62.463 | -28.1485 | -1189.7302 | 1191.9195 | 11.2892 | 2.1744 | -5.9723 | 13.4041 | 1189.7302 | -2.1744 | 18.1666
0.1972 | -60.7651 | -28.2081 | -1187.4664 | 1189.5515 | 11.3041 | 2.2489 | -5.9276 | 13.4041 | 1187.4664 | -2.2489 | 18.334
0.199 | -56.446 | -26.7039 | -1186.6323 | 1188.5685 | 11.3786 | 2.2191 | -5.8382 | 13.419 | 1186.6323 | -2.2191 | 18.5013
0.2008 | -52.5142 | -25.2741 | -1185.8579 | 1187.7494 | 11.4828 | 2.2191 | -5.7489 | 13.4785 | 1185.8579 | -2.2191 | 18.6687
0.2026 | -48.9546 | -24.0231 | -1186.8111 | 1188.6727 | 11.602 | 2.2638 | -5.6297 | 13.5232 | 1186.8111 | -2.2638 | 18.836
0.2044 | -49.3121 | -25.4379 | -1191.7408 | 1193.5876 | 11.6318 | 2.5617 | -5.5255 | 13.4487 | 1191.7408 | -2.5617 | 19.0033
0.2062 | -51.0099 | -27.7166 | -1195.6875 | 1197.5343 | 11.6318 | 2.9042 | -5.4212 | 13.4041 | 1195.6875 | -2.9042 | 19.1707
0.208 | -52.7971 | -30.5464 | -1194.9875 | 1196.9088 | 11.6913 | 3.217 | -5.3616 | 13.4339 | 1194.9875 | -3.217 | 19.338
0.2097 | -55.0758 | -34.3889 | -1194.4961 | 1196.5216 | 11.7509 | 3.5893 | -5.2872 | 13.4785 | 1194.4961 | -3.5893 | 19.4961
0.2115 | -58.8736 | -37.8442 | -1196.2535 | 1198.5024 | 11.7956 | 3.9616 | -5.2127 | 13.5679 | 1196.2535 | -3.9616 | 19.6634
0.2133 | -64.5034 | -43.593 | -1204.6087 | 1207.2895 | 11.8105 | 4.4531 | -5.1829 | 13.7168 | 1204.6087 | -4.4531 | 19.8308
0.2151 | -71.5479 | -49.6695 | -1214.2298 | 1217.4468 | 11.8552 | 4.9893 | -5.1233 | 13.8956 | 1214.2298 | -4.9893 | 19.9981
0.2169 | -78.5478 | -56.2524 | -1229.0041 | 1232.9062 | 11.7807 | 5.4361 | -5.1084 | 13.9998 | 1229.0041 | -5.4361 | 20.1655
0.2187 | -86.5158 | -64.0119 | -1249.3932 | 1254.1442 | 11.6764 | 5.8978 | -5.0935 | 14.0892 | 1249.3932 | -5.8978 | 20.3328
0.2205 | -94.1859 | -71.5181 | -1275.5311 | 1281.1906 | 11.6615 | 6.3893 | -5.0638 | 14.3126 | 1275.5311 | -6.3893 | 20.5002
0.2223 | -97.2391 | -74.8543 | -1301.118 | 1307.1051 | 11.7807 | 6.6574 | -4.9893 | 14.5211 | 1301.118 | -6.6574 | 20.6675
0.2241 | -96.1221 | -73.8713 | -1316.8156 | 1322.5943 | 12.0786 | 6.7765 | -4.9595 | 14.804 | 1316.8156 | -6.7765 | 20.8349
0.2259 | -97.5071 | -76.4628 | -1334.7622 | 1340.7196 | 12.1977 | 6.9254 | -4.9446 | 14.9679 | 1334.7622 | -6.9254 | 21.0022
0.2277 | -98.5646 | -78.6223 | -1354.8534 | 1360.9448 | 12.3317 | 7.0595 | -4.9297 | 15.1317 | 1354.8534 | -7.0595 | 21.1696
0.2295 | -97.3433 | -78.1904 | -1375.9425 | 1381.8254 | 12.5403 | 7.1042 | -4.9148 | 15.3253 | 1375.9425 | -7.1042 | 21.3369
0.2313 | -95.4519 | -77.3415 | -1391.7891 | 1397.4188 | 12.7488 | 7.1488 | -4.8999 | 15.504 | 1391.7891 | -7.1488 | 21.5043
0.2331 | -95.6455 | -78.9946 | -1411.2845 | 1416.9589 | 12.8381 | 7.2531 | -4.8701 | 15.6083 | 1411.2845 | -7.2531 | 21.6716
0.2348 | -96.554 | -81.9137 | -1437.9437 | 1443.7522 | 12.8977 | 7.3425 | -4.8404 | 15.6679 | 1437.9437 | -7.3425 | 21.8297
0.2366 | -97.1348 | -84.9818 | -1469.7114 | 1475.6092 | 13.0019 | 7.3871 | -4.8255 | 15.7721 | 1469.7114 | -7.3871 | 21.997
0.2384 | -96.2263 | -86.7839 | -1500.6152 | 1506.4236 | 13.2105 | 7.402 | -4.7957 | 15.9359 | 1500.6152 | -7.402 | 22.1644
0.2402 | -93.7391 | -87.4094 | -1526.3808 | 1531.9658 | 13.419 | 7.4318 | -4.7659 | 16.1147 | 1526.3808 | -7.4318 | 22.3317
0.242 | -90.9243 | -87.6626 | -1550.7315 | 1556.0782 | 13.6275 | 7.4914 | -4.751 | 16.3083 | 1550.7315 | -7.4914 | 22.4991
0.2438 | -87.2009 | -85.7414 | -1566.9355 | 1571.9248 | 13.8509 | 7.5659 | -4.751 | 16.5466 | 1566.9355 | -7.5659 | 22.6664
0.2456 | -82.4201 | -83.5967 | -1579.2375 | 1583.7502 | 14.1934 | 7.5659 | -4.7212 | 16.8147 | 1579.2375 | -7.5659 | 22.8338
0.2474 | -77.9223 | -79.7542 | -1587.5629 | 1591.599 | 14.5806 | 7.5807 | -4.7361 | 17.1721 | 1587.5629 | -7.5807 | 23.0011
0.2492 | -73.916 | -76.15 | -1598.0032 | 1601.6372 | 14.9232 | 7.6403 | -4.7361 | 17.5146 | 1598.0032 | -7.6403 | 23.1685
0.251 | -70.6245 | -73.916 | -1615.2498 | 1618.6008 | 15.1317 | 7.7446 | -4.7212 | 17.738 | 1615.2498 | -7.7446 | 23.3358
0.2528 | -66.8565 | -71.9798 | -1633.4793 | 1636.5473 | 15.3551 | 7.7595 | -4.7063 | 17.9317 | 1633.4793 | -7.7595 | 23.5032
0.2546 | -63.0289 | -70.8182 | -1648.8642 | 1651.6939 | 15.5934 | 7.7595 | -4.7212 | 18.1551 | 1648.8642 | -7.7595 | 23.6705
0.2564 | -60.3928 | -70.148 | -1665.515 | 1668.2256 | 15.7721 | 7.7744 | -4.7361 | 18.3189 | 1665.515 | -7.7744 | 23.8379
0.2581 | -56.4311 | -67.7501 | -1683.2382 | 1685.77 | 15.9359 | 7.7297 | -4.7361 | 18.438 | 1683.2382 | -7.7297 | 23.9959
0.2599 | -53.6163 | -66.648 | -1699.7103 | 1702.1081 | 16.0849 | 7.7744 | -4.7361 | 18.5721 | 1699.7103 | -7.7744 | 24.1633
0.2617 | -53.2886 | -68.2565 | -1715.8696 | 1718.2973 | 16.1742 | 7.9978 | -4.7957 | 18.7508 | 1715.8696 | -7.9978 | 24.3306
0.2635 | -53.661 | -70.4607 | -1735.3502 | 1737.8821 | 16.2785 | 8.2956 | -4.8404 | 18.9593 | 1735.3502 | -8.2956 | 24.498
0.2653 | -54.0482 | -72.6649 | -1754.9797 | 1757.6159 | 16.353 | 8.5935 | -4.885 | 19.1529 | 1754.9797 | -8.5935 | 24.6653
0.2671 | -54.7482 | -75.2415 | -1778.7645 | 1781.5198 | 16.3381 | 8.8914 | -4.9148 | 19.3019 | 1778.7645 | -8.8914 | 24.8327
0.2689 | -54.495 | -75.9713 | -1801.0301 | 1803.8152 | 16.4572 | 9.1148 | -4.9744 | 19.5253 | 1801.0301 | -9.1148 | 25
0.2707 | -54.4652 | -75.867 | -1824.8894 | 1827.6596 | 16.5615 | 9.368 | -5.0042 | 19.7636 | 1824.8894 | -9.368 | 25.1673
0.2725 | -54.8077 | -74.3032 | -1851.5486 | 1854.2592 | 16.6359 | 9.6807 | -5.0191 | 20.0018 | 1851.5486 | -9.6807 | 25.3347
0.2743 | -56.4162 | -74.616 | -1883.2864 | 1885.9673 | 16.6657 | 9.9786 | -5.034 | 20.1955 | 1883.2864 | -9.9786 | 25.502
0.2761 | -57.6077 | -75.9415 | -1917.0349 | 1919.6859 | 16.6955 | 10.202 | -5.0787 | 20.3295 | 1917.0349 | -10.202 | 25.6694
0.2779 | -56.9524 | -75.9415 | -1949.8451 | 1952.377 | 16.77 | 10.3658 | -5.0935 | 20.4784 | 1949.8451 | -10.3658 | 25.8367
0.2797 | -55.2843 | -75.9266 | -1975.2532 | 1977.8 | 16.904 | 10.4552 | -5.1382 | 20.6274 | 1975.2532 | -10.4552 | 26.0041
0.2815 | -53.2737 | -76.4181 | -2002.7614 | 2005.2932 | 17.0976 | 10.5743 | -5.2276 | 20.8806 | 2002.7614 | -10.5743 | 26.1714
0.2832 | -52.7822 | -78.3393 | -2036.8822 | 2039.4438 | 17.2615 | 10.7381 | -5.2872 | 21.1189 | 2036.8822 | -10.7381 | 26.3295
0.285 | -51.2929 | -79.3074 | -2066.9966 | 2069.5285 | 17.4402 | 10.8573 | -5.3318 | 21.3125 | 2066.9966 | -10.8573 | 26.4968
0.2868 | -48.0163 | -78.1159 | -2088.8602 | 2091.2431 | 17.5891 | 10.902 | -5.3318 | 21.4614 | 2088.8602 | -10.902 | 26.6642
0.2886 | -45.3057 | -76.5074 | -2110.679 | 2112.913 | 17.738 | 11.0211 | -5.3169 | 21.655 | 2110.679 | -11.0211 | 26.8315
0.2904 | -42.5951 | -74.899 | -2132.4978 | 2134.5978 | 17.8721 | 11.1403 | -5.3021 | 21.8337 | 2132.4978 | -11.1403 | 26.9989
0.2922 | -39.6909 | -74.1245 | -2147.57 | 2149.5955 | 17.9615 | 11.2743 | -5.3169 | 21.9827 | 2147.57 | -11.2743 | 27.1662
0.294 | -35.9527 | -71.1607 | -2154.2422 | 2156.1039 | 18.0657 | 11.4679 | -5.3467 | 22.1912 | 2154.2422 | -11.4679 | 27.3336
0.2958 | -32.4825 | -64.6374 | -2154.7784 | 2156.4464 | 18.1551 | 11.6764 | -5.3467 | 22.3997 | 2154.7784 | -11.6764 | 27.5009
0.2976 | -30.6357 | -56.9077 | -2146.3338 | 2147.853 | 18.0508 | 11.8105 | -5.2872 | 22.3848 | 2146.3338 | -11.8105 | 27.6683
0.2994 | -29.9506 | -51.561 | -2132.0957 | 2133.5851 | 17.7678 | 11.9296 | -5.2276 | 22.2359 | 2132.0957 | -11.9296 | 27.8356
0.3012 | -29.8464 | -47.2419 | -2119.1534 | 2120.6129 | 17.4402 | 12.0786 | -5.1829 | 22.1018 | 2119.1534 | -12.0786 | 28.003
0.303 | -31.4102 | -47.5993 | -2124.2618 | 2125.7362 | 17.2466 | 12.302 | -5.2127 | 22.1167 | 2124.2618 | -12.302 | 28.1703
0.3048 | -31.723 | -46.11 | -2129.8915 | 2131.3659 | 17.1572 | 12.4956 | -5.2127 | 22.2061 | 2129.8915 | -12.4956 | 28.3377
0.3066 | -32.8549 | -45.5143 | -2136.2957 | 2137.785 | 17.1274 | 12.6743 | -5.2574 | 22.3252 | 2136.2957 | -12.6743 | 28.505
0.3083 | -34.3591 | -45.4547 | -2143.1466 | 2144.6509 | 17.0976 | 12.853 | -5.3021 | 22.4295 | 2143.1466 | -12.853 | 28.6631
0.3101 | -35.2378 | -45.9611 | -2149.3125 | 2150.8763 | 17.0678 | 13.0168 | -5.3616 | 22.5337 | 2149.3125 | -13.0168 | 28.8304
0.3119 | -37.6208 | -49.2078 | -2155.2401 | 2156.9081 | 16.9934 | 13.2105 | -5.4361 | 22.638 | 2155.2401 | -13.2105 | 28.9978
0.3137 | -41.1654 | -53.2886 | -2161.5698 | 2163.3272 | 16.8444 | 13.4339 | -5.4957 | 22.7422 | 2161.5698 | -13.4339 | 29.1651
0.3155 | -43.4441 | -55.3737 | -2167.8101 | 2169.642 | 16.6955 | 13.7317 | -5.5106 | 22.8316 | 2167.8101 | -13.7317 | 29.3325
0.3173 | -43.6973 | -54.1375 | -2168.0782 | 2169.8952 | 16.6508 | 13.97 | -5.5255 | 22.9358 | 2168.0782 | -13.97 | 29.4998
0.3191 | -43.7717 | -51.561 | -2169.359 | 2171.1909 | 16.6657 | 14.1487 | -5.5701 | 23.055 | 2169.359 | -14.1487 | 29.6672
0.3209 | -43.8313 | -48.9248 | -2170.759 | 2172.6058 | 16.6955 | 14.3275 | -5.5999 | 23.1593 | 2170.759 | -14.3275 | 29.8345
0.3227 | -41.8952 | -43.7419 | -2167.2591 | 2169.0761 | 16.7998 | 14.4019 | -5.6595 | 23.2635 | 2167.2591 | -14.4019 | 30.0019
0.3245 | -40.1228 | -37.725 | -2162.5974 | 2164.3995 | 16.8891 | 14.3572 | -5.7042 | 23.2784 | 2162.5974 | -14.3572 | 30.1692
0.3263 | -39.6314 | -34.2251 | -2160.7655 | 2162.5825 | 16.8593 | 14.387 | -5.7786 | 23.3082 | 2160.7655 | -14.387 | 30.3366
0.3281 | -39.4526 | -32.7804 | -2146.0806 | 2147.9125 | 16.6955 | 14.2083 | -5.8531 | 23.0848 | 2146.0806 | -14.2083 | 30.5039
0.3299 | -39.5271 | -32.691 | -2142.3573 | 2144.1743 | 16.6062 | 14.1041 | -5.9425 | 22.9805 | 2142.3573 | -14.1041 | 30.6713
0.3316 | -40.7484 | -33.4506 | -2147.6891 | 2149.6104 | 16.5615 | 14.1338 | -6.0318 | 22.9805 | 2147.6891 | -14.1338 | 30.8293
0.3334 | -41.5824 | -33.5549 | -2157.8762 | 2159.9017 | 16.5913 | 14.1785 | -6.1212 | 23.0401 | 2157.8762 | -14.1785 | 30.9967
0.3352 | -42.2675 | -34.5527 | -2171.9207 | 2174.0654 | 16.6062 | 14.2381 | -6.1957 | 23.0848 | 2171.9207 | -14.2381 | 31.164
0.337 | -43.0271 | -37.0548 | -2189.2566 | 2191.5055 | 16.6359 | 14.2828 | -6.2999 | 23.1295 | 2189.2566 | -14.2828 | 31.3313
0.3388 | -43.459 | -40.1228 | -2209.1542 | 2211.4478 | 16.7104 | 14.2977 | -6.4042 | 23.189 | 2209.1542 | -14.2977 | 31.4987
0.3406 | -42.7888 | -40.4803 | -2234.4581 | 2236.6326 | 16.8147 | 14.4466 | -6.4042 | 23.3529 | 2234.4581 | -14.4466 | 31.666
0.3424 | -40.1526 | -39.4973 | -2259.1067 | 2261.0429 | 17.0083 | 14.5062 | -6.4042 | 23.5465 | 2259.1067 | -14.5062 | 31.8334
0.3442 | -36.1612 | -37.4718 | -2284.917 | 2286.5702 | 17.321 | 14.4615 | -6.4191 | 23.7997 | 2284.917 | -14.4615 | 32.0007
0.346 | -32.1698 | -35.4612 | -2310.7273 | 2312.0975 | 17.6487 | 14.4168 | -6.434 | 24.0529 | 2310.7273 | -14.4168 | 32.1681
0.3478 | -32.6315 | -35.5208 | -2342.2566 | 2343.5226 | 17.887 | 14.4317 | -6.4637 | 24.2911 | 2342.2566 | -14.4317 | 32.3354
0.3496 | -32.6464 | -34.7016 | -2369.0201 | 2370.1371 | 18.0806 | 14.4019 | -6.4935 | 24.4699 | 2369.0201 | -14.4019 | 32.5028
0.3514 | -32.2442 | -34.508 | -2388.4411 | 2389.4538 | 18.1849 | 14.3721 | -6.5382 | 24.5741 | 2388.4411 | -14.3721 | 32.6701
0.3532 | -32.0804 | -34.8506 | -2403.7515 | 2404.6898 | 18.2146 | 14.3721 | -6.5829 | 24.6486 | 2403.7515 | -14.3721 | 32.8375
0.355 | -32.8102 | -36.1761 | -2416.1428 | 2417.0513 | 18.2146 | 14.4168 | -6.6276 | 24.6784 | 2416.1428 | -14.4168 | 33.0048
0.3567 | -31.8123 | -36.6676 | -2425.9874 | 2426.8363 | 18.2593 | 14.387 | -6.6871 | 24.7082 | 2425.9874 | -14.387 | 33.1629
0.3585 | -29.1166 | -36.1612 | -2432.4362 | 2433.2107 | 18.3338 | 14.2977 | -6.7169 | 24.738 | 2432.4362 | -14.2977 | 33.3302
0.3603 | -26.7635 | -35.6101 | -2437.2021 | 2437.917 | 18.4231 | 14.1934 | -6.702 | 24.7231 | 2437.2021 | -14.1934 | 33.4976
0.3621 | -26.6443 | -34.6421 | -2444.8722 | 2445.5722 | 18.4976 | 14.1487 | -6.6425 | 24.7231 | 2444.8722 | -14.1487 | 33.6649
0.3639 | -25.1699 | -33.1527 | -2448.3722 | 2449.0573 | 18.5572 | 14.0743 | -6.6425 | 24.7082 | 2448.3722 | -14.0743 | 33.8323
0.3657 | -23.6358 | -31.5591 | -2452.2444 | 2452.8998 | 18.6168 | 13.97 | -6.6425 | 24.6784 | 2452.2444 | -13.97 | 33.9996
0.3675 | -23.7699 | -30.0847 | -2465.0826 | 2465.7379 | 18.6317 | 13.9104 | -6.6276 | 24.6337 | 2465.0826 | -13.9104 | 34.167
0.3693 | -23.8741 | -28.0145 | -2478.9483 | 2479.5887 | 18.7359 | 13.8062 | -6.6722 | 24.6337 | 2478.9483 | -13.8062 | 34.3343
0.3711 | -23.7699 | -25.006 | -2488.778 | 2489.4035 | 18.8699 | 13.7019 | -6.7318 | 24.6635 | 2488.778 | -13.7019 | 34.5017
0.3729 | -23.055 | -21.372 | -2488.4652 | 2489.0907 | 19.0189 | 13.5083 | -6.851 | 24.6784 | 2488.4652 | -13.5083 | 34.669
0.3747 | -21.238 | -17.9168 | -2485.1291 | 2485.7546 | 19.138 | 13.27 | -6.9403 | 24.6337 | 2485.1291 | -13.27 | 34.8364
0.3765 | -20.8806 | -17.9317 | -2482.3589 | 2482.9695 | 19.0189 | 13.1509 | -6.9552 | 24.4848 | 2482.3589 | -13.1509 | 35.0037
0.3783 | -21.2976 | -19.138 | -2484.5334 | 2485.1291 | 18.8551 | 13.0168 | -6.9701 | 24.2763 | 2484.5334 | -13.0168 | 35.1711
0.3801 | -22.1763 | -20.5827 | -2492.7843 | 2493.3651 | 18.7359 | 12.8828 | -6.9552 | 24.0975 | 2492.7843 | -12.8828 | 35.3384
0.3818 | -22.7273 | -21.0891 | -2497.4013 | 2497.9821 | 18.6614 | 12.7637 | -6.9701 | 23.9635 | 2497.4013 | -12.7637 | 35.4965
0.3836 | -23.0252 | -22.221 | -2494.7055 | 2495.2864 | 18.587 | 12.7041 | -7.0297 | 23.9039 | 2494.7055 | -12.7041 | 35.6638
0.3854 | -23.9933 | -24.9465 | -2496.0906 | 2496.6566 | 18.4827 | 12.6743 | -7.0893 | 23.8295 | 2496.0906 | -12.6743 | 35.8312
0.3872 | -26.3762 | -28.8932 | -2508.8841 | 2509.4798 | 18.3487 | 12.719 | -7.119 | 23.7848 | 2508.8841 | -12.719 | 35.9985
0.389 | -28.9081 | -32.3038 | -2524.0307 | 2524.6562 | 18.304 | 12.7934 | -7.1935 | 23.8295 | 2524.0307 | -12.7934 | 36.1659
0.3908 | -31.7676 | -36.5782 | -2538.656 | 2539.3709 | 18.2295 | 12.9275 | -7.268 | 23.8592 | 2538.656 | -12.9275 | 36.3332
0.3926 | -34.7612 | -42.2228 | -2544.0921 | 2544.9559 | 18.1402 | 13.1658 | -7.4169 | 23.9933 | 2544.0921 | -13.1658 | 36.5006
0.3944 | -38.142 | -48.5376 | -2552.075 | 2553.1175 | 18.021 | 13.3892 | -7.551 | 24.0826 | 2552.075 | -13.3892 | 36.6679
0.3962 | -42.6845 | -55.7311 | -2567.0875 | 2568.4428 | 17.8274 | 13.6721 | -7.5807 | 24.172 | 2567.0875 | -13.6721 | 36.8353
0.398 | -47.3908 | -62.5225 | -2582.6213 | 2584.3341 | 17.6487 | 13.9998 | -7.5956 | 24.2614 | 2582.6213 | -13.9998 | 37.0026
0.3998 | -50.444 | -66.8714 | -2597.4403 | 2599.3615 | 17.5742 | 14.253 | -7.5956 | 24.3954 | 2597.4403 | -14.253 | 37.17
0.4016 | -52.112 | -69.1799 | -2618.6187 | 2620.6293 | 17.6487 | 14.4168 | -7.6254 | 24.5592 | 2618.6187 | -14.4168 | 37.3373
0.4034 | -53.2141 | -72.2032 | -2644.012 | 2646.0375 | 17.7083 | 14.5062 | -7.6999 | 24.6635 | 2644.012 | -14.5062 | 37.5046
0.4051 | -55.0163 | -75.0032 | -2673.5754 | 2675.6307 | 17.7678 | 14.4913 | -7.7595 | 24.7231 | 2673.5754 | -14.4913 | 37.6627
0.4069 | -57.2949 | -77.9521 | -2704.1218 | 2706.2366 | 17.8274 | 14.4615 | -7.8339 | 24.7677 | 2704.1218 | -14.4615 | 37.83
0.4087 | -58.6354 | -80.1265 | -2730.6469 | 2732.8214 | 17.9317 | 14.4466 | -7.9084 | 24.8422 | 2730.6469 | -14.4466 | 37.9974
0.4105 | -57.7715 | -79.4117 | -2752.5849 | 2754.7296 | 18.1551 | 14.4019 | -7.9829 | 25.006 | 2752.5849 | -14.4019 | 38.1647
0.4123 | -54.8375 | -75.5989 | -2778.425 | 2780.4356 | 18.5572 | 14.3275 | -7.9978 | 25.2294 | 2778.425 | -14.3275 | 38.3321
0.4141 | -51.6354 | -72.8586 | -2809.3884 | 2811.2352 | 18.9295 | 14.2679 | -7.9829 | 25.423 | 2809.3884 | -14.2679 | 38.4994
0.4159 | -49.9674 | -73.216 | -2845.1177 | 2846.9049 | 19.1827 | 14.253 | -7.9531 | 25.5869 | 2845.1177 | -14.253 | 38.6668
0.4177 | -51.7695 | -77.4159 | -2881.7555 | 2883.6022 | 19.2572 | 14.3275 | -7.9531 | 25.6911 | 2881.7555 | -14.3275 | 38.8341
0.4195 | -49.5504 | -76.984 | -2915.9358 | 2917.7379 | 19.5848 | 14.2828 | -7.9978 | 25.9145 | 2915.9358 | -14.2828 | 39.0015
0.4213 | -47.6738 | -76.6117 | -2951.5758 | 2953.3183 | 19.9125 | 14.253 | -8.0573 | 26.1677 | 2951.5758 | -14.253 | 39.1688
0.4231 | -45.2313 | -75.3606 | -2983.0009 | 2984.6838 | 20.2699 | 14.1934 | -8.1467 | 26.4209 | 2983.0009 | -14.1934 | 39.3362
0.4249 | -41.5675 | -72.1884 | -3009.1537 | 3010.643 | 20.6869 | 14.1487 | -8.2361 | 26.7188 | 3009.1537 | -14.1487 | 39.5035
0.4267 | -38.0229 | -70.0139 | -3029.2598 | 3030.6598 | 20.9401 | 13.9551 | -8.2658 | 26.8379 | 3029.2598 | -13.9551 | 39.6709
0.4285 | -37.1739 | -70.3416 | -3050.0659 | 3051.4212 | 21.104 | 13.9253 | -8.3254 | 26.972 | 3050.0659 | -13.9253 | 39.8382
0.4302 | -35.3421 | -68.0927 | -3062.3827 | 3063.604 | 21.3125 | 13.9253 | -8.3403 | 27.1507 | 3062.3827 | -13.9253 | 39.9963
0.432 | -31.6336 | -64.3842 | -3065.0337 | 3066.0763 | 21.5508 | 13.8807 | -8.3552 | 27.2996 | 3065.0337 | -13.8807 | 40.1636
0.4338 | -28.8634 | -61.6587 | -3064.7508 | 3065.7039 | 21.7146 | 13.8807 | -8.385 | 27.4337 | 3064.7508 | -13.8807 | 40.331
0.4356 | -26.1528 | -58.1737 | -3058.7487 | 3059.6125 | 21.8784 | 13.8807 | -8.3999 | 27.5826 | 3058.7487 | -13.8807 | 40.4983
0.4374 | -24.6784 | -55.7609 | -3049.0233 | 3049.8574 | 21.9231 | 13.836 | -8.4595 | 27.6124 | 3049.0233 | -13.836 | 40.6657
0.4392 | -22.7422 | -54.3609 | -3043.1702 | 3043.9894 | 21.9231 | 13.836 | -8.4595 | 27.6124 | 3043.1702 | -13.836 | 40.833
0.441 | -20.4933 | -53.3035 | -3039.983 | 3040.7724 | 21.9082 | 13.9402 | -8.4595 | 27.6422 | 3039.983 | -13.9402 | 41.0004
0.4428 | -17.5891 | -50.3248 | -3031.1363 | 3031.881 | 21.9976 | 14.0445 | -8.5041 | 27.7762 | 3031.1363 | -14.0445 | 41.1677
0.4446 | -14.9381 | -47.5398 | -3024.0471 | 3024.7471 | 22.1167 | 14.1338 | -8.5339 | 27.8954 | 3024.0471 | -14.1338 | 41.3351
0.4464 | -15.6083 | -48.3589 | -3025.313 | 3026.0428 | 22.072 | 14.3721 | -8.5488 | 27.9698 | 3025.313 | -14.3721 | 41.5024
0.4482 | -19.4359 | -53.0503 | -3033.1767 | 3034.0555 | 21.8635 | 14.7147 | -8.5488 | 27.9698 | 3033.1767 | -14.7147 | 41.6698
0.45 | -21.6848 | -55.5077 | -3045.7468 | 3046.6851 | 21.7742 | 14.953 | -8.5041 | 27.9847 | 3045.7468 | -14.953 | 41.8371
0.4518 | -25.1252 | -59.1268 | -3062.7253 | 3063.7976 | 21.6103 | 15.236 | -8.4595 | 27.9847 | 3062.7253 | -15.236 | 42.0045
0.4536 | -26.8379 | -59.4098 | -3078.7804 | 3079.8527 | 21.5359 | 15.4594 | -8.4148 | 28.0145 | 3078.7804 | -15.4594 | 42.1718
0.4553 | -27.4634 | -58.6354 | -3091.8717 | 3092.9142 | 21.521 | 15.5785 | -8.385 | 28.089 | 3091.8717 | -15.5785 | 42.3299
0.4571 | -28.1932 | -58.1588 | -3100.3013 | 3101.2992 | 21.5061 | 15.7125 | -8.385 | 28.1485 | 3100.3013 | -15.7125 | 42.4972
0.4589 | -27.6273 | -56.1184 | -3102.8481 | 3103.7417 | 21.5657 | 15.787 | -8.385 | 28.223 | 3102.8481 | -15.787 | 42.6646
0.4607 | -26.6592 | -53.9588 | -3102.7141 | 3103.5183 | 21.6252 | 15.8615 | -8.3701 | 28.2975 | 3102.7141 | -15.8615 | 42.8319
0.4625 | -26.5401 | -53.1993 | -3101.7013 | 3102.4609 | 21.6401 | 15.9657 | -8.3701 | 28.3422 | 3101.7013 | -15.9657 | 42.9993
0.4643 | -26.9124 | -52.812 | -3103.563 | 3104.2928 | 21.6252 | 16.07 | -8.385 | 28.4017 | 3103.563 | -16.07 | 43.1666
0.4661 | -28.5953 | -53.6163 | -3113.1543 | 3113.9139 | 21.5657 | 16.204 | -8.3999 | 28.4315 | 3113.1543 | -16.204 | 43.334
0.4679 | -28.2975 | -52.4397 | -3126.901 | 3127.6456 | 21.5359 | 16.2189 | -8.4743 | 28.4464 | 3126.901 | -16.2189 | 43.5013
0.4697 | -28.2528 | -50.3844 | -3137.1178 | 3137.8476 | 21.5657 | 16.2338 | -8.5339 | 28.4762 | 3137.1178 | -16.2338 | 43.6687
0.4715 | -27.7464 | -49.4163 | -3141.6603 | 3142.4497 | 21.521 | 16.1742 | -8.519 | 28.4315 | 3141.6603 | -16.1742 | 43.836
0.4733 | -28.1336 | -48.9397 | -3142.822 | 3143.6263 | 21.4465 | 16.1742 | -8.519 | 28.3868 | 3142.822 | -16.1742 | 44.0033
0.4751 | -31.6634 | -51.9631 | -3143.8199 | 3144.6837 | 21.238 | 16.3232 | -8.5041 | 28.2975 | 3143.8199 | -16.3232 | 44.1707
0.4769 | -34.374 | -53.229 | -3138.488 | 3139.3667 | 21.1189 | 16.4572 | -8.5637 | 28.2677 | 3138.488 | -16.4572 | 44.338
0.4786 | -34.2102 | -52.3801 | -3123.9521 | 3124.8159 | 21.0742 | 16.5168 | -8.5935 | 28.2826 | 3123.9521 | -16.5168 | 44.4961
0.4804 | -33.3166 | -50.3993 | -3106.8991 | 3107.7183 | 21.0593 | 16.5764 | -8.6233 | 28.2975 | 3106.8991 | -16.5764 | 44.6634
0.4822 | -32.5868 | -47.8227 | -3086.5696 | 3087.3441 | 21.0742 | 16.6359 | -8.6382 | 28.3422 | 3086.5696 | -16.6359 | 44.8308
0.484 | -31.8272 | -45.2313 | -3066.2997 | 3067.0295 | 21.0742 | 16.7104 | -8.6531 | 28.3868 | 3066.2997 | -16.7104 | 44.9981
0.4858 | -31.9166 | -44.1143 | -3054.1169 | 3054.8615 | 21.0444 | 16.7998 | -8.668 | 28.4166 | 3054.1169 | -16.7998 | 45.1655
0.4876 | -31.2315 | -42.9377 | -3046.4766 | 3047.2063 | 21.0295 | 16.8444 | -8.6531 | 28.4315 | 3046.4766 | -16.8444 | 45.3328
0.4894 | -30.9932 | -43.0866 | -3043.1553 | 3043.9 | 20.9848 | 16.9189 | -8.6233 | 28.4315 | 3043.1553 | -16.9189 | 45.5002
0.4912 | -32.5123 | -44.993 | -3045.2255 | 3046.0595 | 20.8806 | 17.053 | -8.5935 | 28.4315 | 3045.2255 | -17.053 | 45.6675
0.493 | -34.508 | -47.5993 | -3050.1999 | 3051.1084 | 20.7614 | 17.2168 | -8.5488 | 28.4315 | 3050.1999 | -17.2168 | 45.8349
0.4948 | -35.0144 | -48.6419 | -3055.5318 | 3056.4402 | 20.7167 | 17.3061 | -8.4892 | 28.4315 | 3055.5318 | -17.3061 | 46.0022
0.4966 | -37.6505 | -51.7695 | -3062.6657 | 3063.5891 | 20.5827 | 17.4551 | -8.4595 | 28.3868 | 3062.6657 | -17.4551 | 46.1696
0.4984 | -37.6803 | -52.0674 | -3060.4764 | 3061.3849 | 20.5678 | 17.5295 | -8.4446 | 28.4017 | 3060.4764 | -17.5295 | 46.3369
0.5002 | -40.1377 | -53.1099 | -3054.6381 | 3055.5615 | 20.4933 | 17.738 | -8.4892 | 28.4762 | 3054.6381 | -17.738 | 46.5043
0.502 | -38.4548 | -49.2674 | -3047.0723 | 3047.9361 | 20.5976 | 17.738 | -8.4743 | 28.5507 | 3047.0723 | -17.738 | 46.6716
0.5037 | -38.8867 | -48.9546 | -3043.8106 | 3044.6745 | 20.538 | 17.7827 | -8.4595 | 28.5358 | 3043.8106 | -17.7827 | 46.8297
0.5055 | -44.0547 | -54.8673 | -3046.3276 | 3047.3553 | 20.1955 | 18.0061 | -8.4892 | 28.4315 | 3046.3276 | -18.0061 | 46.997
0.5073 | -47.227 | -58.9332 | -3045.7766 | 3046.9234 | 19.9721 | 18.1253 | -8.5041 | 28.3571 | 3045.7766 | -18.1253 | 47.1644
0.5091 | -46.8845 | -58.9481 | -3038.6575 | 3039.8043 | 19.9125 | 18.1253 | -8.5339 | 28.3273 | 3038.6575 | -18.1253 | 47.3317
0.5109 | -47.1227 | -58.6056 | -3030.7044 | 3031.8363 | 19.838 | 18.1253 | -8.5488 | 28.2677 | 3030.7044 | -18.1253 | 47.4991
0.5127 | -47.5844 | -58.0396 | -3021.3663 | 3022.4982 | 19.7636 | 18.1253 | -8.5935 | 28.2379 | 3021.3663 | -18.1253 | 47.6664
0.5145 | -44.9185 | -53.7354 | -3007.3069 | 3008.3345 | 19.8529 | 18.021 | -8.5935 | 28.223 | 3007.3069 | -18.021 | 47.8338
0.5163 | -43.4888 | -51.844 | -2990.0901 | 2991.0731 | 19.8082 | 17.9466 | -8.5786 | 28.1485 | 2990.0901 | -17.9466 | 48.0011
0.5181 | -43.9504 | -52.4546 | -2971.9053 | 2972.9329 | 19.6593 | 17.9615 | -8.5935 | 28.0443 | 2971.9053 | -17.9615 | 48.1685
0.5199 | -43.7122 | -51.6801 | -2954.8523 | 2955.88 | 19.5699 | 17.9615 | -8.5935 | 27.9847 | 2954.8523 | -17.9615 | 48.3358
0.5217 | -41.7015 | -46.661 | -2941.8206 | 2942.8035 | 19.6742 | 17.9019 | -8.5935 | 28.0294 | 2941.8206 | -17.9019 | 48.5032
0.5235 | -38.708 | -43.0122 | -2933.7186 | 2934.6568 | 19.704 | 17.738 | -8.5637 | 27.94 | 2933.7186 | -17.738 | 48.6705
0.5253 | -40.3314 | -45.2164 | -2929.742 | 2930.7846 | 19.5253 | 17.7976 | -8.5637 | 27.8656 | 2929.742 | -17.7976 | 48.8379
0.5271 | -40.0782 | -45.1568 | -2927.5825 | 2928.625 | 19.5104 | 17.8274 | -8.5339 | 27.8656 | 2927.5825 | -17.8274 | 49.0052
0.5288 | -39.3633 | -44.427 | -2924.8421 | 2925.8697 | 19.5402 | 17.8423 | -8.5041 | 27.8805 | 2924.8421 | -17.8423 | 49.1633
0.5306 | -36.6676 | -40.6888 | -2919.5698 | 2920.4932 | 19.6891 | 17.7976 | -8.4595 | 27.94 | 2919.5698 | -17.7976 | 49.3306
0.5324 | -33.9719 | -36.9505 | -2914.2976 | 2915.1316 | 19.8231 | 17.738 | -8.4148 | 27.9996 | 2914.2976 | -17.738 | 49.498
0.5342 | -31.6336 | -34.2697 | -2910.6785 | 2911.3933 | 19.8678 | 17.5891 | -8.385 | 27.8954 | 2910.6785 | -17.5891 | 49.6653
0.536 | -31.7825 | -33.5102 | -2911.7806 | 2912.4955 | 19.8529 | 17.5891 | -8.3403 | 27.8805 | 2911.7806 | -17.5891 | 49.8327
0.5378 | -32.9442 | -33.1676 | -2912.9274 | 2913.6422 | 19.8529 | 17.7083 | -8.2956 | 27.9251 | 2912.9274 | -17.7083 | 50
0.5396 | -33.1825 | -32.6464 | -2914.5358 | 2915.2358 | 19.8529 | 17.738 | -8.2509 | 27.94 | 2914.5358 | -17.738 | 50.1673
0.5414 | -32.423 | -32.3783 | -2918.3486 | 2919.0486 | 19.838 | 17.6785 | -8.2212 | 27.8805 | 2918.3486 | -17.6785 | 50.3347
0.5432 | -35.2676 | -36.6229 | -2928.8484 | 2929.6527 | 19.6444 | 17.738 | -8.2212 | 27.7911 | 2928.8484 | -17.738 | 50.502
0.545 | -36.8165 | -37.9186 | -2940.078 | 2940.8823 | 19.5848 | 17.7827 | -8.1914 | 27.7613 | 2940.078 | -17.7827 | 50.6694
0.5468 | -38.3356 | -39.4377 | -2952.5289 | 2953.3481 | 19.5253 | 17.7976 | -8.1765 | 27.7166 | 2952.5289 | -17.7976 | 50.8367
0.5486 | -43.3845 | -46.7653 | -2972.6648 | 2973.6031 | 19.2274 | 17.9168 | -8.1765 | 27.5677 | 2972.6648 | -17.9168 | 51.0041
0.5504 | -48.1951 | -53.929 | -2993.3369 | 2994.3794 | 18.9444 | 18.0061 | -8.1914 | 27.4337 | 2993.3369 | -18.0061 | 51.1714
0.5521 | -50.6823 | -58.1439 | -3010.2558 | 3011.3728 | 18.8104 | 18.0359 | -8.1914 | 27.3443 | 3010.2558 | -18.0359 | 51.3295
0.5539 | -52.8716 | -62.3438 | -3026.1471 | 3027.3683 | 18.6465 | 18.0657 | -8.1616 | 27.2549 | 3026.1471 | -18.0657 | 51.4968
0.5557 | -54.1971 | -66.1565 | -3040.8171 | 3042.0979 | 18.5423 | 18.0657 | -8.1914 | 27.1954 | 3040.8171 | -18.0657 | 51.6642
0.5575 | -55.9248 | -69.5224 | -3055.9041 | 3057.2743 | 18.4976 | 18.1402 | -8.2956 | 27.24 | 3055.9041 | -18.1402 | 51.8315
0.5593 | -56.8481 | -71.5628 | -3074.2528 | 3075.7123 | 18.5125 | 18.17 | -8.3701 | 27.2996 | 3074.2528 | -18.17 | 51.9989
0.5611 | -55.9396 | -71.2203 | -3094.9993 | 3096.4737 | 18.6763 | 18.0806 | -8.4595 | 27.4039 | 3094.9993 | -18.0806 | 52.1662
0.5629 | -56.9226 | -73.633 | -3121.3159 | 3122.8797 | 18.721 | 18.0955 | -8.519 | 27.4783 | 3121.3159 | -18.0955 | 52.3336
0.5647 | -58.7992 | -76.1649 | -3156.375 | 3157.9984 | 18.8402 | 18.0806 | -8.6233 | 27.6124 | 3156.375 | -18.0806 | 52.5009
0.5665 | -58.963 | -76.2989 | -3189.7065 | 3191.315 | 19.0189 | 18.0359 | -8.668 | 27.7315 | 3189.7065 | -18.0359 | 52.6683
0.5683 | -57.6971 | -75.1968 | -3212.7913 | 3214.3551 | 19.1827 | 17.9466 | -8.6829 | 27.806 | 3212.7913 | -17.9466 | 52.8356
0.5701 | -55.1205 | -73.3351 | -3232.8527 | 3234.3569 | 19.3614 | 17.7976 | -8.6978 | 27.8656 | 3232.8527 | -17.7976 | 53.003
0.5719 | -53.0801 | -72.0245 | -3250.4716 | 3251.946 | 19.5104 | 17.6636 | -8.7275 | 27.9102 | 3250.4716 | -17.6636 | 53.1703
0.5737 | -50.2355 | -68.9714 | -3260.5097 | 3261.865 | 19.704 | 17.5593 | -8.7573 | 27.9847 | 3260.5097 | -17.5593 | 53.3377
0.5755 | -47.2717 | -66.365 | -3265.1416 | 3266.348 | 19.8827 | 17.3955 | -8.7871 | 28.0145 | 3265.1416 | -17.3955 | 53.505
0.5772 | -44.9334 | -64.4736 | -3270.5926 | 3271.6947 | 20.0316 | 17.2615 | -8.8318 | 28.0443 | 3270.5926 | -17.2615 | 53.6631
0.579 | -44.3973 | -63.58 | -3282.9094 | 3283.9817 | 20.0912 | 17.2466 | -8.8318 | 28.0741 | 3282.9094 | -17.2466 | 53.8304
0.5808 | -46.4674 | -65.9033 | -3296.7752 | 3297.9518 | 20.0018 | 17.3359 | -8.8467 | 28.0741 | 3296.7752 | -17.3359 | 53.9978
0.5826 | -45.9313 | -65.7693 | -3306.4559 | 3307.6623 | 20.0614 | 17.2615 | -8.8765 | 28.1039 | 3306.4559 | -17.2615 | 54.1651
0.5844 | -45.0377 | -65.2033 | -3315.2132 | 3316.4643 | 20.1508 | 17.187 | -8.9212 | 28.1336 | 3315.2132 | -17.187 | 54.3325
0.5862 | -43.0717 | -63.014 | -3321.3493 | 3322.5557 | 20.2848 | 17.0827 | -8.9658 | 28.1932 | 3321.3493 | -17.0827 | 54.4998
0.588 | -40.0186 | -58.8439 | -3332.6087 | 3333.6513 | 20.4487 | 16.9934 | -8.9509 | 28.2528 | 3332.6087 | -16.9934 | 54.6672
0.5898 | -41.9994 | -61.4204 | -3344.8809 | 3346.0277 | 20.3295 | 17.0232 | -8.9658 | 28.2081 | 3344.8809 | -17.0232 | 54.8345
0.5916 | -44.1143 | -63.714 | -3351.1361 | 3352.417 | 20.2104 | 17.053 | -8.9807 | 28.1634 | 3351.1361 | -17.053 | 55.0019
0.5934 | -44.3973 | -62.9544 | -3345.3873 | 3346.683 | 20.1508 | 17.0083 | -8.9956 | 28.089 | 3345.3873 | -17.0083 | 55.1692
0.5952 | -44.3824 | -61.8672 | -3338.1491 | 3339.4001 | 20.0614 | 16.9636 | -8.9807 | 27.9847 | 3338.1491 | -16.9636 | 55.3366
0.597 | -44.3675 | -61.2715 | -3327.7684 | 3328.9896 | 19.9572 | 16.904 | -8.9956 | 27.8656 | 3327.7684 | -16.904 | 55.5039
0.5988 | -44.3675 | -61.5545 | -3312.0707 | 3313.3069 | 19.7933 | 16.8296 | -9.0105 | 27.7166 | 3312.0707 | -16.8296 | 55.6713
0.6006 | -44.6951 | -62.18 | -3294.5263 | 3295.7624 | 19.6146 | 16.8147 | -9.0105 | 27.5677 | 3294.5263 | -16.8147 | 55.8386
0.6023 | -43.9504 | -61.4949 | -3274.2712 | 3275.4776 | 19.4657 | 16.8147 | -8.9956 | 27.4337 | 3274.2712 | -16.8147 | 55.9967
0.6041 | -45.7972 | -62.3736 | -3255.997 | 3257.1885 | 19.2721 | 16.9189 | -8.9807 | 27.3145 | 3255.997 | -16.9189 | 56.164
0.6059 | -49.3568 | -62.9991 | -3237.0675 | 3238.2441 | 19.1231 | 17.053 | -8.9807 | 27.2698 | 3237.0675 | -17.053 | 56.3313
0.6077 | -51.144 | -62.6566 | -3223.3805 | 3224.5124 | 19.0487 | 17.1125 | -8.9509 | 27.2252 | 3223.3805 | -17.1125 | 56.4987
0.6095 | -50.6674 | -62.2991 | -3223.9613 | 3225.0634 | 19.0338 | 17.0232 | -8.8914 | 27.1507 | 3223.9613 | -17.0232 | 56.666
0.6113 | -53.7652 | -66.3055 | -3233.2548 | 3234.4761 | 18.8848 | 17.0827 | -8.8616 | 27.0464 | 3233.2548 | -17.0827 | 56.8334
0.6131 | -57.9056 | -69.582 | -3241.4164 | 3242.7717 | 18.7955 | 17.187 | -8.9063 | 27.0613 | 3241.4164 | -17.187 | 57.0007
0.6149 | -58.963 | -69.4629 | -3243.844 | 3245.1993 | 18.8253 | 17.2317 | -8.9509 | 27.1358 | 3243.844 | -17.2317 | 57.1681
0.6167 | -56.6992 | -65.3523 | -3243.144 | 3244.3653 | 19.0189 | 17.1572 | -8.9509 | 27.2549 | 3243.144 | -17.1572 | 57.3354
0.6185 | -52.8269 | -59.2162 | -3240.0611 | 3241.0738 | 19.3465 | 17.0381 | -8.9509 | 27.4039 | 3240.0611 | -17.0381 | 57.5028
0.6203 | -47.7781 | -53.1099 | -3247.8652 | 3248.6695 | 19.6593 | 16.8147 | -8.8765 | 27.4932 | 3247.8652 | -16.8147 | 57.6701
0.6221 | -48.6866 | -54.0035 | -3262.3118 | 3263.1459 | 19.6295 | 16.8891 | -8.8467 | 27.5081 | 3262.3118 | -16.8891 | 57.8375
0.6239 | -50.2355 | -56.3567 | -3277.5776 | 3278.4712 | 19.6146 | 16.9338 | -8.8914 | 27.5379 | 3277.5776 | -16.9338 | 58.0048
0.6256 | -48.0908 | -55.0907 | -3290.3412 | 3291.1753 | 19.7784 | 16.8296 | -8.936 | 27.6124 | 3290.3412 | -16.8296 | 58.1629
0.6274 | -45.4398 | -54.2567 | -3306.3814 | 3307.1559 | 19.9125 | 16.6508 | -8.9509 | 27.6273 | 3306.3814 | -16.6508 | 58.3302
0.6292 | -44.0696 | -56.3716 | -3327.6343 | 3328.4237 | 19.8976 | 16.4572 | -8.9063 | 27.4783 | 3327.6343 | -16.4572 | 58.4976
0.631 | -46.6759 | -62.2247 | -3351.5531 | 3352.4765 | 19.7338 | 16.4274 | -8.8914 | 27.3443 | 3351.5531 | -16.4274 | 58.6649
0.6328 | -48.061 | -65.665 | -3367.4891 | 3368.4721 | 19.6146 | 16.4572 | -8.8765 | 27.2996 | 3367.4891 | -16.4572 | 58.8323
0.6346 | -48.2546 | -66.8565 | -3371.2571 | 3372.2997 | 19.5848 | 16.4274 | -8.9063 | 27.2549 | 3371.2571 | -16.4274 | 58.9996
0.6364 | -45.7972 | -64.8906 | -3372.8805 | 3373.8337 | 19.6593 | 16.3083 | -8.9063 | 27.2549 | 3372.8805 | -16.3083 | 59.167
0.6382 | -42.893 | -62.5374 | -3374.3401 | 3375.2039 | 19.7487 | 16.1445 | -8.8765 | 27.24 | 3374.3401 | -16.1445 | 59.3343
0.64 | -42.0143 | -62.8353 | -3384.6314 | 3385.4952 | 19.6742 | 16.07 | -8.802 | 27.1209 | 3384.6314 | -16.07 | 59.5017
0.6418 | -43.3845 | -65.4416 | -3395.3249 | 3396.2632 | 19.4955 | 16.07 | -8.7126 | 26.9571 | 3395.3249 | -16.07 | 59.669
0.6436 | -43.1015 | -64.5182 | -3394.2972 | 3395.2206 | 19.4955 | 16.0998 | -8.7126 | 26.972 | 3394.2972 | -16.0998 | 59.8364
0.6454 | -40.7335 | -61.0779 | -3387.1335 | 3387.9824 | 19.6295 | 16.0402 | -8.6978 | 27.0464 | 3387.1335 | -16.0402 | 60.0037
0.6472 | -37.9335 | -59.4992 | -3380.5059 | 3381.2804 | 19.6593 | 15.9806 | -8.668 | 27.0315 | 3380.5059 | -15.9806 | 60.1711
0.649 | -37.0846 | -59.8268 | -3382.2931 | 3383.0825 | 19.5848 | 16.0104 | -8.5935 | 26.9571 | 3382.2931 | -16.0104 | 60.3384
0.6507 | -36.1612 | -58.9332 | -3378.3464 | 3379.1357 | 19.555 | 16.0551 | -8.5488 | 26.9273 | 3378.3464 | -16.0551 | 60.4965
0.6525 | -33.6293 | -54.1078 | -3358.7318 | 3359.4318 | 19.6891 | 16.1147 | -8.519 | 27.0613 | 3358.7318 | -16.1147 | 60.6638
0.6543 | -31.2613 | -50.5631 | -3342.3788 | 3342.9596 | 19.7338 | 16.2189 | -8.4446 | 27.1209 | 3342.3788 | -16.2189 | 60.8312
0.6561 | -31.1272 | -48.9248 | -3328.9896 | 3329.5258 | 19.704 | 16.3976 | -8.3999 | 27.1805 | 3328.9896 | -16.3976 | 60.9985
0.6579 | -32.2442 | -48.061 | -3316.3302 | 3316.8664 | 19.6444 | 16.621 | -8.3701 | 27.24 | 3316.3302 | -16.621 | 61.1659
0.6597 | -30.9485 | -44.4121 | -3301.69 | 3302.1517 | 19.6146 | 16.6955 | -8.2658 | 27.2252 | 3301.69 | -16.6955 | 61.3332
0.6615 | -29.3996 | -40.7335 | -3287.9732 | 3288.3604 | 19.5848 | 16.7402 | -8.1616 | 27.1805 | 3287.9732 | -16.7402 | 61.5006
0.6633 | -30.2485 | -39.5569 | -3265.1416 | 3265.5288 | 19.4061 | 16.8593 | -8.0871 | 27.0911 | 3265.1416 | -16.8593 | 61.6679
0.6651 | -32.84 | -41.076 | -3244.1419 | 3244.5887 | 19.1231 | 17.0381 | -8.0573 | 26.9869 | 3244.1419 | -17.0381 | 61.8353
0.6669 | -37.0846 | -45.5589 | -3236.6505 | 3237.2016 | 18.8104 | 17.2764 | -8.0275 | 26.8826 | 3236.6505 | -17.2764 | 62.0026
0.6687 | -38.708 | -47.4653 | -3234.908 | 3235.5186 | 18.6465 | 17.4104 | -7.9829 | 26.823 | 3234.908 | -17.4104 | 62.17
0.6705 | -41.1207 | -50.3546 | -3237.0973 | 3237.7824 | 18.4678 | 17.604 | -7.9829 | 26.8081 | 3237.0973 | -17.604 | 62.3373
0.6723 | -40.9867 | -51.2482 | -3239.1824 | 3239.8973 | 18.3934 | 17.6934 | -7.9531 | 26.8081 | 3239.1824 | -17.6934 | 62.5046
0.6741 | -39.3484 | -48.4929 | -3238.0952 | 3238.7505 | 18.5125 | 17.7529 | -7.9382 | 26.9273 | 3238.0952 | -17.7529 | 62.672
0.6758 | -37.6803 | -45.6632 | -3236.9782 | 3237.5739 | 18.6465 | 17.8125 | -7.9233 | 27.0613 | 3236.9782 | -17.8125 | 62.83
0.6776 | -36.1314 | -43.3696 | -3237.7377 | 3238.2739 | 18.7359 | 17.8572 | -7.9084 | 27.1507 | 3237.7377 | -17.8572 | 62.9974
0.6794 | -34.7761 | -42.2824 | -3242.7866 | 3243.293 | 18.7508 | 17.8572 | -7.8339 | 27.1358 | 3242.7866 | -17.8572 | 63.1647
0.6812 | -37.5463 | -45.8866 | -3256.7566 | 3257.3672 | 18.5423 | 18.0061 | -7.7744 | 27.0613 | 3256.7566 | -18.0061 | 63.3321
0.683 | -40.4207 | -49.5951 | -3275.8648 | 3276.5648 | 18.4678 | 18.2146 | -7.7893 | 27.1507 | 3275.8648 | -18.2146 | 63.4994
0.6848 | -42.4462 | -52.5439 | -3295.4795 | 3296.239 | 18.438 | 18.4083 | -7.8041 | 27.2549 | 3295.4795 | -18.4083 | 63.6668
0.6866 | -41.9249 | -53.0354 | -3315.4962 | 3316.2558 | 18.4976 | 18.5125 | -7.819 | 27.3741 | 3315.4962 | -18.5125 | 63.8341
0.6884 | -39.9441 | -50.6525 | -3332.3704 | 3333.0555 | 18.7061 | 18.6019 | -7.8339 | 27.5677 | 3332.3704 | -18.6019 | 64.0015
0.6902 | -36.3548 | -46.9589 | -3342.8256 | 3343.4065 | 18.8699 | 18.5274 | -7.8339 | 27.6422 | 3342.8256 | -18.5274 | 64.1688
0.692 | -34.4633 | -44.1143 | -3355.4105 | 3355.9318 | 18.9593 | 18.4678 | -7.8041 | 27.672 | 3355.4105 | -18.4678 | 64.3362
0.6938 | -37.0101 | -46.5568 | -3368.2933 | 3368.8593 | 18.8104 | 18.5721 | -7.8041 | 27.6422 | 3368.2933 | -18.5721 | 64.5035
0.6956 | -37.5016 | -46.6759 | -3370.3635 | 3370.9593 | 18.7359 | 18.6019 | -7.819 | 27.6124 | 3370.3635 | -18.6019 | 64.6709
0.6974 | -32.423 | -40.242 | -3360.2211 | 3360.653 | 18.8848 | 18.4827 | -7.7446 | 27.6124 | 3360.2211 | -18.4827 | 64.8382
0.6992 | -31.6187 | -39.7803 | -3348.8574 | 3349.2745 | 18.6614 | 18.4678 | -7.6403 | 27.4188 | 3348.8574 | -18.4678 | 65.0056
0.7009 | -31.991 | -41.4186 | -3337.2555 | 3337.6874 | 18.304 | 18.4827 | -7.551 | 27.1507 | 3337.2555 | -18.4827 | 65.1636
0.7027 | -29.7421 | -39.3633 | -3325.2961 | 3325.6833 | 18.1849 | 18.4231 | -7.4914 | 27.0166 | 3325.2961 | -18.4231 | 65.331
0.7045 | -26.9869 | -36.5633 | -3317.1941 | 3317.5217 | 18.1402 | 18.304 | -7.4467 | 26.9124 | 3317.1941 | -18.304 | 65.4983
0.7063 | -25.1996 | -35.0889 | -3316.1813 | 3316.4792 | 18.0657 | 18.2444 | -7.4169 | 26.7932 | 3316.1813 | -18.2444 | 65.6657
0.7081 | -24.2018 | -34.4782 | -3316.2111 | 3316.509 | 17.9317 | 18.2593 | -7.3871 | 26.7039 | 3316.2111 | -18.2593 | 65.833
0.7099 | -25.6762 | -35.9527 | -3321.1557 | 3321.4685 | 17.7678 | 18.2444 | -7.3573 | 26.5847 | 3321.1557 | -18.2444 | 66.0004
0.7117 | -27.24 | -36.9357 | -3325.2514 | 3325.6088 | 17.6487 | 18.2444 | -7.3276 | 26.4954 | 3325.2514 | -18.2444 | 66.1677
0.7135 | -24.8273 | -34.0017 | -3321.0514 | 3321.3493 | 17.738 | 18.0359 | -7.3127 | 26.4209 | 3321.0514 | -18.0359 | 66.3351
0.7153 | -23.2784 | -32.7208 | -3321.0068 | 3321.2748 | 17.7827 | 17.8423 | -7.3276 | 26.3167 | 3321.0068 | -17.8423 | 66.5024
0.7171 | -22.5784 | -32.8995 | -3324.194 | 3324.4918 | 17.7678 | 17.6487 | -7.3425 | 26.1826 | 3324.194 | -17.6487 | 66.6698
0.7189 | -22.4593 | -35.1931 | -3332.6236 | 3332.9215 | 17.6934 | 17.4551 | -7.3573 | 26.0039 | 3332.6236 | -17.4551 | 66.8371
0.7207 | -21.2231 | -36.6825 | -3339.7278 | 3340.0256 | 17.6636 | 17.2317 | -7.3573 | 25.8252 | 3339.7278 | -17.2317 | 67.0045
0.7225 | -20.4189 | -38.008 | -3340.6958 | 3341.0235 | 17.6934 | 17.0678 | -7.3871 | 25.7507 | 3340.6958 | -17.0678 | 67.1718
0.7242 | -16.621 | -36.042 | -3335.4087 | 3335.6768 | 17.887 | 16.7849 | -7.402 | 25.7358 | 3335.4087 | -16.7849 | 67.3299
0.726 | -16.487 | -38.6633 | -3334.8427 | 3335.1555 | 17.8721 | 16.6657 | -7.4318 | 25.6464 | 3334.8427 | -16.6657 | 67.4972
0.7278 | -15.3402 | -39.7356 | -3330.8513 | 3331.179 | 17.8274 | 16.5019 | -7.402 | 25.5124 | 3330.8513 | -16.5019 | 67.6646
0.7296 | -13.9104 | -40.2271 | -3326.0407 | 3326.3833 | 17.8125 | 16.3232 | -7.3871 | 25.3784 | 3326.0407 | -16.3232 | 67.8319
0.7314 | -10.6935 | -36.7718 | -3315.8536 | 3316.1515 | 17.9168 | 16.07 | -7.3573 | 25.2741 | 3315.8536 | -16.07 | 67.9993
0.7332 | -8.8914 | -33.7634 | -3306.9176 | 3307.1708 | 17.9466 | 15.8317 | -7.3276 | 25.1401 | 3306.9176 | -15.8317 | 68.1666
0.735 | -6.434 | -29.7272 | -3294.3773 | 3294.6156 | 18.0359 | 15.5487 | -7.2978 | 25.0209 | 3294.3773 | -15.5487 | 68.334
0.7368 | -2.5021 | -24.1273 | -3276.5202 | 3276.7584 | 18.1997 | 15.2062 | -7.2978 | 24.9167 | 3276.5202 | -15.2062 | 68.5013
0.7386 | 0.61063 | -18.4678 | -3259.1991 | 3259.4523 | 18.304 | 14.8487 | -7.2531 | 24.738 | 3259.1991 | -14.8487 | 68.6687
0.7404 | 1.4596 | -16.5764 | -3243.5461 | 3243.8142 | 18.304 | 14.5658 | -7.2084 | 24.5592 | 3243.5461 | -14.5658 | 68.836
0.7422 | 3.3212 | -13.3743 | -3228.3549 | 3228.6378 | 18.2891 | 14.1934 | -7.1042 | 24.306 | 3228.3549 | -14.1934 | 69.0033
0.744 | 5.168 | -10.1722 | -3213.1636 | 3213.4615 | 18.2742 | 13.8211 | -7.0148 | 24.0529 | 3213.1636 | -13.8211 | 69.1707
0.7458 | 5.2872 | -9.7105 | -3194.6362 | 3194.9787 | 18.1253 | 13.4934 | -6.9701 | 23.7401 | 3194.6362 | -13.4934 | 69.338
0.7476 | 6.1063 | -7.5956 | -3172.311 | 3172.6982 | 18.0508 | 13.1062 | -6.9254 | 23.4422 | 3172.311 | -13.1062 | 69.5054
0.7493 | 8.6084 | -2.1893 | -3149.8517 | 3150.373 | 17.9912 | 12.5998 | -6.9105 | 23.0848 | 3149.8517 | -12.5998 | 69.6634
0.7511 | 11.453 | 3.8723 | -3127.5265 | 3128.1967 | 17.9317 | 12.0488 | -6.8956 | 22.7124 | 3127.5265 | -12.0488 | 69.8308
0.7529 | 8.2212 | 0.38723 | -3106.8693 | 3107.5544 | 17.6338 | 11.7658 | -6.9254 | 22.3252 | 3106.8693 | -11.7658 | 69.9981
0.7547 | 7.1339 | -0.81914 | -3080.225 | 3080.9399 | 17.4402 | 11.453 | -6.9403 | 22.0274 | 3080.225 | -11.453 | 70.1655
0.7565 | 8.7424 | 1.5489 | -3047.5936 | 3048.3084 | 17.3508 | 11.0211 | -6.9552 | 21.7295 | 3047.5936 | -11.0211 | 70.3328
0.7583 | 9.8892 | 4.8701 | -3009.3026 | 3010.0771 | 17.2764 | 10.619 | -6.9999 | 21.4912 | 3009.3026 | -10.619 | 70.5002
0.7601 | 11.0956 | 7.9084 | -2970.7138 | 2971.5925 | 17.0827 | 10.1126 | -7.0297 | 21.104 | 2970.7138 | -10.1126 | 70.6675
0.7619 | 12.3764 | 10.9913 | -2932.0803 | 2933.0484 | 16.8891 | 9.6211 | -7.0595 | 20.7018 | 2932.0803 | -9.6211 | 70.8349
0.7637 | 13.7913 | 13.9551 | -2892.9851 | 2894.0723 | 16.6806 | 9.0999 | -7.0595 | 20.2997 | 2892.9851 | -9.0999 | 71.0022
0.7655 | 16.204 | 17.0083 | -2851.0899 | 2852.3261 | 16.5168 | 8.519 | -6.9999 | 19.8827 | 2851.0899 | -8.519 | 71.1696
0.7673 | 19.3465 | 20.3742 | -2803.7438 | 2805.0097 | 16.3679 | 7.9382 | -6.9105 | 19.5104 | 2803.7438 | -7.9382 | 71.3369
0.7691 | 21.4167 | 22.8614 | -2757.8274 | 2759.0933 | 16.1742 | 7.4467 | -6.8361 | 19.1231 | 2757.8274 | -7.4467 | 71.5043
0.7709 | 21.9976 | 23.4869 | -2716.9301 | 2718.1663 | 15.8913 | 7.0297 | -6.7765 | 18.7061 | 2716.9301 | -7.0297 | 71.6716
0.7727 | 21.7742 | 22.9805 | -2681.1561 | 2682.3774 | 15.5636 | 6.6722 | -6.7169 | 18.2742 | 2681.1561 | -6.6722 | 71.839
0.7744 | 21.9976 | 23.338 | -2648.7928 | 2650.0289 | 15.3104 | 6.434 | -6.6871 | 17.9466 | 2648.7928 | -6.434 | 71.997
0.7762 | 22.9061 | 24.6039 | -2607.6869 | 2608.8784 | 15.0423 | 6.0169 | -6.568 | 17.5444 | 2607.6869 | -6.0169 | 72.1644
0.778 | 21.7146 | 23.3678 | -2570.796 | 2571.9874 | 14.6998 | 5.6446 | -6.4786 | 17.0976 | 2570.796 | -5.6446 | 72.3317
0.7798 | 20.6721 | 21.9082 | -2538.522 | 2539.6985 | 14.4019 | 5.3318 | -6.434 | 16.7402 | 2538.522 | -5.3318 | 72.4991
0.7816 | 19.1976 | 19.704 | -2507.7671 | 2508.9139 | 14.0743 | 5.0489 | -6.3893 | 16.3679 | 2507.7671 | -5.0489 | 72.6664
0.7834 | 13.7615 | 12.8977 | -2486.1418 | 2487.0354 | 13.5977 | 5.0191 | -6.3744 | 15.9657 | 2486.1418 | -5.0191 | 72.8338
0.7852 | 7.3722 | 6.9105 | -2463.8315 | 2464.5166 | 13.1807 | 5.0489 | -6.3595 | 15.6232 | 2463.8315 | -5.0489 | 73.0011
0.787 | 3.9914 | 5.868 | -2434.6851 | 2435.4447 | 12.9126 | 4.9595 | -6.2999 | 15.37 | 2434.6851 | -4.9595 | 73.1685
0.7888 | 1.4298 | 5.2872 | -2404.3472 | 2405.1962 | 12.6594 | 4.8255 | -6.2106 | 15.1019 | 2404.3472 | -4.8255 | 73.3358
0.7906 | 1.0723 | 7.0893 | -2362.5861 | 2363.5095 | 12.436 | 4.6319 | -6.1361 | 14.804 | 2362.5861 | -4.6319 | 73.5032
0.7924 | 0.93828 | 10.7828 | -2317.5931 | 2318.5612 | 12.2573 | 4.4382 | -6.0616 | 14.5806 | 2317.5931 | -4.4382 | 73.6705
0.7942 | 1.8915 | 20.1061 | -2268.6534 | 2269.6512 | 12.2126 | 4.2148 | -5.9425 | 14.4317 | 2268.6534 | -4.2148 | 73.8379
0.796 | 1.9808 | 26.3464 | -2227.5774 | 2228.6497 | 12.0934 | 3.9468 | -5.7935 | 14.1785 | 2227.5774 | -3.9468 | 74.0052
0.7977 | 2.3085 | 31.0379 | -2187.6035 | 2188.7354 | 11.9892 | 3.6787 | -5.6744 | 13.97 | 2187.6035 | -3.6787 | 74.1633
0.7995 | 2.1744 | 33.7485 | -2153.0954 | 2154.3167 | 11.8254 | 3.4553 | -5.5552 | 13.7168 | 2153.0954 | -3.4553 | 74.3306
0.8013 | 2.8446 | 36.2059 | -2119.4661 | 2120.8065 | 11.602 | 3.2021 | -5.451 | 13.4041 | 2119.4661 | -3.2021 | 74.498
0.8031 | 3.5893 | 37.7995 | -2087.7878 | 2089.1878 | 11.3637 | 2.934 | -5.3467 | 13.0764 | 2087.7878 | -2.934 | 74.6653
0.8049 | 4.185 | 37.2186 | -2062.7371 | 2064.1222 | 11.0658 | 2.651 | -5.2425 | 12.7041 | 2062.7371 | -2.651 | 74.8327
0.8067 | 4.2297 | 36.9654 | -2041.5885 | 2042.9885 | 10.7084 | 2.3829 | -5.1233 | 12.2573 | 2041.5885 | -2.3829 | 75
0.8085 | 7.9382 | 41.3441 | -2011.7123 | 2013.2314 | 10.6488 | 1.9957 | -4.9893 | 12.0786 | 2011.7123 | -1.9957 | 75.1673
0.8103 | 12.5998 | 47.7483 | -1977.7702 | 1979.2894 | 10.7084 | 1.5638 | -4.8255 | 12.0488 | 1977.7702 | -1.5638 | 75.3347
0.8121 | 16.4125 | 54.2567 | -1941.7431 | 1943.4558 | 10.7679 | 1.2213 | -4.6914 | 12.0339 | 1941.7431 | -1.2213 | 75.502
0.8139 | 19.0189 | 59.4098 | -1904.8521 | 1906.7882 | 10.753 | 0.96807 | -4.5425 | 11.9743 | 1904.8521 | -0.96807 | 75.6694
0.8157 | 21.0742 | 64.1906 | -1867.6633 | 1869.659 | 10.7233 | 0.78935 | -4.4084 | 11.8998 | 1867.6633 | -0.78935 | 75.8367
0.8175 | 21.0742 | 66.1565 | -1838.9339 | 1841.0636 | 10.5743 | 0.6702 | -4.3042 | 11.6913 | 1838.9339 | -0.6702 | 76.0041
0.8193 | 19.2274 | 65.0395 | -1814.9108 | 1816.9959 | 10.3062 | 0.61063 | -4.1999 | 11.4084 | 1814.9108 | -0.61063 | 76.1714
0.8211 | 17.6189 | 64.2353 | -1791.0367 | 1793.1217 | 10.0382 | 0.53616 | -4.0957 | 11.1254 | 1791.0367 | -0.53616 | 76.3388
0.8228 | 16.3381 | 63.2672 | -1768.637 | 1770.6774 | 9.7254 | 0.47659 | -4.0361 | 10.7977 | 1768.637 | -0.47659 | 76.4968
0.8246 | 15.37 | 63.3715 | -1742.6182 | 1744.6735 | 9.4126 | 0.43191 | -3.9765 | 10.485 | 1742.6182 | -0.43191 | 76.6642
0.8264 | 14.7594 | 66.1267 | -1717.3143 | 1719.6079 | 9.2041 | 0.41702 | -3.917 | 10.2616 | 1717.3143 | -0.41702 | 76.8315
0.8282 | 12.5403 | 65.2033 | -1698.6975 | 1701.0209 | 8.936 | 0.4617 | -3.8425 | 9.9935 | 1698.6975 | -0.4617 | 76.9989
0.83 | 10.6786 | 65.3821 | -1682.0169 | 1684.3552 | 8.7573 | 0.55106 | -3.8276 | 9.8296 | 1682.0169 | -0.55106 | 77.1662
0.8318 | 8.9956 | 65.6353 | -1666.1405 | 1668.5235 | 8.5786 | 0.64042 | -3.8276 | 9.6807 | 1666.1405 | -0.64042 | 77.3336
0.8336 | 7.5063 | 64.5629 | -1655.1343 | 1657.547 | 8.3701 | 0.77446 | -3.8127 | 9.4871 | 1655.1343 | -0.77446 | 77.5009
0.8354 | 5.9723 | 64.2502 | -1644.2472 | 1646.7195 | 8.1616 | 0.84892 | -3.7829 | 9.2935 | 1644.2472 | -0.84892 | 77.6683
0.8372 | 3.6936 | 64.3097 | -1634.5963 | 1637.292 | 7.968 | 0.95318 | -3.7829 | 9.0999 | 1634.5963 | -0.95318 | 77.8356
0.839 | 0.14893 | 62.5672 | -1626.6432 | 1629.607 | 7.7297 | 1.1319 | -3.7829 | 8.8914 | 1626.6432 | -1.1319 | 78.003
0.8408 | -0.2234 | 64.4587 | -1616.6051 | 1619.8816 | 7.5956 | 1.2064 | -3.8127 | 8.7871 | 1616.6051 | -1.2064 | 78.1703
0.8426 | -0.19361 | 67.5267 | -1605.9712 | 1609.7094 | 7.5807 | 1.3553 | -3.8723 | 8.802 | 1605.9712 | -1.3553 | 78.3377
0.8444 | -2.1893 | 68.7182 | -1596.5139 | 1600.5946 | 7.4318 | 1.5489 | -3.9319 | 8.7424 | 1596.5139 | -1.5489 | 78.505
0.8462 | -5.2574 | 68.748 | -1589.3799 | 1593.699 | 7.1935 | 1.7574 | -3.9914 | 8.6531 | 1589.3799 | -1.7574 | 78.6724
0.8479 | -9.0701 | 68.8075 | -1585.0757 | 1589.5288 | 6.9403 | 2.0255 | -4.0808 | 8.5935 | 1585.0757 | -2.0255 | 78.8304
0.8497 | -11.1254 | 70.5203 | -1583.3332 | 1587.9204 | 6.8659 | 2.2638 | -4.1999 | 8.668 | 1583.3332 | -2.2638 | 78.9978
0.8515 | -16.5168 | 67.4969 | -1584.8821 | 1589.6331 | 6.6276 | 2.6361 | -4.2595 | 8.668 | 1584.8821 | -2.6361 | 79.1651
0.8533 | -20.5827 | 64.9353 | -1588.6501 | 1593.2522 | 6.434 | 2.9489 | -4.3191 | 8.6829 | 1588.6501 | -2.9489 | 79.3325
0.8551 | -24.3507 | 62.2396 | -1591.6884 | 1596.0522 | 6.2403 | 3.2765 | -4.3638 | 8.7424 | 1591.6884 | -3.2765 | 79.4998
0.8569 | -28.6996 | 58.6056 | -1591.7182 | 1596.1117 | 6.0616 | 3.6489 | -4.4084 | 8.8169 | 1591.7182 | -3.6489 | 79.6672
0.8587 | -32.423 | 55.6716 | -1590.4224 | 1594.7862 | 5.8829 | 4.0361 | -4.4233 | 8.936 | 1590.4224 | -4.0361 | 79.8345
0.8605 | -35.3421 | 53.7801 | -1588.7842 | 1593.0437 | 5.7637 | 4.4233 | -4.4531 | 9.1446 | 1588.7842 | -4.4233 | 80.0019
0.8623 | -37.7697 | 52.4099 | -1586.3565 | 1590.5714 | 5.6893 | 4.7808 | -4.5276 | 9.4275 | 1586.3565 | -4.7808 | 80.1692
0.8641 | -40.1973 | 50.9206 | -1583.8396 | 1588.0097 | 5.6148 | 5.1531 | -4.5872 | 9.7105 | 1583.8396 | -5.1531 | 80.3366
0.8659 | -42.5058 | 47.7334 | -1579.9375 | 1583.914 | 5.4957 | 5.451 | -4.5276 | 9.8296 | 1579.9375 | -5.451 | 80.5039
0.8677 | -45.4547 | 44.159 | -1573.9354 | 1577.8077 | 5.3914 | 5.734 | -4.468 | 9.919 | 1573.9354 | -5.734 | 80.6713
0.8695 | -47.4951 | 41.8058 | -1566.1015 | 1569.9589 | 5.3765 | 5.9723 | -4.4382 | 10.0233 | 1566.1015 | -5.9723 | 80.8386
0.8712 | -47.4504 | 41.5377 | -1560.2186 | 1564.076 | 5.4659 | 6.0616 | -4.4233 | 10.1126 | 1560.2186 | -6.0616 | 80.9967
0.873 | -45.3653 | 42.8186 | -1551.2975 | 1555.2293 | 5.585 | 6.1063 | -4.4084 | 10.1871 | 1551.2975 | -6.1063 | 81.164
0.8748 | -46.0057 | 41.493 | -1546.3231 | 1550.389 | 5.6595 | 6.2106 | -4.3936 | 10.2318 | 1546.3231 | -6.2106 | 81.3313
0.8766 | -48.21 | 36.5335 | -1545.1465 | 1548.9592 | 5.585 | 6.2403 | -4.4084 | 10.2467 | 1545.1465 | -6.2403 | 81.4987
0.8784 | -51.1589 | 30.6804 | -1545.6975 | 1549.5103 | 5.4957 | 6.3297 | -4.3936 | 10.202 | 1545.6975 | -6.3297 | 81.666
0.8802 | -51.6503 | 26.8826 | -1541.4678 | 1545.2061 | 5.451 | 6.3148 | -4.3787 | 10.1722 | 1541.4678 | -6.3148 | 81.8334
0.882 | -51.7248 | 22.9507 | -1536.8657 | 1540.4551 | 5.4659 | 6.1808 | -4.4382 | 10.1275 | 1536.8657 | -6.1808 | 82.0007
0.8838 | -52.2908 | 18.8848 | -1532.1148 | 1535.5551 | 5.4361 | 6.0467 | -4.4531 | 10.0828 | 1532.1148 | -6.0467 | 82.1681
0.8856 | -55.0014 | 14.536 | -1527.2446 | 1530.4914 | 5.2276 | 6.0765 | -4.2744 | 10.0084 | 1527.2446 | -6.0765 | 82.3354
0.8874 | -55.3439 | 11.7509 | -1518.383 | 1521.466 | 5.168 | 5.9723 | -4.1702 | 9.8743 | 1518.383 | -5.9723 | 82.5028
0.8892 | -53.7056 | 10.2913 | -1506.1258 | 1509.0746 | 5.2276 | 5.734 | -4.1106 | 9.6807 | 1506.1258 | -5.734 | 82.6701
0.891 | -51.0844 | 12.5551 | -1489.3558 | 1492.0217 | 5.3467 | 5.5999 | -4.0659 | 9.6807 | 1489.3558 | -5.5999 | 82.8375
0.8928 | -47.6887 | 15.9508 | -1475.5645 | 1477.9176 | 5.3616 | 5.4063 | -3.9616 | 9.5616 | 1475.5645 | -5.4063 | 83.0048
0.8946 | -41.8356 | 19.3912 | -1456.1137 | 1458.02 | 5.4063 | 5.034 | -3.8574 | 9.3382 | 1456.1137 | -5.034 | 83.1722
0.8963 | -34.5378 | 25.9741 | -1427.6077 | 1429.0673 | 5.4361 | 4.6914 | -3.8276 | 9.2041 | 1427.6077 | -4.6914 | 83.3302
0.8981 | -30.4123 | 29.6379 | -1403.5846 | 1404.8506 | 5.3765 | 4.468 | -3.8425 | 8.9956 | 1403.5846 | -4.468 | 83.4976
0.8999 | -29.0124 | 31.0081 | -1385.5338 | 1386.6955 | 5.2872 | 4.2595 | -3.8574 | 8.6531 | 1385.5338 | -4.2595 | 83.6649
0.9017 | -30.0549 | 30.4868 | -1372.6659 | 1373.768 | 5.168 | 4.0957 | -3.8425 | 8.2807 | 1372.6659 | -4.0957 | 83.8323
0.9035 | -32.3783 | 26.9571 | -1365.9192 | 1366.9022 | 5.1829 | 3.8276 | -3.8872 | 8.0127 | 1365.9192 | -3.8276 | 83.9996
0.9053 | -33.4655 | 23.7252 | -1360.081 | 1361.0044 | 5.2872 | 3.4851 | -3.9616 | 7.8637 | 1360.081 | -3.4851 | 84.167
0.9071 | -34.5527 | 20.4933 | -1354.2428 | 1355.0917 | 5.4063 | 3.1574 | -4.0212 | 7.7297 | 1354.2428 | -3.1574 | 84.3343
0.9089 | -37.3825 | 18.5274 | -1352.0534 | 1352.947 | 5.3467 | 3.0085 | -4.0212 | 7.5807 | 1352.0534 | -3.0085 | 84.5017
0.9107 | -38.9612 | 18.6763 | -1349.6556 | 1350.5939 | 5.2723 | 2.8595 | -3.9914 | 7.4467 | 1349.6556 | -2.8595 | 84.669
0.9125 | -40.4058 | 16.6508 | -1351.4279 | 1352.3662 | 5.1829 | 2.651 | -4.0063 | 7.3127 | 1351.4279 | -2.651 | 84.8364
0.9143 | -39.0058 | 17.4253 | -1350.7577 | 1351.5917 | 5.1829 | 2.234 | -4.0361 | 7.2233 | 1350.7577 | -2.234 | 85.0037
0.9161 | -37.6803 | 21.8784 | -1349.462 | 1350.2811 | 5.3318 | 1.6681 | -4.0808 | 7.3722 | 1349.462 | -1.6681 | 85.1711
0.9179 | -36.3548 | 26.3464 | -1348.1663 | 1348.9705 | 5.4957 | 1.1021 | -4.1404 | 7.5212 | 1348.1663 | -1.1021 | 85.3384
0.9197 | -34.6867 | 28.2975 | -1344.5621 | 1345.3216 | 5.585 | 0.53616 | -4.2744 | 7.7297 | 1344.5621 | -0.53616 | 85.5058
0.9214 | -33.2421 | 29.4145 | -1341.7323 | 1342.477 | 5.6893 | -0.059574 | -4.4233 | 8.0424 | 1341.7323 | 0.059574 | 85.6638
0.9232 | -32.0655 | 30.2634 | -1340.094 | 1340.8238 | 5.7786 | -0.62552 | -4.5723 | 8.3403 | 1340.094 | 0.62552 | 85.8312
0.925 | -32.4825 | 29.8911 | -1342.7897 | 1343.5344 | 5.868 | -1.1021 | -4.7212 | 8.5786 | 1342.7897 | 1.1021 | 85.9985
0.9268 | -35.1931 | 28.4613 | -1348.3003 | 1349.0897 | 5.9127 | -1.534 | -4.8404 | 8.7871 | 1348.3003 | 1.534 | 86.1659
0.9286 | -38.6037 | 23.5763 | -1357.4746 | 1358.2789 | 5.7637 | -1.9361 | -5.0042 | 8.936 | 1357.4746 | 1.9361 | 86.3332
0.9304 | -39.1995 | 21.6103 | -1362.6278 | 1363.4171 | 5.7191 | -2.3978 | -5.1829 | 9.1743 | 1362.6278 | 2.3978 | 86.5006
0.9322 | -35.9229 | 24.455 | -1361.7044 | 1362.4639 | 5.8829 | -2.9489 | -5.3467 | 9.5616 | 1361.7044 | 2.9489 | 86.6679
0.934 | -32.5421 | 30.3081 | -1358.5321 | 1359.3363 | 6.1361 | -3.4553 | -5.5403 | 9.9786 | 1358.5321 | 3.4553 | 86.8353
0.9358 | -29.5336 | 36.891 | -1357.281 | 1358.1895 | 6.4191 | -3.917 | -5.734 | 10.4105 | 1357.281 | 3.917 | 87.0026
0.9376 | -26.123 | 42.5058 | -1358.6363 | 1359.6342 | 6.6425 | -4.4084 | -5.9276 | 10.8871 | 1358.6363 | 4.4084 | 87.17
0.9394 | -22.8465 | 47.0334 | -1362.0171 | 1363.0895 | 6.8212 | -4.8999 | -6.1212 | 11.3488 | 1362.0171 | 4.8999 | 87.3373
0.9412 | -19.421 | 49.3121 | -1369.7021 | 1370.8191 | 6.8808 | -5.3318 | -6.3148 | 11.7658 | 1369.7021 | 5.3318 | 87.5046
0.943 | -18.7061 | 49.4014 | -1378.5637 | 1379.6509 | 6.8212 | -5.4957 | -6.5978 | 11.9743 | 1378.5637 | 5.4957 | 87.672
0.9447 | -17.2764 | 52.9163 | -1378.3552 | 1379.5914 | 6.8808 | -5.5403 | -6.9254 | 12.1679 | 1378.3552 | 5.5403 | 87.83
0.9465 | -15.7274 | 56.6247 | -1385.6828 | 1387.0678 | 6.9701 | -5.7489 | -7.119 | 12.3615 | 1385.6828 | 5.7489 | 87.9974
0.9483 | -15.2509 | 58.7396 | -1395.8252 | 1397.2996 | 6.985 | -6.0914 | -7.2978 | 12.5998 | 1395.8252 | 6.0914 | 88.1647
0.9501 | -15.2955 | 60.6609 | -1406.8016 | 1408.3356 | 6.9701 | -6.434 | -7.5212 | 12.7934 | 1406.8016 | 6.434 | 88.3321
0.9519 | -14.0147 | 63.4161 | -1418.6419 | 1420.2653 | 6.9999 | -6.7914 | -7.7148 | 13.0168 | 1418.6419 | 6.7914 | 88.4994
0.9537 | -13.6424 | 64.831 | -1431.9119 | 1433.5651 | 6.9999 | -7.0446 | -7.9084 | 13.1956 | 1431.9119 | 7.0446 | 88.6668
0.9555 | -14.6551 | 64.0566 | -1447.5649 | 1449.1585 | 6.9403 | -7.1339 | -8.102 | 13.2849 | 1447.5649 | 7.1339 | 88.8341
0.9573 | -11.4828 | 64.5331 | -1466.4199 | 1467.9688 | 6.9701 | -7.551 | -8.1765 | 13.6275 | 1466.4199 | 7.551 | 89.0015
0.9591 | -10.1275 | 62.7906 | -1490.1302 | 1491.5749 | 6.9254 | -7.8041 | -8.2361 | 13.8211 | 1490.1302 | 7.8041 | 89.1688
0.9609 | -8.2212 | 62.5821 | -1522.3149 | 1523.7149 | 6.9403 | -8.1765 | -8.2212 | 14.1041 | 1522.3149 | 8.1765 | 89.3362
0.9627 | -9.2637 | 61.3162 | -1545.2954 | 1546.6507 | 6.8956 | -8.2807 | -8.3105 | 14.1934 | 1545.2954 | 8.2807 | 89.5035
0.9645 | -11.185 | 58.5013 | -1567.5015 | 1568.7972 | 6.8361 | -8.3552 | -8.3701 | 14.253 | 1567.5015 | 8.3552 | 89.6709
0.9663 | -14.2679 | 51.5312 | -1591.0629 | 1592.2394 | 6.6127 | -8.4892 | -8.4595 | 14.2977 | 1591.0629 | 8.4892 | 89.8382
0.9681 | -17.3508 | 44.5462 | -1614.6242 | 1615.6966 | 6.3893 | -8.6233 | -8.5339 | 14.3424 | 1614.6242 | 8.6233 | 90.0056
0.9698 | -18.2891 | 41.4186 | -1631.3942 | 1632.4368 | 6.2999 | -8.802 | -8.6531 | 14.4615 | 1631.3942 | 8.802 | 90.1636
0.9716 | -17.8572 | 40.7037 | -1642.6834 | 1643.7408 | 6.2999 | -8.9807 | -8.7722 | 14.6402 | 1642.6834 | 8.9807 | 90.331
0.9734 | -17.9466 | 38.9612 | -1652.6471 | 1653.645 | 6.2701 | -9.1892 | -8.8914 | 14.804 | 1652.6471 | 9.1892 | 90.4983
0.9752 | -17.4104 | 37.2633 | -1661.6427 | 1662.581 | 6.2552 | -9.3977 | -8.9807 | 14.9977 | 1661.6427 | 9.3977 | 90.6657
0.977 | -16.6508 | 37.2931 | -1667.7788 | 1668.6873 | 6.3446 | -9.502 | -9.0552 | 15.1466 | 1667.7788 | 9.502 | 90.833
0.9788 | -15.1764 | 38.1569 | -1669.6256 | 1670.5192 | 6.4637 | -9.5914 | -9.1594 | 15.2955 | 1669.6256 | 9.5914 | 91.0004
0.9806 | -12.5551 | 40.7633 | -1671.1447 | 1672.083 | 6.6276 | -9.7105 | -9.219 | 15.4743 | 1671.1447 | 9.7105 | 91.1677
0.9824 | -10.9318 | 41.3441 | -1672.0234 | 1672.9617 | 6.6574 | -9.8148 | -9.2786 | 15.5785 | 1672.0234 | 9.8148 | 91.3351
0.9842 | -11.2296 | 39.825 | -1676.834 | 1677.7127 | 6.5829 | -9.7552 | -9.3531 | 15.5934 | 1676.834 | 9.7552 | 91.5024
0.986 | -9.4126 | 42.1483 | -1678.4276 | 1679.3063 | 6.7169 | -9.785 | -9.4275 | 15.7274 | 1678.4276 | 9.785 | 91.6698
0.9878 | -5.8382 | 48.2249 | -1673.6617 | 1674.6298 | 7.0297 | -9.8594 | -9.5169 | 15.9955 | 1673.6617 | 9.8594 | 91.8371
0.9896 | -2.8893 | 54.8524 | -1665.2171 | 1666.3639 | 7.3871 | -9.8743 | -9.6211 | 16.2487 | 1665.2171 | 9.8743 | 92.0045
0.9914 | -1.683 | 58.6502 | -1655.2386 | 1656.5194 | 7.685 | -9.8743 | -9.7552 | 16.4572 | 1655.2386 | 9.8743 | 92.1718
0.9932 | -1.6085 | 60.4226 | -1643.4281 | 1644.7834 | 7.968 | -9.8594 | -9.9339 | 16.6508 | 1643.4281 | 9.8594 | 92.3392
0.9949 | -0.4617 | 64.2948 | -1627.9389 | 1629.3985 | 8.385 | -9.919 | -10.0679 | 16.9636 | 1627.9389 | 9.919 | 92.4972
0.9967 | 1.2808 | 69.0012 | -1613.3881 | 1615.0264 | 8.8318 | -10.0084 | -10.1871 | 17.3061 | 1613.3881 | 10.0084 | 92.6646
0.9985 | 4.9148 | 73.0969 | -1603.9903 | 1605.8073 | 9.1743 | -10.2169 | -10.2169 | 17.6785 | 1603.9903 | 10.2169 | 92.8319
1.0003 | 8.5488 | 77.1776 | -1594.5926 | 1596.5883 | 9.5318 | -10.4105 | -10.2616 | 18.0359 | 1594.5926 | 10.4105 | 92.9993
1.0021 | 9.8296 | 80.5584 | -1588.948 | 1591.1671 | 9.7999 | -10.4552 | -10.3062 | 18.2742 | 1588.948 | 10.4552 | 93.1666
1.0039 | 10.0531 | 80.767 | -1590.4076 | 1592.6565 | 9.9041 | -10.4701 | -10.336 | 18.3785 | 1590.4076 | 10.4701 | 93.334
1.0057 | 9.8148 | 79.918 | -1593.7139 | 1595.933 | 9.9637 | -10.4552 | -10.3509 | 18.4231 | 1593.7139 | 10.4552 | 93.5013
1.0075 | 7.5361 | 76.2989 | -1599.9691 | 1601.9946 | 9.9339 | -10.336 | -10.4254 | 18.3934 | 1599.9691 | 10.336 | 93.6687
1.0093 | 3.8425 | 71.0565 | -1606.7158 | 1608.5328 | 9.8445 | -10.0977 | -10.5594 | 18.3189 | 1606.7158 | 10.0977 | 93.836
1.0111 | 2.517 | 69.448 | -1615.5029 | 1617.2753 | 9.9041 | -9.9637 | -10.6041 | 18.304 | 1615.5029 | 9.9637 | 94.0033
1.0129 | 0.56595 | 67.8991 | -1623.8581 | 1625.6007 | 9.9488 | -9.7403 | -10.6935 | 18.2593 | 1623.8581 | 9.7403 | 94.1707
1.0147 | 0.58084 | 67.8693 | -1628.2815 | 1629.9942 | 10.0679 | -9.6956 | -10.7233 | 18.3189 | 1628.2815 | 9.6956 | 94.338
1.0165 | -0.08936 | 66.7225 | -1633.2261 | 1634.8942 | 10.1573 | -9.6211 | -10.7679 | 18.3487 | 1633.2261 | 9.6211 | 94.5054
1.0182 | -0.99786 | 65.9331 | -1635.6686 | 1637.292 | 10.2318 | -9.6509 | -10.753 | 18.4083 | 1635.6686 | 9.6509 | 94.6634
1.02 | 1.1319 | 69.999 | -1633.5389 | 1635.2963 | 10.5147 | -9.7403 | -10.7679 | 18.6465 | 1633.5389 | 9.7403 | 94.8308
1.0218 | 2.0404 | 73.1415 | -1630.4112 | 1632.3176 | 10.7233 | -9.8148 | -10.8126 | 18.8699 | 1630.4112 | 9.8148 | 94.9981
1.0236 | 0.65531 | 72.099 | -1635.5495 | 1637.3814 | 10.7381 | -9.8296 | -10.8424 | 18.8997 | 1635.5495 | 9.8296 | 95.1655
1.0254 | -2.7404 | 67.0501 | -1644.8579 | 1646.4515 | 10.5445 | -9.7999 | -10.8573 | 18.8104 | 1644.8579 | 9.7999 | 95.3328
1.0272 | -6.285 | 61.9268 | -1655.9088 | 1657.279 | 10.336 | -9.7701 | -10.8573 | 18.6763 | 1655.9088 | 9.7701 | 95.5002
1.029 | -10.0531 | 55.7907 | -1667.2278 | 1668.3894 | 10.1275 | -9.6956 | -10.8424 | 18.5274 | 1667.2278 | 9.6956 | 95.6675
1.0308 | -13.8956 | 49.461 | -1678.0255 | 1678.9787 | 9.9041 | -9.6062 | -10.8275 | 18.3785 | 1678.0255 | 9.6062 | 95.8349
1.0326 | -15.6977 | 46.0951 | -1685.1147 | 1685.9786 | 9.8148 | -9.5765 | -10.8275 | 18.304 | 1685.1147 | 9.5765 | 96.0022
1.0344 | -16.9487 | 43.0271 | -1689.6572 | 1690.4317 | 9.7105 | -9.5467 | -10.8424 | 18.1997 | 1689.6572 | 9.5467 | 96.1696
1.0362 | -17.9615 | 40.5994 | -1691.0274 | 1691.7572 | 9.6509 | -9.4871 | -10.8573 | 18.1402 | 1691.0274 | 9.4871 | 96.3369
1.038 | -18.6912 | 39.0654 | -1689.0764 | 1689.7764 | 9.6956 | -9.4424 | -10.8871 | 18.1551 | 1689.0764 | 9.4424 | 96.5043
1.0398 | -18.7508 | 37.9484 | -1685.9935 | 1686.6934 | 9.7999 | -9.3828 | -10.8871 | 18.17 | 1685.9935 | 9.3828 | 96.6716
1.0416 | -19.7784 | 36.7123 | -1680.9446 | 1681.6595 | 9.9488 | -9.3084 | -10.8722 | 18.1997 | 1680.9446 | 9.3084 | 96.839
1.0433 | -22.0125 | 33.5846 | -1680.7957 | 1681.5254 | 10.0233 | -9.2339 | -10.8275 | 18.1849 | 1680.7957 | 9.2339 | 96.997
1.0451 | -23.2337 | 31.7527 | -1676.3127 | 1677.0872 | 10.1275 | -9.1743 | -10.8275 | 18.1849 | 1676.3127 | 9.1743 | 97.1644
1.0469 | -21.9231 | 33.2421 | -1668.5831 | 1669.3724 | 10.336 | -9.1743 | -10.7977 | 18.2742 | 1668.5831 | 9.1743 | 97.3317
1.0487 | -22.1912 | 31.991 | -1661.4491 | 1662.298 | 10.4701 | -9.2637 | -10.7084 | 18.3189 | 1661.4491 | 9.2637 | 97.4991
1.0505 | -19.8529 | 34.9548 | -1648.477 | 1649.2961 | 10.7381 | -9.4424 | -10.619 | 18.4976 | 1648.477 | 9.4424 | 97.6664
1.0523 | -16.6359 | 39.8845 | -1633.8516 | 1634.641 | 11.0062 | -9.6211 | -10.5594 | 18.721 | 1633.8516 | 9.6211 | 97.8338
1.0541 | -14.6551 | 41.4633 | -1624.9007 | 1625.6751 | 11.1254 | -9.7105 | -10.5147 | 18.7955 | 1624.9007 | 9.7105 | 98.0011
1.0559 | -14.8934 | 39.9441 | -1620.3433 | 1621.088 | 11.1105 | -9.7403 | -10.485 | 18.7806 | 1620.3433 | 9.7403 | 98.1685
1.0577 | -17.2019 | 34.5974 | -1616.9625 | 1617.7221 | 10.8722 | -9.6956 | -10.4403 | 18.5721 | 1616.9625 | 9.6956 | 98.3358
1.0595 | -17.9019 | 33.9719 | -1613.1498 | 1613.9094 | 10.8424 | -9.6509 | -10.3956 | 18.4976 | 1613.1498 | 9.6509 | 98.5032
1.0613 | -18.438 | 33.674 | -1608.5179 | 1609.2775 | 10.8275 | -9.5914 | -10.3658 | 18.438 | 1608.5179 | 9.5914 | 98.6705
1.0631 | -19.5848 | 31.3953 | -1607.6243 | 1608.369 | 10.7828 | -9.5318 | -10.336 | 18.3338 | 1607.6243 | 9.5318 | 98.8379
1.0649 | -21.804 | 26.823 | -1611.9732 | 1612.6583 | 10.6488 | -9.4424 | -10.2765 | 18.1402 | 1611.9732 | 9.4424 | 99.0052
1.0667 | -24.5741 | 23.3231 | -1623.456 | 1624.0964 | 10.4999 | -9.2339 | -10.1871 | 17.8721 | 1623.456 | 9.2339 | 99.1726
1.0684 | -25.7954 | 22.6827 | -1634.2686 | 1634.8793 | 10.4999 | -9.0403 | -10.0828 | 17.6785 | 1634.2686 | 9.0403 | 99.3306
1.0702 | -25.8401 | 24.738 | -1631.8708 | 1632.4963 | 10.5892 | -9.0105 | -10.0233 | 17.6785 | 1631.8708 | 9.0105 | 99.498
1.072 | -24.7231 | 25.7954 | -1627.2985 | 1627.8794 | 10.6339 | -9.085 | -9.9488 | 17.7232 | 1627.2985 | 9.085 | 99.6653
1.0738 | -23.5912 | 26.8528 | -1622.7263 | 1623.2624 | 10.6637 | -9.1743 | -9.8892 | 17.7529 | 1622.7263 | 9.1743 | 99.8327
1.0756 | -23.6209 | 27.5826 | -1623.1433 | 1623.6943 | 10.7679 | -9.1446 | -9.8148 | 17.7827 | 1623.1433 | 9.1446 | 100
| Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My
Minimum | -98.5646 | -87.6626 | -3395.3249 | 1109.7228 | 5.168 | -10.4701 | -10.8871 | 7.2233 | 1107.9207 | -18.6019
Maximum | 22.9061 | 80.767 | -1107.9207 | 3396.2632 | 22.1167 | 18.6019 | -3.7829 | 28.5507 | 3395.3249 | 10.4701

## Diagr_HIGH100
HIGH100 = Loads at High Level in Subject with 100kg Body Weight.
Contact forces and moments, measured telemetrically in knee implants of 8 subjects.
Averages of one individual subject.
Fx  lateral, Fy  anterior, -Fz  distal, Fres = Resultant Force
Mx around x, My around y, Mz around z, Mres = Resultant Moment
z-axis = axis of implant stem!
From Bergmann et al, PlosOne (2014) "Standardized Loads Acting in Knee Implants"
Subject K3R
Subject K3R
G.Bergmann et al., PlosOne, 2014
G.Bergmann et al., PlosOne, 2014