## Dat_Individ
Data File: | k8l_2266_0_6_i1.akf
Curve Alignment Program: | 1.5.7
------------------------MEASURING INFOS------------------------
Diagram Title #1 | Forces and Moments at Knee Joint
Diagram Title #2 | Measured Relative to Implant
Comment #1 | Knee Joint; Gaitanalysis; Stairs up;
Comment #2 | K8L, 25 Months PO *
BodyWeight [N]: | 775
Patient(en): | K8L | [7 empty] | Subject | K8L
Implant Type: | Knee Joint
Varus Angle Leg [°] | 4 | [7 empty] | Individual loads in N (forces) and Nm (moments)
Axes Directions X, Y, Z: | Lateral, Anterior, Superior
Axes Origin: | Tibial Plateau at Stem Axis | [7 empty] | BW [kg] = | 79.001
Transformation X Y Z (°): | 11.5 | -1 | 12
Displacement X, Y, Z (mm): | 0 | 0 | 0
Activity Code: | KneeActivities | 1P_1_15_4_1
---------------------------DATA--------------------------------
Number of Data Sets = | 601
Number of analog channels = | 0
Time-Offset(sec) = | 0
Time | Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My | Time
[s] | [N] | [N] | [N] | [N] | [Nm] | [Nm] | [Nm] | [Nm] | [N] | (Nm] | [%Cycle]
0 | -2.67 | 23.32 | -172.56 | 174.24 | -0.39 | -1.12 | -0.73 | 1.4 | 172.56 | 1.12 | 0
0.0019 | -2.95 | 22.97 | -172.45 | 174.09 | -0.39 | -1.08 | -0.73 | 1.37 | 172.45 | 1.08 | 0.16678
0.0038 | -3.03 | 22.75 | -172.16 | 173.78 | -0.4 | -1.06 | -0.72 | 1.35 | 172.16 | 1.06 | 0.33357
0.0057 | -3.01 | 22.61 | -171.85 | 173.45 | -0.4 | -1.05 | -0.71 | 1.34 | 171.85 | 1.05 | 0.50035
0.0076 | -2.92 | 22.8 | -174.08 | 175.68 | -0.39 | -1.05 | -0.7 | 1.33 | 174.08 | 1.05 | 0.66713
0.0095 | -2.97 | 22.63 | -176.1 | 177.68 | -0.4 | -1.03 | -0.69 | 1.31 | 176.1 | 1.03 | 0.83392
0.0114 | -3 | 22.64 | -177.66 | 179.25 | -0.39 | -1.02 | -0.68 | 1.3 | 177.66 | 1.02 | 1.0007
0.0133 | -2.93 | 22.97 | -178.56 | 180.18 | -0.37 | -1.01 | -0.67 | 1.28 | 178.56 | 1.01 | 1.1675
0.0152 | -2.89 | 23.23 | -178.61 | 180.26 | -0.36 | -1 | -0.67 | 1.26 | 178.61 | 1 | 1.3343
0.0171 | -3.02 | 23.3 | -177.59 | 179.25 | -0.35 | -0.97 | -0.66 | 1.24 | 177.59 | 0.97 | 1.5011
0.019 | -3.38 | 23.1 | -176.05 | 177.7 | -0.35 | -0.93 | -0.66 | 1.2 | 176.05 | 0.93 | 1.6678
0.0209 | -3.78 | 22.77 | -174.43 | 176.06 | -0.36 | -0.89 | -0.65 | 1.17 | 174.43 | 0.89 | 1.8346
0.0228 | -3.84 | 22.76 | -174.24 | 175.88 | -0.37 | -0.87 | -0.63 | 1.15 | 174.24 | 0.87 | 2.0014
0.0247 | -3.75 | 23.28 | -175.2 | 176.91 | -0.37 | -0.86 | -0.62 | 1.14 | 175.2 | 0.86 | 2.1682
0.0266 | -3.67 | 23.84 | -176.23 | 178.02 | -0.38 | -0.86 | -0.61 | 1.13 | 176.23 | 0.86 | 2.335
0.0285 | -3.4 | 24.43 | -176.93 | 178.8 | -0.38 | -0.86 | -0.6 | 1.13 | 176.93 | 0.86 | 2.5018
0.0304 | -3.21 | 25.02 | -176.46 | 178.41 | -0.37 | -0.87 | -0.59 | 1.12 | 176.46 | 0.87 | 2.6685
0.0323 | -2.91 | 25.76 | -175.81 | 177.87 | -0.35 | -0.88 | -0.58 | 1.12 | 175.81 | 0.88 | 2.8353
0.0342 | -2.58 | 26.58 | -174.77 | 176.91 | -0.36 | -0.9 | -0.58 | 1.14 | 174.77 | 0.9 | 3.0021
0.0361 | -2.09 | 27.67 | -174.04 | 176.33 | -0.37 | -0.93 | -0.58 | 1.16 | 174.04 | 0.93 | 3.1689
0.038 | -2.04 | 27.9 | -172.34 | 174.7 | -0.38 | -0.92 | -0.57 | 1.15 | 172.34 | 0.92 | 3.3357
0.0399 | -2.14 | 27.95 | -170.23 | 172.61 | -0.41 | -0.91 | -0.56 | 1.15 | 170.23 | 0.91 | 3.5025
0.0418 | -2.16 | 28.16 | -168.86 | 171.29 | -0.43 | -0.9 | -0.55 | 1.15 | 168.86 | 0.9 | 3.6692
0.0437 | -1.89 | 28.69 | -169.98 | 172.49 | -0.45 | -0.92 | -0.54 | 1.17 | 169.98 | 0.92 | 3.836
0.0456 | -1.62 | 29.26 | -171.85 | 174.43 | -0.48 | -0.94 | -0.53 | 1.19 | 171.85 | 0.94 | 4.0028
0.0475 | -1.33 | 29.9 | -173.55 | 176.21 | -0.5 | -0.97 | -0.53 | 1.21 | 173.55 | 0.97 | 4.1696
0.0494 | -1.01 | 30.76 | -174.8 | 177.58 | -0.51 | -0.99 | -0.52 | 1.23 | 174.8 | 0.99 | 4.3364
0.0513 | -0.69 | 31.63 | -175.4 | 178.32 | -0.52 | -1.01 | -0.52 | 1.26 | 175.4 | 1.01 | 4.5032
0.0532 | -0.45 | 32.77 | -175 | 178.14 | -0.53 | -1.04 | -0.53 | 1.28 | 175 | 1.04 | 4.6699
0.0551 | -0.22 | 33.51 | -174.29 | 177.6 | -0.55 | -1.05 | -0.54 | 1.31 | 174.29 | 1.05 | 4.8367
0.057 | 0.01 | 34.03 | -173.76 | 177.19 | -0.57 | -1.07 | -0.54 | 1.33 | 173.76 | 1.07 | 5.0035
0.0589 | 0.24 | 34.61 | -174.37 | 177.89 | -0.6 | -1.09 | -0.54 | 1.36 | 174.37 | 1.09 | 5.1703
0.0608 | 0.47 | 34.86 | -175.47 | 179.02 | -0.64 | -1.11 | -0.54 | 1.39 | 175.47 | 1.11 | 5.3371
0.0627 | 0.71 | 35.27 | -177.35 | 180.94 | -0.69 | -1.13 | -0.54 | 1.43 | 177.35 | 1.13 | 5.5039
0.0646 | 1.18 | 36.23 | -180.05 | 183.78 | -0.7 | -1.17 | -0.54 | 1.48 | 180.05 | 1.17 | 5.6706
0.0665 | 1.74 | 37.27 | -182.96 | 186.83 | -0.72 | -1.21 | -0.55 | 1.51 | 182.96 | 1.21 | 5.8374
0.0684 | 2.13 | 38.03 | -184.75 | 188.73 | -0.75 | -1.24 | -0.57 | 1.55 | 184.75 | 1.24 | 6.0042
0.0703 | 2.48 | 38.77 | -186.33 | 190.41 | -0.78 | -1.26 | -0.58 | 1.59 | 186.33 | 1.26 | 6.171
0.0721 | 2.85 | 39.48 | -188.85 | 193.02 | -0.82 | -1.29 | -0.59 | 1.64 | 188.85 | 1.29 | 6.329
0.074 | 3.51 | 40.23 | -193.15 | 197.39 | -0.87 | -1.33 | -0.6 | 1.7 | 193.15 | 1.33 | 6.4958
0.0759 | 4.21 | 41.25 | -197.8 | 202.15 | -0.9 | -1.38 | -0.61 | 1.76 | 197.8 | 1.38 | 6.6626
0.0778 | 4.91 | 42.27 | -202.32 | 206.81 | -0.93 | -1.43 | -0.62 | 1.82 | 202.32 | 1.43 | 6.8294
0.0797 | 5.58 | 43.52 | -206.4 | 211.07 | -0.95 | -1.48 | -0.64 | 1.87 | 206.4 | 1.48 | 6.9961
0.0816 | 6.07 | 44.73 | -209.38 | 214.25 | -0.98 | -1.52 | -0.66 | 1.93 | 209.38 | 1.52 | 7.1629
0.0835 | 6.33 | 45.68 | -211.23 | 216.27 | -1 | -1.54 | -0.68 | 1.96 | 211.23 | 1.54 | 7.3297
0.0854 | 6.49 | 46.33 | -212.6 | 217.74 | -1.03 | -1.56 | -0.7 | 1.99 | 212.6 | 1.56 | 7.4965
0.0873 | 6.65 | 47 | -214.16 | 219.42 | -1.06 | -1.57 | -0.71 | 2.02 | 214.16 | 1.57 | 7.6633
0.0892 | 7 | 48.06 | -217.07 | 222.5 | -1.09 | -1.59 | -0.72 | 2.06 | 217.07 | 1.59 | 7.8301
0.0911 | 7.44 | 49.42 | -220.99 | 226.65 | -1.12 | -1.62 | -0.73 | 2.1 | 220.99 | 1.62 | 7.9968
0.093 | 7.86 | 50.99 | -225.68 | 231.58 | -1.13 | -1.65 | -0.75 | 2.14 | 225.68 | 1.65 | 8.1636
0.0949 | 8.59 | 52.95 | -230.95 | 237.16 | -1.13 | -1.7 | -0.77 | 2.18 | 230.95 | 1.7 | 8.3304
0.0968 | 9.04 | 54.6 | -234.76 | 241.25 | -1.14 | -1.71 | -0.78 | 2.2 | 234.76 | 1.71 | 8.4972
0.0987 | 9.16 | 55.87 | -237.48 | 244.2 | -1.16 | -1.69 | -0.79 | 2.2 | 237.48 | 1.69 | 8.664
0.1006 | 9.25 | 57.05 | -240.04 | 246.96 | -1.17 | -1.67 | -0.8 | 2.19 | 240.04 | 1.67 | 8.8308
0.1025 | 9.49 | 58.34 | -243.36 | 250.5 | -1.2 | -1.65 | -0.81 | 2.2 | 243.36 | 1.65 | 8.9975
0.1044 | 9.93 | 59.72 | -247.89 | 255.24 | -1.23 | -1.66 | -0.82 | 2.22 | 247.89 | 1.66 | 9.1643
0.1063 | 10.5 | 61.64 | -253.85 | 261.5 | -1.24 | -1.66 | -0.83 | 2.23 | 253.85 | 1.66 | 9.3311
0.1082 | 11.09 | 63.77 | -260.31 | 268.29 | -1.25 | -1.66 | -0.83 | 2.24 | 260.31 | 1.66 | 9.4979
0.1101 | 11.56 | 65.71 | -266.38 | 274.66 | -1.26 | -1.65 | -0.83 | 2.24 | 266.38 | 1.65 | 9.6647
0.112 | 11.82 | 67.41 | -271.46 | 280.01 | -1.27 | -1.61 | -0.84 | 2.22 | 271.46 | 1.61 | 9.8315
0.1139 | 12.04 | 68.92 | -275.63 | 284.42 | -1.29 | -1.57 | -0.84 | 2.21 | 275.63 | 1.57 | 9.9982
0.1158 | 12.05 | 69.92 | -278.73 | 287.67 | -1.31 | -1.52 | -0.84 | 2.19 | 278.73 | 1.52 | 10.165
0.1177 | 11.89 | 70.82 | -281.88 | 290.94 | -1.35 | -1.46 | -0.83 | 2.18 | 281.88 | 1.46 | 10.3318
0.1196 | 11.78 | 71.82 | -285.44 | 294.64 | -1.38 | -1.41 | -0.83 | 2.16 | 285.44 | 1.41 | 10.4986
0.1215 | 11.65 | 73.14 | -290.16 | 299.53 | -1.39 | -1.35 | -0.82 | 2.13 | 290.16 | 1.35 | 10.6654
0.1234 | 11.67 | 74.61 | -295.67 | 305.22 | -1.41 | -1.29 | -0.8 | 2.1 | 295.67 | 1.29 | 10.8322
0.1253 | 11.8 | 76.11 | -301.61 | 311.36 | -1.42 | -1.24 | -0.79 | 2.08 | 301.61 | 1.24 | 10.9989
0.1272 | 11.96 | 77.58 | -307.11 | 317.05 | -1.43 | -1.19 | -0.77 | 2.06 | 307.11 | 1.19 | 11.1657
0.1291 | 12.09 | 78.81 | -310.88 | 321.01 | -1.44 | -1.14 | -0.76 | 2.05 | 310.88 | 1.14 | 11.3325
0.131 | 12.05 | 79.85 | -313.73 | 324.03 | -1.45 | -1.08 | -0.75 | 2.03 | 313.73 | 1.08 | 11.4993
0.1329 | 11.64 | 80.43 | -314.98 | 325.36 | -1.47 | -1.01 | -0.74 | 2 | 314.98 | 1.01 | 11.6661
0.1348 | 11.23 | 80.88 | -316.63 | 327.06 | -1.49 | -0.95 | -0.73 | 1.98 | 316.63 | 0.95 | 11.8329
0.1367 | 11.03 | 81.46 | -320.26 | 330.71 | -1.51 | -0.89 | -0.71 | 1.96 | 320.26 | 0.89 | 11.9996
0.1386 | 10.98 | 82.1 | -323.8 | 334.3 | -1.53 | -0.84 | -0.69 | 1.94 | 323.8 | 0.84 | 12.1664
0.1405 | 11.12 | 83.03 | -326.9 | 337.53 | -1.54 | -0.78 | -0.67 | 1.92 | 326.9 | 0.78 | 12.3332
0.1424 | 11.28 | 83.94 | -329.76 | 340.52 | -1.54 | -0.73 | -0.65 | 1.9 | 329.76 | 0.73 | 12.5
0.1443 | 11.37 | 84.6 | -331.15 | 342.03 | -1.55 | -0.69 | -0.64 | 1.89 | 331.15 | 0.69 | 12.6668
0.1462 | 11.14 | 84.85 | -331.66 | 342.57 | -1.57 | -0.64 | -0.62 | 1.89 | 331.66 | 0.64 | 12.8336
0.1481 | 10.57 | 84.61 | -331.34 | 342.18 | -1.6 | -0.58 | -0.61 | 1.9 | 331.34 | 0.58 | 13.0004
0.15 | 9.91 | 84.28 | -331.24 | 341.97 | -1.63 | -0.51 | -0.6 | 1.9 | 331.24 | 0.51 | 13.1671
0.1519 | 9.48 | 84.57 | -333.65 | 344.36 | -1.65 | -0.45 | -0.58 | 1.91 | 333.65 | 0.45 | 13.3339
0.1538 | 9.07 | 85.06 | -337.21 | 347.92 | -1.67 | -0.39 | -0.56 | 1.92 | 337.21 | 0.39 | 13.5007
0.1557 | 8.74 | 85.73 | -342.19 | 352.9 | -1.69 | -0.33 | -0.53 | 1.96 | 342.19 | 0.33 | 13.6675
0.1576 | 8.59 | 86.56 | -346.72 | 357.49 | -1.71 | -0.29 | -0.52 | 2 | 346.72 | 0.29 | 13.8343
0.1595 | 8.39 | 87.09 | -350.22 | 361.01 | -1.72 | -0.25 | -0.5 | 2.02 | 350.22 | 0.25 | 14.0011
0.1614 | 8.26 | 87.48 | -352.99 | 363.78 | -1.72 | -0.21 | -0.49 | 2.03 | 352.99 | 0.21 | 14.1678
0.1633 | 8.32 | 87.75 | -354.82 | 365.62 | -1.73 | -0.2 | -0.48 | 2.05 | 354.82 | 0.2 | 14.3346
0.1652 | 8.43 | 87.64 | -355.67 | 366.42 | -1.73 | -0.18 | -0.48 | 2.06 | 355.67 | 0.18 | 14.5014
0.1671 | 8.34 | 87.12 | -355.52 | 366.15 | -1.73 | -0.14 | -0.47 | 2.06 | 355.52 | 0.14 | 14.6682
0.169 | 8.08 | 86.23 | -354.72 | 365.16 | -1.73 | -0.08 | -0.45 | 2.09 | 354.72 | 0.08 | 14.835
0.1709 | 7.73 | 85.05 | -353.21 | 363.41 | -1.74 | -0.03 | -0.44 | 2.12 | 353.21 | 0.03 | 15.0018
0.1728 | 7.33 | 83.58 | -351.99 | 361.89 | -1.76 | 0.02 | -0.43 | 2.15 | 351.99 | -0.02 | 15.1685
0.1747 | 6.88 | 82.08 | -351.73 | 361.3 | -1.79 | 0.06 | -0.42 | 2.19 | 351.73 | -0.06 | 15.3353
0.1766 | 6.45 | 80.71 | -352.08 | 361.34 | -1.81 | 0.1 | -0.4 | 2.22 | 352.08 | -0.1 | 15.5021
0.1785 | 6.08 | 79.51 | -353.05 | 362.02 | -1.83 | 0.14 | -0.39 | 2.24 | 353.05 | -0.14 | 15.6689
0.1804 | 6.06 | 78.61 | -356.14 | 364.86 | -1.82 | 0.15 | -0.39 | 2.24 | 356.14 | -0.15 | 15.8357
0.1823 | 6 | 77.9 | -359.69 | 368.2 | -1.81 | 0.15 | -0.38 | 2.25 | 359.69 | -0.15 | 16.0025
0.1842 | 6.05 | 77.19 | -363.15 | 371.45 | -1.8 | 0.15 | -0.38 | 2.26 | 363.15 | -0.15 | 16.1692
0.1861 | 6.34 | 76.68 | -366.34 | 374.48 | -1.76 | 0.16 | -0.39 | 2.24 | 366.34 | -0.16 | 16.336
0.188 | 6.7 | 75.93 | -368.82 | 376.78 | -1.72 | 0.17 | -0.39 | 2.21 | 368.82 | -0.17 | 16.5028
0.1899 | 7.02 | 75.11 | -370.25 | 378.04 | -1.67 | 0.19 | -0.39 | 2.18 | 370.25 | -0.19 | 16.6696
0.1918 | 7.5 | 74.21 | -370.98 | 378.59 | -1.61 | 0.21 | -0.4 | 2.14 | 370.98 | -0.21 | 16.8364
0.1937 | 7.8 | 73.02 | -371.01 | 378.4 | -1.56 | 0.24 | -0.4 | 2.11 | 371.01 | -0.24 | 17.0032
0.1956 | 7.96 | 71.54 | -370.26 | 377.39 | -1.52 | 0.3 | -0.4 | 2.08 | 370.26 | -0.3 | 17.1699
0.1975 | 7.73 | 69.78 | -368.68 | 375.48 | -1.49 | 0.38 | -0.39 | 2.05 | 368.68 | -0.38 | 17.3367
0.1994 | 7.47 | 67.78 | -367.5 | 373.92 | -1.47 | 0.47 | -0.38 | 2.04 | 367.5 | -0.47 | 17.5035
0.2013 | 7.57 | 65.74 | -367.68 | 373.72 | -1.43 | 0.55 | -0.38 | 2.02 | 367.68 | -0.55 | 17.6703
0.2032 | 7.74 | 63.63 | -368.75 | 374.4 | -1.39 | 0.63 | -0.37 | 2.01 | 368.75 | -0.63 | 17.8371
0.2051 | 7.99 | 61.53 | -371.34 | 376.62 | -1.35 | 0.72 | -0.36 | 2 | 371.34 | -0.72 | 18.0039
0.207 | 8.28 | 59.51 | -375 | 379.92 | -1.29 | 0.83 | -0.35 | 1.98 | 375 | -0.83 | 18.1706
0.2089 | 8.57 | 57.6 | -379.66 | 384.23 | -1.22 | 0.93 | -0.34 | 1.96 | 379.66 | -0.93 | 18.3374
0.2108 | 9.16 | 55.85 | -385.5 | 389.77 | -1.15 | 1.01 | -0.34 | 1.96 | 385.5 | -1.01 | 18.5042
0.2126 | 10.16 | 54.49 | -391.94 | 395.93 | -1.09 | 1.05 | -0.34 | 1.9 | 391.94 | -1.05 | 18.6622
0.2145 | 11.2 | 52.21 | -396.88 | 400.54 | -1.02 | 1.12 | -0.34 | 1.84 | 396.88 | -1.12 | 18.829
0.2164 | 12.19 | 49.88 | -402.28 | 405.65 | -0.94 | 1.21 | -0.34 | 1.82 | 402.28 | -1.21 | 18.9958
0.2183 | 13.03 | 47.84 | -409.29 | 412.4 | -0.84 | 1.32 | -0.33 | 1.8 | 409.29 | -1.32 | 19.1626
0.2202 | 13.56 | 45.88 | -418.2 | 421.06 | -0.75 | 1.45 | -0.33 | 1.82 | 418.2 | -1.45 | 19.3294
0.2221 | 14.03 | 43.83 | -428.23 | 430.82 | -0.67 | 1.61 | -0.33 | 1.88 | 428.23 | -1.61 | 19.4961
0.224 | 14.77 | 41.89 | -440.37 | 442.73 | -0.57 | 1.73 | -0.33 | 1.94 | 440.37 | -1.73 | 19.6629
0.2259 | 15.53 | 39.87 | -454.32 | 456.47 | -0.47 | 1.88 | -0.34 | 2.04 | 454.32 | -1.88 | 19.8297
0.2278 | 16.35 | 37.92 | -469.36 | 471.33 | -0.37 | 2.01 | -0.35 | 2.13 | 469.36 | -2.01 | 19.9965
0.2297 | 17.09 | 35.95 | -486.12 | 487.9 | -0.27 | 2.14 | -0.36 | 2.23 | 486.12 | -2.14 | 20.1633
0.2316 | 17.74 | 33.68 | -504.93 | 506.53 | -0.16 | 2.3 | -0.38 | 2.37 | 504.93 | -2.3 | 20.3301
0.2335 | 18.56 | 31.38 | -525.3 | 526.72 | -0.04 | 2.44 | -0.4 | 2.51 | 525.3 | -2.44 | 20.4968
0.2354 | 19.57 | 29.03 | -546.65 | 547.95 | 0.09 | 2.58 | -0.43 | 2.66 | 546.65 | -2.58 | 20.6636
0.2373 | 20.64 | 26.57 | -568.13 | 569.32 | 0.21 | 2.71 | -0.45 | 2.79 | 568.13 | -2.71 | 20.8304
0.2392 | 21.87 | 23.96 | -590.12 | 591.2 | 0.34 | 2.82 | -0.48 | 2.92 | 590.12 | -2.82 | 20.9972
0.2411 | 23.56 | 21.27 | -614.69 | 615.71 | 0.49 | 2.88 | -0.52 | 3 | 614.69 | -2.88 | 21.164
0.243 | 25.33 | 18.5 | -640.12 | 641.08 | 0.64 | 2.93 | -0.55 | 3.09 | 640.12 | -2.93 | 21.3308
0.2449 | 27.27 | 15.62 | -666.75 | 667.66 | 0.8 | 2.96 | -0.59 | 3.16 | 666.75 | -2.96 | 21.4975
0.2468 | 29.55 | 12.69 | -695.13 | 696.05 | 0.98 | 2.96 | -0.63 | 3.23 | 695.13 | -2.96 | 21.6643
0.2487 | 32.04 | 9.71 | -723.58 | 724.53 | 1.17 | 2.93 | -0.67 | 3.3 | 723.58 | -2.93 | 21.8311
0.2506 | 34.55 | 6.73 | -752.19 | 753.19 | 1.35 | 2.89 | -0.72 | 3.36 | 752.19 | -2.89 | 21.9979
0.2525 | 37.1 | 3.8 | -781.76 | 782.81 | 1.54 | 2.84 | -0.77 | 3.43 | 781.76 | -2.84 | 22.1647
0.2544 | 39.88 | 1 | -812.99 | 814.13 | 1.74 | 2.74 | -0.82 | 3.53 | 812.99 | -2.74 | 22.3315
0.2563 | 42.78 | -1.51 | -845.78 | 847.01 | 1.96 | 2.61 | -0.87 | 3.64 | 845.78 | -2.61 | 22.4982
0.2582 | 45.61 | -3.87 | -879.08 | 880.43 | 2.18 | 2.49 | -0.91 | 3.76 | 879.08 | -2.49 | 22.665
0.2601 | 48.47 | -6.04 | -911.6 | 913.06 | 2.41 | 2.35 | -0.96 | 3.89 | 911.6 | -2.35 | 22.8318
0.262 | 51.44 | -8.35 | -945.33 | 946.92 | 2.63 | 2.2 | -1.01 | 4.03 | 945.33 | -2.2 | 22.9986
0.2639 | 54.37 | -10.8 | -979.67 | 981.37 | 2.86 | 2.04 | -1.06 | 4.18 | 979.67 | -2.04 | 23.1654
0.2658 | 57.14 | -13.05 | -1013.3 | 1015.13 | 3.09 | 1.86 | -1.11 | 4.35 | 1013.3 | -1.86 | 23.3322
0.2677 | 59.68 | -15.52 | -1047.17 | 1049.11 | 3.3 | 1.69 | -1.16 | 4.54 | 1047.17 | -1.69 | 23.4989
0.2696 | 62.2 | -18.05 | -1080.36 | 1082.42 | 3.5 | 1.5 | -1.21 | 4.74 | 1080.36 | -1.5 | 23.6657
0.2715 | 65.07 | -20.44 | -1117.88 | 1120.08 | 3.74 | 1.31 | -1.27 | 4.95 | 1117.88 | -1.31 | 23.8325
0.2734 | 67.83 | -22.77 | -1156.56 | 1158.88 | 3.99 | 1.15 | -1.33 | 5.13 | 1156.56 | -1.15 | 23.9993
0.2753 | 71.09 | -25.45 | -1196.03 | 1198.52 | 4.25 | 0.98 | -1.39 | 5.33 | 1196.03 | -0.98 | 24.1661
0.2772 | 74.08 | -28.03 | -1233.95 | 1236.59 | 4.5 | 0.79 | -1.45 | 5.56 | 1233.95 | -0.79 | 24.3329
0.2791 | 77.07 | -30.67 | -1271.43 | 1274.22 | 4.75 | 0.59 | -1.5 | 5.79 | 1271.43 | -0.59 | 24.4996
0.281 | 80 | -33.27 | -1308.15 | 1311.11 | 5 | 0.39 | -1.56 | 6.04 | 1308.15 | -0.39 | 24.6664
0.2829 | 82.25 | -35.74 | -1340.78 | 1343.86 | 5.22 | 0.23 | -1.61 | 6.26 | 1340.78 | -0.23 | 24.8332
0.2848 | 84.26 | -38.17 | -1371.57 | 1374.78 | 5.42 | 0.09 | -1.67 | 6.47 | 1371.57 | -0.09 | 25
0.2867 | 86.52 | -41.21 | -1408.34 | 1411.69 | 5.66 | -0.09 | -1.72 | 6.72 | 1408.34 | 0.09 | 25.1668
0.2886 | 88.78 | -44.34 | -1446.72 | 1450.21 | 5.91 | -0.29 | -1.78 | 6.99 | 1446.72 | 0.29 | 25.3336
0.2905 | 91.17 | -47.57 | -1486.85 | 1490.49 | 6.17 | -0.49 | -1.84 | 7.26 | 1486.85 | 0.49 | 25.5004
0.2924 | 93.47 | -50.56 | -1526.86 | 1530.64 | 6.44 | -0.71 | -1.9 | 7.54 | 1526.86 | 0.71 | 25.6671
0.2943 | 95.71 | -53.32 | -1567.1 | 1571.01 | 6.71 | -0.93 | -1.96 | 7.84 | 1567.1 | 0.93 | 25.8339
0.2962 | 97.96 | -56.03 | -1607.74 | 1611.79 | 7 | -1.16 | -2.03 | 8.14 | 1607.74 | 1.16 | 26.0007
0.2981 | 100.26 | -58.75 | -1648.64 | 1652.82 | 7.29 | -1.4 | -2.09 | 8.45 | 1648.64 | 1.4 | 26.1675
0.3 | 102.59 | -61.39 | -1689.6 | 1693.92 | 7.58 | -1.65 | -2.16 | 8.77 | 1689.6 | 1.65 | 26.3343
0.3019 | 104.99 | -64.11 | -1731.66 | 1736.11 | 7.89 | -1.9 | -2.23 | 9.11 | 1731.66 | 1.9 | 26.5011
0.3038 | 107.29 | -66.58 | -1772.46 | 1777.04 | 8.2 | -2.17 | -2.3 | 9.45 | 1772.46 | 2.17 | 26.6678
0.3057 | 109.44 | -69.06 | -1812.25 | 1816.95 | 8.51 | -2.42 | -2.37 | 9.79 | 1812.25 | 2.42 | 26.8346
0.3076 | 111.43 | -71.62 | -1850.33 | 1855.14 | 8.79 | -2.69 | -2.42 | 10.13 | 1850.33 | 2.69 | 27.0014
0.3095 | 113.34 | -74.16 | -1887.95 | 1892.88 | 9.08 | -2.96 | -2.48 | 10.47 | 1887.95 | 2.96 | 27.1682
0.3114 | 114.95 | -76.66 | -1925.43 | 1930.45 | 9.36 | -3.22 | -2.54 | 10.81 | 1925.43 | 3.22 | 27.335
0.3133 | 116.27 | -79.02 | -1961.11 | 1966.21 | 9.64 | -3.48 | -2.59 | 11.14 | 1961.11 | 3.48 | 27.5018
0.3152 | 117.5 | -81.32 | -1991.72 | 1996.91 | 9.87 | -3.75 | -2.63 | 11.43 | 1991.72 | 3.75 | 27.6685
0.3171 | 118.48 | -83.53 | -2021 | 2026.26 | 10.08 | -4.01 | -2.66 | 11.71 | 2021 | 4.01 | 27.8353
0.319 | 119.75 | -85.35 | -2049.05 | 2054.38 | 10.31 | -4.31 | -2.7 | 12.02 | 2049.05 | 4.31 | 28.0021
0.3209 | 120.94 | -86.65 | -2075.75 | 2081.14 | 10.55 | -4.6 | -2.74 | 12.34 | 2075.75 | 4.6 | 28.1689
0.3228 | 122.14 | -87.94 | -2101.93 | 2107.37 | 10.79 | -4.9 | -2.79 | 12.67 | 2101.93 | 4.9 | 28.3357
0.3247 | 123.36 | -89.5 | -2127.2 | 2132.72 | 11.02 | -5.24 | -2.82 | 12.99 | 2127.2 | 5.24 | 28.5025
0.3266 | 124.25 | -91.21 | -2150.79 | 2156.37 | 11.23 | -5.57 | -2.86 | 13.28 | 2150.79 | 5.57 | 28.6692
0.3285 | 124.67 | -92.96 | -2172.5 | 2178.12 | 11.4 | -5.89 | -2.89 | 13.55 | 2172.5 | 5.89 | 28.836
0.3304 | 124.93 | -94.71 | -2193.38 | 2199.05 | 11.57 | -6.2 | -2.92 | 13.84 | 2193.38 | 6.2 | 29.0028
0.3323 | 125.28 | -96.68 | -2213.02 | 2218.74 | 11.73 | -6.52 | -2.95 | 14.11 | 2213.02 | 6.52 | 29.1696
0.3342 | 125.39 | -98.58 | -2230.75 | 2236.51 | 11.88 | -6.85 | -2.98 | 14.39 | 2230.75 | 6.85 | 29.3364
0.3361 | 125.11 | -100.21 | -2246.93 | 2252.71 | 12.04 | -7.17 | -3.01 | 14.7 | 2246.93 | 7.17 | 29.5032
0.338 | 124.94 | -102.07 | -2263.81 | 2269.62 | 12.18 | -7.5 | -3.05 | 14.98 | 2263.81 | 7.5 | 29.6699
0.3399 | 124.75 | -103.82 | -2279.79 | 2285.63 | 12.31 | -7.83 | -3.08 | 15.23 | 2279.79 | 7.83 | 29.8367
0.3418 | 124.17 | -105.07 | -2294.2 | 2300.04 | 12.44 | -8.15 | -3.1 | 15.49 | 2294.2 | 8.15 | 30.0035
0.3437 | 123.15 | -106.12 | -2307.24 | 2313.04 | 12.57 | -8.45 | -3.13 | 15.75 | 2307.24 | 8.45 | 30.1703
0.3456 | 122.23 | -107.19 | -2320.05 | 2325.83 | 12.69 | -8.74 | -3.16 | 16.02 | 2320.05 | 8.74 | 30.3371
0.3475 | 121.42 | -107.91 | -2332.46 | 2338.2 | 12.82 | -9.05 | -3.18 | 16.27 | 2332.46 | 9.05 | 30.5039
0.3494 | 120.09 | -108.63 | -2344.5 | 2350.19 | 12.94 | -9.35 | -3.2 | 16.52 | 2344.5 | 9.35 | 30.6706
0.3513 | 118.42 | -109.31 | -2355.42 | 2361.03 | 13.06 | -9.66 | -3.23 | 16.79 | 2355.42 | 9.66 | 30.8374
0.3531 | 116.77 | -110.12 | -2365.62 | 2371.16 | 13.17 | -9.95 | -3.25 | 17.05 | 2365.62 | 9.95 | 30.9954
0.355 | 115.14 | -110.98 | -2375.26 | 2380.75 | 13.28 | -10.25 | -3.28 | 17.32 | 2375.26 | 10.25 | 31.1622
0.3569 | 113.5 | -111.85 | -2384.45 | 2389.88 | 13.38 | -10.53 | -3.3 | 17.56 | 2384.45 | 10.53 | 31.329
0.3588 | 111.98 | -112.94 | -2392.94 | 2398.34 | 13.46 | -10.79 | -3.32 | 17.78 | 2392.94 | 10.79 | 31.4958
0.3607 | 110.54 | -114.28 | -2400.43 | 2405.81 | 13.53 | -11.07 | -3.34 | 18.01 | 2400.43 | 11.07 | 31.6626
0.3626 | 109.21 | -115.86 | -2406.9 | 2412.29 | 13.58 | -11.35 | -3.36 | 18.22 | 2406.9 | 11.35 | 31.8294
0.3645 | 107.84 | -117.4 | -2412.86 | 2418.27 | 13.62 | -11.6 | -3.37 | 18.42 | 2412.86 | 11.6 | 31.9961
0.3664 | 106.1 | -118.92 | -2418.36 | 2423.76 | 13.66 | -11.8 | -3.37 | 18.59 | 2418.36 | 11.8 | 32.1629
0.3683 | 104.22 | -120.21 | -2423.74 | 2429.12 | 13.69 | -12.02 | -3.38 | 18.76 | 2423.74 | 12.02 | 32.3297
0.3702 | 102.32 | -121.46 | -2429.13 | 2434.48 | 13.72 | -12.25 | -3.39 | 18.94 | 2429.13 | 12.25 | 32.4965
0.3721 | 100.13 | -122.58 | -2434.26 | 2439.56 | 13.76 | -12.46 | -3.4 | 19.1 | 2434.26 | 12.46 | 32.6633
0.374 | 97.49 | -123.45 | -2438.19 | 2443.43 | 13.81 | -12.62 | -3.42 | 19.25 | 2438.19 | 12.62 | 32.8301
0.3759 | 95.18 | -124.57 | -2441.37 | 2446.57 | 13.85 | -12.8 | -3.43 | 19.38 | 2441.37 | 12.8 | 32.9968
0.3778 | 92.92 | -125.88 | -2443.52 | 2448.7 | 13.88 | -12.93 | -3.45 | 19.49 | 2443.52 | 12.93 | 33.1636
0.3797 | 90.38 | -126.71 | -2445.04 | 2450.17 | 13.93 | -13.07 | -3.47 | 19.62 | 2445.04 | 13.07 | 33.3304
0.3816 | 87.72 | -127.33 | -2446.35 | 2451.41 | 13.99 | -13.19 | -3.5 | 19.75 | 2446.35 | 13.19 | 33.4972
0.3835 | 84.98 | -128.58 | -2447.17 | 2452.21 | 14.03 | -13.21 | -3.54 | 19.8 | 2447.17 | 13.21 | 33.664
0.3854 | 82.22 | -129.84 | -2447.88 | 2452.92 | 14.08 | -13.24 | -3.58 | 19.86 | 2447.88 | 13.24 | 33.8308
0.3873 | 79.49 | -130.94 | -2448.33 | 2453.33 | 14.11 | -13.33 | -3.6 | 19.96 | 2448.33 | 13.33 | 33.9975
0.3892 | 76.8 | -132.09 | -2448.53 | 2453.51 | 14.11 | -13.42 | -3.62 | 20.02 | 2448.53 | 13.42 | 34.1643
0.3911 | 74.15 | -132.76 | -2448.32 | 2453.25 | 14.14 | -13.55 | -3.64 | 20.12 | 2448.32 | 13.55 | 34.3311
0.393 | 71.48 | -133.13 | -2447.18 | 2452.06 | 14.17 | -13.69 | -3.66 | 20.24 | 2447.18 | 13.69 | 34.4979
0.3949 | 68.93 | -133.9 | -2445.14 | 2449.99 | 14.18 | -13.85 | -3.68 | 20.36 | 2445.14 | 13.85 | 34.6647
0.3968 | 66.24 | -134.18 | -2442.94 | 2447.71 | 14.17 | -14.02 | -3.69 | 20.47 | 2442.94 | 14.02 | 34.8315
0.3987 | 63.36 | -134.33 | -2440.12 | 2444.83 | 14.17 | -14.07 | -3.72 | 20.51 | 2440.12 | 14.07 | 34.9982
0.4006 | 60.85 | -134.97 | -2436.88 | 2441.57 | 14.17 | -14.2 | -3.73 | 20.6 | 2436.88 | 14.2 | 35.165
0.4025 | 58.16 | -136 | -2433.12 | 2437.81 | 14.16 | -14.28 | -3.74 | 20.65 | 2433.12 | 14.28 | 35.3318
0.4044 | 55.69 | -136.97 | -2428.79 | 2433.48 | 14.13 | -14.44 | -3.72 | 20.74 | 2428.79 | 14.44 | 35.4986
0.4063 | 52.68 | -137.9 | -2424.22 | 2428.91 | 14.11 | -14.56 | -3.72 | 20.81 | 2424.22 | 14.56 | 35.6654
0.4082 | 49.79 | -138.7 | -2419.39 | 2424.07 | 14.08 | -14.67 | -3.73 | 20.87 | 2419.39 | 14.67 | 35.8322
0.4101 | 47.34 | -139.79 | -2413.81 | 2418.54 | 14.07 | -14.73 | -3.74 | 20.9 | 2413.81 | 14.73 | 35.9989
0.412 | 45 | -140.44 | -2408.64 | 2413.37 | 14.04 | -14.88 | -3.75 | 20.96 | 2408.64 | 14.88 | 36.1657
0.4139 | 42.62 | -140.84 | -2404.54 | 2409.26 | 14.02 | -14.93 | -3.78 | 20.98 | 2404.54 | 14.93 | 36.3325
0.4158 | 40.31 | -141.13 | -2400.61 | 2405.31 | 14.01 | -14.98 | -3.81 | 21 | 2400.61 | 14.98 | 36.4993
0.4177 | 38.1 | -141.3 | -2396.12 | 2400.8 | 14 | -15.04 | -3.83 | 21.03 | 2396.12 | 15.04 | 36.6661
0.4196 | 35.79 | -141.03 | -2390.45 | 2395.1 | 13.97 | -15.11 | -3.84 | 21.07 | 2390.45 | 15.11 | 36.8329
0.4215 | 32.99 | -140.41 | -2384.38 | 2388.96 | 13.92 | -15.2 | -3.86 | 21.11 | 2384.38 | 15.2 | 36.9996
0.4234 | 30.66 | -139.98 | -2377.88 | 2382.42 | 13.92 | -15.24 | -3.87 | 21.14 | 2377.88 | 15.24 | 37.1664
0.4253 | 28.29 | -139.42 | -2370.33 | 2374.82 | 13.9 | -15.26 | -3.89 | 21.15 | 2370.33 | 15.26 | 37.3332
0.4272 | 25.98 | -139.07 | -2362.04 | 2366.5 | 13.87 | -15.26 | -3.89 | 21.14 | 2362.04 | 15.26 | 37.5
0.4291 | 23.49 | -138.54 | -2353.51 | 2357.92 | 13.84 | -15.27 | -3.9 | 21.14 | 2353.51 | 15.27 | 37.6668
0.431 | 21.39 | -137.58 | -2344.48 | 2348.83 | 13.81 | -15.24 | -3.92 | 21.11 | 2344.48 | 15.24 | 37.8336
0.4329 | 19.51 | -136.73 | -2335.17 | 2339.47 | 13.77 | -15.21 | -3.95 | 21.07 | 2335.17 | 15.21 | 38.0004
0.4348 | 17.56 | -136.01 | -2325.43 | 2329.68 | 13.73 | -15.15 | -3.97 | 21.01 | 2325.43 | 15.15 | 38.1671
0.4367 | 15.5 | -135.35 | -2316 | 2320.21 | 13.67 | -15.1 | -4 | 20.95 | 2316 | 15.1 | 38.3339
0.4386 | 13.2 | -134.74 | -2306.64 | 2310.81 | 13.6 | -15.05 | -4.04 | 20.88 | 2306.64 | 15.05 | 38.5007
0.4405 | 11.16 | -133.63 | -2297.49 | 2301.61 | 13.54 | -14.96 | -4.08 | 20.79 | 2297.49 | 14.96 | 38.6675
0.4424 | 9.2 | -132.23 | -2288.75 | 2292.8 | 13.48 | -14.88 | -4.13 | 20.7 | 2288.75 | 14.88 | 38.8343
0.4443 | 7.9 | -130.67 | -2278.39 | 2282.36 | 13.44 | -14.79 | -4.18 | 20.63 | 2278.39 | 14.79 | 39.0011
0.4462 | 6.61 | -128.84 | -2267.67 | 2271.55 | 13.4 | -14.71 | -4.23 | 20.57 | 2267.67 | 14.71 | 39.1678
0.4481 | 6.09 | -127.49 | -2256.23 | 2260.05 | 13.33 | -14.61 | -4.29 | 20.46 | 2256.23 | 14.61 | 39.3346
0.45 | 5.17 | -126.46 | -2246.14 | 2249.92 | 13.26 | -14.51 | -4.33 | 20.35 | 2246.14 | 14.51 | 39.5014
0.4519 | 4.28 | -125.64 | -2236.92 | 2240.67 | 13.18 | -14.41 | -4.36 | 20.24 | 2236.92 | 14.41 | 39.6682
0.4538 | 3.57 | -124.56 | -2226.99 | 2230.7 | 13.09 | -14.32 | -4.4 | 20.12 | 2226.99 | 14.32 | 39.835
0.4557 | 3.17 | -123.04 | -2215.86 | 2219.51 | 13 | -14.23 | -4.44 | 20.01 | 2215.86 | 14.23 | 40.0018
0.4576 | 2.89 | -121.64 | -2204.04 | 2207.63 | 12.91 | -14.12 | -4.47 | 19.88 | 2204.04 | 14.12 | 40.1685
0.4595 | 2.39 | -119.63 | -2192.44 | 2195.95 | 12.84 | -13.97 | -4.52 | 19.74 | 2192.44 | 13.97 | 40.3353
0.4614 | 1.84 | -117.9 | -2180.35 | 2183.78 | 12.75 | -13.81 | -4.57 | 19.58 | 2180.35 | 13.81 | 40.5021
0.4633 | 1.35 | -116.43 | -2168.41 | 2171.78 | 12.64 | -13.65 | -4.61 | 19.41 | 2168.41 | 13.65 | 40.6689
0.4652 | 1.07 | -114.99 | -2156.39 | 2159.69 | 12.53 | -13.49 | -4.65 | 19.23 | 2156.39 | 13.49 | 40.8357
0.4671 | 0.95 | -113.3 | -2145.2 | 2148.43 | 12.43 | -13.36 | -4.7 | 19.09 | 2145.2 | 13.36 | 41.0025
0.469 | 0.97 | -111.47 | -2134.68 | 2137.83 | 12.35 | -13.25 | -4.74 | 18.97 | 2134.68 | 13.25 | 41.1692
0.4709 | 1.13 | -109.45 | -2121.73 | 2124.78 | 12.25 | -13.12 | -4.79 | 18.82 | 2121.73 | 13.12 | 41.336
0.4728 | 1.41 | -107.72 | -2107.79 | 2110.77 | 12.12 | -12.99 | -4.83 | 18.66 | 2107.79 | 12.99 | 41.5028
0.4747 | 1.63 | -106.31 | -2093.7 | 2096.63 | 11.99 | -12.83 | -4.86 | 18.48 | 2093.7 | 12.83 | 41.6696
0.4766 | 2.06 | -104.82 | -2079.46 | 2082.33 | 11.86 | -12.7 | -4.9 | 18.32 | 2079.46 | 12.7 | 41.8364
0.4785 | 3.02 | -103.18 | -2064.15 | 2066.97 | 11.72 | -12.59 | -4.93 | 18.17 | 2064.15 | 12.59 | 42.0032
0.4804 | 4.07 | -101.51 | -2048.7 | 2051.45 | 11.59 | -12.49 | -4.95 | 18.02 | 2048.7 | 12.49 | 42.1699
0.4823 | 4.98 | -99.73 | -2033 | 2035.68 | 11.45 | -12.39 | -4.98 | 17.87 | 2033 | 12.39 | 42.3367
0.4842 | 5.53 | -97.84 | -2017.37 | 2019.97 | 11.32 | -12.27 | -5.01 | 17.72 | 2017.37 | 12.27 | 42.5035
0.4861 | 5.84 | -95.86 | -2001.9 | 2004.42 | 11.19 | -12.15 | -5.03 | 17.57 | 2001.9 | 12.15 | 42.6703
0.488 | 6.09 | -93.96 | -1986.56 | 1989.01 | 11.07 | -12.01 | -5.06 | 17.41 | 1986.56 | 12.01 | 42.8371
0.4899 | 6.42 | -92.26 | -1970.47 | 1972.86 | 10.92 | -11.94 | -5.08 | 17.27 | 1970.47 | 11.94 | 43.0039
0.4918 | 6.78 | -90.49 | -1954.11 | 1956.43 | 10.78 | -11.83 | -5.09 | 17.12 | 1954.11 | 11.83 | 43.1706
0.4936 | 7.46 | -88.81 | -1937.45 | 1939.71 | 10.65 | -11.73 | -5.1 | 16.97 | 1937.45 | 11.73 | 43.3287
0.4955 | 8.38 | -87.18 | -1920.55 | 1922.76 | 10.51 | -11.65 | -5.11 | 16.83 | 1920.55 | 11.65 | 43.4954
0.4974 | 9.64 | -85.58 | -1902.76 | 1904.91 | 10.37 | -11.55 | -5.12 | 16.68 | 1902.76 | 11.55 | 43.6622
0.4993 | 10.93 | -84 | -1885.77 | 1887.88 | 10.23 | -11.49 | -5.12 | 16.55 | 1885.77 | 11.49 | 43.829
0.5012 | 11.97 | -82.36 | -1869.29 | 1871.34 | 10.09 | -11.43 | -5.13 | 16.43 | 1869.29 | 11.43 | 43.9958
0.5031 | 12.8 | -80.72 | -1853.06 | 1855.07 | 9.96 | -11.38 | -5.13 | 16.3 | 1853.06 | 11.38 | 44.1626
0.505 | 13.74 | -79.28 | -1837.14 | 1839.11 | 9.82 | -11.33 | -5.13 | 16.19 | 1837.14 | 11.33 | 44.3294
0.5069 | 14.91 | -77.9 | -1822.17 | 1824.11 | 9.7 | -11.29 | -5.14 | 16.08 | 1822.17 | 11.29 | 44.4961
0.5088 | 15.71 | -76.57 | -1808.72 | 1810.63 | 9.59 | -11.27 | -5.14 | 16 | 1808.72 | 11.27 | 44.6629
0.5107 | 15.99 | -75.04 | -1796.74 | 1798.61 | 9.53 | -11.21 | -5.15 | 15.92 | 1796.74 | 11.21 | 44.8297
0.5126 | 16 | -73.32 | -1784.29 | 1786.1 | 9.45 | -11.17 | -5.15 | 15.84 | 1784.29 | 11.17 | 44.9965
0.5145 | 16.21 | -71.77 | -1770.42 | 1772.18 | 9.35 | -11.11 | -5.15 | 15.74 | 1770.42 | 11.11 | 45.1633
0.5164 | 16.39 | -70.48 | -1756.6 | 1758.32 | 9.24 | -11.06 | -5.16 | 15.63 | 1756.6 | 11.06 | 45.3301
0.5183 | 16.6 | -69.35 | -1743.09 | 1744.78 | 9.14 | -11 | -5.16 | 15.53 | 1743.09 | 11 | 45.4968
0.5202 | 17.14 | -68.45 | -1730.04 | 1731.72 | 9.03 | -10.96 | -5.17 | 15.43 | 1730.04 | 10.96 | 45.6636
0.5221 | 18.02 | -67.79 | -1717.28 | 1718.96 | 8.92 | -10.92 | -5.17 | 15.34 | 1717.28 | 10.92 | 45.8304
0.524 | 18.92 | -67.19 | -1703.78 | 1705.46 | 8.81 | -10.9 | -5.17 | 15.25 | 1703.78 | 10.9 | 45.9972
0.5259 | 19.8 | -66.63 | -1689.74 | 1691.43 | 8.69 | -10.87 | -5.17 | 15.15 | 1689.74 | 10.87 | 46.164
0.5278 | 20.48 | -65.87 | -1675.71 | 1677.39 | 8.58 | -10.86 | -5.16 | 15.06 | 1675.71 | 10.86 | 46.3308
0.5297 | 21.18 | -65.25 | -1661.77 | 1663.43 | 8.48 | -10.83 | -5.16 | 14.97 | 1661.77 | 10.83 | 46.4975
0.5316 | 21.92 | -64.54 | -1647.64 | 1649.3 | 8.37 | -10.81 | -5.15 | 14.89 | 1647.64 | 10.81 | 46.6643
0.5335 | 22.69 | -63.91 | -1633.55 | 1635.2 | 8.27 | -10.8 | -5.14 | 14.82 | 1633.55 | 10.8 | 46.8311
0.5354 | 23.51 | -63.14 | -1618.85 | 1620.49 | 8.17 | -10.81 | -5.13 | 14.74 | 1618.85 | 10.81 | 46.9979
0.5373 | 24.07 | -62.34 | -1604.91 | 1606.53 | 8.08 | -10.8 | -5.12 | 14.67 | 1604.91 | 10.8 | 47.1647
0.5392 | 24.55 | -61.44 | -1591.69 | 1593.29 | 7.99 | -10.8 | -5.11 | 14.6 | 1591.69 | 10.8 | 47.3315
0.5411 | 25.02 | -60.94 | -1578.96 | 1580.54 | 7.91 | -10.8 | -5.11 | 14.54 | 1578.96 | 10.8 | 47.4982
0.543 | 25.6 | -60.5 | -1566.59 | 1568.17 | 7.83 | -10.8 | -5.1 | 14.48 | 1566.59 | 10.8 | 47.665
0.5449 | 26.09 | -59.77 | -1554.36 | 1555.93 | 7.76 | -10.8 | -5.09 | 14.42 | 1554.36 | 10.8 | 47.8318
0.5468 | 26.5 | -58.96 | -1542.13 | 1543.68 | 7.67 | -10.81 | -5.08 | 14.37 | 1542.13 | 10.81 | 47.9986
0.5487 | 26.76 | -58.29 | -1530.85 | 1532.38 | 7.59 | -10.83 | -5.07 | 14.32 | 1530.85 | 10.83 | 48.1654
0.5506 | 26.76 | -57.63 | -1520.64 | 1522.15 | 7.52 | -10.83 | -5.06 | 14.28 | 1520.64 | 10.83 | 48.3322
0.5525 | 26.59 | -57 | -1511.16 | 1512.65 | 7.46 | -10.83 | -5.05 | 14.23 | 1511.16 | 10.83 | 48.4989
0.5544 | 26.41 | -56.57 | -1502.77 | 1504.24 | 7.4 | -10.86 | -5.04 | 14.21 | 1502.77 | 10.86 | 48.6657
0.5563 | 26.31 | -56.18 | -1494.54 | 1496 | 7.35 | -10.9 | -5.02 | 14.21 | 1494.54 | 10.9 | 48.8325
0.5582 | 26.14 | -55.61 | -1486.02 | 1487.46 | 7.28 | -10.97 | -4.99 | 14.21 | 1486.02 | 10.97 | 48.9993
0.5601 | 26.05 | -55.38 | -1478.28 | 1479.71 | 7.22 | -11.01 | -4.98 | 14.19 | 1478.28 | 11.01 | 49.1661
0.562 | 26.21 | -55.27 | -1470.42 | 1471.85 | 7.13 | -11.07 | -4.97 | 14.18 | 1470.42 | 11.07 | 49.3329
0.5639 | 26.24 | -55.22 | -1463.08 | 1464.52 | 7.07 | -11.1 | -4.96 | 14.17 | 1463.08 | 11.1 | 49.4996
0.5658 | 26.18 | -55.15 | -1456.62 | 1458.06 | 7.05 | -11.15 | -4.95 | 14.19 | 1456.62 | 11.15 | 49.6664
0.5677 | 26.12 | -55.02 | -1450.33 | 1451.76 | 7.02 | -11.21 | -4.94 | 14.22 | 1450.33 | 11.21 | 49.8332
0.5696 | 25.81 | -54.84 | -1444.71 | 1446.14 | 7 | -11.26 | -4.92 | 14.25 | 1444.71 | 11.26 | 50
0.5715 | 25.06 | -54.57 | -1440.59 | 1441.98 | 6.99 | -11.32 | -4.91 | 14.28 | 1440.59 | 11.32 | 50.1668
0.5734 | 24.3 | -54.42 | -1436.74 | 1438.1 | 6.98 | -11.38 | -4.89 | 14.31 | 1436.74 | 11.38 | 50.3336
0.5753 | 23.89 | -54.62 | -1432.7 | 1434.05 | 6.96 | -11.46 | -4.88 | 14.36 | 1432.7 | 11.46 | 50.5004
0.5772 | 23.7 | -54.86 | -1427.75 | 1429.12 | 6.94 | -11.55 | -4.85 | 14.42 | 1427.75 | 11.55 | 50.6671
0.5791 | 23.44 | -55.17 | -1423.41 | 1424.79 | 6.94 | -11.63 | -4.83 | 14.47 | 1423.41 | 11.63 | 50.8339
0.581 | 23.08 | -55.21 | -1418.88 | 1420.26 | 6.94 | -11.72 | -4.8 | 14.54 | 1418.88 | 11.72 | 51.0007
0.5829 | 22.25 | -55.21 | -1416.04 | 1417.41 | 6.96 | -11.8 | -4.78 | 14.61 | 1416.04 | 11.8 | 51.1675
0.5848 | 21.56 | -55.56 | -1414.28 | 1415.65 | 6.99 | -11.88 | -4.76 | 14.68 | 1414.28 | 11.88 | 51.3343
0.5867 | 21.06 | -55.96 | -1412.16 | 1413.54 | 7.02 | -11.99 | -4.74 | 14.78 | 1412.16 | 11.99 | 51.5011
0.5886 | 20.75 | -56.59 | -1409.98 | 1411.39 | 7.04 | -12.09 | -4.72 | 14.87 | 1409.98 | 12.09 | 51.6678
0.5905 | 20.39 | -57.25 | -1408.05 | 1409.48 | 7.06 | -12.2 | -4.7 | 14.96 | 1408.05 | 12.2 | 51.8346
0.5924 | 19.7 | -58.09 | -1408 | 1409.46 | 7.1 | -12.29 | -4.68 | 15.04 | 1408 | 12.29 | 52.0014
0.5943 | 18.96 | -59.09 | -1408.61 | 1410.1 | 7.13 | -12.38 | -4.65 | 15.13 | 1408.61 | 12.38 | 52.1682
0.5962 | 18.64 | -60.25 | -1408.93 | 1410.48 | 7.16 | -12.53 | -4.63 | 15.26 | 1408.93 | 12.53 | 52.335
0.5981 | 18.35 | -61.42 | -1409.24 | 1410.84 | 7.18 | -12.69 | -4.6 | 15.39 | 1409.24 | 12.69 | 52.5018
0.6 | 17.55 | -62.36 | -1410.16 | 1411.79 | 7.22 | -12.81 | -4.57 | 15.51 | 1410.16 | 12.81 | 52.6685
0.6019 | 16.33 | -63.06 | -1411.73 | 1413.38 | 7.28 | -12.92 | -4.55 | 15.63 | 1411.73 | 12.92 | 52.8353
0.6038 | 15.76 | -63.91 | -1413.86 | 1415.54 | 7.34 | -13.08 | -4.52 | 15.78 | 1413.86 | 13.08 | 53.0021
0.6057 | 15.28 | -64.78 | -1416.01 | 1417.72 | 7.39 | -13.25 | -4.49 | 15.94 | 1416.01 | 13.25 | 53.1689
0.6076 | 15.13 | -65.69 | -1417.34 | 1419.1 | 7.45 | -13.4 | -4.46 | 16.08 | 1417.34 | 13.4 | 53.3357
0.6095 | 14.75 | -66.73 | -1417.74 | 1419.55 | 7.5 | -13.52 | -4.44 | 16.21 | 1417.74 | 13.52 | 53.5025
0.6114 | 14.11 | -67.76 | -1418.69 | 1420.54 | 7.56 | -13.63 | -4.41 | 16.33 | 1418.69 | 13.63 | 53.6692
0.6133 | 12.81 | -68.51 | -1421.79 | 1423.67 | 7.64 | -13.72 | -4.38 | 16.44 | 1421.79 | 13.72 | 53.836
0.6152 | 11.93 | -69.38 | -1425.56 | 1427.48 | 7.72 | -13.83 | -4.33 | 16.56 | 1425.56 | 13.83 | 54.0028
0.6171 | 11.27 | -70.37 | -1429.36 | 1431.33 | 7.8 | -13.95 | -4.28 | 16.69 | 1429.36 | 13.95 | 54.1696
0.619 | 10.91 | -71.52 | -1433.29 | 1435.32 | 7.87 | -14.07 | -4.27 | 16.82 | 1433.29 | 14.07 | 54.3364
0.6209 | 10.19 | -72.6 | -1436.66 | 1438.74 | 7.93 | -14.15 | -4.24 | 16.92 | 1436.66 | 14.15 | 54.5032
0.6228 | 9.32 | -73.55 | -1439.71 | 1441.83 | 8.01 | -14.22 | -4.22 | 17.01 | 1439.71 | 14.22 | 54.6699
0.6247 | 8.71 | -74.4 | -1443.34 | 1445.5 | 8.08 | -14.32 | -4.2 | 17.13 | 1443.34 | 14.32 | 54.8367
0.6266 | 8.21 | -75.26 | -1447.82 | 1450.01 | 8.16 | -14.43 | -4.19 | 17.26 | 1447.82 | 14.43 | 55.0035
0.6285 | 8.6 | -76.56 | -1452.34 | 1454.6 | 8.22 | -14.59 | -4.18 | 17.43 | 1452.34 | 14.59 | 55.1703
0.6304 | 8.93 | -77.99 | -1455.09 | 1457.43 | 8.27 | -14.73 | -4.16 | 17.57 | 1455.09 | 14.73 | 55.3371
0.6323 | 9.19 | -79.51 | -1456.48 | 1458.9 | 8.32 | -14.87 | -4.14 | 17.7 | 1456.48 | 14.87 | 55.5039
0.6341 | 9.18 | -81.13 | -1457.95 | 1460.46 | 8.36 | -15.02 | -4.11 | 17.85 | 1457.95 | 15.02 | 55.6619
0.636 | 9.13 | -82.93 | -1459.57 | 1462.18 | 8.41 | -15.19 | -4.08 | 18.01 | 1459.57 | 15.19 | 55.8287
0.6379 | 8.92 | -84.38 | -1462.39 | 1465.07 | 8.47 | -15.36 | -4.05 | 18.17 | 1462.39 | 15.36 | 55.9954
0.6398 | 8.58 | -85.43 | -1466.9 | 1469.63 | 8.54 | -15.52 | -4.03 | 18.34 | 1466.9 | 15.52 | 56.1622
0.6417 | 8.24 | -86.2 | -1471.92 | 1474.68 | 8.61 | -15.67 | -4.01 | 18.5 | 1471.92 | 15.67 | 56.329
0.6436 | 8.15 | -87.14 | -1477.15 | 1479.96 | 8.7 | -15.86 | -3.99 | 18.69 | 1477.15 | 15.86 | 56.4958
0.6455 | 8.24 | -88.35 | -1482.31 | 1485.18 | 8.78 | -16.05 | -3.97 | 18.9 | 1482.31 | 16.05 | 56.6626
0.6474 | 7.82 | -89.55 | -1487.67 | 1490.6 | 8.86 | -16.24 | -3.94 | 19.1 | 1487.67 | 16.24 | 56.8294
0.6493 | 7.05 | -90.62 | -1493.47 | 1496.45 | 8.95 | -16.43 | -3.9 | 19.3 | 1493.47 | 16.43 | 56.9961
0.6512 | 6.34 | -91.48 | -1500.27 | 1503.29 | 9.03 | -16.62 | -3.88 | 19.5 | 1500.27 | 16.62 | 57.1629
0.6531 | 5.89 | -92.55 | -1507.63 | 1510.71 | 9.09 | -16.82 | -3.85 | 19.7 | 1507.63 | 16.82 | 57.3297
0.655 | 5.74 | -93.82 | -1516.22 | 1519.37 | 9.16 | -17.05 | -3.84 | 19.93 | 1516.22 | 17.05 | 57.4965
0.6569 | 5.82 | -95.32 | -1525.08 | 1528.32 | 9.22 | -17.26 | -3.83 | 20.15 | 1525.08 | 17.26 | 57.6633
0.6588 | 5.58 | -96.68 | -1532.94 | 1536.25 | 9.28 | -17.43 | -3.83 | 20.32 | 1532.94 | 17.43 | 57.8301
0.6607 | 5.7 | -97.89 | -1539.69 | 1543.06 | 9.36 | -17.6 | -3.82 | 20.51 | 1539.69 | 17.6 | 57.9968
0.6626 | 5.88 | -99.35 | -1546.1 | 1549.54 | 9.42 | -17.78 | -3.81 | 20.69 | 1546.1 | 17.78 | 58.1636
0.6645 | 6.07 | -100.81 | -1552.5 | 1556.03 | 9.49 | -17.95 | -3.79 | 20.87 | 1552.5 | 17.95 | 58.3304
0.6664 | 5.68 | -102.1 | -1560.36 | 1563.96 | 9.55 | -18.13 | -3.78 | 21.06 | 1560.36 | 18.13 | 58.4972
0.6683 | 5.27 | -103.22 | -1567.64 | 1571.3 | 9.61 | -18.32 | -3.76 | 21.25 | 1567.64 | 18.32 | 58.664
0.6702 | 4.96 | -104.19 | -1575.96 | 1579.67 | 9.68 | -18.53 | -3.74 | 21.46 | 1575.96 | 18.53 | 58.8308
0.6721 | 4.58 | -105.18 | -1583.81 | 1587.57 | 9.75 | -18.72 | -3.73 | 21.66 | 1583.81 | 18.72 | 58.9975
0.674 | 4.2 | -106.12 | -1591.96 | 1595.76 | 9.82 | -18.91 | -3.72 | 21.86 | 1591.96 | 18.91 | 59.1643
0.6759 | 4.08 | -106.67 | -1601.67 | 1605.48 | 9.9 | -19.11 | -3.72 | 22.07 | 1601.67 | 19.11 | 59.3311
0.6778 | 3.96 | -107.46 | -1611.15 | 1614.98 | 9.99 | -19.28 | -3.73 | 22.25 | 1611.15 | 19.28 | 59.4979
0.6797 | 3.98 | -108.06 | -1621.53 | 1625.37 | 10.07 | -19.48 | -3.73 | 22.47 | 1621.53 | 19.48 | 59.6647
0.6816 | 3.87 | -108.99 | -1630.75 | 1634.61 | 10.16 | -19.65 | -3.74 | 22.65 | 1630.75 | 19.65 | 59.8315
0.6835 | 3.93 | -110.14 | -1639.79 | 1643.7 | 10.24 | -19.81 | -3.75 | 22.83 | 1639.79 | 19.81 | 59.9982
0.6854 | 4.08 | -111.22 | -1648.77 | 1652.72 | 10.33 | -19.98 | -3.77 | 23.01 | 1648.77 | 19.98 | 60.165
0.6873 | 4.36 | -112.07 | -1657.78 | 1661.76 | 10.42 | -20.16 | -3.78 | 23.2 | 1657.78 | 20.16 | 60.3318
0.6892 | 4.68 | -113 | -1666.6 | 1670.61 | 10.51 | -20.34 | -3.79 | 23.39 | 1666.6 | 20.34 | 60.4986
0.6911 | 5.04 | -114.03 | -1674.86 | 1678.91 | 10.59 | -20.51 | -3.8 | 23.58 | 1674.86 | 20.51 | 60.6654
0.693 | 5.61 | -115.45 | -1681.65 | 1685.76 | 10.66 | -20.71 | -3.8 | 23.77 | 1681.65 | 20.71 | 60.8322
0.6949 | 5.81 | -116.8 | -1688.45 | 1692.63 | 10.74 | -20.88 | -3.8 | 23.95 | 1688.45 | 20.88 | 60.9989
0.6968 | 5.94 | -117.69 | -1695.8 | 1700 | 10.83 | -21.04 | -3.81 | 24.13 | 1695.8 | 21.04 | 61.1657
0.6987 | 6.07 | -118.28 | -1703.49 | 1707.71 | 10.93 | -21.21 | -3.82 | 24.32 | 1703.49 | 21.21 | 61.3325
0.7006 | 6.12 | -118.77 | -1711.38 | 1715.61 | 11.04 | -21.37 | -3.82 | 24.51 | 1711.38 | 21.37 | 61.4993
0.7025 | 6.17 | -119.24 | -1719.19 | 1723.43 | 11.14 | -21.52 | -3.83 | 24.69 | 1719.19 | 21.52 | 61.6661
0.7044 | 6.15 | -119.79 | -1726.36 | 1730.62 | 11.25 | -21.66 | -3.84 | 24.86 | 1726.36 | 21.66 | 61.8329
0.7063 | 6.04 | -120.51 | -1731.49 | 1735.78 | 11.35 | -21.83 | -3.83 | 25.03 | 1731.49 | 21.83 | 61.9996
0.7082 | 5.86 | -121.27 | -1735.69 | 1740.01 | 11.44 | -22.01 | -3.8 | 25.22 | 1735.69 | 22.01 | 62.1664
0.7101 | 5.43 | -121.67 | -1740.28 | 1744.61 | 11.55 | -22.18 | -3.77 | 25.4 | 1740.28 | 22.18 | 62.3332
0.712 | 5.04 | -121.96 | -1745.19 | 1749.52 | 11.65 | -22.33 | -3.76 | 25.57 | 1745.19 | 22.33 | 62.5
0.7139 | 4.67 | -122.21 | -1750.45 | 1754.79 | 11.74 | -22.47 | -3.75 | 25.73 | 1750.45 | 22.47 | 62.6668
0.7158 | 4.24 | -122.55 | -1756.87 | 1761.21 | 11.85 | -22.6 | -3.76 | 25.89 | 1756.87 | 22.6 | 62.8336
0.7177 | 3.72 | -122.97 | -1764.34 | 1768.69 | 11.98 | -22.72 | -3.77 | 26.06 | 1764.34 | 22.72 | 63.0004
0.7196 | 3.33 | -123.58 | -1771.59 | 1775.97 | 12.12 | -22.82 | -3.8 | 26.22 | 1771.59 | 22.82 | 63.1671
0.7215 | 3.06 | -124.28 | -1775.79 | 1780.2 | 12.22 | -22.9 | -3.8 | 26.34 | 1775.79 | 22.9 | 63.3339
0.7234 | 2.79 | -124.81 | -1779.92 | 1784.35 | 12.36 | -22.99 | -3.81 | 26.49 | 1779.92 | 22.99 | 63.5007
0.7253 | 2.72 | -125.09 | -1782.5 | 1786.95 | 12.49 | -23.06 | -3.82 | 26.6 | 1782.5 | 23.06 | 63.6675
0.7272 | 2.68 | -125.43 | -1785.25 | 1789.72 | 12.6 | -23.1 | -3.83 | 26.7 | 1785.25 | 23.1 | 63.8343
0.7291 | 2.45 | -125.96 | -1787.79 | 1792.28 | 12.63 | -23.12 | -3.85 | 26.74 | 1787.79 | 23.12 | 64.0011
0.731 | 2.08 | -126.62 | -1790.55 | 1795.08 | 12.64 | -23.13 | -3.86 | 26.75 | 1790.55 | 23.13 | 64.1678
0.7329 | 1.69 | -127.26 | -1794.37 | 1798.93 | 12.68 | -23.15 | -3.88 | 26.79 | 1794.37 | 23.15 | 64.3346
0.7348 | 1.41 | -127.73 | -1799.09 | 1803.68 | 12.79 | -23.17 | -3.91 | 26.87 | 1799.09 | 23.17 | 64.5014
0.7367 | 1.14 | -128.03 | -1804.21 | 1808.8 | 12.95 | -23.21 | -3.93 | 26.99 | 1804.21 | 23.21 | 64.6682
0.7386 | 0.93 | -128.22 | -1809.22 | 1813.82 | 13.1 | -23.24 | -3.95 | 27.09 | 1809.22 | 23.24 | 64.835
0.7405 | 0.68 | -128.3 | -1813.49 | 1818.08 | 13.25 | -23.25 | -3.98 | 27.17 | 1813.49 | 23.25 | 65.0018
0.7424 | -0.18 | -127.94 | -1816.23 | 1820.79 | 13.47 | -23.17 | -3.99 | 27.21 | 1816.23 | 23.17 | 65.1685
0.7443 | -1 | -127.41 | -1818.34 | 1822.86 | 13.66 | -23.07 | -4.02 | 27.23 | 1818.34 | 23.07 | 65.3353
0.7462 | -1.68 | -126.81 | -1820.27 | 1824.74 | 13.81 | -22.97 | -4.05 | 27.23 | 1820.27 | 22.97 | 65.5021
0.7481 | -1.21 | -126.71 | -1820.93 | 1825.38 | 13.97 | -22.9 | -4.08 | 27.27 | 1820.93 | 22.9 | 65.6689
0.75 | -0.54 | -126.93 | -1821.18 | 1825.65 | 14.09 | -22.84 | -4.11 | 27.3 | 1821.18 | 22.84 | 65.8357
0.7519 | -1.46 | -127.25 | -1822.62 | 1827.11 | 14.15 | -22.74 | -4.15 | 27.26 | 1822.62 | 22.74 | 66.0025
0.7538 | -2.6 | -127.52 | -1824.81 | 1829.32 | 14.22 | -22.63 | -4.19 | 27.22 | 1824.81 | 22.63 | 66.1692
0.7557 | -4.1 | -127.71 | -1828.7 | 1833.22 | 14.32 | -22.48 | -4.24 | 27.17 | 1828.7 | 22.48 | 66.336
0.7576 | -5.34 | -128.14 | -1833.95 | 1838.49 | 14.47 | -22.34 | -4.28 | 27.15 | 1833.95 | 22.34 | 66.5028
0.7595 | -6.62 | -128.94 | -1839.22 | 1843.81 | 14.6 | -22.19 | -4.32 | 27.12 | 1839.22 | 22.19 | 66.6696
0.7614 | -7.01 | -129.17 | -1846.1 | 1850.68 | 14.73 | -22.17 | -4.36 | 27.19 | 1846.1 | 22.17 | 66.8364
0.7633 | -7.2 | -129.65 | -1853.71 | 1858.32 | 14.84 | -22.2 | -4.39 | 27.27 | 1853.71 | 22.2 | 67.0032
0.7652 | -7.45 | -129.8 | -1861.83 | 1866.42 | 15.03 | -22.25 | -4.41 | 27.42 | 1861.83 | 22.25 | 67.1699
0.7671 | -7.76 | -129.86 | -1870.07 | 1874.65 | 15.23 | -22.29 | -4.42 | 27.57 | 1870.07 | 22.29 | 67.3367
0.769 | -8.44 | -130.64 | -1878.56 | 1883.18 | 15.37 | -22.28 | -4.45 | 27.64 | 1878.56 | 22.28 | 67.5035
0.7709 | -9.4 | -131.89 | -1886.94 | 1891.63 | 15.5 | -22.19 | -4.48 | 27.66 | 1886.94 | 22.19 | 67.6703
0.7728 | -10.19 | -132.38 | -1896.74 | 1901.46 | 15.63 | -22.11 | -4.53 | 27.68 | 1896.74 | 22.11 | 67.8371
0.7746 | -10.43 | -133.21 | -1905.52 | 1910.28 | 15.79 | -22.08 | -4.58 | 27.75 | 1905.52 | 22.08 | 67.9951
0.7765 | -10.7 | -134.29 | -1913.28 | 1918.1 | 15.97 | -22.04 | -4.62 | 27.82 | 1913.28 | 22.04 | 68.1619
0.7784 | -11.46 | -135.69 | -1919.49 | 1924.39 | 16.13 | -21.97 | -4.65 | 27.87 | 1919.49 | 21.97 | 68.3287
0.7803 | -12.26 | -136.94 | -1926.12 | 1931.1 | 16.3 | -21.92 | -4.67 | 27.92 | 1926.12 | 21.92 | 68.4954
0.7822 | -13.29 | -137.63 | -1934.35 | 1939.37 | 16.49 | -21.84 | -4.69 | 27.98 | 1934.35 | 21.84 | 68.6622
0.7841 | -14.82 | -138.74 | -1941.71 | 1946.81 | 16.67 | -21.67 | -4.74 | 27.97 | 1941.71 | 21.67 | 68.829
0.786 | -15.47 | -139.62 | -1949.05 | 1954.2 | 16.83 | -21.56 | -4.78 | 27.99 | 1949.05 | 21.56 | 68.9958
0.7879 | -15.77 | -140.07 | -1956.53 | 1961.71 | 16.98 | -21.5 | -4.81 | 28.03 | 1956.53 | 21.5 | 69.1626
0.7898 | -16.68 | -140.64 | -1963.51 | 1968.72 | 17.21 | -21.39 | -4.87 | 28.07 | 1963.51 | 21.39 | 69.3294
0.7917 | -17.24 | -141.53 | -1970.31 | 1975.58 | 17.4 | -21.28 | -4.9 | 28.1 | 1970.31 | 21.28 | 69.4961
0.7936 | -17.73 | -142.32 | -1975.9 | 1981.22 | 17.48 | -21.2 | -4.91 | 28.09 | 1975.9 | 21.2 | 69.6629
0.7955 | -18.44 | -143.12 | -1982.24 | 1987.6 | 17.56 | -21.12 | -4.93 | 28.07 | 1982.24 | 21.12 | 69.8297
0.7974 | -20.39 | -143.7 | -1990.96 | 1996.35 | 17.73 | -20.92 | -5.01 | 28.02 | 1990.96 | 20.92 | 69.9965
0.7993 | -22.69 | -144.17 | -2000.1 | 2005.51 | 17.93 | -20.68 | -5.11 | 27.95 | 2000.1 | 20.68 | 70.1633
0.8012 | -24.36 | -144.66 | -2007.9 | 2013.34 | 18.03 | -20.46 | -5.18 | 27.86 | 2007.9 | 20.46 | 70.3301
0.8031 | -25.61 | -145.24 | -2014.74 | 2020.21 | 18.05 | -20.26 | -5.25 | 27.73 | 2014.74 | 20.26 | 70.4968
0.805 | -26.96 | -146.59 | -2021.13 | 2026.68 | 17.99 | -20.07 | -5.3 | 27.56 | 2021.13 | 20.07 | 70.6636
0.8069 | -28.78 | -148.42 | -2027.84 | 2033.51 | 17.98 | -19.82 | -5.37 | 27.39 | 2027.84 | 19.82 | 70.8304
0.8088 | -30.63 | -150.11 | -2033.59 | 2039.38 | 18.02 | -19.54 | -5.45 | 27.23 | 2033.59 | 19.54 | 70.9972
0.8107 | -32.54 | -151.64 | -2038.33 | 2044.24 | 18.11 | -19.22 | -5.54 | 27.07 | 2038.33 | 19.22 | 71.164
0.8126 | -34.54 | -152.65 | -2042.09 | 2048.1 | 18.24 | -18.84 | -5.63 | 26.91 | 2042.09 | 18.84 | 71.3308
0.8145 | -36.65 | -153.15 | -2044.9 | 2050.97 | 18.4 | -18.41 | -5.74 | 26.72 | 2044.9 | 18.41 | 71.4975
0.8164 | -38.6 | -153.56 | -2047.17 | 2053.3 | 18.55 | -17.99 | -5.84 | 26.55 | 2047.17 | 17.99 | 71.6643
0.8183 | -39.95 | -153.65 | -2047.26 | 2053.42 | 18.68 | -17.59 | -5.9 | 26.39 | 2047.26 | 17.59 | 71.8311
0.8202 | -41.09 | -153.87 | -2045.81 | 2052.01 | 18.82 | -17.2 | -5.95 | 26.24 | 2045.81 | 17.2 | 71.9979
0.8221 | -42.31 | -153.95 | -2043.13 | 2049.37 | 18.94 | -16.75 | -6.01 | 26.05 | 2043.13 | 16.75 | 72.1647
0.824 | -43.36 | -153.3 | -2038.72 | 2044.94 | 19.09 | -16.34 | -6.04 | 25.9 | 2038.72 | 16.34 | 72.3315
0.8259 | -44.45 | -152.29 | -2032.67 | 2038.86 | 19.22 | -15.92 | -6.06 | 25.75 | 2032.67 | 15.92 | 72.4982
0.8278 | -45.51 | -151.2 | -2025.16 | 2031.32 | 19.35 | -15.49 | -6.07 | 25.58 | 2025.16 | 15.49 | 72.665
0.8297 | -46.49 | -149.86 | -2015.36 | 2021.47 | 19.47 | -15.05 | -6.07 | 25.41 | 2015.36 | 15.05 | 72.8318
0.8316 | -47.41 | -148.31 | -2004.1 | 2010.15 | 19.59 | -14.63 | -6.06 | 25.25 | 2004.1 | 14.63 | 72.9986
0.8335 | -48.15 | -146.53 | -1991.09 | 1997.07 | 19.69 | -14.21 | -6.03 | 25.09 | 1991.09 | 14.21 | 73.1654
0.8354 | -48.92 | -144.84 | -1976.24 | 1982.16 | 19.78 | -13.81 | -5.99 | 24.92 | 1976.24 | 13.81 | 73.3322
0.8373 | -49.91 | -143.48 | -1959.69 | 1965.58 | 19.81 | -13.42 | -5.93 | 24.72 | 1959.69 | 13.42 | 73.4989
0.8392 | -50.94 | -142.05 | -1942.57 | 1948.44 | 19.83 | -13.04 | -5.87 | 24.52 | 1942.57 | 13.04 | 73.6657
0.8411 | -52 | -140.26 | -1924.27 | 1930.09 | 19.86 | -12.67 | -5.8 | 24.34 | 1924.27 | 12.67 | 73.8325
0.843 | -52.81 | -137.83 | -1903.69 | 1909.42 | 19.88 | -12.33 | -5.72 | 24.17 | 1903.69 | 12.33 | 73.9993
0.8449 | -53.39 | -135 | -1881.59 | 1887.2 | 19.9 | -12 | -5.65 | 24.01 | 1881.59 | 12 | 74.1661
0.8468 | -53.87 | -131.96 | -1857.65 | 1863.13 | 19.91 | -11.67 | -5.57 | 23.84 | 1857.65 | 11.67 | 74.3329
0.8487 | -54.21 | -128.72 | -1830.92 | 1836.26 | 19.89 | -11.37 | -5.48 | 23.68 | 1830.92 | 11.37 | 74.4996
0.8506 | -54.8 | -125.46 | -1801.81 | 1807.03 | 19.83 | -11.05 | -5.4 | 23.47 | 1801.81 | 11.05 | 74.6664
0.8525 | -55.43 | -122.12 | -1771.61 | 1776.7 | 19.76 | -10.75 | -5.31 | 23.24 | 1771.61 | 10.75 | 74.8332
0.8544 | -56.17 | -118.51 | -1739.84 | 1744.81 | 19.66 | -10.45 | -5.22 | 23.02 | 1739.84 | 10.45 | 75
0.8563 | -56.85 | -114.73 | -1706.54 | 1711.37 | 19.55 | -10.17 | -5.12 | 22.78 | 1706.54 | 10.17 | 75.1668
0.8582 | -57.56 | -110.72 | -1671.42 | 1676.11 | 19.41 | -9.87 | -5.03 | 22.52 | 1671.42 | 9.87 | 75.3336
0.8601 | -58.12 | -106.51 | -1635.11 | 1639.65 | 19.26 | -9.58 | -4.94 | 22.24 | 1635.11 | 9.58 | 75.5004
0.862 | -58.28 | -101.75 | -1596.77 | 1601.11 | 19.07 | -9.29 | -4.84 | 21.94 | 1596.77 | 9.29 | 75.6671
0.8639 | -57.94 | -96.3 | -1555.87 | 1559.97 | 18.85 | -9.01 | -4.73 | 21.62 | 1555.87 | 9.01 | 75.8339
0.8658 | -57.43 | -90.62 | -1514.29 | 1518.14 | 18.61 | -8.74 | -4.63 | 21.28 | 1514.29 | 8.74 | 76.0007
0.8677 | -56.82 | -85.12 | -1472.6 | 1476.22 | 18.32 | -8.46 | -4.52 | 20.9 | 1472.6 | 8.46 | 76.1675
0.8696 | -56.18 | -79.67 | -1431.14 | 1434.51 | 17.99 | -8.18 | -4.41 | 20.47 | 1431.14 | 8.18 | 76.3343
0.8715 | -55.3 | -73.76 | -1388.04 | 1391.16 | 17.64 | -7.87 | -4.31 | 20.01 | 1388.04 | 7.87 | 76.5011
0.8734 | -54.32 | -67.55 | -1343.19 | 1346.04 | 17.25 | -7.55 | -4.19 | 19.52 | 1343.19 | 7.55 | 76.6678
0.8753 | -53.3 | -61.2 | -1296.78 | 1299.39 | 16.82 | -7.22 | -4.07 | 18.99 | 1296.78 | 7.22 | 76.8346
0.8772 | -52.07 | -54.59 | -1249.46 | 1251.82 | 16.35 | -6.86 | -3.94 | 18.42 | 1249.46 | 6.86 | 77.0014
0.8791 | -50.86 | -48.07 | -1202.98 | 1205.07 | 15.88 | -6.49 | -3.82 | 17.83 | 1202.98 | 6.49 | 77.1682
0.881 | -49.56 | -41.83 | -1155.45 | 1157.35 | 15.35 | -6.12 | -3.69 | 17.19 | 1155.45 | 6.12 | 77.335
0.8829 | -48.21 | -35.57 | -1107.48 | 1109.19 | 14.79 | -5.72 | -3.55 | 16.52 | 1107.48 | 5.72 | 77.5018
0.8848 | -46.85 | -29.41 | -1060.27 | 1061.79 | 14.22 | -5.34 | -3.41 | 15.84 | 1060.27 | 5.34 | 77.6685
0.8867 | -45.52 | -23.48 | -1014.14 | 1015.52 | 13.64 | -4.94 | -3.28 | 15.15 | 1014.14 | 4.94 | 77.8353
0.8886 | -44.08 | -17.83 | -968.05 | 969.32 | 13.03 | -4.55 | -3.14 | 14.43 | 968.05 | 4.55 | 78.0021
0.8905 | -42.56 | -12.33 | -921.99 | 923.16 | 12.41 | -4.16 | -2.99 | 13.71 | 921.99 | 4.16 | 78.1689
0.8924 | -41.02 | -6.79 | -875.08 | 876.16 | 11.78 | -3.72 | -2.85 | 12.96 | 875.08 | 3.72 | 78.3357
0.8943 | -39.31 | -1.78 | -829.03 | 830.07 | 11.13 | -3.3 | -2.72 | 12.2 | 829.03 | 3.3 | 78.5025
0.8962 | -37.66 | 2.76 | -784.62 | 785.64 | 10.47 | -2.91 | -2.58 | 11.45 | 784.62 | 2.91 | 78.6692
0.8981 | -36.29 | 6.72 | -742.5 | 743.51 | 9.82 | -2.51 | -2.46 | 10.7 | 742.5 | 2.51 | 78.836
0.9 | -35.11 | 10.13 | -702.48 | 703.5 | 9.17 | -2.1 | -2.33 | 9.96 | 702.48 | 2.1 | 79.0028
0.9019 | -34.01 | 12.71 | -665.96 | 667.02 | 8.54 | -1.75 | -2.22 | 9.26 | 665.96 | 1.75 | 79.1696
0.9038 | -32.99 | 14.9 | -631.63 | 632.74 | 7.92 | -1.42 | -2.11 | 8.59 | 631.63 | 1.42 | 79.3364
0.9057 | -32.1 | 16.66 | -600.12 | 601.28 | 7.34 | -1.11 | -2.01 | 7.97 | 600.12 | 1.11 | 79.5032
0.9076 | -31.31 | 17.8 | -570.56 | 571.77 | 6.78 | -0.82 | -1.92 | 7.38 | 570.56 | 0.82 | 79.6699
0.9095 | -30.58 | 18.85 | -542.1 | 543.36 | 6.24 | -0.55 | -1.83 | 6.81 | 542.1 | 0.55 | 79.8367
0.9114 | -30.07 | 19.82 | -516.28 | 517.62 | 5.73 | -0.3 | -1.75 | 6.29 | 516.28 | 0.3 | 80.0035
0.9133 | -29.82 | 20.7 | -493.68 | 495.1 | 5.28 | -0.08 | -1.67 | 5.82 | 493.68 | 0.08 | 80.1703
0.9151 | -29.71 | 21.21 | -473.54 | 475.04 | 4.88 | 0.08 | -1.61 | 5.42 | 473.54 | -0.08 | 80.3283
0.917 | -29.71 | 21.43 | -455.11 | 456.67 | 4.5 | 0.24 | -1.55 | 5.05 | 455.11 | -0.24 | 80.4951
0.9189 | -29.8 | 21.53 | -439.09 | 440.71 | 4.13 | 0.43 | -1.5 | 4.68 | 439.09 | -0.43 | 80.6619
0.9208 | -29.8 | 21.65 | -424.43 | 426.1 | 3.79 | 0.61 | -1.45 | 4.35 | 424.43 | -0.61 | 80.8287
0.9227 | -29.87 | 21.62 | -411.52 | 413.23 | 3.51 | 0.74 | -1.4 | 4.1 | 411.52 | -0.74 | 80.9954
0.9246 | -30.07 | 21.62 | -400.59 | 402.36 | 3.28 | 0.86 | -1.37 | 3.9 | 400.59 | -0.86 | 81.1622
0.9265 | -30.4 | 22.05 | -392.2 | 394.05 | 3.11 | 0.93 | -1.35 | 3.77 | 392.2 | -0.93 | 81.329
0.9284 | -30.44 | 22.85 | -383.81 | 385.75 | 3.03 | 0.96 | -1.34 | 3.73 | 383.81 | -0.96 | 81.4958
0.9303 | -30.07 | 23.14 | -374.33 | 376.33 | 2.97 | 0.99 | -1.3 | 3.7 | 374.33 | -0.99 | 81.6626
0.9322 | -29.87 | 22.93 | -364.77 | 366.8 | 2.87 | 1.06 | -1.24 | 3.63 | 364.77 | -1.06 | 81.8294
0.9341 | -29.93 | 22.5 | -356.26 | 358.31 | 2.72 | 1.17 | -1.19 | 3.52 | 356.26 | -1.17 | 81.9961
0.936 | -29.89 | 22.04 | -349.79 | 351.84 | 2.58 | 1.27 | -1.14 | 3.44 | 349.79 | -1.27 | 82.1629
0.9379 | -29.73 | 21.73 | -344.47 | 346.5 | 2.48 | 1.37 | -1.09 | 3.36 | 344.47 | -1.37 | 82.3297
0.9398 | -29.69 | 21.88 | -340.19 | 342.24 | 2.39 | 1.49 | -1.06 | 3.28 | 340.19 | -1.49 | 82.4965
0.9417 | -29.8 | 21.73 | -336.6 | 338.66 | 2.3 | 1.62 | -1.02 | 3.24 | 336.6 | -1.62 | 82.6633
0.9436 | -29.98 | 21.39 | -333.57 | 335.64 | 2.21 | 1.75 | -0.98 | 3.21 | 333.57 | -1.75 | 82.8301
0.9455 | -30.21 | 20.71 | -331.32 | 333.38 | 2.11 | 1.87 | -0.94 | 3.18 | 331.32 | -1.87 | 82.9968
0.9474 | -30.4 | 20.57 | -330.61 | 332.68 | 2.04 | 1.97 | -0.92 | 3.18 | 330.61 | -1.97 | 83.1636
0.9493 | -30.43 | 20.92 | -330.22 | 332.32 | 2.01 | 2.08 | -0.89 | 3.23 | 330.22 | -2.08 | 83.3304
0.9512 | -29.96 | 20.82 | -327.43 | 329.48 | 1.95 | 2.2 | -0.86 | 3.27 | 327.43 | -2.2 | 83.4972
0.9531 | -29.97 | 20.64 | -326.85 | 328.9 | 1.89 | 2.32 | -0.83 | 3.28 | 326.85 | -2.32 | 83.664
0.955 | -29.83 | 20.49 | -327.35 | 329.37 | 1.82 | 2.48 | -0.79 | 3.34 | 327.35 | -2.48 | 83.8308
0.9569 | -29.86 | 20.61 | -329.87 | 331.9 | 1.78 | 2.66 | -0.73 | 3.44 | 329.87 | -2.66 | 83.9975
0.9588 | -29.86 | 20.86 | -331.97 | 333.99 | 1.74 | 2.83 | -0.68 | 3.54 | 331.97 | -2.83 | 84.1643
0.9607 | -29.58 | 21.25 | -331.6 | 333.63 | 1.7 | 2.94 | -0.64 | 3.59 | 331.6 | -2.94 | 84.3311
0.9626 | -29.41 | 21.41 | -330.91 | 332.94 | 1.65 | 3 | -0.59 | 3.62 | 330.91 | -3 | 84.4979
0.9645 | -29.35 | 21.48 | -330.01 | 332.05 | 1.58 | 3.04 | -0.55 | 3.62 | 330.01 | -3.04 | 84.6647
0.9664 | -29.16 | 21.96 | -327.94 | 330.01 | 1.54 | 3.07 | -0.51 | 3.62 | 327.94 | -3.07 | 84.8315
0.9683 | -28.91 | 22.62 | -325.09 | 327.2 | 1.49 | 3.13 | -0.47 | 3.64 | 325.09 | -3.13 | 84.9982
0.9702 | -28.47 | 23.21 | -322.1 | 324.23 | 1.45 | 3.18 | -0.43 | 3.65 | 322.1 | -3.18 | 85.165
0.9721 | -27.96 | 24.04 | -318.61 | 320.79 | 1.42 | 3.19 | -0.39 | 3.64 | 318.61 | -3.19 | 85.3318
0.974 | -27.36 | 24.98 | -314.67 | 316.9 | 1.39 | 3.19 | -0.37 | 3.61 | 314.67 | -3.19 | 85.4986
0.9759 | -26.68 | 25.49 | -310.69 | 312.93 | 1.36 | 3.2 | -0.32 | 3.6 | 310.69 | -3.2 | 85.6654
0.9778 | -25.98 | 26.03 | -306.37 | 308.62 | 1.32 | 3.19 | -0.28 | 3.58 | 306.37 | -3.19 | 85.8322
0.9797 | -25.26 | 26.63 | -301.16 | 303.44 | 1.29 | 3.17 | -0.26 | 3.53 | 301.16 | -3.17 | 85.9989
0.9816 | -24.46 | 26.93 | -295.16 | 297.45 | 1.24 | 3.15 | -0.21 | 3.49 | 295.16 | -3.15 | 86.1657
0.9835 | -23.62 | 27.16 | -288.56 | 290.87 | 1.19 | 3.14 | -0.17 | 3.45 | 288.56 | -3.14 | 86.3325
0.9854 | -22.76 | 27.34 | -280.79 | 283.13 | 1.15 | 3.15 | -0.11 | 3.43 | 280.79 | -3.15 | 86.4993
0.9873 | -21.96 | 27.32 | -272.96 | 275.32 | 1.09 | 3.19 | -0.04 | 3.43 | 272.96 | -3.19 | 86.6661
0.9892 | -21.27 | 26.85 | -266.78 | 269.07 | 1.03 | 3.17 | 0.02 | 3.37 | 266.78 | -3.17 | 86.8329
0.9911 | -20.62 | 26.49 | -260.39 | 262.65 | 0.97 | 3.13 | 0.06 | 3.31 | 260.39 | -3.13 | 86.9996
0.993 | -20.03 | 25.93 | -253.38 | 255.6 | 0.92 | 3.05 | 0.07 | 3.21 | 253.38 | -3.05 | 87.1664
0.9949 | -19.42 | 25.15 | -247.34 | 249.48 | 0.86 | 2.97 | 0.09 | 3.12 | 247.34 | -2.97 | 87.3332
0.9968 | -18.68 | 24.4 | -242.07 | 244.11 | 0.79 | 2.89 | 0.11 | 3.02 | 242.07 | -2.89 | 87.5
0.9987 | -17.88 | 23.59 | -237.1 | 239.03 | 0.71 | 2.8 | 0.13 | 2.91 | 237.1 | -2.8 | 87.6668
1.0006 | -17.22 | 22.5 | -232.93 | 234.74 | 0.63 | 2.72 | 0.14 | 2.81 | 232.93 | -2.72 | 87.8336
1.0025 | -16.73 | 21.48 | -229.55 | 231.24 | 0.56 | 2.62 | 0.14 | 2.71 | 229.55 | -2.62 | 88.0004
1.0044 | -16.36 | 20.25 | -227 | 228.56 | 0.47 | 2.55 | 0.15 | 2.62 | 227 | -2.55 | 88.1671
1.0063 | -16.01 | 19.43 | -224.71 | 226.17 | 0.41 | 2.45 | 0.16 | 2.51 | 224.71 | -2.45 | 88.3339
1.0082 | -15.8 | 18.35 | -223.46 | 224.82 | 0.33 | 2.38 | 0.17 | 2.43 | 223.46 | -2.38 | 88.5007
1.0101 | -15.7 | 17.55 | -222.96 | 224.25 | 0.27 | 2.31 | 0.19 | 2.35 | 222.96 | -2.31 | 88.6675
1.012 | -15.67 | 17.1 | -222.74 | 223.99 | 0.21 | 2.23 | 0.2 | 2.27 | 222.74 | -2.23 | 88.8343
1.0139 | -15.49 | 16.66 | -223.07 | 224.27 | 0.17 | 2.14 | 0.21 | 2.17 | 223.07 | -2.14 | 89.0011
1.0158 | -15.48 | 16.33 | -223.79 | 224.97 | 0.12 | 2.05 | 0.21 | 2.08 | 223.79 | -2.05 | 89.1678
1.0177 | -15.52 | 16 | -224.59 | 225.74 | 0.08 | 1.95 | 0.21 | 1.98 | 224.59 | -1.95 | 89.3346
1.0196 | -15.69 | 16.03 | -224.56 | 225.72 | 0.05 | 1.84 | 0.2 | 1.86 | 224.56 | -1.84 | 89.5014
1.0215 | -15.8 | 16.28 | -223.92 | 225.11 | 0.04 | 1.73 | 0.18 | 1.76 | 223.92 | -1.73 | 89.6682
1.0234 | -15.82 | 16.69 | -222.38 | 223.6 | 0.04 | 1.62 | 0.16 | 1.64 | 222.38 | -1.62 | 89.835
1.0253 | -15.53 | 17.15 | -219.39 | 220.65 | 0.06 | 1.51 | 0.14 | 1.53 | 219.39 | -1.51 | 90.0018
1.0272 | -15.12 | 17.53 | -216.04 | 217.32 | 0.07 | 1.39 | 0.12 | 1.41 | 216.04 | -1.39 | 90.1685
1.0291 | -14.72 | 17.88 | -212.34 | 213.64 | 0.08 | 1.28 | 0.1 | 1.3 | 212.34 | -1.28 | 90.3353
1.031 | -14.37 | 18.17 | -208.07 | 209.4 | 0.09 | 1.17 | 0.07 | 1.2 | 208.07 | -1.17 | 90.5021
1.0329 | -14.08 | 18.21 | -202.55 | 203.9 | 0.09 | 1.07 | 0.04 | 1.1 | 202.55 | -1.07 | 90.6689
1.0348 | -13.87 | 18.06 | -196.11 | 197.46 | 0.09 | 0.96 | 0.01 | 0.99 | 196.11 | -0.96 | 90.8357
1.0367 | -13.54 | 17.73 | -189.7 | 191.05 | 0.09 | 0.86 | -0.02 | 0.88 | 189.7 | -0.86 | 91.0025
1.0386 | -13.27 | 17.06 | -184.5 | 185.81 | 0.08 | 0.77 | -0.04 | 0.8 | 184.5 | -0.77 | 91.1692
1.0405 | -13.03 | 15.78 | -180.66 | 181.88 | 0.04 | 0.7 | -0.07 | 0.73 | 180.66 | -0.7 | 91.336
1.0424 | -12.78 | 14.49 | -176.91 | 178.02 | 0.01 | 0.63 | -0.09 | 0.66 | 176.91 | -0.63 | 91.5028
1.0443 | -12.48 | 13.27 | -174.24 | 175.25 | -0.04 | 0.56 | -0.11 | 0.61 | 174.24 | -0.56 | 91.6696
1.0462 | -12.17 | 12.25 | -172.84 | 173.77 | -0.08 | 0.49 | -0.12 | 0.57 | 172.84 | -0.49 | 91.8364
1.0481 | -11.84 | 11.47 | -171.99 | 172.84 | -0.11 | 0.41 | -0.14 | 0.53 | 171.99 | -0.41 | 92.0032
1.05 | -11.46 | 11.15 | -171.96 | 172.77 | -0.12 | 0.32 | -0.16 | 0.48 | 171.96 | -0.32 | 92.1699
1.0519 | -10.99 | 11.18 | -173.2 | 173.97 | -0.14 | 0.22 | -0.18 | 0.47 | 173.2 | -0.22 | 92.3367
1.0538 | -10.45 | 11.32 | -174.64 | 175.4 | -0.15 | 0.11 | -0.21 | 0.5 | 174.64 | -0.11 | 92.5035
1.0556 | -10.04 | 11.77 | -174.67 | 175.46 | -0.14 | 0.03 | -0.23 | 0.53 | 174.67 | -0.03 | 92.6615
1.0575 | -9.62 | 12.21 | -174.33 | 175.14 | -0.12 | -0.05 | -0.26 | 0.56 | 174.33 | 0.05 | 92.8283
1.0594 | -9.25 | 12.43 | -173.75 | 174.55 | -0.11 | -0.11 | -0.28 | 0.59 | 173.75 | 0.11 | 92.9951
1.0613 | -8.98 | 12.64 | -172.5 | 173.31 | -0.1 | -0.17 | -0.31 | 0.58 | 172.5 | 0.17 | 93.1619
1.0632 | -8.81 | 12.77 | -170.36 | 171.18 | -0.09 | -0.23 | -0.33 | 0.59 | 170.36 | 0.23 | 93.3287
1.0651 | -8.63 | 12.42 | -167.42 | 168.22 | -0.1 | -0.26 | -0.35 | 0.6 | 167.42 | 0.26 | 93.4954
1.067 | -8.53 | 12.05 | -164.09 | 164.87 | -0.1 | -0.3 | -0.37 | 0.6 | 164.09 | 0.3 | 93.6622
1.0689 | -8.45 | 11.35 | -161.63 | 162.36 | -0.13 | -0.32 | -0.39 | 0.63 | 161.63 | 0.32 | 93.829
1.0708 | -8.25 | 10.66 | -160.59 | 161.29 | -0.15 | -0.35 | -0.4 | 0.65 | 160.59 | 0.35 | 93.9958
1.0727 | -8.02 | 10.19 | -160.32 | 160.98 | -0.17 | -0.38 | -0.41 | 0.66 | 160.32 | 0.38 | 94.1626
1.0746 | -7.84 | 9.73 | -160.62 | 161.24 | -0.19 | -0.41 | -0.42 | 0.67 | 160.62 | 0.41 | 94.3294
1.0765 | -7.67 | 9.36 | -161.92 | 162.5 | -0.21 | -0.43 | -0.42 | 0.69 | 161.92 | 0.43 | 94.4961
1.0784 | -7.39 | 9.06 | -163.75 | 164.28 | -0.23 | -0.46 | -0.43 | 0.72 | 163.75 | 0.46 | 94.6629
1.0803 | -7.02 | 9.35 | -166.12 | 166.66 | -0.22 | -0.49 | -0.43 | 0.73 | 166.12 | 0.49 | 94.8297
1.0822 | -6.65 | 9.86 | -168.3 | 168.87 | -0.19 | -0.52 | -0.44 | 0.75 | 168.3 | 0.52 | 94.9965
1.0841 | -6.23 | 10.09 | -169.5 | 170.07 | -0.17 | -0.54 | -0.45 | 0.76 | 169.5 | 0.54 | 95.1633
1.086 | -6.05 | 10.34 | -169.6 | 170.18 | -0.16 | -0.55 | -0.46 | 0.77 | 169.6 | 0.55 | 95.3301
1.0879 | -5.91 | 10.34 | -169.47 | 170.05 | -0.14 | -0.56 | -0.47 | 0.78 | 169.47 | 0.56 | 95.4968
1.0898 | -6.02 | 10.15 | -168.26 | 168.83 | -0.14 | -0.55 | -0.48 | 0.78 | 168.26 | 0.55 | 95.6636
1.0917 | -6.22 | 9.56 | -166.71 | 167.25 | -0.15 | -0.52 | -0.49 | 0.77 | 166.71 | 0.52 | 95.8304
1.0936 | -6.37 | 8.86 | -166.33 | 166.83 | -0.16 | -0.5 | -0.49 | 0.76 | 166.33 | 0.5 | 95.9972
1.0955 | -6.5 | 8.05 | -167.6 | 168.05 | -0.19 | -0.47 | -0.49 | 0.74 | 167.6 | 0.47 | 96.164
1.0974 | -6.67 | 7.3 | -169.48 | 169.89 | -0.21 | -0.44 | -0.49 | 0.73 | 169.48 | 0.44 | 96.3308
1.0993 | -6.78 | 7.09 | -172.31 | 172.7 | -0.21 | -0.43 | -0.49 | 0.72 | 172.31 | 0.43 | 96.4975
1.1012 | -6.85 | 7.21 | -176.07 | 176.45 | -0.2 | -0.43 | -0.49 | 0.71 | 176.07 | 0.43 | 96.6643
1.1031 | -6.85 | 7.2 | -179.28 | 179.66 | -0.19 | -0.42 | -0.49 | 0.7 | 179.28 | 0.42 | 96.8311
1.105 | -6.81 | 7.21 | -181.67 | 182.04 | -0.18 | -0.41 | -0.49 | 0.7 | 181.67 | 0.41 | 96.9979
1.1069 | -6.63 | 7.23 | -183.45 | 183.81 | -0.17 | -0.41 | -0.5 | 0.69 | 183.45 | 0.41 | 97.1647
1.1088 | -6.37 | 7.11 | -184.3 | 184.65 | -0.18 | -0.39 | -0.5 | 0.69 | 184.3 | 0.39 | 97.3315
1.1107 | -6.2 | 7.19 | -185.28 | 185.61 | -0.17 | -0.39 | -0.5 | 0.68 | 185.28 | 0.39 | 97.4982
1.1126 | -6.02 | 7.31 | -186.17 | 186.5 | -0.17 | -0.38 | -0.5 | 0.68 | 186.17 | 0.38 | 97.665
1.1145 | -5.81 | 7.51 | -186.86 | 187.2 | -0.17 | -0.36 | -0.49 | 0.66 | 186.86 | 0.36 | 97.8318
1.1164 | -5.63 | 7.84 | -186.69 | 187.04 | -0.16 | -0.36 | -0.49 | 0.65 | 186.69 | 0.36 | 97.9986
1.1183 | -5.45 | 7.83 | -185.92 | 186.26 | -0.17 | -0.35 | -0.48 | 0.65 | 185.92 | 0.35 | 98.1654
1.1202 | -5.19 | 8.37 | -185.02 | 185.39 | -0.17 | -0.35 | -0.47 | 0.64 | 185.02 | 0.35 | 98.3322
1.1221 | -4.96 | 8.86 | -183.57 | 183.96 | -0.16 | -0.35 | -0.47 | 0.63 | 183.57 | 0.35 | 98.4989
1.124 | -4.8 | 8.95 | -182.12 | 182.52 | -0.18 | -0.34 | -0.46 | 0.63 | 182.12 | 0.34 | 98.6657
1.1259 | -4.66 | 8.86 | -182.08 | 182.48 | -0.19 | -0.34 | -0.45 | 0.63 | 182.08 | 0.34 | 98.8325
1.1278 | -4.45 | 8.84 | -182.53 | 182.94 | -0.21 | -0.34 | -0.44 | 0.63 | 182.53 | 0.34 | 98.9993
1.1297 | -4.07 | 9.35 | -184.1 | 184.56 | -0.21 | -0.36 | -0.43 | 0.63 | 184.1 | 0.36 | 99.1661
1.1316 | -3.74 | 9.78 | -185.93 | 186.42 | -0.22 | -0.37 | -0.42 | 0.63 | 185.93 | 0.37 | 99.3329
1.1335 | -3.45 | 10.34 | -186.33 | 186.86 | -0.22 | -0.38 | -0.41 | 0.63 | 186.33 | 0.38 | 99.4996
1.1354 | -3.23 | 11.09 | -185.23 | 185.8 | -0.21 | -0.38 | -0.41 | 0.62 | 185.23 | 0.38 | 99.6664
1.1373 | -3.08 | 11.68 | -184.1 | 184.7 | -0.2 | -0.37 | -0.39 | 0.61 | 184.1 | 0.37 | 99.8332
1.1392 | -2.78 | 12.38 | -183.63 | 184.31 | -0.2 | -0.38 | -0.39 | 0.61 | 183.63 | 0.38 | 100
| Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My
Minimum | -58.28 | -153.95 | -2448.53 | 160.98 | -1.83 | -23.25 | -6.07 | 0.47 | 160.32 | -3.2
Maximum | 125.39 | 87.75 | -160.32 | 2453.51 | 19.91 | 3.2 | 0.21 | 28.1 | 2448.53 | 23.25

## Diagr_Individ
Individual Loads
Contact forces and moments, measured telemetrically in knee implants of 8 subjects.
Averages of one individual subject.
Fx  lateral, Fy  anterior, -Fz  distal, Fres = Resultant Force
Mx around x, My around y, Mz around z, Mres = Resultant Moment
z-axis = axis of implant stem!
From Bergmann et al, PlosOne (2014) "Standardized Loads Acting in Knee Implants"
Subject K8L
Subject K8L

## Dat_BW
Data File: | k8l_2266_0_6_i1.akf
Curve Alignment Program: | 1.5.7
------------------------MEASURING INFOS------------------------
Diagram Title #1 | Forces and Moments at Knee Joint
Diagram Title #2 | Measured Relative to Implant
Comment #1 | Knee Joint; Gaitanalysis; Stairs up;
Comment #2 | K8L, 25 Months PO *
BodyWeight [N]: | 775
Patient(en): | K8L | [7 empty] | Subject | K8L
Implant Type: | Knee Joint
Varus Angle Leg [°] | 4 | [7 empty] | Normalized Loads in %BW (Forces) and %BWm (moments)
Axes Directions X, Y, Z: | Lateral, Anterior, Superior
Axes Origin: | Tibial Plateau at Stem Axis | [7 empty] | BW [kg] = | 79.001
Transformation X Y Z (°): | 11.5 | -1 | 12
Displacement X, Y, Z (mm): | 0 | 0 | 0
Activity Code: | KneeActivities | 1P_1_15_4_1
---------------------------DATA--------------------------------
Number of Data Sets = | 601
Number of analog channels = | 0
Time-Offset(sec) = | 0
Time | Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My | Time
[s] | [%BW] | [%BW] | [%BW] | [%BW] | [%BWm] | [%BWm] | [%BWm] | [%BWm] | [%BW] | [%BWm] | [%Cycle]
0 | -0.34452 | 3.009 | -22.2658 | 22.4826 | -0.050323 | -0.14452 | -0.094194 | 0.18065 | 22.2658 | 0.14452 | 0
0.0018 | -0.38065 | 2.9639 | -22.2516 | 22.4632 | -0.050323 | -0.13935 | -0.094194 | 0.17677 | 22.2516 | 0.13935 | 0.16735
0.0036 | -0.39097 | 2.9355 | -22.2142 | 22.4232 | -0.051613 | -0.13677 | -0.092903 | 0.17419 | 22.2142 | 0.13677 | 0.3347
0.0054 | -0.38839 | 2.9174 | -22.1742 | 22.3806 | -0.051613 | -0.13548 | -0.091613 | 0.1729 | 22.1742 | 0.13548 | 0.50205
0.0072 | -0.37677 | 2.9419 | -22.4619 | 22.6684 | -0.050323 | -0.13548 | -0.090323 | 0.17161 | 22.4619 | 0.13548 | 0.66939
0.009 | -0.38323 | 2.92 | -22.7226 | 22.9265 | -0.051613 | -0.1329 | -0.089032 | 0.16903 | 22.7226 | 0.1329 | 0.83674
0.0108 | -0.3871 | 2.9213 | -22.9239 | 23.129 | -0.050323 | -0.13161 | -0.087742 | 0.16774 | 22.9239 | 0.13161 | 1.0041
0.0125 | -0.37806 | 2.9639 | -23.04 | 23.249 | -0.047742 | -0.13032 | -0.086452 | 0.16516 | 23.04 | 0.13032 | 1.1621
0.0143 | -0.3729 | 2.9974 | -23.0465 | 23.2594 | -0.046452 | -0.12903 | -0.086452 | 0.16258 | 23.0465 | 0.12903 | 1.3295
0.0161 | -0.38968 | 3.0065 | -22.9148 | 23.129 | -0.045161 | -0.12516 | -0.085161 | 0.16 | 22.9148 | 0.12516 | 1.4968
0.0179 | -0.43613 | 2.9806 | -22.7161 | 22.929 | -0.045161 | -0.12 | -0.085161 | 0.15484 | 22.7161 | 0.12 | 1.6642
0.0197 | -0.48774 | 2.9381 | -22.5071 | 22.7174 | -0.046452 | -0.11484 | -0.083871 | 0.15097 | 22.5071 | 0.11484 | 1.8315
0.0215 | -0.49548 | 2.9368 | -22.4826 | 22.6942 | -0.047742 | -0.11226 | -0.08129 | 0.14839 | 22.4826 | 0.11226 | 1.9989
0.0233 | -0.48387 | 3.0039 | -22.6065 | 22.8271 | -0.047742 | -0.11097 | -0.08 | 0.1471 | 22.6065 | 0.11097 | 2.1662
0.0251 | -0.47355 | 3.0761 | -22.7394 | 22.9703 | -0.049032 | -0.11097 | -0.07871 | 0.14581 | 22.7394 | 0.11097 | 2.3336
0.0269 | -0.43871 | 3.1523 | -22.8297 | 23.071 | -0.049032 | -0.11097 | -0.077419 | 0.14581 | 22.8297 | 0.11097 | 2.5009
0.0287 | -0.41419 | 3.2284 | -22.769 | 23.0206 | -0.047742 | -0.11226 | -0.076129 | 0.14452 | 22.769 | 0.11226 | 2.6683
0.0305 | -0.37548 | 3.3239 | -22.6852 | 22.951 | -0.045161 | -0.11355 | -0.074839 | 0.14452 | 22.6852 | 0.11355 | 2.8356
0.0323 | -0.3329 | 3.4297 | -22.551 | 22.8271 | -0.046452 | -0.11613 | -0.074839 | 0.1471 | 22.551 | 0.11613 | 3.003
0.0341 | -0.26968 | 3.5703 | -22.4568 | 22.7523 | -0.047742 | -0.12 | -0.074839 | 0.14968 | 22.4568 | 0.12 | 3.1703
0.0359 | -0.26323 | 3.6 | -22.2374 | 22.5419 | -0.049032 | -0.11871 | -0.073548 | 0.14839 | 22.2374 | 0.11871 | 3.3377
0.0376 | -0.27613 | 3.6065 | -21.9652 | 22.2723 | -0.052903 | -0.11742 | -0.072258 | 0.14839 | 21.9652 | 0.11742 | 3.4957
0.0394 | -0.27871 | 3.6335 | -21.7884 | 22.1019 | -0.055484 | -0.11613 | -0.070968 | 0.14839 | 21.7884 | 0.11613 | 3.6631
0.0412 | -0.24387 | 3.7019 | -21.9329 | 22.2568 | -0.058065 | -0.11871 | -0.069677 | 0.15097 | 21.9329 | 0.11871 | 3.8304
0.043 | -0.20903 | 3.7755 | -22.1742 | 22.5071 | -0.061935 | -0.12129 | -0.068387 | 0.15355 | 22.1742 | 0.12129 | 3.9978
0.0448 | -0.17161 | 3.8581 | -22.3935 | 22.7368 | -0.064516 | -0.12516 | -0.068387 | 0.15613 | 22.3935 | 0.12516 | 4.1651
0.0466 | -0.13032 | 3.969 | -22.5548 | 22.9135 | -0.065806 | -0.12774 | -0.067097 | 0.15871 | 22.5548 | 0.12774 | 4.3325
0.0484 | -0.089032 | 4.0813 | -22.6323 | 23.009 | -0.067097 | -0.13032 | -0.067097 | 0.16258 | 22.6323 | 0.13032 | 4.4998
0.0502 | -0.058065 | 4.2284 | -22.5806 | 22.9858 | -0.068387 | -0.13419 | -0.068387 | 0.16516 | 22.5806 | 0.13419 | 4.6672
0.052 | -0.028387 | 4.3239 | -22.489 | 22.9161 | -0.070968 | -0.13548 | -0.069677 | 0.16903 | 22.489 | 0.13548 | 4.8345
0.0538 | 0.0012903 | 4.391 | -22.4206 | 22.8632 | -0.073548 | -0.13806 | -0.069677 | 0.17161 | 22.4206 | 0.13806 | 5.0019
0.0556 | 0.030968 | 4.4658 | -22.4994 | 22.9535 | -0.077419 | -0.14065 | -0.069677 | 0.17548 | 22.4994 | 0.14065 | 5.1692
0.0574 | 0.060645 | 4.4981 | -22.6413 | 23.0994 | -0.082581 | -0.14323 | -0.069677 | 0.17935 | 22.6413 | 0.14323 | 5.3366
0.0592 | 0.091613 | 4.551 | -22.8839 | 23.3471 | -0.089032 | -0.14581 | -0.069677 | 0.18452 | 22.8839 | 0.14581 | 5.5039
0.061 | 0.15226 | 4.6748 | -23.2323 | 23.7135 | -0.090323 | -0.15097 | -0.069677 | 0.19097 | 23.2323 | 0.15097 | 5.6713
0.0627 | 0.22452 | 4.809 | -23.6077 | 24.1071 | -0.092903 | -0.15613 | -0.070968 | 0.19484 | 23.6077 | 0.15613 | 5.8293
0.0645 | 0.27484 | 4.9071 | -23.8387 | 24.3523 | -0.096774 | -0.16 | -0.073548 | 0.2 | 23.8387 | 0.16 | 5.9967
0.0663 | 0.32 | 5.0026 | -24.0426 | 24.569 | -0.10065 | -0.16258 | -0.074839 | 0.20516 | 24.0426 | 0.16258 | 6.164
0.0681 | 0.36774 | 5.0942 | -24.3677 | 24.9058 | -0.10581 | -0.16645 | -0.076129 | 0.21161 | 24.3677 | 0.16645 | 6.3313
0.0699 | 0.4529 | 5.191 | -24.9226 | 25.4697 | -0.11226 | -0.17161 | -0.077419 | 0.21935 | 24.9226 | 0.17161 | 6.4987
0.0717 | 0.54323 | 5.3226 | -25.5226 | 26.0839 | -0.11613 | -0.17806 | -0.07871 | 0.2271 | 25.5226 | 0.17806 | 6.666
0.0735 | 0.63355 | 5.4542 | -26.1058 | 26.6852 | -0.12 | -0.18452 | -0.08 | 0.23484 | 26.1058 | 0.18452 | 6.8334
0.0753 | 0.72 | 5.6155 | -26.6323 | 27.2348 | -0.12258 | -0.19097 | -0.082581 | 0.24129 | 26.6323 | 0.19097 | 7.0007
0.0771 | 0.78323 | 5.7716 | -27.0168 | 27.6452 | -0.12645 | -0.19613 | -0.085161 | 0.24903 | 27.0168 | 0.19613 | 7.1681
0.0789 | 0.81677 | 5.8942 | -27.2555 | 27.9058 | -0.12903 | -0.19871 | -0.087742 | 0.2529 | 27.2555 | 0.19871 | 7.3354
0.0807 | 0.83742 | 5.9781 | -27.4323 | 28.0955 | -0.1329 | -0.20129 | -0.090323 | 0.25677 | 27.4323 | 0.20129 | 7.5028
0.0825 | 0.85806 | 6.0645 | -27.6335 | 28.3123 | -0.13677 | -0.20258 | -0.091613 | 0.26065 | 27.6335 | 0.20258 | 7.6701
0.0843 | 0.90323 | 6.2013 | -28.009 | 28.7097 | -0.14065 | -0.20516 | -0.092903 | 0.26581 | 28.009 | 0.20516 | 7.8375
0.086 | 0.96 | 6.3768 | -28.5148 | 29.2452 | -0.14452 | -0.20903 | -0.094194 | 0.27097 | 28.5148 | 0.20903 | 7.9955
0.0878 | 1.0142 | 6.5794 | -29.12 | 29.8813 | -0.14581 | -0.2129 | -0.096774 | 0.27613 | 29.12 | 0.2129 | 8.1629
0.0896 | 1.1084 | 6.8323 | -29.8 | 30.6013 | -0.14581 | -0.21935 | -0.099355 | 0.28129 | 29.8 | 0.21935 | 8.3302
0.0914 | 1.1665 | 7.0452 | -30.2916 | 31.129 | -0.1471 | -0.22065 | -0.10065 | 0.28387 | 30.2916 | 0.22065 | 8.4976
0.0932 | 1.1819 | 7.209 | -30.6426 | 31.5097 | -0.14968 | -0.21806 | -0.10194 | 0.28387 | 30.6426 | 0.21806 | 8.6649
0.095 | 1.1935 | 7.3613 | -30.9729 | 31.8658 | -0.15097 | -0.21548 | -0.10323 | 0.28258 | 30.9729 | 0.21548 | 8.8323
0.0968 | 1.2245 | 7.5277 | -31.4013 | 32.3226 | -0.15484 | -0.2129 | -0.10452 | 0.28387 | 31.4013 | 0.2129 | 8.9996
0.0986 | 1.2813 | 7.7058 | -31.9858 | 32.9342 | -0.15871 | -0.21419 | -0.10581 | 0.28645 | 31.9858 | 0.21419 | 9.167
0.1004 | 1.3548 | 7.9535 | -32.7548 | 33.7419 | -0.16 | -0.21419 | -0.1071 | 0.28774 | 32.7548 | 0.21419 | 9.3343
0.1022 | 1.431 | 8.2284 | -33.5884 | 34.6181 | -0.16129 | -0.21419 | -0.1071 | 0.28903 | 33.5884 | 0.21419 | 9.5017
0.104 | 1.4916 | 8.4787 | -34.3716 | 35.44 | -0.16258 | -0.2129 | -0.1071 | 0.28903 | 34.3716 | 0.2129 | 9.669
0.1058 | 1.5252 | 8.6981 | -35.0271 | 36.1303 | -0.16387 | -0.20774 | -0.10839 | 0.28645 | 35.0271 | 0.20774 | 9.8364
0.1076 | 1.5535 | 8.8929 | -35.5652 | 36.6994 | -0.16645 | -0.20258 | -0.10839 | 0.28516 | 35.5652 | 0.20258 | 10.0037
0.1094 | 1.5548 | 9.0219 | -35.9652 | 37.1187 | -0.16903 | -0.19613 | -0.10839 | 0.28258 | 35.9652 | 0.19613 | 10.1711
0.1111 | 1.5342 | 9.1381 | -36.3716 | 37.5406 | -0.17419 | -0.18839 | -0.1071 | 0.28129 | 36.3716 | 0.18839 | 10.3291
0.1129 | 1.52 | 9.2671 | -36.831 | 38.0181 | -0.17806 | -0.18194 | -0.1071 | 0.27871 | 36.831 | 0.18194 | 10.4965
0.1147 | 1.5032 | 9.4374 | -37.44 | 38.649 | -0.17935 | -0.17419 | -0.10581 | 0.27484 | 37.44 | 0.17419 | 10.6638
0.1165 | 1.5058 | 9.6271 | -38.151 | 39.3832 | -0.18194 | -0.16645 | -0.10323 | 0.27097 | 38.151 | 0.16645 | 10.8312
0.1183 | 1.5226 | 9.8206 | -38.9174 | 40.1755 | -0.18323 | -0.16 | -0.10194 | 0.26839 | 38.9174 | 0.16 | 10.9985
0.1201 | 1.5432 | 10.0103 | -39.6271 | 40.9097 | -0.18452 | -0.15355 | -0.099355 | 0.26581 | 39.6271 | 0.15355 | 11.1659
0.1219 | 1.56 | 10.169 | -40.1135 | 41.4206 | -0.18581 | -0.1471 | -0.098065 | 0.26452 | 40.1135 | 0.1471 | 11.3332
0.1237 | 1.5548 | 10.3032 | -40.4813 | 41.8103 | -0.1871 | -0.13935 | -0.096774 | 0.26194 | 40.4813 | 0.13935 | 11.5006
0.1255 | 1.5019 | 10.3781 | -40.6426 | 41.9819 | -0.18968 | -0.13032 | -0.095484 | 0.25806 | 40.6426 | 0.13032 | 11.6679
0.1273 | 1.449 | 10.4361 | -40.8555 | 42.2013 | -0.19226 | -0.12258 | -0.094194 | 0.25548 | 40.8555 | 0.12258 | 11.8353
0.1291 | 1.4232 | 10.511 | -41.3239 | 42.6723 | -0.19484 | -0.11484 | -0.091613 | 0.2529 | 41.3239 | 0.11484 | 12.0026
0.1309 | 1.4168 | 10.5935 | -41.7806 | 43.1355 | -0.19742 | -0.10839 | -0.089032 | 0.25032 | 41.7806 | 0.10839 | 12.17
0.1327 | 1.4348 | 10.7135 | -42.1806 | 43.5523 | -0.19871 | -0.10065 | -0.086452 | 0.24774 | 42.1806 | 0.10065 | 12.3373
0.1345 | 1.4555 | 10.831 | -42.5497 | 43.9381 | -0.19871 | -0.094194 | -0.083871 | 0.24516 | 42.5497 | 0.094194 | 12.5046
0.1362 | 1.4671 | 10.9161 | -42.729 | 44.1329 | -0.2 | -0.089032 | -0.082581 | 0.24387 | 42.729 | 0.089032 | 12.6627
0.138 | 1.4374 | 10.9484 | -42.7948 | 44.2026 | -0.20258 | -0.082581 | -0.08 | 0.24387 | 42.7948 | 0.082581 | 12.83
0.1398 | 1.3639 | 10.9174 | -42.7535 | 44.1523 | -0.20645 | -0.074839 | -0.07871 | 0.24516 | 42.7535 | 0.074839 | 12.9974
0.1416 | 1.2787 | 10.8748 | -42.7406 | 44.1252 | -0.21032 | -0.065806 | -0.077419 | 0.24516 | 42.7406 | 0.065806 | 13.1647
0.1434 | 1.2232 | 10.9123 | -43.0516 | 44.4335 | -0.2129 | -0.058065 | -0.074839 | 0.24645 | 43.0516 | 0.058065 | 13.3321
0.1452 | 1.1703 | 10.9755 | -43.511 | 44.8929 | -0.21548 | -0.050323 | -0.072258 | 0.24774 | 43.511 | 0.050323 | 13.4994
0.147 | 1.1277 | 11.0619 | -44.1535 | 45.5355 | -0.21806 | -0.042581 | -0.068387 | 0.2529 | 44.1535 | 0.042581 | 13.6668
0.1488 | 1.1084 | 11.169 | -44.7381 | 46.1277 | -0.22065 | -0.037419 | -0.067097 | 0.25806 | 44.7381 | 0.037419 | 13.8341
0.1506 | 1.0826 | 11.2374 | -45.1897 | 46.5819 | -0.22194 | -0.032258 | -0.064516 | 0.26065 | 45.1897 | 0.032258 | 14.0015
0.1524 | 1.0658 | 11.2877 | -45.5471 | 46.9394 | -0.22194 | -0.027097 | -0.063226 | 0.26194 | 45.5471 | 0.027097 | 14.1688
0.1542 | 1.0735 | 11.3226 | -45.7832 | 47.1768 | -0.22323 | -0.025806 | -0.061935 | 0.26452 | 45.7832 | 0.025806 | 14.3362
0.156 | 1.0877 | 11.3084 | -45.8929 | 47.28 | -0.22323 | -0.023226 | -0.061935 | 0.26581 | 45.8929 | 0.023226 | 14.5035
0.1578 | 1.0761 | 11.2413 | -45.8735 | 47.2452 | -0.22323 | -0.018065 | -0.060645 | 0.26581 | 45.8735 | 0.018065 | 14.6709
0.1595 | 1.0426 | 11.1265 | -45.7703 | 47.1174 | -0.22323 | -0.010323 | -0.058065 | 0.26968 | 45.7703 | 0.010323 | 14.8289
0.1613 | 0.99742 | 10.9742 | -45.5755 | 46.8916 | -0.22452 | -0.003871 | -0.056774 | 0.27355 | 45.5755 | 0.003871 | 14.9963
0.1631 | 0.94581 | 10.7845 | -45.4181 | 46.6955 | -0.2271 | 0.0025806 | -0.055484 | 0.27742 | 45.4181 | -0.0025806 | 15.1636
0.1649 | 0.88774 | 10.591 | -45.3845 | 46.6194 | -0.23097 | 0.0077419 | -0.054194 | 0.28258 | 45.3845 | -0.0077419 | 15.331
0.1667 | 0.83226 | 10.4142 | -45.4297 | 46.6245 | -0.23355 | 0.012903 | -0.051613 | 0.28645 | 45.4297 | -0.012903 | 15.4983
0.1685 | 0.78452 | 10.2594 | -45.5548 | 46.7123 | -0.23613 | 0.018065 | -0.050323 | 0.28903 | 45.5548 | -0.018065 | 15.6657
0.1703 | 0.78194 | 10.1432 | -45.9535 | 47.0787 | -0.23484 | 0.019355 | -0.050323 | 0.28903 | 45.9535 | -0.019355 | 15.833
0.1721 | 0.77419 | 10.0516 | -46.4116 | 47.5097 | -0.23355 | 0.019355 | -0.049032 | 0.29032 | 46.4116 | -0.019355 | 16.0004
0.1739 | 0.78065 | 9.96 | -46.8581 | 47.929 | -0.23226 | 0.019355 | -0.049032 | 0.29161 | 46.8581 | -0.019355 | 16.1677
0.1757 | 0.81806 | 9.8942 | -47.2697 | 48.32 | -0.2271 | 0.020645 | -0.050323 | 0.28903 | 47.2697 | -0.020645 | 16.3351
0.1775 | 0.86452 | 9.7974 | -47.5897 | 48.6168 | -0.22194 | 0.021935 | -0.050323 | 0.28516 | 47.5897 | -0.021935 | 16.5024
0.1793 | 0.90581 | 9.6916 | -47.7742 | 48.7794 | -0.21548 | 0.024516 | -0.050323 | 0.28129 | 47.7742 | -0.024516 | 16.6698
0.1811 | 0.96774 | 9.5755 | -47.8684 | 48.8503 | -0.20774 | 0.027097 | -0.051613 | 0.27613 | 47.8684 | -0.027097 | 16.8371
0.1829 | 1.0065 | 9.4219 | -47.8723 | 48.8258 | -0.20129 | 0.030968 | -0.051613 | 0.27226 | 47.8723 | -0.030968 | 17.0045
0.1846 | 1.0271 | 9.231 | -47.7755 | 48.6955 | -0.19613 | 0.03871 | -0.051613 | 0.26839 | 47.7755 | -0.03871 | 17.1625
0.1864 | 0.99742 | 9.0039 | -47.5716 | 48.449 | -0.19226 | 0.049032 | -0.050323 | 0.26452 | 47.5716 | -0.049032 | 17.3299
0.1882 | 0.96387 | 8.7458 | -47.4194 | 48.2477 | -0.18968 | 0.060645 | -0.049032 | 0.26323 | 47.4194 | -0.060645 | 17.4972
0.19 | 0.97677 | 8.4826 | -47.4426 | 48.2219 | -0.18452 | 0.070968 | -0.049032 | 0.26065 | 47.4426 | -0.070968 | 17.6646
0.1918 | 0.99871 | 8.2103 | -47.5806 | 48.3097 | -0.17935 | 0.08129 | -0.047742 | 0.25935 | 47.5806 | -0.08129 | 17.8319
0.1936 | 1.031 | 7.9394 | -47.9148 | 48.5961 | -0.17419 | 0.092903 | -0.046452 | 0.25806 | 47.9148 | -0.092903 | 17.9993
0.1954 | 1.0684 | 7.6787 | -48.3871 | 49.0219 | -0.16645 | 0.1071 | -0.045161 | 0.25548 | 48.3871 | -0.1071 | 18.1666
0.1972 | 1.1058 | 7.4323 | -48.9884 | 49.5781 | -0.15742 | 0.12 | -0.043871 | 0.2529 | 48.9884 | -0.12 | 18.334
0.199 | 1.1819 | 7.2065 | -49.7419 | 50.2929 | -0.14839 | 0.13032 | -0.043871 | 0.2529 | 49.7419 | -0.13032 | 18.5013
0.2008 | 1.311 | 7.031 | -50.5729 | 51.0877 | -0.14065 | 0.13548 | -0.043871 | 0.24516 | 50.5729 | -0.13548 | 18.6687
0.2026 | 1.4452 | 6.7368 | -51.2103 | 51.6826 | -0.13161 | 0.14452 | -0.043871 | 0.23742 | 51.2103 | -0.14452 | 18.836
0.2044 | 1.5729 | 6.4361 | -51.9071 | 52.3419 | -0.12129 | 0.15613 | -0.043871 | 0.23484 | 51.9071 | -0.15613 | 19.0033
0.2062 | 1.6813 | 6.1729 | -52.8116 | 53.2129 | -0.10839 | 0.17032 | -0.042581 | 0.23226 | 52.8116 | -0.17032 | 19.1707
0.208 | 1.7497 | 5.92 | -53.9613 | 54.3303 | -0.096774 | 0.1871 | -0.042581 | 0.23484 | 53.9613 | -0.1871 | 19.338
0.2097 | 1.8103 | 5.6555 | -55.2555 | 55.5897 | -0.086452 | 0.20774 | -0.042581 | 0.24258 | 55.2555 | -0.20774 | 19.4961
0.2115 | 1.9058 | 5.4052 | -56.8219 | 57.1265 | -0.073548 | 0.22323 | -0.042581 | 0.25032 | 56.8219 | -0.22323 | 19.6634
0.2133 | 2.0039 | 5.1445 | -58.6219 | 58.8994 | -0.060645 | 0.24258 | -0.043871 | 0.26323 | 58.6219 | -0.24258 | 19.8308
0.2151 | 2.1097 | 4.8929 | -60.5626 | 60.8168 | -0.047742 | 0.25935 | -0.045161 | 0.27484 | 60.5626 | -0.25935 | 19.9981
0.2169 | 2.2052 | 4.6387 | -62.7252 | 62.9548 | -0.034839 | 0.27613 | -0.046452 | 0.28774 | 62.7252 | -0.27613 | 20.1655
0.2187 | 2.289 | 4.3458 | -65.1523 | 65.3587 | -0.020645 | 0.29677 | -0.049032 | 0.30581 | 65.1523 | -0.29677 | 20.3328
0.2205 | 2.3948 | 4.049 | -67.7806 | 67.9639 | -0.0051613 | 0.31484 | -0.051613 | 0.32387 | 67.7806 | -0.31484 | 20.5002
0.2223 | 2.5252 | 3.7458 | -70.5355 | 70.7032 | 0.011613 | 0.3329 | -0.055484 | 0.34323 | 70.5355 | -0.3329 | 20.6675
0.2241 | 2.6632 | 3.4284 | -73.3071 | 73.4606 | 0.027097 | 0.34968 | -0.058065 | 0.36 | 73.3071 | -0.34968 | 20.8349
0.2259 | 2.8219 | 3.0916 | -76.1445 | 76.2839 | 0.043871 | 0.36387 | -0.061935 | 0.37677 | 76.1445 | -0.36387 | 21.0022
0.2277 | 3.04 | 2.7445 | -79.3148 | 79.4465 | 0.063226 | 0.37161 | -0.067097 | 0.3871 | 79.3148 | -0.37161 | 21.1696
0.2295 | 3.2684 | 2.3871 | -82.5961 | 82.72 | 0.082581 | 0.37806 | -0.070968 | 0.39871 | 82.5961 | -0.37806 | 21.3369
0.2313 | 3.5187 | 2.0155 | -86.0323 | 86.1497 | 0.10323 | 0.38194 | -0.076129 | 0.40774 | 86.0323 | -0.38194 | 21.5043
0.2331 | 3.8129 | 1.6374 | -89.6942 | 89.8129 | 0.12645 | 0.38194 | -0.08129 | 0.41677 | 89.6942 | -0.38194 | 21.6716
0.2348 | 4.1342 | 1.2529 | -93.3652 | 93.4877 | 0.15097 | 0.37806 | -0.086452 | 0.42581 | 93.3652 | -0.37806 | 21.8297
0.2366 | 4.4581 | 0.86839 | -97.0568 | 97.1858 | 0.17419 | 0.3729 | -0.092903 | 0.43355 | 97.0568 | -0.3729 | 21.997
0.2384 | 4.7871 | 0.49032 | -100.8723 | 101.0077 | 0.19871 | 0.36645 | -0.099355 | 0.44258 | 100.8723 | -0.36645 | 22.1644
0.2402 | 5.1458 | 0.12903 | -104.9019 | 105.049 | 0.22452 | 0.35355 | -0.10581 | 0.45548 | 104.9019 | -0.35355 | 22.3317
0.242 | 5.52 | -0.19484 | -109.1329 | 109.2916 | 0.2529 | 0.33677 | -0.11226 | 0.46968 | 109.1329 | -0.33677 | 22.4991
0.2438 | 5.8852 | -0.49935 | -113.4297 | 113.6039 | 0.28129 | 0.32129 | -0.11742 | 0.48516 | 113.4297 | -0.32129 | 22.6664
0.2456 | 6.2542 | -0.77935 | -117.6258 | 117.8142 | 0.31097 | 0.30323 | -0.12387 | 0.50194 | 117.6258 | -0.30323 | 22.8338
0.2474 | 6.6374 | -1.0774 | -121.9781 | 122.1832 | 0.33935 | 0.28387 | -0.13032 | 0.52 | 121.9781 | -0.28387 | 23.0011
0.2492 | 7.0155 | -1.3935 | -126.409 | 126.6284 | 0.36903 | 0.26323 | -0.13677 | 0.53935 | 126.409 | -0.26323 | 23.1685
0.251 | 7.3729 | -1.6839 | -130.7484 | 130.9845 | 0.39871 | 0.24 | -0.14323 | 0.56129 | 130.7484 | -0.24 | 23.3358
0.2528 | 7.7006 | -2.0026 | -135.1187 | 135.369 | 0.42581 | 0.21806 | -0.14968 | 0.58581 | 135.1187 | -0.21806 | 23.5032
0.2546 | 8.0258 | -2.329 | -139.4013 | 139.6671 | 0.45161 | 0.19355 | -0.15613 | 0.61161 | 139.4013 | -0.19355 | 23.6705
0.2564 | 8.3961 | -2.6374 | -144.2426 | 144.5265 | 0.48258 | 0.16903 | -0.16387 | 0.63871 | 144.2426 | -0.16903 | 23.8379
0.2581 | 8.7523 | -2.9381 | -149.2335 | 149.5329 | 0.51484 | 0.14839 | -0.17161 | 0.66194 | 149.2335 | -0.14839 | 23.9959
0.2599 | 9.1729 | -3.2839 | -154.3265 | 154.6477 | 0.54839 | 0.12645 | -0.17935 | 0.68774 | 154.3265 | -0.12645 | 24.1633
0.2617 | 9.5587 | -3.6168 | -159.2194 | 159.56 | 0.58065 | 0.10194 | -0.1871 | 0.71742 | 159.2194 | -0.10194 | 24.3306
0.2635 | 9.9445 | -3.9574 | -164.0555 | 164.4155 | 0.6129 | 0.076129 | -0.19355 | 0.7471 | 164.0555 | -0.076129 | 24.498
0.2653 | 10.3226 | -4.2929 | -168.7935 | 169.1755 | 0.64516 | 0.050323 | -0.20129 | 0.77935 | 168.7935 | -0.050323 | 24.6653
0.2671 | 10.6129 | -4.6116 | -173.0039 | 173.4013 | 0.67355 | 0.029677 | -0.20774 | 0.80774 | 173.0039 | -0.029677 | 24.8327
0.2689 | 10.8723 | -4.9252 | -176.9768 | 177.391 | 0.69935 | 0.011613 | -0.21548 | 0.83484 | 176.9768 | -0.011613 | 25
0.2707 | 11.1639 | -5.3174 | -181.7213 | 182.1535 | 0.73032 | -0.011613 | -0.22194 | 0.8671 | 181.7213 | 0.011613 | 25.1673
0.2725 | 11.4555 | -5.7213 | -186.6735 | 187.1239 | 0.76258 | -0.037419 | -0.22968 | 0.90194 | 186.6735 | 0.037419 | 25.3347
0.2743 | 11.7639 | -6.1381 | -191.8516 | 192.3213 | 0.79613 | -0.063226 | -0.23742 | 0.93677 | 191.8516 | 0.063226 | 25.502
0.2761 | 12.0606 | -6.5239 | -197.0142 | 197.5019 | 0.83097 | -0.091613 | -0.24516 | 0.9729 | 197.0142 | 0.091613 | 25.6694
0.2779 | 12.3497 | -6.88 | -202.2065 | 202.711 | 0.86581 | -0.12 | -0.2529 | 1.0116 | 202.2065 | 0.12 | 25.8367
0.2797 | 12.64 | -7.2297 | -207.4503 | 207.9729 | 0.90323 | -0.14968 | -0.26194 | 1.0503 | 207.4503 | 0.14968 | 26.0041
0.2815 | 12.9368 | -7.5806 | -212.7277 | 213.2671 | 0.94065 | -0.18065 | -0.26968 | 1.0903 | 212.7277 | 0.18065 | 26.1714
0.2832 | 13.2374 | -7.9213 | -218.0129 | 218.5703 | 0.97806 | -0.2129 | -0.27871 | 1.1316 | 218.0129 | 0.2129 | 26.3295
0.285 | 13.5471 | -8.2723 | -223.44 | 224.0142 | 1.0181 | -0.24516 | -0.28774 | 1.1755 | 223.44 | 0.24516 | 26.4968
0.2868 | 13.8439 | -8.591 | -228.7045 | 229.2955 | 1.0581 | -0.28 | -0.29677 | 1.2194 | 228.7045 | 0.28 | 26.6642
0.2886 | 14.1213 | -8.911 | -233.8387 | 234.4452 | 1.0981 | -0.31226 | -0.30581 | 1.2632 | 233.8387 | 0.31226 | 26.8315
0.2904 | 14.3781 | -9.2413 | -238.7523 | 239.3729 | 1.1342 | -0.3471 | -0.31226 | 1.3071 | 238.7523 | 0.3471 | 26.9989
0.2922 | 14.6245 | -9.569 | -243.6065 | 244.2426 | 1.1716 | -0.38194 | -0.32 | 1.351 | 243.6065 | 0.38194 | 27.1662
0.294 | 14.8323 | -9.8916 | -248.4426 | 249.0903 | 1.2077 | -0.41548 | -0.32774 | 1.3948 | 248.4426 | 0.41548 | 27.3336
0.2958 | 15.0026 | -10.1961 | -253.0465 | 253.7045 | 1.2439 | -0.44903 | -0.33419 | 1.4374 | 253.0465 | 0.44903 | 27.5009
0.2976 | 15.1613 | -10.4929 | -256.9961 | 257.6658 | 1.2735 | -0.48387 | -0.33935 | 1.4748 | 256.9961 | 0.48387 | 27.6683
0.2994 | 15.2877 | -10.7781 | -260.7742 | 261.4529 | 1.3006 | -0.51742 | -0.34323 | 1.511 | 260.7742 | 0.51742 | 27.8356
0.3012 | 15.4516 | -11.0129 | -264.3935 | 265.0813 | 1.3303 | -0.55613 | -0.34839 | 1.551 | 264.3935 | 0.55613 | 28.003
0.303 | 15.6052 | -11.1806 | -267.8387 | 268.5342 | 1.3613 | -0.59355 | -0.35355 | 1.5923 | 267.8387 | 0.59355 | 28.1703
0.3048 | 15.76 | -11.3471 | -271.2168 | 271.9187 | 1.3923 | -0.63226 | -0.36 | 1.6348 | 271.2168 | 0.63226 | 28.3377
0.3066 | 15.9174 | -11.5484 | -274.4774 | 275.1897 | 1.4219 | -0.67613 | -0.36387 | 1.6761 | 274.4774 | 0.67613 | 28.505
0.3083 | 16.0323 | -11.769 | -277.5213 | 278.2413 | 1.449 | -0.71871 | -0.36903 | 1.7135 | 277.5213 | 0.71871 | 28.6631
0.3101 | 16.0865 | -11.9948 | -280.3226 | 281.0477 | 1.471 | -0.76 | -0.3729 | 1.7484 | 280.3226 | 0.76 | 28.8304
0.3119 | 16.12 | -12.2206 | -283.0168 | 283.7484 | 1.4929 | -0.8 | -0.37677 | 1.7858 | 283.0168 | 0.8 | 28.9978
0.3137 | 16.1652 | -12.4748 | -285.551 | 286.289 | 1.5135 | -0.84129 | -0.38065 | 1.8206 | 285.551 | 0.84129 | 29.1651
0.3155 | 16.1794 | -12.72 | -287.8387 | 288.5819 | 1.5329 | -0.88387 | -0.38452 | 1.8568 | 287.8387 | 0.88387 | 29.3325
0.3173 | 16.1432 | -12.9303 | -289.9265 | 290.6723 | 1.5535 | -0.92516 | -0.38839 | 1.8968 | 289.9265 | 0.92516 | 29.4998
0.3191 | 16.1213 | -13.1703 | -292.1045 | 292.8542 | 1.5716 | -0.96774 | -0.39355 | 1.9329 | 292.1045 | 0.96774 | 29.6672
0.3209 | 16.0968 | -13.3961 | -294.1665 | 294.92 | 1.5884 | -1.0103 | -0.39742 | 1.9652 | 294.1665 | 1.0103 | 29.8345
0.3227 | 16.0219 | -13.5574 | -296.0258 | 296.7794 | 1.6052 | -1.0516 | -0.4 | 1.9987 | 296.0258 | 1.0516 | 30.0019
0.3245 | 15.8903 | -13.6929 | -297.7084 | 298.4568 | 1.6219 | -1.0903 | -0.40387 | 2.0323 | 297.7084 | 1.0903 | 30.1692
0.3263 | 15.7716 | -13.831 | -299.3613 | 300.1071 | 1.6374 | -1.1277 | -0.40774 | 2.0671 | 299.3613 | 1.1277 | 30.3366
0.3281 | 15.6671 | -13.9239 | -300.9626 | 301.7032 | 1.6542 | -1.1677 | -0.41032 | 2.0994 | 300.9626 | 1.1677 | 30.5039
0.3299 | 15.4955 | -14.0168 | -302.5161 | 303.2503 | 1.6697 | -1.2065 | -0.4129 | 2.1316 | 302.5161 | 1.2065 | 30.6713
0.3316 | 15.28 | -14.1045 | -303.9252 | 304.649 | 1.6852 | -1.2465 | -0.41677 | 2.1665 | 303.9252 | 1.2465 | 30.8293
0.3334 | 15.0671 | -14.209 | -305.2413 | 305.9561 | 1.6994 | -1.2839 | -0.41935 | 2.2 | 305.2413 | 1.2839 | 30.9967
0.3352 | 14.8568 | -14.32 | -306.4852 | 307.1935 | 1.7135 | -1.3226 | -0.42323 | 2.2348 | 306.4852 | 1.3226 | 31.164
0.337 | 14.6452 | -14.4323 | -307.671 | 308.3716 | 1.7265 | -1.3587 | -0.42581 | 2.2658 | 307.671 | 1.3587 | 31.3313
0.3388 | 14.449 | -14.5729 | -308.7665 | 309.4632 | 1.7368 | -1.3923 | -0.42839 | 2.2942 | 308.7665 | 1.3923 | 31.4987
0.3406 | 14.2632 | -14.7458 | -309.7329 | 310.4271 | 1.7458 | -1.4284 | -0.43097 | 2.3239 | 309.7329 | 1.4284 | 31.666
0.3424 | 14.0916 | -14.9497 | -310.5677 | 311.2632 | 1.7523 | -1.4645 | -0.43355 | 2.351 | 310.5677 | 1.4645 | 31.8334
0.3442 | 13.9148 | -15.1484 | -311.3368 | 312.0348 | 1.7574 | -1.4968 | -0.43484 | 2.3768 | 311.3368 | 1.4968 | 32.0007
0.346 | 13.6903 | -15.3445 | -312.0465 | 312.7432 | 1.7626 | -1.5226 | -0.43484 | 2.3987 | 312.0465 | 1.5226 | 32.1681
0.3478 | 13.4477 | -15.511 | -312.7406 | 313.4348 | 1.7665 | -1.551 | -0.43613 | 2.4206 | 312.7406 | 1.551 | 32.3354
0.3496 | 13.2026 | -15.6723 | -313.4361 | 314.1265 | 1.7703 | -1.5806 | -0.43742 | 2.4439 | 313.4361 | 1.5806 | 32.5028
0.3514 | 12.92 | -15.8168 | -314.0981 | 314.7819 | 1.7755 | -1.6077 | -0.43871 | 2.4645 | 314.0981 | 1.6077 | 32.6701
0.3532 | 12.5794 | -15.929 | -314.6052 | 315.2813 | 1.7819 | -1.6284 | -0.44129 | 2.4839 | 314.6052 | 1.6284 | 32.8375
0.355 | 12.2813 | -16.0735 | -315.0155 | 315.6865 | 1.7871 | -1.6516 | -0.44258 | 2.5006 | 315.0155 | 1.6516 | 33.0048
0.3567 | 11.9897 | -16.2426 | -315.2929 | 315.9613 | 1.791 | -1.6684 | -0.44516 | 2.5148 | 315.2929 | 1.6684 | 33.1629
0.3585 | 11.6619 | -16.3497 | -315.489 | 316.151 | 1.7974 | -1.6865 | -0.44774 | 2.5316 | 315.489 | 1.6865 | 33.3302
0.3603 | 11.3187 | -16.4297 | -315.6581 | 316.311 | 1.8052 | -1.7019 | -0.45161 | 2.5484 | 315.6581 | 1.7019 | 33.4976
0.3621 | 10.9652 | -16.591 | -315.7639 | 316.4142 | 1.8103 | -1.7045 | -0.45677 | 2.5548 | 315.7639 | 1.7045 | 33.6649
0.3639 | 10.609 | -16.7535 | -315.8555 | 316.5058 | 1.8168 | -1.7084 | -0.46194 | 2.5626 | 315.8555 | 1.7084 | 33.8323
0.3657 | 10.2568 | -16.8955 | -315.9135 | 316.5587 | 1.8206 | -1.72 | -0.46452 | 2.5755 | 315.9135 | 1.72 | 33.9996
0.3675 | 9.9097 | -17.0439 | -315.9394 | 316.5819 | 1.8206 | -1.7316 | -0.4671 | 2.5832 | 315.9394 | 1.7316 | 34.167
0.3693 | 9.5677 | -17.1303 | -315.9123 | 316.5484 | 1.8245 | -1.7484 | -0.46968 | 2.5961 | 315.9123 | 1.7484 | 34.3343
0.3711 | 9.2232 | -17.1781 | -315.7652 | 316.3948 | 1.8284 | -1.7665 | -0.47226 | 2.6116 | 315.7652 | 1.7665 | 34.5017
0.3729 | 8.8942 | -17.2774 | -315.5019 | 316.1277 | 1.8297 | -1.7871 | -0.47484 | 2.6271 | 315.5019 | 1.7871 | 34.669
0.3747 | 8.5471 | -17.3135 | -315.2181 | 315.8335 | 1.8284 | -1.809 | -0.47613 | 2.6413 | 315.2181 | 1.809 | 34.8364
0.3765 | 8.1755 | -17.3329 | -314.8542 | 315.4619 | 1.8284 | -1.8155 | -0.48 | 2.6465 | 314.8542 | 1.8155 | 35.0037
0.3783 | 7.8516 | -17.4155 | -314.4361 | 315.0413 | 1.8284 | -1.8323 | -0.48129 | 2.6581 | 314.4361 | 1.8323 | 35.1711
0.3801 | 7.5045 | -17.5484 | -313.951 | 314.5561 | 1.8271 | -1.8426 | -0.48258 | 2.6645 | 313.951 | 1.8426 | 35.3384
0.3818 | 7.1858 | -17.6735 | -313.3923 | 313.9974 | 1.8232 | -1.8632 | -0.48 | 2.6761 | 313.3923 | 1.8632 | 35.4965
0.3836 | 6.7974 | -17.7935 | -312.8026 | 313.4077 | 1.8206 | -1.8787 | -0.48 | 2.6852 | 312.8026 | 1.8787 | 35.6638
0.3854 | 6.4245 | -17.8968 | -312.1794 | 312.7832 | 1.8168 | -1.8929 | -0.48129 | 2.6929 | 312.1794 | 1.8929 | 35.8312
0.3872 | 6.1084 | -18.0374 | -311.4594 | 312.0697 | 1.8155 | -1.9006 | -0.48258 | 2.6968 | 311.4594 | 1.9006 | 35.9985
0.389 | 5.8065 | -18.1213 | -310.7923 | 311.4026 | 1.8116 | -1.92 | -0.48387 | 2.7045 | 310.7923 | 1.92 | 36.1659
0.3908 | 5.4994 | -18.1729 | -310.2632 | 310.8723 | 1.809 | -1.9265 | -0.48774 | 2.7071 | 310.2632 | 1.9265 | 36.3332
0.3926 | 5.2013 | -18.2103 | -309.7561 | 310.3626 | 1.8077 | -1.9329 | -0.49161 | 2.7097 | 309.7561 | 1.9329 | 36.5006
0.3944 | 4.9161 | -18.2323 | -309.1768 | 309.7806 | 1.8065 | -1.9406 | -0.49419 | 2.7135 | 309.1768 | 1.9406 | 36.6679
0.3962 | 4.6181 | -18.1974 | -308.4452 | 309.0452 | 1.8026 | -1.9497 | -0.49548 | 2.7187 | 308.4452 | 1.9497 | 36.8353
0.398 | 4.2568 | -18.1174 | -307.6619 | 308.2529 | 1.7961 | -1.9613 | -0.49806 | 2.7239 | 307.6619 | 1.9613 | 37.0026
0.3998 | 3.9561 | -18.0619 | -306.8232 | 307.409 | 1.7961 | -1.9665 | -0.49935 | 2.7277 | 306.8232 | 1.9665 | 37.17
0.4016 | 3.6503 | -17.9897 | -305.849 | 306.4284 | 1.7935 | -1.969 | -0.50194 | 2.729 | 305.849 | 1.969 | 37.3373
0.4034 | 3.3523 | -17.9445 | -304.7794 | 305.3548 | 1.7897 | -1.969 | -0.50194 | 2.7277 | 304.7794 | 1.969 | 37.5046
0.4051 | 3.031 | -17.8761 | -303.6787 | 304.2477 | 1.7858 | -1.9703 | -0.50323 | 2.7277 | 303.6787 | 1.9703 | 37.6627
0.4069 | 2.76 | -17.7523 | -302.5135 | 303.0748 | 1.7819 | -1.9665 | -0.50581 | 2.7239 | 302.5135 | 1.9665 | 37.83
0.4087 | 2.5174 | -17.6426 | -301.3123 | 301.8671 | 1.7768 | -1.9626 | -0.50968 | 2.7187 | 301.3123 | 1.9626 | 37.9974
0.4105 | 2.2658 | -17.5497 | -300.0555 | 300.6039 | 1.7716 | -1.9548 | -0.51226 | 2.711 | 300.0555 | 1.9548 | 38.1647
0.4123 | 2 | -17.4645 | -298.8387 | 299.3819 | 1.7639 | -1.9484 | -0.51613 | 2.7032 | 298.8387 | 1.9484 | 38.3321
0.4141 | 1.7032 | -17.3858 | -297.631 | 298.169 | 1.7548 | -1.9419 | -0.52129 | 2.6942 | 297.631 | 1.9419 | 38.4994
0.4159 | 1.44 | -17.2426 | -296.4503 | 296.9819 | 1.7471 | -1.9303 | -0.52645 | 2.6826 | 296.4503 | 1.9303 | 38.6668
0.4177 | 1.1871 | -17.0619 | -295.3226 | 295.8452 | 1.7394 | -1.92 | -0.5329 | 2.671 | 295.3226 | 1.92 | 38.8341
0.4195 | 1.0194 | -16.8606 | -293.9858 | 294.4981 | 1.7342 | -1.9084 | -0.53935 | 2.6619 | 293.9858 | 1.9084 | 39.0015
0.4213 | 0.8529 | -16.6245 | -292.6026 | 293.1032 | 1.729 | -1.8981 | -0.54581 | 2.6542 | 292.6026 | 1.8981 | 39.1688
0.4231 | 0.78581 | -16.4503 | -291.1265 | 291.6194 | 1.72 | -1.8852 | -0.55355 | 2.64 | 291.1265 | 1.8852 | 39.3362
0.4249 | 0.6671 | -16.3174 | -289.8245 | 290.3123 | 1.711 | -1.8723 | -0.55871 | 2.6258 | 289.8245 | 1.8723 | 39.5035
0.4267 | 0.55226 | -16.2116 | -288.6348 | 289.1187 | 1.7006 | -1.8594 | -0.56258 | 2.6116 | 288.6348 | 1.8594 | 39.6709
0.4285 | 0.46065 | -16.0723 | -287.3535 | 287.8323 | 1.689 | -1.8477 | -0.56774 | 2.5961 | 287.3535 | 1.8477 | 39.8382
0.4302 | 0.40903 | -15.8761 | -285.9174 | 286.3884 | 1.6774 | -1.8361 | -0.5729 | 2.5819 | 285.9174 | 1.8361 | 39.9963
0.432 | 0.3729 | -15.6955 | -284.3923 | 284.8555 | 1.6658 | -1.8219 | -0.57677 | 2.5652 | 284.3923 | 1.8219 | 40.1636
0.4338 | 0.30839 | -15.4361 | -282.8955 | 283.3484 | 1.6568 | -1.8026 | -0.58323 | 2.5471 | 282.8955 | 1.8026 | 40.331
0.4356 | 0.23742 | -15.2129 | -281.3355 | 281.7781 | 1.6452 | -1.7819 | -0.58968 | 2.5265 | 281.3355 | 1.7819 | 40.4983
0.4374 | 0.17419 | -15.0232 | -279.7948 | 280.2297 | 1.631 | -1.7613 | -0.59484 | 2.5045 | 279.7948 | 1.7613 | 40.6657
0.4392 | 0.13806 | -14.8374 | -278.2439 | 278.6697 | 1.6168 | -1.7406 | -0.6 | 2.4813 | 278.2439 | 1.7406 | 40.833
0.441 | 0.12258 | -14.6194 | -276.8 | 277.2168 | 1.6039 | -1.7239 | -0.60645 | 2.4632 | 276.8 | 1.7239 | 41.0004
0.4428 | 0.12516 | -14.3832 | -275.4426 | 275.849 | 1.5935 | -1.7097 | -0.61161 | 2.4477 | 275.4426 | 1.7097 | 41.1677
0.4446 | 0.14581 | -14.1226 | -273.7716 | 274.1652 | 1.5806 | -1.6929 | -0.61806 | 2.4284 | 273.7716 | 1.6929 | 41.3351
0.4464 | 0.18194 | -13.8994 | -271.9729 | 272.3574 | 1.5639 | -1.6761 | -0.62323 | 2.4077 | 271.9729 | 1.6761 | 41.5024
0.4482 | 0.21032 | -13.7174 | -270.1548 | 270.5329 | 1.5471 | -1.6555 | -0.6271 | 2.3845 | 270.1548 | 1.6555 | 41.6698
0.45 | 0.26581 | -13.5252 | -268.3174 | 268.6877 | 1.5303 | -1.6387 | -0.63226 | 2.3639 | 268.3174 | 1.6387 | 41.8371
0.4518 | 0.38968 | -13.3135 | -266.3419 | 266.7058 | 1.5123 | -1.6245 | -0.63613 | 2.3445 | 266.3419 | 1.6245 | 42.0045
0.4536 | 0.52516 | -13.0981 | -264.3484 | 264.7032 | 1.4955 | -1.6116 | -0.63871 | 2.3252 | 264.3484 | 1.6116 | 42.1718
0.4553 | 0.64258 | -12.8684 | -262.3226 | 262.6684 | 1.4774 | -1.5987 | -0.64258 | 2.3058 | 262.3226 | 1.5987 | 42.3299
0.4571 | 0.71355 | -12.6245 | -260.3058 | 260.6413 | 1.4606 | -1.5832 | -0.64645 | 2.2865 | 260.3058 | 1.5832 | 42.4972
0.4589 | 0.75355 | -12.369 | -258.3097 | 258.6348 | 1.4439 | -1.5677 | -0.64903 | 2.2671 | 258.3097 | 1.5677 | 42.6646
0.4607 | 0.78581 | -12.1239 | -256.3303 | 256.6465 | 1.4284 | -1.5497 | -0.6529 | 2.2465 | 256.3303 | 1.5497 | 42.8319
0.4625 | 0.82839 | -11.9045 | -254.2542 | 254.5626 | 1.409 | -1.5406 | -0.65548 | 2.2284 | 254.2542 | 1.5406 | 42.9993
0.4643 | 0.87484 | -11.6761 | -252.1432 | 252.4426 | 1.391 | -1.5265 | -0.65677 | 2.209 | 252.1432 | 1.5265 | 43.1666
0.4661 | 0.96258 | -11.4594 | -249.9935 | 250.2852 | 1.3742 | -1.5135 | -0.65806 | 2.1897 | 249.9935 | 1.5135 | 43.334
0.4679 | 1.0813 | -11.249 | -247.8129 | 248.0981 | 1.3561 | -1.5032 | -0.65935 | 2.1716 | 247.8129 | 1.5032 | 43.5013
0.4697 | 1.2439 | -11.0426 | -245.5174 | 245.7948 | 1.3381 | -1.4903 | -0.66065 | 2.1523 | 245.5174 | 1.4903 | 43.6687
0.4715 | 1.4103 | -10.8387 | -243.3252 | 243.5974 | 1.32 | -1.4826 | -0.66065 | 2.1355 | 243.3252 | 1.4826 | 43.836
0.4733 | 1.5445 | -10.6271 | -241.1987 | 241.4632 | 1.3019 | -1.4748 | -0.66194 | 2.12 | 241.1987 | 1.4748 | 44.0033
0.4751 | 1.6516 | -10.4155 | -239.1045 | 239.3639 | 1.2852 | -1.4684 | -0.66194 | 2.1032 | 239.1045 | 1.4684 | 44.1707
0.4769 | 1.7729 | -10.2297 | -237.0503 | 237.3045 | 1.2671 | -1.4619 | -0.66194 | 2.089 | 237.0503 | 1.4619 | 44.338
0.4786 | 1.9239 | -10.0516 | -235.1187 | 235.369 | 1.2516 | -1.4568 | -0.66323 | 2.0748 | 235.1187 | 1.4568 | 44.4961
0.4804 | 2.0271 | -9.88 | -233.3832 | 233.6297 | 1.2374 | -1.4542 | -0.66323 | 2.0645 | 233.3832 | 1.4542 | 44.6634
0.4822 | 2.0632 | -9.6826 | -231.8374 | 232.0787 | 1.2297 | -1.4465 | -0.66452 | 2.0542 | 231.8374 | 1.4465 | 44.8308
0.484 | 2.0645 | -9.4606 | -230.231 | 230.4645 | 1.2194 | -1.4413 | -0.66452 | 2.0439 | 230.231 | 1.4413 | 44.9981
0.4858 | 2.0916 | -9.2606 | -228.4413 | 228.6684 | 1.2065 | -1.4335 | -0.66452 | 2.031 | 228.4413 | 1.4335 | 45.1655
0.4876 | 2.1148 | -9.0942 | -226.6581 | 226.88 | 1.1923 | -1.4271 | -0.66581 | 2.0168 | 226.6581 | 1.4271 | 45.3328
0.4894 | 2.1419 | -8.9484 | -224.9148 | 225.1329 | 1.1794 | -1.4194 | -0.66581 | 2.0039 | 224.9148 | 1.4194 | 45.5002
0.4912 | 2.2116 | -8.8323 | -223.231 | 223.4477 | 1.1652 | -1.4142 | -0.6671 | 1.991 | 223.231 | 1.4142 | 45.6675
0.493 | 2.3252 | -8.7471 | -221.5845 | 221.8013 | 1.151 | -1.409 | -0.6671 | 1.9794 | 221.5845 | 1.409 | 45.8349
0.4948 | 2.4413 | -8.6697 | -219.8426 | 220.0594 | 1.1368 | -1.4065 | -0.6671 | 1.9677 | 219.8426 | 1.4065 | 46.0022
0.4966 | 2.5548 | -8.5974 | -218.031 | 218.249 | 1.1213 | -1.4026 | -0.6671 | 1.9548 | 218.031 | 1.4026 | 46.1696
0.4984 | 2.6426 | -8.4994 | -216.2206 | 216.4374 | 1.1071 | -1.4013 | -0.66581 | 1.9432 | 216.2206 | 1.4013 | 46.3369
0.5002 | 2.7329 | -8.4194 | -214.4219 | 214.6361 | 1.0942 | -1.3974 | -0.66581 | 1.9316 | 214.4219 | 1.3974 | 46.5043
0.502 | 2.8284 | -8.3277 | -212.5987 | 212.8129 | 1.08 | -1.3948 | -0.66452 | 1.9213 | 212.5987 | 1.3948 | 46.6716
0.5037 | 2.9277 | -8.2465 | -210.7806 | 210.9935 | 1.0671 | -1.3935 | -0.66323 | 1.9123 | 210.7806 | 1.3935 | 46.8297
0.5055 | 3.0335 | -8.1471 | -208.8839 | 209.0955 | 1.0542 | -1.3948 | -0.66194 | 1.9019 | 208.8839 | 1.3948 | 46.997
0.5073 | 3.1058 | -8.0439 | -207.0852 | 207.2942 | 1.0426 | -1.3935 | -0.66065 | 1.8929 | 207.0852 | 1.3935 | 47.1644
0.5091 | 3.1677 | -7.9277 | -205.3794 | 205.5858 | 1.031 | -1.3935 | -0.65935 | 1.8839 | 205.3794 | 1.3935 | 47.3317
0.5109 | 3.2284 | -7.8632 | -203.7368 | 203.9406 | 1.0206 | -1.3935 | -0.65935 | 1.8761 | 203.7368 | 1.3935 | 47.4991
0.5127 | 3.3032 | -7.8065 | -202.1406 | 202.3445 | 1.0103 | -1.3935 | -0.65806 | 1.8684 | 202.1406 | 1.3935 | 47.6664
0.5145 | 3.3665 | -7.7123 | -200.5626 | 200.7652 | 1.0013 | -1.3935 | -0.65677 | 1.8606 | 200.5626 | 1.3935 | 47.8338
0.5163 | 3.4194 | -7.6077 | -198.9845 | 199.1845 | 0.98968 | -1.3948 | -0.65548 | 1.8542 | 198.9845 | 1.3948 | 48.0011
0.5181 | 3.4529 | -7.5213 | -197.529 | 197.7265 | 0.97935 | -1.3974 | -0.65419 | 1.8477 | 197.529 | 1.3974 | 48.1685
0.5199 | 3.4529 | -7.4361 | -196.2116 | 196.4065 | 0.97032 | -1.3974 | -0.6529 | 1.8426 | 196.2116 | 1.3974 | 48.3358
0.5217 | 3.431 | -7.3548 | -194.9884 | 195.1806 | 0.96258 | -1.3974 | -0.65161 | 1.8361 | 194.9884 | 1.3974 | 48.5032
0.5235 | 3.4077 | -7.2994 | -193.9058 | 194.0955 | 0.95484 | -1.4013 | -0.65032 | 1.8335 | 193.9058 | 1.4013 | 48.6705
0.5253 | 3.3948 | -7.249 | -192.8439 | 193.0323 | 0.94839 | -1.4065 | -0.64774 | 1.8335 | 192.8439 | 1.4065 | 48.8379
0.5271 | 3.3729 | -7.1755 | -191.7445 | 191.9303 | 0.93935 | -1.4155 | -0.64387 | 1.8335 | 191.7445 | 1.4155 | 49.0052
0.5288 | 3.3613 | -7.1458 | -190.7458 | 190.9303 | 0.93161 | -1.4206 | -0.64258 | 1.831 | 190.7458 | 1.4206 | 49.1633
0.5306 | 3.3819 | -7.1316 | -189.7316 | 189.9161 | 0.92 | -1.4284 | -0.64129 | 1.8297 | 189.7316 | 1.4284 | 49.3306
0.5324 | 3.3858 | -7.1252 | -188.7845 | 188.9703 | 0.91226 | -1.4323 | -0.64 | 1.8284 | 188.7845 | 1.4323 | 49.498
0.5342 | 3.3781 | -7.1161 | -187.951 | 188.1368 | 0.90968 | -1.4387 | -0.63871 | 1.831 | 187.951 | 1.4387 | 49.6653
0.536 | 3.3703 | -7.0994 | -187.1394 | 187.3239 | 0.90581 | -1.4465 | -0.63742 | 1.8348 | 187.1394 | 1.4465 | 49.8327
0.5378 | 3.3303 | -7.0761 | -186.4142 | 186.5987 | 0.90323 | -1.4529 | -0.63484 | 1.8387 | 186.4142 | 1.4529 | 50
0.5396 | 3.2335 | -7.0413 | -185.8826 | 186.0619 | 0.90194 | -1.4606 | -0.63355 | 1.8426 | 185.8826 | 1.4606 | 50.1673
0.5414 | 3.1355 | -7.0219 | -185.3858 | 185.5613 | 0.90065 | -1.4684 | -0.63097 | 1.8465 | 185.3858 | 1.4684 | 50.3347
0.5432 | 3.0826 | -7.0477 | -184.8645 | 185.0387 | 0.89806 | -1.4787 | -0.62968 | 1.8529 | 184.8645 | 1.4787 | 50.502
0.545 | 3.0581 | -7.0787 | -184.2258 | 184.4026 | 0.89548 | -1.4903 | -0.62581 | 1.8606 | 184.2258 | 1.4903 | 50.6694
0.5468 | 3.0245 | -7.1187 | -183.6658 | 183.8439 | 0.89548 | -1.5006 | -0.62323 | 1.8671 | 183.6658 | 1.5006 | 50.8367
0.5486 | 2.9781 | -7.1239 | -183.0813 | 183.2594 | 0.89548 | -1.5123 | -0.61935 | 1.8761 | 183.0813 | 1.5123 | 51.0041
0.5504 | 2.871 | -7.1239 | -182.7148 | 182.8916 | 0.89806 | -1.5226 | -0.61677 | 1.8852 | 182.7148 | 1.5226 | 51.1714
0.5521 | 2.7819 | -7.169 | -182.4877 | 182.6645 | 0.90194 | -1.5329 | -0.61419 | 1.8942 | 182.4877 | 1.5329 | 51.3295
0.5539 | 2.7174 | -7.2206 | -182.2142 | 182.3923 | 0.90581 | -1.5471 | -0.61161 | 1.9071 | 182.2142 | 1.5471 | 51.4968
0.5557 | 2.6774 | -7.3019 | -181.9329 | 182.1148 | 0.90839 | -1.56 | -0.60903 | 1.9187 | 181.9329 | 1.56 | 51.6642
0.5575 | 2.631 | -7.3871 | -181.6839 | 181.8684 | 0.91097 | -1.5742 | -0.60645 | 1.9303 | 181.6839 | 1.5742 | 51.8315
0.5593 | 2.5419 | -7.4955 | -181.6774 | 181.8658 | 0.91613 | -1.5858 | -0.60387 | 1.9406 | 181.6774 | 1.5858 | 51.9989
0.5611 | 2.4465 | -7.6245 | -181.7561 | 181.9484 | 0.92 | -1.5974 | -0.6 | 1.9523 | 181.7561 | 1.5974 | 52.1662
0.5629 | 2.4052 | -7.7742 | -181.7974 | 181.9974 | 0.92387 | -1.6168 | -0.59742 | 1.969 | 181.7974 | 1.6168 | 52.3336
0.5647 | 2.3677 | -7.9252 | -181.8374 | 182.0439 | 0.92645 | -1.6374 | -0.59355 | 1.9858 | 181.8374 | 1.6374 | 52.5009
0.5665 | 2.2645 | -8.0465 | -181.9561 | 182.1665 | 0.93161 | -1.6529 | -0.58968 | 2.0013 | 181.9561 | 1.6529 | 52.6683
0.5683 | 2.1071 | -8.1368 | -182.1587 | 182.3716 | 0.93935 | -1.6671 | -0.5871 | 2.0168 | 182.1587 | 1.6671 | 52.8356
0.5701 | 2.0335 | -8.2465 | -182.4335 | 182.6503 | 0.9471 | -1.6877 | -0.58323 | 2.0361 | 182.4335 | 1.6877 | 53.003
0.5719 | 1.9716 | -8.3587 | -182.711 | 182.9316 | 0.95355 | -1.7097 | -0.57935 | 2.0568 | 182.711 | 1.7097 | 53.1703
0.5737 | 1.9523 | -8.4761 | -182.8826 | 183.1097 | 0.96129 | -1.729 | -0.57548 | 2.0748 | 182.8826 | 1.729 | 53.3377
0.5755 | 1.9032 | -8.6103 | -182.9342 | 183.1677 | 0.96774 | -1.7445 | -0.5729 | 2.0916 | 182.9342 | 1.7445 | 53.505
0.5772 | 1.8206 | -8.7432 | -183.0568 | 183.2955 | 0.97548 | -1.7587 | -0.56903 | 2.1071 | 183.0568 | 1.7587 | 53.6631
0.579 | 1.6529 | -8.84 | -183.4568 | 183.6994 | 0.98581 | -1.7703 | -0.56516 | 2.1213 | 183.4568 | 1.7703 | 53.8304
0.5808 | 1.5394 | -8.9523 | -183.9432 | 184.191 | 0.99613 | -1.7845 | -0.55871 | 2.1368 | 183.9432 | 1.7845 | 53.9978
0.5826 | 1.4542 | -9.08 | -184.4335 | 184.6877 | 1.0065 | -1.8 | -0.55226 | 2.1535 | 184.4335 | 1.8 | 54.1651
0.5844 | 1.4077 | -9.2284 | -184.9406 | 185.2026 | 1.0155 | -1.8155 | -0.55097 | 2.1703 | 184.9406 | 1.8155 | 54.3325
0.5862 | 1.3148 | -9.3677 | -185.3755 | 185.6439 | 1.0232 | -1.8258 | -0.5471 | 2.1832 | 185.3755 | 1.8258 | 54.4998
0.588 | 1.2026 | -9.4903 | -185.769 | 186.0426 | 1.0335 | -1.8348 | -0.54452 | 2.1948 | 185.769 | 1.8348 | 54.6672
0.5898 | 1.1239 | -9.6 | -186.2374 | 186.5161 | 1.0426 | -1.8477 | -0.54194 | 2.2103 | 186.2374 | 1.8477 | 54.8345
0.5916 | 1.0594 | -9.711 | -186.8155 | 187.0981 | 1.0529 | -1.8619 | -0.54065 | 2.2271 | 186.8155 | 1.8619 | 55.0019
0.5934 | 1.1097 | -9.8787 | -187.3987 | 187.6903 | 1.0606 | -1.8826 | -0.53935 | 2.249 | 187.3987 | 1.8826 | 55.1692
0.5952 | 1.1523 | -10.0632 | -187.7535 | 188.0555 | 1.0671 | -1.9006 | -0.53677 | 2.2671 | 187.7535 | 1.9006 | 55.3366
0.597 | 1.1858 | -10.2594 | -187.9329 | 188.2452 | 1.0735 | -1.9187 | -0.53419 | 2.2839 | 187.9329 | 1.9187 | 55.5039
0.5988 | 1.1845 | -10.4684 | -188.1226 | 188.4465 | 1.0787 | -1.9381 | -0.53032 | 2.3032 | 188.1226 | 1.9381 | 55.6713
0.6006 | 1.1781 | -10.7006 | -188.3316 | 188.6684 | 1.0852 | -1.96 | -0.52645 | 2.3239 | 188.3316 | 1.96 | 55.8386
0.6023 | 1.151 | -10.8877 | -188.6955 | 189.0413 | 1.0929 | -1.9819 | -0.52258 | 2.3445 | 188.6955 | 1.9819 | 55.9967
0.6041 | 1.1071 | -11.0232 | -189.2774 | 189.6297 | 1.1019 | -2.0026 | -0.52 | 2.3665 | 189.2774 | 2.0026 | 56.164
0.6059 | 1.0632 | -11.1226 | -189.9252 | 190.2813 | 1.111 | -2.0219 | -0.51742 | 2.3871 | 189.9252 | 2.0219 | 56.3313
0.6077 | 1.0516 | -11.2439 | -190.6 | 190.9626 | 1.1226 | -2.0465 | -0.51484 | 2.4116 | 190.6 | 2.0465 | 56.4987
0.6095 | 1.0632 | -11.4 | -191.2658 | 191.6361 | 1.1329 | -2.071 | -0.51226 | 2.4387 | 191.2658 | 2.071 | 56.666
0.6113 | 1.009 | -11.5548 | -191.9574 | 192.3355 | 1.1432 | -2.0955 | -0.50839 | 2.4645 | 191.9574 | 2.0955 | 56.8334
0.6131 | 0.90968 | -11.6929 | -192.7058 | 193.0903 | 1.1548 | -2.12 | -0.50323 | 2.4903 | 192.7058 | 2.12 | 57.0007
0.6149 | 0.81806 | -11.8039 | -193.5832 | 193.9729 | 1.1652 | -2.1445 | -0.50065 | 2.5161 | 193.5832 | 2.1445 | 57.1681
0.6167 | 0.76 | -11.9419 | -194.5329 | 194.9303 | 1.1729 | -2.1703 | -0.49677 | 2.5419 | 194.5329 | 2.1703 | 57.3354
0.6185 | 0.74065 | -12.1058 | -195.6413 | 196.0477 | 1.1819 | -2.2 | -0.49548 | 2.5716 | 195.6413 | 2.2 | 57.5028
0.6203 | 0.75097 | -12.2994 | -196.7845 | 197.2026 | 1.1897 | -2.2271 | -0.49419 | 2.6 | 196.7845 | 2.2271 | 57.6701
0.6221 | 0.72 | -12.4748 | -197.7987 | 198.2258 | 1.1974 | -2.249 | -0.49419 | 2.6219 | 197.7987 | 2.249 | 57.8375
0.6239 | 0.73548 | -12.631 | -198.6697 | 199.1045 | 1.2077 | -2.271 | -0.4929 | 2.6465 | 198.6697 | 2.271 | 58.0048
0.6256 | 0.75871 | -12.8194 | -199.4968 | 199.9406 | 1.2155 | -2.2942 | -0.49161 | 2.6697 | 199.4968 | 2.2942 | 58.1629
0.6274 | 0.78323 | -13.0077 | -200.3226 | 200.7781 | 1.2245 | -2.3161 | -0.48903 | 2.6929 | 200.3226 | 2.3161 | 58.3302
0.6292 | 0.7329 | -13.1742 | -201.3368 | 201.8013 | 1.2323 | -2.3394 | -0.48774 | 2.7174 | 201.3368 | 2.3394 | 58.4976
0.631 | 0.68 | -13.3187 | -202.2761 | 202.7484 | 1.24 | -2.3639 | -0.48516 | 2.7419 | 202.2761 | 2.3639 | 58.6649
0.6328 | 0.64 | -13.4439 | -203.3497 | 203.8284 | 1.249 | -2.391 | -0.48258 | 2.769 | 203.3497 | 2.391 | 58.8323
0.6346 | 0.59097 | -13.5716 | -204.3626 | 204.8477 | 1.2581 | -2.4155 | -0.48129 | 2.7948 | 204.3626 | 2.4155 | 58.9996
0.6364 | 0.54194 | -13.6929 | -205.4142 | 205.9045 | 1.2671 | -2.44 | -0.48 | 2.8206 | 205.4142 | 2.44 | 59.167
0.6382 | 0.52645 | -13.7639 | -206.6671 | 207.1587 | 1.2774 | -2.4658 | -0.48 | 2.8477 | 206.6671 | 2.4658 | 59.3343
0.64 | 0.51097 | -13.8658 | -207.8903 | 208.3845 | 1.289 | -2.4877 | -0.48129 | 2.871 | 207.8903 | 2.4877 | 59.5017
0.6418 | 0.51355 | -13.9432 | -209.2297 | 209.7252 | 1.2994 | -2.5135 | -0.48129 | 2.8994 | 209.2297 | 2.5135 | 59.669
0.6436 | 0.49935 | -14.0632 | -210.4194 | 210.9174 | 1.311 | -2.5355 | -0.48258 | 2.9226 | 210.4194 | 2.5355 | 59.8364
0.6454 | 0.5071 | -14.2116 | -211.5858 | 212.0903 | 1.3213 | -2.5561 | -0.48387 | 2.9458 | 211.5858 | 2.5561 | 60.0037
0.6472 | 0.52645 | -14.351 | -212.7445 | 213.2542 | 1.3329 | -2.5781 | -0.48645 | 2.969 | 212.7445 | 2.5781 | 60.1711
0.649 | 0.56258 | -14.4606 | -213.9071 | 214.4206 | 1.3445 | -2.6013 | -0.48774 | 2.9935 | 213.9071 | 2.6013 | 60.3384
0.6507 | 0.60387 | -14.5806 | -215.0452 | 215.5626 | 1.3561 | -2.6245 | -0.48903 | 3.0181 | 215.0452 | 2.6245 | 60.4965
0.6525 | 0.65032 | -14.7135 | -216.111 | 216.6335 | 1.3665 | -2.6465 | -0.49032 | 3.0426 | 216.111 | 2.6465 | 60.6638
0.6543 | 0.72387 | -14.8968 | -216.9871 | 217.5174 | 1.3755 | -2.6723 | -0.49032 | 3.0671 | 216.9871 | 2.6723 | 60.8312
0.6561 | 0.74968 | -15.071 | -217.8645 | 218.4039 | 1.3858 | -2.6942 | -0.49032 | 3.0903 | 217.8645 | 2.6942 | 60.9985
0.6579 | 0.76645 | -15.1858 | -218.8129 | 219.3548 | 1.3974 | -2.7148 | -0.49161 | 3.1135 | 218.8129 | 2.7148 | 61.1659
0.6597 | 0.78323 | -15.2619 | -219.8052 | 220.3497 | 1.4103 | -2.7368 | -0.4929 | 3.1381 | 219.8052 | 2.7368 | 61.3332
0.6615 | 0.78968 | -15.3252 | -220.8232 | 221.369 | 1.4245 | -2.7574 | -0.4929 | 3.1626 | 220.8232 | 2.7574 | 61.5006
0.6633 | 0.79613 | -15.3858 | -221.831 | 222.3781 | 1.4374 | -2.7768 | -0.49419 | 3.1858 | 221.831 | 2.7768 | 61.6679
0.6651 | 0.79355 | -15.4568 | -222.7561 | 223.3058 | 1.4516 | -2.7948 | -0.49548 | 3.2077 | 222.7561 | 2.7948 | 61.8353
0.6669 | 0.77935 | -15.5497 | -223.4181 | 223.9716 | 1.4645 | -2.8168 | -0.49419 | 3.2297 | 223.4181 | 2.8168 | 62.0026
0.6687 | 0.75613 | -15.6477 | -223.96 | 224.5174 | 1.4761 | -2.84 | -0.49032 | 3.2542 | 223.96 | 2.84 | 62.17
0.6705 | 0.70065 | -15.6994 | -224.5523 | 225.111 | 1.4903 | -2.8619 | -0.48645 | 3.2774 | 224.5523 | 2.8619 | 62.3373
0.6723 | 0.65032 | -15.7368 | -225.1858 | 225.7445 | 1.5032 | -2.8813 | -0.48516 | 3.2994 | 225.1858 | 2.8813 | 62.5046
0.6741 | 0.60258 | -15.769 | -225.8645 | 226.4245 | 1.5148 | -2.8994 | -0.48387 | 3.32 | 225.8645 | 2.8994 | 62.672
0.6758 | 0.5471 | -15.8129 | -226.6929 | 227.2529 | 1.529 | -2.9161 | -0.48516 | 3.3406 | 226.6929 | 2.9161 | 62.83
0.6776 | 0.48 | -15.8671 | -227.6568 | 228.2181 | 1.5458 | -2.9316 | -0.48645 | 3.3626 | 227.6568 | 2.9316 | 62.9974
0.6794 | 0.42968 | -15.9458 | -228.5923 | 229.1574 | 1.5639 | -2.9445 | -0.49032 | 3.3832 | 228.5923 | 2.9445 | 63.1647
0.6812 | 0.39484 | -16.0361 | -229.1342 | 229.7032 | 1.5768 | -2.9548 | -0.49032 | 3.3987 | 229.1342 | 2.9548 | 63.3321
0.683 | 0.36 | -16.1045 | -229.6671 | 230.2387 | 1.5948 | -2.9665 | -0.49161 | 3.4181 | 229.6671 | 2.9665 | 63.4994
0.6848 | 0.35097 | -16.1406 | -230 | 230.5742 | 1.6116 | -2.9755 | -0.4929 | 3.4323 | 230 | 2.9755 | 63.6668
0.6866 | 0.34581 | -16.1845 | -230.3548 | 230.9316 | 1.6258 | -2.9806 | -0.49419 | 3.4452 | 230.3548 | 2.9806 | 63.8341
0.6884 | 0.31613 | -16.2529 | -230.6826 | 231.2619 | 1.6297 | -2.9832 | -0.49677 | 3.4503 | 230.6826 | 2.9832 | 64.0015
0.6902 | 0.26839 | -16.3381 | -231.0387 | 231.6232 | 1.631 | -2.9845 | -0.49806 | 3.4516 | 231.0387 | 2.9845 | 64.1688
0.692 | 0.21806 | -16.4206 | -231.5316 | 232.12 | 1.6361 | -2.9871 | -0.50065 | 3.4568 | 231.5316 | 2.9871 | 64.3362
0.6938 | 0.18194 | -16.4813 | -232.1406 | 232.7329 | 1.6503 | -2.9897 | -0.50452 | 3.4671 | 232.1406 | 2.9897 | 64.5035
0.6956 | 0.1471 | -16.52 | -232.8013 | 233.3935 | 1.671 | -2.9948 | -0.5071 | 3.4826 | 232.8013 | 2.9948 | 64.6709
0.6974 | 0.12 | -16.5445 | -233.4477 | 234.0413 | 1.6903 | -2.9987 | -0.50968 | 3.4955 | 233.4477 | 2.9987 | 64.8382
0.6992 | 0.087742 | -16.5548 | -233.9987 | 234.591 | 1.7097 | -3 | -0.51355 | 3.5058 | 233.9987 | 3 | 65.0056
0.7009 | -0.023226 | -16.5084 | -234.3523 | 234.9406 | 1.7381 | -2.9897 | -0.51484 | 3.511 | 234.3523 | 2.9897 | 65.1636
0.7027 | -0.12903 | -16.44 | -234.6245 | 235.2077 | 1.7626 | -2.9768 | -0.51871 | 3.5135 | 234.6245 | 2.9768 | 65.331
0.7045 | -0.21677 | -16.3626 | -234.8735 | 235.4503 | 1.7819 | -2.9639 | -0.52258 | 3.5135 | 234.8735 | 2.9639 | 65.4983
0.7063 | -0.15613 | -16.3497 | -234.9587 | 235.5329 | 1.8026 | -2.9548 | -0.52645 | 3.5187 | 234.9587 | 2.9548 | 65.6657
0.7081 | -0.069677 | -16.3781 | -234.991 | 235.5677 | 1.8181 | -2.9471 | -0.53032 | 3.5226 | 234.991 | 2.9471 | 65.833
0.7099 | -0.18839 | -16.4194 | -235.1768 | 235.7561 | 1.8258 | -2.9342 | -0.53548 | 3.5174 | 235.1768 | 2.9342 | 66.0004
0.7117 | -0.33548 | -16.4542 | -235.4594 | 236.0413 | 1.8348 | -2.92 | -0.54065 | 3.5123 | 235.4594 | 2.92 | 66.1677
0.7135 | -0.52903 | -16.4787 | -235.9613 | 236.5445 | 1.8477 | -2.9006 | -0.5471 | 3.5058 | 235.9613 | 2.9006 | 66.3351
0.7153 | -0.68903 | -16.5342 | -236.6387 | 237.2245 | 1.8671 | -2.8826 | -0.55226 | 3.5032 | 236.6387 | 2.8826 | 66.5024
0.7171 | -0.85419 | -16.6374 | -237.3187 | 237.911 | 1.8839 | -2.8632 | -0.55742 | 3.4994 | 237.3187 | 2.8632 | 66.6698
0.7189 | -0.90452 | -16.6671 | -238.2065 | 238.7974 | 1.9006 | -2.8606 | -0.56258 | 3.5084 | 238.2065 | 2.8606 | 66.8371
0.7207 | -0.92903 | -16.729 | -239.1884 | 239.7832 | 1.9148 | -2.8645 | -0.56645 | 3.5187 | 239.1884 | 2.8645 | 67.0045
0.7225 | -0.96129 | -16.7484 | -240.2361 | 240.8284 | 1.9394 | -2.871 | -0.56903 | 3.5381 | 240.2361 | 2.871 | 67.1718
0.7242 | -1.0013 | -16.7561 | -241.2994 | 241.8903 | 1.9652 | -2.8761 | -0.57032 | 3.5574 | 241.2994 | 2.8761 | 67.3299
0.726 | -1.089 | -16.8568 | -242.3948 | 242.991 | 1.9832 | -2.8748 | -0.57419 | 3.5665 | 242.3948 | 2.8748 | 67.4972
0.7278 | -1.2129 | -17.0181 | -243.4761 | 244.0813 | 2 | -2.8632 | -0.57806 | 3.569 | 243.4761 | 2.8632 | 67.6646
0.7296 | -1.3148 | -17.0813 | -244.7406 | 245.3497 | 2.0168 | -2.8529 | -0.58452 | 3.5716 | 244.7406 | 2.8529 | 67.8319
0.7314 | -1.3458 | -17.1884 | -245.8735 | 246.4877 | 2.0374 | -2.849 | -0.59097 | 3.5806 | 245.8735 | 2.849 | 67.9993
0.7332 | -1.3806 | -17.3277 | -246.8748 | 247.4968 | 2.0606 | -2.8439 | -0.59613 | 3.5897 | 246.8748 | 2.8439 | 68.1666
0.735 | -1.4787 | -17.5084 | -247.6761 | 248.3084 | 2.0813 | -2.8348 | -0.6 | 3.5961 | 247.6761 | 2.8348 | 68.334
0.7368 | -1.5819 | -17.6697 | -248.5316 | 249.1742 | 2.1032 | -2.8284 | -0.60258 | 3.6026 | 248.5316 | 2.8284 | 68.5013
0.7386 | -1.7148 | -17.7587 | -249.5935 | 250.2413 | 2.1277 | -2.8181 | -0.60516 | 3.6103 | 249.5935 | 2.8181 | 68.6687
0.7404 | -1.9123 | -17.9019 | -250.5432 | 251.2013 | 2.151 | -2.7961 | -0.61161 | 3.609 | 250.5432 | 2.7961 | 68.836
0.7422 | -1.9961 | -18.0155 | -251.4903 | 252.1548 | 2.1716 | -2.7819 | -0.61677 | 3.6116 | 251.4903 | 2.7819 | 69.0033
0.744 | -2.0348 | -18.0735 | -252.4555 | 253.1239 | 2.191 | -2.7742 | -0.62065 | 3.6168 | 252.4555 | 2.7742 | 69.1707
0.7458 | -2.1523 | -18.1471 | -253.3561 | 254.0284 | 2.2206 | -2.76 | -0.62839 | 3.6219 | 253.3561 | 2.76 | 69.338
0.7476 | -2.2245 | -18.2619 | -254.2335 | 254.9135 | 2.2452 | -2.7458 | -0.63226 | 3.6258 | 254.2335 | 2.7458 | 69.5054
0.7493 | -2.2877 | -18.3639 | -254.9548 | 255.6413 | 2.2555 | -2.7355 | -0.63355 | 3.6245 | 254.9548 | 2.7355 | 69.6634
0.7511 | -2.3794 | -18.4671 | -255.7729 | 256.4645 | 2.2658 | -2.7252 | -0.63613 | 3.6219 | 255.7729 | 2.7252 | 69.8308
0.7529 | -2.631 | -18.5419 | -256.8981 | 257.5935 | 2.2877 | -2.6994 | -0.64645 | 3.6155 | 256.8981 | 2.6994 | 69.9981
0.7547 | -2.9277 | -18.6026 | -258.0774 | 258.7755 | 2.3135 | -2.6684 | -0.65935 | 3.6065 | 258.0774 | 2.6684 | 70.1655
0.7565 | -3.1432 | -18.6658 | -259.0839 | 259.7858 | 2.3265 | -2.64 | -0.66839 | 3.5948 | 259.0839 | 2.64 | 70.3328
0.7583 | -3.3045 | -18.7406 | -259.9665 | 260.6723 | 2.329 | -2.6142 | -0.67742 | 3.5781 | 259.9665 | 2.6142 | 70.5002
0.7601 | -3.4787 | -18.9148 | -260.791 | 261.5071 | 2.3213 | -2.5897 | -0.68387 | 3.5561 | 260.791 | 2.5897 | 70.6675
0.7619 | -3.7135 | -19.151 | -261.6568 | 262.3884 | 2.32 | -2.5574 | -0.6929 | 3.5342 | 261.6568 | 2.5574 | 70.8349
0.7637 | -3.9523 | -19.369 | -262.3987 | 263.1458 | 2.3252 | -2.5213 | -0.70323 | 3.5135 | 262.3987 | 2.5213 | 71.0022
0.7655 | -4.1987 | -19.5665 | -263.0103 | 263.7729 | 2.3368 | -2.48 | -0.71484 | 3.4929 | 263.0103 | 2.48 | 71.1696
0.7673 | -4.4568 | -19.6968 | -263.4955 | 264.271 | 2.3535 | -2.431 | -0.72645 | 3.4723 | 263.4955 | 2.431 | 71.3369
0.7691 | -4.729 | -19.7613 | -263.8581 | 264.6413 | 2.3742 | -2.3755 | -0.74065 | 3.4477 | 263.8581 | 2.3755 | 71.5043
0.7709 | -4.9806 | -19.8142 | -264.151 | 264.9419 | 2.3935 | -2.3213 | -0.75355 | 3.4258 | 264.151 | 2.3213 | 71.6716
0.7727 | -5.1548 | -19.8258 | -264.1626 | 264.9574 | 2.4103 | -2.2697 | -0.76129 | 3.4052 | 264.1626 | 2.2697 | 71.839
0.7744 | -5.3019 | -19.8542 | -263.9755 | 264.7755 | 2.4284 | -2.2194 | -0.76774 | 3.3858 | 263.9755 | 2.2194 | 71.997
0.7762 | -5.4594 | -19.8645 | -263.6297 | 264.4348 | 2.4439 | -2.1613 | -0.77548 | 3.3613 | 263.6297 | 2.1613 | 72.1644
0.778 | -5.5948 | -19.7806 | -263.0606 | 263.8632 | 2.4632 | -2.1084 | -0.77935 | 3.3419 | 263.0606 | 2.1084 | 72.3317
0.7798 | -5.7355 | -19.6503 | -262.28 | 263.0787 | 2.48 | -2.0542 | -0.78194 | 3.3226 | 262.28 | 2.0542 | 72.4991
0.7816 | -5.8723 | -19.5097 | -261.311 | 262.1058 | 2.4968 | -1.9987 | -0.78323 | 3.3006 | 261.311 | 1.9987 | 72.6664
0.7834 | -5.9987 | -19.3368 | -260.0465 | 260.8348 | 2.5123 | -1.9419 | -0.78323 | 3.2787 | 260.0465 | 1.9419 | 72.8338
0.7852 | -6.1174 | -19.1368 | -258.5935 | 259.3742 | 2.5277 | -1.8877 | -0.78194 | 3.2581 | 258.5935 | 1.8877 | 73.0011
0.787 | -6.2129 | -18.9071 | -256.9148 | 257.6865 | 2.5406 | -1.8335 | -0.77806 | 3.2374 | 256.9148 | 1.8335 | 73.1685
0.7888 | -6.3123 | -18.689 | -254.9987 | 255.7626 | 2.5523 | -1.7819 | -0.7729 | 3.2155 | 254.9987 | 1.7819 | 73.3358
0.7906 | -6.44 | -18.5135 | -252.8632 | 253.6232 | 2.5561 | -1.7316 | -0.76516 | 3.1897 | 252.8632 | 1.7316 | 73.5032
0.7924 | -6.5729 | -18.329 | -250.6542 | 251.4116 | 2.5587 | -1.6826 | -0.75742 | 3.1639 | 250.6542 | 1.6826 | 73.6705
0.7942 | -6.7097 | -18.0981 | -248.2929 | 249.0439 | 2.5626 | -1.6348 | -0.74839 | 3.1406 | 248.2929 | 1.6348 | 73.8379
0.796 | -6.8142 | -17.7845 | -245.6374 | 246.3768 | 2.5652 | -1.591 | -0.73806 | 3.1187 | 245.6374 | 1.591 | 74.0052
0.7977 | -6.889 | -17.4194 | -242.7858 | 243.5097 | 2.5677 | -1.5484 | -0.72903 | 3.0981 | 242.7858 | 1.5484 | 74.1633
0.7995 | -6.951 | -17.0271 | -239.6968 | 240.4039 | 2.569 | -1.5058 | -0.71871 | 3.0761 | 239.6968 | 1.5058 | 74.3306
0.8013 | -6.9948 | -16.609 | -236.2477 | 236.9368 | 2.5665 | -1.4671 | -0.7071 | 3.0555 | 236.2477 | 1.4671 | 74.498
0.8031 | -7.071 | -16.1884 | -232.4916 | 233.1652 | 2.5587 | -1.4258 | -0.69677 | 3.0284 | 232.4916 | 1.4258 | 74.6653
0.8049 | -7.1523 | -15.7574 | -228.5948 | 229.2516 | 2.5497 | -1.3871 | -0.68516 | 2.9987 | 228.5948 | 1.3871 | 74.8327
0.8067 | -7.2477 | -15.2916 | -224.4955 | 225.1368 | 2.5368 | -1.3484 | -0.67355 | 2.9703 | 224.4955 | 1.3484 | 75
0.8085 | -7.3355 | -14.8039 | -220.1987 | 220.8219 | 2.5226 | -1.3123 | -0.66065 | 2.9394 | 220.1987 | 1.3123 | 75.1673
0.8103 | -7.4271 | -14.2865 | -215.6671 | 216.2723 | 2.5045 | -1.2735 | -0.64903 | 2.9058 | 215.6671 | 1.2735 | 75.3347
0.8121 | -7.4994 | -13.7432 | -210.9819 | 211.5677 | 2.4852 | -1.2361 | -0.63742 | 2.8697 | 210.9819 | 1.2361 | 75.502
0.8139 | -7.52 | -13.129 | -206.0348 | 206.5948 | 2.4606 | -1.1987 | -0.62452 | 2.831 | 206.0348 | 1.1987 | 75.6694
0.8157 | -7.4761 | -12.4258 | -200.7574 | 201.2865 | 2.4323 | -1.1626 | -0.61032 | 2.7897 | 200.7574 | 1.1626 | 75.8367
0.8175 | -7.4103 | -11.6929 | -195.3923 | 195.889 | 2.4013 | -1.1277 | -0.59742 | 2.7458 | 195.3923 | 1.1277 | 76.0041
0.8193 | -7.3316 | -10.9832 | -190.0129 | 190.48 | 2.3639 | -1.0916 | -0.58323 | 2.6968 | 190.0129 | 1.0916 | 76.1714
0.8211 | -7.249 | -10.28 | -184.6632 | 185.0981 | 2.3213 | -1.0555 | -0.56903 | 2.6413 | 184.6632 | 1.0555 | 76.3388
0.8228 | -7.1355 | -9.5174 | -179.1019 | 179.5045 | 2.2761 | -1.0155 | -0.55613 | 2.5819 | 179.1019 | 1.0155 | 76.4968
0.8246 | -7.009 | -8.7161 | -173.3148 | 173.6826 | 2.2258 | -0.97419 | -0.54065 | 2.5187 | 173.3148 | 0.97419 | 76.6642
0.8264 | -6.8774 | -7.8968 | -167.3265 | 167.6632 | 2.1703 | -0.93161 | -0.52516 | 2.4503 | 167.3265 | 0.93161 | 76.8315
0.8282 | -6.7187 | -7.0439 | -161.2206 | 161.5252 | 2.1097 | -0.88516 | -0.50839 | 2.3768 | 161.2206 | 0.88516 | 76.9989
0.83 | -6.5626 | -6.2026 | -155.2232 | 155.4929 | 2.049 | -0.83742 | -0.4929 | 2.3006 | 155.2232 | 0.83742 | 77.1662
0.8318 | -6.3948 | -5.3974 | -149.0903 | 149.3355 | 1.9806 | -0.78968 | -0.47613 | 2.2181 | 149.0903 | 0.78968 | 77.3336
0.8336 | -6.2206 | -4.5897 | -142.9006 | 143.1213 | 1.9084 | -0.73806 | -0.45806 | 2.1316 | 142.9006 | 0.73806 | 77.5009
0.8354 | -6.0452 | -3.7948 | -136.809 | 137.0052 | 1.8348 | -0.68903 | -0.44 | 2.0439 | 136.809 | 0.68903 | 77.6683
0.8372 | -5.8735 | -3.0297 | -130.8568 | 131.0348 | 1.76 | -0.63742 | -0.42323 | 1.9548 | 130.8568 | 0.63742 | 77.8356
0.839 | -5.6877 | -2.3006 | -124.9097 | 125.0735 | 1.6813 | -0.5871 | -0.40516 | 1.8619 | 124.9097 | 0.5871 | 78.003
0.8408 | -5.4916 | -1.591 | -118.9665 | 119.1174 | 1.6013 | -0.53677 | -0.38581 | 1.769 | 118.9665 | 0.53677 | 78.1703
0.8426 | -5.2929 | -0.87613 | -112.9135 | 113.0529 | 1.52 | -0.48 | -0.36774 | 1.6723 | 112.9135 | 0.48 | 78.3377
0.8444 | -5.0723 | -0.22968 | -106.9716 | 107.1058 | 1.4361 | -0.42581 | -0.35097 | 1.5742 | 106.9716 | 0.42581 | 78.505
0.8462 | -4.8594 | 0.35613 | -101.2413 | 101.3729 | 1.351 | -0.37548 | -0.3329 | 1.4774 | 101.2413 | 0.37548 | 78.6724
0.8479 | -4.6826 | 0.8671 | -95.8065 | 95.9368 | 1.2671 | -0.32387 | -0.31742 | 1.3806 | 95.8065 | 0.32387 | 78.8304
0.8497 | -4.5303 | 1.3071 | -90.6426 | 90.7742 | 1.1832 | -0.27097 | -0.30065 | 1.2852 | 90.6426 | 0.27097 | 78.9978
0.8515 | -4.3884 | 1.64 | -85.9303 | 86.0671 | 1.1019 | -0.22581 | -0.28645 | 1.1948 | 85.9303 | 0.22581 | 79.1651
0.8533 | -4.2568 | 1.9226 | -81.5006 | 81.6439 | 1.0219 | -0.18323 | -0.27226 | 1.1084 | 81.5006 | 0.18323 | 79.3325
0.8551 | -4.1419 | 2.1497 | -77.4348 | 77.5845 | 0.9471 | -0.14323 | -0.25935 | 1.0284 | 77.4348 | 0.14323 | 79.4998
0.8569 | -4.04 | 2.2968 | -73.6206 | 73.7768 | 0.87484 | -0.10581 | -0.24774 | 0.95226 | 73.6206 | 0.10581 | 79.6672
0.8587 | -3.9458 | 2.4323 | -69.9484 | 70.111 | 0.80516 | -0.070968 | -0.23613 | 0.87871 | 69.9484 | 0.070968 | 79.8345
0.8605 | -3.88 | 2.5574 | -66.6168 | 66.7897 | 0.73935 | -0.03871 | -0.22581 | 0.81161 | 66.6168 | 0.03871 | 80.0019
0.8623 | -3.8477 | 2.671 | -63.7006 | 63.8839 | 0.68129 | -0.010323 | -0.21548 | 0.75097 | 63.7006 | 0.010323 | 80.1692
0.8641 | -3.8335 | 2.7368 | -61.1019 | 61.2955 | 0.62968 | 0.010323 | -0.20774 | 0.69935 | 61.1019 | -0.010323 | 80.3366
0.8659 | -3.8335 | 2.7652 | -58.7239 | 58.9252 | 0.58065 | 0.030968 | -0.2 | 0.65161 | 58.7239 | -0.030968 | 80.5039
0.8677 | -3.8452 | 2.7781 | -56.6568 | 56.8658 | 0.5329 | 0.055484 | -0.19355 | 0.60387 | 56.6568 | -0.055484 | 80.6713
0.8695 | -3.8452 | 2.7935 | -54.7652 | 54.9806 | 0.48903 | 0.07871 | -0.1871 | 0.56129 | 54.7652 | -0.07871 | 80.8386
0.8712 | -3.8542 | 2.7897 | -53.0994 | 53.32 | 0.4529 | 0.095484 | -0.18065 | 0.52903 | 53.0994 | -0.095484 | 80.9967
0.873 | -3.88 | 2.7897 | -51.689 | 51.9174 | 0.42323 | 0.11097 | -0.17677 | 0.50323 | 51.689 | -0.11097 | 81.164
0.8748 | -3.9226 | 2.8452 | -50.6065 | 50.8452 | 0.40129 | 0.12 | -0.17419 | 0.48645 | 50.6065 | -0.12 | 81.3313
0.8766 | -3.9277 | 2.9484 | -49.5239 | 49.7742 | 0.39097 | 0.12387 | -0.1729 | 0.48129 | 49.5239 | -0.12387 | 81.4987
0.8784 | -3.88 | 2.9858 | -48.3006 | 48.5587 | 0.38323 | 0.12774 | -0.16774 | 0.47742 | 48.3006 | -0.12774 | 81.666
0.8802 | -3.8542 | 2.9587 | -47.0671 | 47.329 | 0.37032 | 0.13677 | -0.16 | 0.46839 | 47.0671 | -0.13677 | 81.8334
0.882 | -3.8619 | 2.9032 | -45.969 | 46.2335 | 0.35097 | 0.15097 | -0.15355 | 0.45419 | 45.969 | -0.15097 | 82.0007
0.8838 | -3.8568 | 2.8439 | -45.1342 | 45.3987 | 0.3329 | 0.16387 | -0.1471 | 0.44387 | 45.1342 | -0.16387 | 82.1681
0.8856 | -3.8361 | 2.8039 | -44.4477 | 44.7097 | 0.32 | 0.17677 | -0.14065 | 0.43355 | 44.4477 | -0.17677 | 82.3354
0.8874 | -3.831 | 2.8232 | -43.8955 | 44.16 | 0.30839 | 0.19226 | -0.13677 | 0.42323 | 43.8955 | -0.19226 | 82.5028
0.8892 | -3.8452 | 2.8039 | -43.4323 | 43.6981 | 0.29677 | 0.20903 | -0.13161 | 0.41806 | 43.4323 | -0.20903 | 82.6701
0.891 | -3.8684 | 2.76 | -43.0413 | 43.3084 | 0.28516 | 0.22581 | -0.12645 | 0.41419 | 43.0413 | -0.22581 | 82.8375
0.8928 | -3.8981 | 2.6723 | -42.751 | 43.0168 | 0.27226 | 0.24129 | -0.12129 | 0.41032 | 42.751 | -0.24129 | 83.0048
0.8946 | -3.9226 | 2.6542 | -42.6594 | 42.9265 | 0.26323 | 0.25419 | -0.11871 | 0.41032 | 42.6594 | -0.25419 | 83.1722
0.8963 | -3.9265 | 2.6994 | -42.609 | 42.88 | 0.25935 | 0.26839 | -0.11484 | 0.41677 | 42.609 | -0.26839 | 83.3302
0.8981 | -3.8658 | 2.6865 | -42.249 | 42.5135 | 0.25161 | 0.28387 | -0.11097 | 0.42194 | 42.249 | -0.28387 | 83.4976
0.8999 | -3.8671 | 2.6632 | -42.1742 | 42.4387 | 0.24387 | 0.29935 | -0.1071 | 0.42323 | 42.1742 | -0.29935 | 83.6649
0.9017 | -3.849 | 2.6439 | -42.2387 | 42.4994 | 0.23484 | 0.32 | -0.10194 | 0.43097 | 42.2387 | -0.32 | 83.8323
0.9035 | -3.8529 | 2.6594 | -42.5639 | 42.8258 | 0.22968 | 0.34323 | -0.094194 | 0.44387 | 42.5639 | -0.34323 | 83.9996
0.9053 | -3.8529 | 2.6916 | -42.8348 | 43.0955 | 0.22452 | 0.36516 | -0.087742 | 0.45677 | 42.8348 | -0.36516 | 84.167
0.9071 | -3.8168 | 2.7419 | -42.7871 | 43.049 | 0.21935 | 0.37935 | -0.082581 | 0.46323 | 42.7871 | -0.37935 | 84.3343
0.9089 | -3.7948 | 2.7626 | -42.6981 | 42.96 | 0.2129 | 0.3871 | -0.076129 | 0.4671 | 42.6981 | -0.3871 | 84.5017
0.9107 | -3.7871 | 2.7716 | -42.5819 | 42.8452 | 0.20387 | 0.39226 | -0.070968 | 0.4671 | 42.5819 | -0.39226 | 84.669
0.9125 | -3.7626 | 2.8335 | -42.3148 | 42.5819 | 0.19871 | 0.39613 | -0.065806 | 0.4671 | 42.3148 | -0.39613 | 84.8364
0.9143 | -3.7303 | 2.9187 | -41.9471 | 42.2194 | 0.19226 | 0.40387 | -0.060645 | 0.46968 | 41.9471 | -0.40387 | 85.0037
0.9161 | -3.6735 | 2.9948 | -41.5613 | 41.8361 | 0.1871 | 0.41032 | -0.055484 | 0.47097 | 41.5613 | -0.41032 | 85.1711
0.9179 | -3.6077 | 3.1019 | -41.111 | 41.3923 | 0.18323 | 0.41161 | -0.050323 | 0.46968 | 41.111 | -0.41161 | 85.3384
0.9197 | -3.5303 | 3.2232 | -40.6026 | 40.8903 | 0.17935 | 0.41161 | -0.047742 | 0.46581 | 40.6026 | -0.41161 | 85.5058
0.9214 | -3.4426 | 3.289 | -40.089 | 40.3781 | 0.17548 | 0.4129 | -0.04129 | 0.46452 | 40.089 | -0.4129 | 85.6638
0.9232 | -3.3523 | 3.3587 | -39.5316 | 39.8219 | 0.17032 | 0.41161 | -0.036129 | 0.46194 | 39.5316 | -0.41161 | 85.8312
0.925 | -3.2594 | 3.4361 | -38.8594 | 39.1535 | 0.16645 | 0.40903 | -0.033548 | 0.45548 | 38.8594 | -0.40903 | 85.9985
0.9268 | -3.1561 | 3.4748 | -38.0852 | 38.3806 | 0.16 | 0.40645 | -0.027097 | 0.45032 | 38.0852 | -0.40645 | 86.1659
0.9286 | -3.0477 | 3.5045 | -37.2335 | 37.5316 | 0.15355 | 0.40516 | -0.021935 | 0.44516 | 37.2335 | -0.40516 | 86.3332
0.9304 | -2.9368 | 3.5277 | -36.231 | 36.5329 | 0.14839 | 0.40645 | -0.014194 | 0.44258 | 36.231 | -0.40645 | 86.5006
0.9322 | -2.8335 | 3.5252 | -35.2206 | 35.5252 | 0.14065 | 0.41161 | -0.0051613 | 0.44258 | 35.2206 | -0.41161 | 86.6679
0.934 | -2.7445 | 3.4645 | -34.4232 | 34.7187 | 0.1329 | 0.40903 | 0.0025806 | 0.43484 | 34.4232 | -0.40903 | 86.8353
0.9358 | -2.6606 | 3.4181 | -33.5987 | 33.8903 | 0.12516 | 0.40387 | 0.0077419 | 0.4271 | 33.5987 | -0.40387 | 87.0026
0.9376 | -2.5845 | 3.3458 | -32.6942 | 32.9806 | 0.11871 | 0.39355 | 0.0090323 | 0.41419 | 32.6942 | -0.39355 | 87.17
0.9394 | -2.5058 | 3.2452 | -31.9148 | 32.191 | 0.11097 | 0.38323 | 0.011613 | 0.40258 | 31.9148 | -0.38323 | 87.3373
0.9412 | -2.4103 | 3.1484 | -31.2348 | 31.4981 | 0.10194 | 0.3729 | 0.014194 | 0.38968 | 31.2348 | -0.3729 | 87.5046
0.943 | -2.3071 | 3.0439 | -30.5935 | 30.8426 | 0.091613 | 0.36129 | 0.016774 | 0.37548 | 30.5935 | -0.36129 | 87.672
0.9447 | -2.2219 | 2.9032 | -30.0555 | 30.289 | 0.08129 | 0.35097 | 0.018065 | 0.36258 | 30.0555 | -0.35097 | 87.83
0.9465 | -2.1587 | 2.7716 | -29.6194 | 29.8374 | 0.072258 | 0.33806 | 0.018065 | 0.34968 | 29.6194 | -0.33806 | 87.9974
0.9483 | -2.111 | 2.6129 | -29.2903 | 29.4916 | 0.060645 | 0.32903 | 0.019355 | 0.33806 | 29.2903 | -0.32903 | 88.1647
0.9501 | -2.0658 | 2.5071 | -28.9948 | 29.1832 | 0.052903 | 0.31613 | 0.020645 | 0.32387 | 28.9948 | -0.31613 | 88.3321
0.9519 | -2.0387 | 2.3677 | -28.8335 | 29.009 | 0.042581 | 0.3071 | 0.021935 | 0.31355 | 28.8335 | -0.3071 | 88.4994
0.9537 | -2.0258 | 2.2645 | -28.769 | 28.9355 | 0.034839 | 0.29806 | 0.024516 | 0.30323 | 28.769 | -0.29806 | 88.6668
0.9555 | -2.0219 | 2.2065 | -28.7406 | 28.9019 | 0.027097 | 0.28774 | 0.025806 | 0.2929 | 28.7406 | -0.28774 | 88.8341
0.9573 | -1.9987 | 2.1497 | -28.7832 | 28.9381 | 0.021935 | 0.27613 | 0.027097 | 0.28 | 28.7832 | -0.27613 | 89.0015
0.9591 | -1.9974 | 2.1071 | -28.8761 | 29.0284 | 0.015484 | 0.26452 | 0.027097 | 0.26839 | 28.8761 | -0.26452 | 89.1688
0.9609 | -2.0026 | 2.0645 | -28.9794 | 29.1277 | 0.010323 | 0.25161 | 0.027097 | 0.25548 | 28.9794 | -0.25161 | 89.3362
0.9627 | -2.0245 | 2.0684 | -28.9755 | 29.1252 | 0.0064516 | 0.23742 | 0.025806 | 0.24 | 28.9755 | -0.23742 | 89.5035
0.9645 | -2.0387 | 2.1006 | -28.8929 | 29.0465 | 0.0051613 | 0.22323 | 0.023226 | 0.2271 | 28.8929 | -0.22323 | 89.6709
0.9663 | -2.0413 | 2.1535 | -28.6942 | 28.8516 | 0.0051613 | 0.20903 | 0.020645 | 0.21161 | 28.6942 | -0.20903 | 89.8382
0.9681 | -2.0039 | 2.2129 | -28.3084 | 28.471 | 0.0077419 | 0.19484 | 0.018065 | 0.19742 | 28.3084 | -0.19484 | 90.0056
0.9698 | -1.951 | 2.2619 | -27.8761 | 28.0413 | 0.0090323 | 0.17935 | 0.015484 | 0.18194 | 27.8761 | -0.17935 | 90.1636
0.9716 | -1.8994 | 2.3071 | -27.3987 | 27.5665 | 0.010323 | 0.16516 | 0.012903 | 0.16774 | 27.3987 | -0.16516 | 90.331
0.9734 | -1.8542 | 2.3445 | -26.8477 | 27.0194 | 0.011613 | 0.15097 | 0.0090323 | 0.15484 | 26.8477 | -0.15097 | 90.4983
0.9752 | -1.8168 | 2.3497 | -26.1355 | 26.3097 | 0.011613 | 0.13806 | 0.0051613 | 0.14194 | 26.1355 | -0.13806 | 90.6657
0.977 | -1.7897 | 2.3303 | -25.3045 | 25.4787 | 0.011613 | 0.12387 | 0.0012903 | 0.12774 | 25.3045 | -0.12387 | 90.833
0.9788 | -1.7471 | 2.2877 | -24.4774 | 24.6516 | 0.011613 | 0.11097 | -0.0025806 | 0.11355 | 24.4774 | -0.11097 | 91.0004
0.9806 | -1.7123 | 2.2013 | -23.8065 | 23.9755 | 0.010323 | 0.099355 | -0.0051613 | 0.10323 | 23.8065 | -0.099355 | 91.1677
0.9824 | -1.6813 | 2.0361 | -23.311 | 23.4684 | 0.0051613 | 0.090323 | -0.0090323 | 0.094194 | 23.311 | -0.090323 | 91.3351
0.9842 | -1.649 | 1.8697 | -22.8271 | 22.9703 | 0.0012903 | 0.08129 | -0.011613 | 0.085161 | 22.8271 | -0.08129 | 91.5024
0.986 | -1.6103 | 1.7123 | -22.4826 | 22.6129 | -0.0051613 | 0.072258 | -0.014194 | 0.07871 | 22.4826 | -0.072258 | 91.6698
0.9878 | -1.5703 | 1.5806 | -22.3019 | 22.4219 | -0.010323 | 0.063226 | -0.015484 | 0.073548 | 22.3019 | -0.063226 | 91.8371
0.9896 | -1.5277 | 1.48 | -22.1923 | 22.3019 | -0.014194 | 0.052903 | -0.018065 | 0.068387 | 22.1923 | -0.052903 | 92.0045
0.9914 | -1.4787 | 1.4387 | -22.1884 | 22.2929 | -0.015484 | 0.04129 | -0.020645 | 0.061935 | 22.1884 | -0.04129 | 92.1718
0.9932 | -1.4181 | 1.4426 | -22.3484 | 22.4477 | -0.018065 | 0.028387 | -0.023226 | 0.060645 | 22.3484 | -0.028387 | 92.3392
0.9949 | -1.3484 | 1.4606 | -22.5342 | 22.6323 | -0.019355 | 0.014194 | -0.027097 | 0.064516 | 22.5342 | -0.014194 | 92.4972
0.9967 | -1.2955 | 1.5187 | -22.5381 | 22.64 | -0.018065 | 0.003871 | -0.029677 | 0.068387 | 22.5381 | -0.003871 | 92.6646
0.9985 | -1.2413 | 1.5755 | -22.4942 | 22.5987 | -0.015484 | -0.0064516 | -0.033548 | 0.072258 | 22.4942 | 0.0064516 | 92.8319
1.0003 | -1.1935 | 1.6039 | -22.4194 | 22.5226 | -0.014194 | -0.014194 | -0.036129 | 0.076129 | 22.4194 | 0.014194 | 92.9993
1.0021 | -1.1587 | 1.631 | -22.2581 | 22.3626 | -0.012903 | -0.021935 | -0.04 | 0.074839 | 22.2581 | 0.021935 | 93.1666
1.0039 | -1.1368 | 1.6477 | -21.9819 | 22.0877 | -0.011613 | -0.029677 | -0.042581 | 0.076129 | 21.9819 | 0.029677 | 93.334
1.0057 | -1.1135 | 1.6026 | -21.6026 | 21.7058 | -0.012903 | -0.033548 | -0.045161 | 0.077419 | 21.6026 | 0.033548 | 93.5013
1.0075 | -1.1006 | 1.5548 | -21.1729 | 21.2735 | -0.012903 | -0.03871 | -0.047742 | 0.077419 | 21.1729 | 0.03871 | 93.6687
1.0093 | -1.0903 | 1.4645 | -20.8555 | 20.9497 | -0.016774 | -0.04129 | -0.050323 | 0.08129 | 20.8555 | 0.04129 | 93.836
1.0111 | -1.0645 | 1.3755 | -20.7213 | 20.8116 | -0.019355 | -0.045161 | -0.051613 | 0.083871 | 20.7213 | 0.045161 | 94.0033
1.0129 | -1.0348 | 1.3148 | -20.6865 | 20.7716 | -0.021935 | -0.049032 | -0.052903 | 0.085161 | 20.6865 | 0.049032 | 94.1707
1.0147 | -1.0116 | 1.2555 | -20.7252 | 20.8052 | -0.024516 | -0.052903 | -0.054194 | 0.086452 | 20.7252 | 0.052903 | 94.338
1.0165 | -0.98968 | 1.2077 | -20.8929 | 20.9677 | -0.027097 | -0.055484 | -0.054194 | 0.089032 | 20.8929 | 0.055484 | 94.5054
1.0182 | -0.95355 | 1.169 | -21.129 | 21.1974 | -0.029677 | -0.059355 | -0.055484 | 0.092903 | 21.129 | 0.059355 | 94.6634
1.02 | -0.90581 | 1.2065 | -21.4348 | 21.5045 | -0.028387 | -0.063226 | -0.055484 | 0.094194 | 21.4348 | 0.063226 | 94.8308
1.0218 | -0.85806 | 1.2723 | -21.7161 | 21.7897 | -0.024516 | -0.067097 | -0.056774 | 0.096774 | 21.7161 | 0.067097 | 94.9981
1.0236 | -0.80387 | 1.3019 | -21.871 | 21.9445 | -0.021935 | -0.069677 | -0.058065 | 0.098065 | 21.871 | 0.069677 | 95.1655
1.0254 | -0.78065 | 1.3342 | -21.8839 | 21.9587 | -0.020645 | -0.070968 | -0.059355 | 0.099355 | 21.8839 | 0.070968 | 95.3328
1.0272 | -0.76258 | 1.3342 | -21.8671 | 21.9419 | -0.018065 | -0.072258 | -0.060645 | 0.10065 | 21.8671 | 0.072258 | 95.5002
1.029 | -0.77677 | 1.3097 | -21.711 | 21.7845 | -0.018065 | -0.070968 | -0.061935 | 0.10065 | 21.711 | 0.070968 | 95.6675
1.0308 | -0.80258 | 1.2335 | -21.511 | 21.5806 | -0.019355 | -0.067097 | -0.063226 | 0.099355 | 21.511 | 0.067097 | 95.8349
1.0326 | -0.82194 | 1.1432 | -21.4619 | 21.5265 | -0.020645 | -0.064516 | -0.063226 | 0.098065 | 21.4619 | 0.064516 | 96.0022
1.0344 | -0.83871 | 1.0387 | -21.6258 | 21.6839 | -0.024516 | -0.060645 | -0.063226 | 0.095484 | 21.6258 | 0.060645 | 96.1696
1.0362 | -0.86065 | 0.94194 | -21.8684 | 21.9213 | -0.027097 | -0.056774 | -0.063226 | 0.094194 | 21.8684 | 0.056774 | 96.3369
1.038 | -0.87484 | 0.91484 | -22.2335 | 22.2839 | -0.027097 | -0.055484 | -0.063226 | 0.092903 | 22.2335 | 0.055484 | 96.5043
1.0398 | -0.88387 | 0.93032 | -22.7187 | 22.7677 | -0.025806 | -0.055484 | -0.063226 | 0.091613 | 22.7187 | 0.055484 | 96.6716
1.0416 | -0.88387 | 0.92903 | -23.1329 | 23.1819 | -0.024516 | -0.054194 | -0.063226 | 0.090323 | 23.1329 | 0.054194 | 96.839
1.0433 | -0.87871 | 0.93032 | -23.4413 | 23.489 | -0.023226 | -0.052903 | -0.063226 | 0.090323 | 23.4413 | 0.052903 | 96.997
1.0451 | -0.85548 | 0.9329 | -23.671 | 23.7174 | -0.021935 | -0.052903 | -0.064516 | 0.089032 | 23.671 | 0.052903 | 97.1644
1.0469 | -0.82194 | 0.91742 | -23.7806 | 23.8258 | -0.023226 | -0.050323 | -0.064516 | 0.089032 | 23.7806 | 0.050323 | 97.3317
1.0487 | -0.8 | 0.92774 | -23.9071 | 23.9497 | -0.021935 | -0.050323 | -0.064516 | 0.087742 | 23.9071 | 0.050323 | 97.4991
1.0505 | -0.77677 | 0.94323 | -24.0219 | 24.0645 | -0.021935 | -0.049032 | -0.064516 | 0.087742 | 24.0219 | 0.049032 | 97.6664
1.0523 | -0.74968 | 0.96903 | -24.111 | 24.1548 | -0.021935 | -0.046452 | -0.063226 | 0.085161 | 24.111 | 0.046452 | 97.8338
1.0541 | -0.72645 | 1.0116 | -24.089 | 24.1342 | -0.020645 | -0.046452 | -0.063226 | 0.083871 | 24.089 | 0.046452 | 98.0011
1.0559 | -0.70323 | 1.0103 | -23.9897 | 24.0335 | -0.021935 | -0.045161 | -0.061935 | 0.083871 | 23.9897 | 0.045161 | 98.1685
1.0577 | -0.66968 | 1.08 | -23.8735 | 23.9213 | -0.021935 | -0.045161 | -0.060645 | 0.082581 | 23.8735 | 0.045161 | 98.3358
1.0595 | -0.64 | 1.1432 | -23.6865 | 23.7368 | -0.020645 | -0.045161 | -0.060645 | 0.08129 | 23.6865 | 0.045161 | 98.5032
1.0613 | -0.61935 | 1.1548 | -23.4994 | 23.551 | -0.023226 | -0.043871 | -0.059355 | 0.08129 | 23.4994 | 0.043871 | 98.6705
1.0631 | -0.60129 | 1.1432 | -23.4942 | 23.5458 | -0.024516 | -0.043871 | -0.058065 | 0.08129 | 23.4942 | 0.043871 | 98.8379
1.0649 | -0.57419 | 1.1406 | -23.5523 | 23.6052 | -0.027097 | -0.043871 | -0.056774 | 0.08129 | 23.5523 | 0.043871 | 99.0052
1.0667 | -0.52516 | 1.2065 | -23.7548 | 23.8142 | -0.027097 | -0.046452 | -0.055484 | 0.08129 | 23.7548 | 0.046452 | 99.1726
1.0684 | -0.48258 | 1.2619 | -23.991 | 24.0542 | -0.028387 | -0.047742 | -0.054194 | 0.08129 | 23.991 | 0.047742 | 99.3306
1.0702 | -0.44516 | 1.3342 | -24.0426 | 24.111 | -0.028387 | -0.049032 | -0.052903 | 0.08129 | 24.0426 | 0.049032 | 99.498
1.072 | -0.41677 | 1.431 | -23.9006 | 23.9742 | -0.027097 | -0.049032 | -0.052903 | 0.08 | 23.9006 | 0.049032 | 99.6653
1.0738 | -0.39742 | 1.5071 | -23.7548 | 23.8323 | -0.025806 | -0.047742 | -0.050323 | 0.07871 | 23.7548 | 0.047742 | 99.8327
1.0756 | -0.35871 | 1.5974 | -23.6942 | 23.7819 | -0.025806 | -0.049032 | -0.050323 | 0.07871 | 23.6942 | 0.049032 | 100
| Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My
Minimum | -7.52 | -19.8645 | -315.9394 | 20.7716 | -0.23613 | -3 | -0.78323 | 0.060645 | 20.6865 | -0.4129
Maximum | 16.1794 | 11.3226 | -20.6865 | 316.5819 | 2.569 | 0.4129 | 0.027097 | 3.6258 | 315.9394 | 3

## Diagr_BW
Loads in %BW / %BWm
Contact forces and moments, measured telemetrically in knee implants of 8 subjects.
Averages of one individual subject.
Fx  lateral, Fy  anterior, -Fz  distal, Fres = Resultant Force
Mx around x, My around y, Mz around z, Mres = Resultant Moment
z-axis = axis of implant stem!
From Bergmann et al, PlosOne (2014) "Standardized Loads Acting in Knee Implants"
Subject K8L
Subject K8L

## Dat_AVER75
Data File: | k8l_2266_0_6_i1.akf
Curve Alignment Program: | 1.5.7
------------------------MEASURING INFOS------------------------
Diagram Title #1 | Forces and Moments at Knee Joint
Diagram Title #2 | Measured Relative to Implant
Comment #1 | Knee Joint; Gaitanalysis; Stairs up;
Comment #2 | K8L, 25 Months PO *
BodyWeight [N]: | 775
Patient(en): | K8L | [7 empty] | AVER75 loads in N (forces) and Nm (moments)
Implant Type: | Knee Joint
Varus Angle Leg [°] | 4 | [7 empty] | Subject | K8L
Axes Directions X, Y, Z: | Lateral, Anterior, Superior
Axes Origin: | Tibial Plateau at Stem Axis | [7 empty] | The individual peak value of Fres was adapted to the average peak of Fres from all subjects!
Transformation X Y Z (°): | 11.5 | -1 | 12 | [5 empty] | The AVER75 forces Fres of all subjects therefore have the same peak values.
Displacement X, Y, Z (mm): | 0 | 0 | 0 | [5 empty] | All 6 load components were then adapted by the same Multiplication Factor.
Activity Code: | KneeActivities | 1P_1_15_4_1
[9 empty] | Multiplication Factor = |  | Max(Fres, AVER75 of 8 Subjects) / |  |  | Max(Fres, Individual Subject)
[9 empty] | Max(Fres, AVER75 of 8 Subjects) = |  |  | 2285.76 | N
[9 empty] | Max(Fres, Individual Subject) = |  |  | 2453.51 | N
---------------------------DATA--------------------------------
Number of Data Sets = | 601
Number of analog channels = | 0
Time-Offset(sec) = | 0
Time | Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My | Time
[s] | [N] | [N] | [N] | [N] | [Nm] | [Nm] | [Nm] | [Nm] | [N] | (Nm] | [%Cycle]
0 | -2.4874 | 21.7256 | -160.7618 | 162.327 | -0.36334 | -1.0434 | -0.68009 | 1.3043 | 160.7618 | 1.0434 | 0
0.0018 | -2.7483 | 21.3995 | -160.6593 | 162.1872 | -0.36334 | -1.0062 | -0.68009 | 1.2763 | 160.6593 | 1.0062 | 0.16735
0.0036 | -2.8228 | 21.1945 | -160.3892 | 161.8984 | -0.37265 | -0.98753 | -0.67077 | 1.2577 | 160.3892 | 0.98753 | 0.3347
0.0054 | -2.8042 | 21.0641 | -160.1004 | 161.591 | -0.37265 | -0.97821 | -0.66146 | 1.2484 | 160.1004 | 0.97821 | 0.50205
0.0072 | -2.7204 | 21.2411 | -162.1779 | 163.6685 | -0.36334 | -0.97821 | -0.65214 | 1.2391 | 162.1779 | 0.97821 | 0.66939
0.009 | -2.7669 | 21.0828 | -164.0598 | 165.5318 | -0.37265 | -0.95958 | -0.64282 | 1.2204 | 164.0598 | 0.95958 | 0.83674
0.0108 | -2.7949 | 21.0921 | -165.5131 | 166.9944 | -0.36334 | -0.95026 | -0.63351 | 1.2111 | 165.5131 | 0.95026 | 1.0041
0.0125 | -2.7297 | 21.3995 | -166.3516 | 167.8608 | -0.3447 | -0.94094 | -0.62419 | 1.1925 | 166.3516 | 0.94094 | 1.1621
0.0143 | -2.6924 | 21.6417 | -166.3982 | 167.9354 | -0.33539 | -0.93163 | -0.62419 | 1.1739 | 166.3982 | 0.93163 | 1.3295
0.0161 | -2.8135 | 21.7069 | -165.4479 | 166.9944 | -0.32607 | -0.90368 | -0.61487 | 1.1552 | 165.4479 | 0.90368 | 1.4968
0.0179 | -3.1489 | 21.5206 | -164.0132 | 165.5504 | -0.32607 | -0.86641 | -0.61487 | 1.118 | 164.0132 | 0.86641 | 1.6642
0.0197 | -3.5216 | 21.2132 | -162.504 | 164.0225 | -0.33539 | -0.82915 | -0.60556 | 1.09 | 162.504 | 0.82915 | 1.8315
0.0215 | -3.5775 | 21.2039 | -162.327 | 163.8548 | -0.3447 | -0.81052 | -0.58693 | 1.0714 | 162.327 | 0.81052 | 1.9989
0.0233 | -3.4936 | 21.6883 | -163.2213 | 164.8144 | -0.3447 | -0.8012 | -0.57761 | 1.0621 | 163.2213 | 0.8012 | 2.1662
0.0251 | -3.4191 | 22.21 | -164.1809 | 165.8485 | -0.35402 | -0.8012 | -0.56829 | 1.0527 | 164.1809 | 0.8012 | 2.3336
0.0269 | -3.1675 | 22.7597 | -164.833 | 166.5752 | -0.35402 | -0.8012 | -0.55898 | 1.0527 | 164.833 | 0.8012 | 2.5009
0.0287 | -2.9905 | 23.3093 | -164.3952 | 166.2119 | -0.3447 | -0.81052 | -0.54966 | 1.0434 | 164.3952 | 0.81052 | 2.6683
0.0305 | -2.711 | 23.9988 | -163.7896 | 165.7088 | -0.32607 | -0.81983 | -0.54034 | 1.0434 | 163.7896 | 0.81983 | 2.8356
0.0323 | -2.4036 | 24.7627 | -162.8207 | 164.8144 | -0.33539 | -0.83847 | -0.54034 | 1.0621 | 162.8207 | 0.83847 | 3.003
0.0341 | -1.9471 | 25.7782 | -162.1406 | 164.2741 | -0.3447 | -0.86641 | -0.54034 | 1.0807 | 162.1406 | 0.86641 | 3.1703
0.0359 | -1.9005 | 25.9924 | -160.5569 | 162.7555 | -0.35402 | -0.8571 | -0.53103 | 1.0714 | 160.5569 | 0.8571 | 3.3377
0.0376 | -1.9937 | 26.039 | -158.5911 | 160.8084 | -0.38197 | -0.84778 | -0.52171 | 1.0714 | 158.5911 | 0.84778 | 3.4957
0.0394 | -2.0123 | 26.2347 | -157.3148 | 159.5787 | -0.4006 | -0.83847 | -0.5124 | 1.0714 | 157.3148 | 0.83847 | 3.6631
0.0412 | -1.7608 | 26.7284 | -158.3582 | 160.6966 | -0.41923 | -0.8571 | -0.50308 | 1.09 | 158.3582 | 0.8571 | 3.8304
0.043 | -1.5092 | 27.2595 | -160.1004 | 162.504 | -0.44718 | -0.87573 | -0.49376 | 1.1086 | 160.1004 | 0.87573 | 3.9978
0.0448 | -1.2391 | 27.8557 | -161.6841 | 164.1623 | -0.46581 | -0.90368 | -0.49376 | 1.1273 | 161.6841 | 0.90368 | 4.1651
0.0466 | -0.94094 | 28.6569 | -162.8487 | 165.4386 | -0.47513 | -0.92231 | -0.48445 | 1.1459 | 162.8487 | 0.92231 | 4.3325
0.0484 | -0.64282 | 29.4674 | -163.4077 | 166.128 | -0.48445 | -0.94094 | -0.48445 | 1.1739 | 163.4077 | 0.94094 | 4.4998
0.0502 | -0.41923 | 30.5295 | -163.035 | 165.9603 | -0.49376 | -0.96889 | -0.49376 | 1.1925 | 163.035 | 0.96889 | 4.6672
0.052 | -0.20496 | 31.2189 | -162.3735 | 165.4572 | -0.5124 | -0.97821 | -0.50308 | 1.2204 | 162.3735 | 0.97821 | 4.8345
0.0538 | 0.0093163 | 31.7033 | -161.8798 | 165.0753 | -0.53103 | -0.99684 | -0.50308 | 1.2391 | 161.8798 | 0.99684 | 5.0019
0.0556 | 0.22359 | 32.2437 | -162.4481 | 165.7274 | -0.55898 | -1.0155 | -0.50308 | 1.267 | 162.4481 | 1.0155 | 5.1692
0.0574 | 0.43787 | 32.4766 | -163.4729 | 166.7801 | -0.59624 | -1.0341 | -0.50308 | 1.295 | 163.4729 | 1.0341 | 5.3366
0.0592 | 0.66146 | 32.8585 | -165.2243 | 168.5689 | -0.64282 | -1.0527 | -0.50308 | 1.3322 | 165.2243 | 1.0527 | 5.5039
0.061 | 1.0993 | 33.7529 | -167.7397 | 171.2147 | -0.65214 | -1.09 | -0.50308 | 1.3788 | 167.7397 | 1.09 | 5.6713
0.0627 | 1.621 | 34.7218 | -170.4508 | 174.0562 | -0.67077 | -1.1273 | -0.5124 | 1.4068 | 170.4508 | 1.1273 | 5.8293
0.0645 | 1.9844 | 35.4298 | -172.1184 | 175.8263 | -0.69872 | -1.1552 | -0.53103 | 1.444 | 172.1184 | 1.1552 | 5.9967
0.0663 | 2.3104 | 36.1192 | -173.5904 | 177.3914 | -0.72667 | -1.1739 | -0.54034 | 1.4813 | 173.5904 | 1.1739 | 6.164
0.0681 | 2.6551 | 36.7807 | -175.9381 | 179.8229 | -0.76394 | -1.2018 | -0.54966 | 1.5279 | 175.9381 | 1.2018 | 6.3313
0.0699 | 3.27 | 37.4794 | -179.9441 | 183.8942 | -0.81052 | -1.2391 | -0.55898 | 1.5838 | 179.9441 | 1.2391 | 6.4987
0.0717 | 3.9222 | 38.4297 | -184.2761 | 188.3287 | -0.83847 | -1.2856 | -0.56829 | 1.6397 | 184.2761 | 1.2856 | 6.666
0.0735 | 4.5743 | 39.3799 | -188.4871 | 192.6701 | -0.86641 | -1.3322 | -0.57761 | 1.6956 | 188.4871 | 1.3322 | 6.8334
0.0753 | 5.1985 | 40.5445 | -192.2881 | 196.6388 | -0.88505 | -1.3788 | -0.59624 | 1.7421 | 192.2881 | 1.3788 | 7.0007
0.0771 | 5.655 | 41.6717 | -195.0644 | 199.6014 | -0.913 | -1.4161 | -0.61487 | 1.798 | 195.0644 | 1.4161 | 7.1681
0.0789 | 5.8972 | 42.5568 | -196.7879 | 201.4833 | -0.93163 | -1.4347 | -0.63351 | 1.826 | 196.7879 | 1.4347 | 7.3354
0.0807 | 6.0463 | 43.1624 | -198.0642 | 202.8528 | -0.95958 | -1.4533 | -0.65214 | 1.8539 | 198.0642 | 1.4533 | 7.5028
0.0825 | 6.1953 | 43.7865 | -199.5176 | 204.4179 | -0.98753 | -1.4627 | -0.66146 | 1.8819 | 199.5176 | 1.4627 | 7.6701
0.0843 | 6.5214 | 44.7741 | -202.2286 | 207.2874 | -1.0155 | -1.4813 | -0.67077 | 1.9192 | 202.2286 | 1.4813 | 7.8375
0.086 | 6.9313 | 46.0411 | -205.8806 | 211.1536 | -1.0434 | -1.5092 | -0.68009 | 1.9564 | 205.8806 | 1.5092 | 7.9955
0.0878 | 7.3226 | 47.5037 | -210.2499 | 215.7465 | -1.0527 | -1.5372 | -0.69872 | 1.9937 | 210.2499 | 1.5372 | 8.1629
0.0896 | 8.0027 | 49.3297 | -215.1596 | 220.945 | -1.0527 | -1.5838 | -0.71735 | 2.031 | 215.1596 | 1.5838 | 8.3302
0.0914 | 8.4219 | 50.8669 | -218.7091 | 224.7554 | -1.0621 | -1.5931 | -0.72667 | 2.0496 | 218.7091 | 1.5931 | 8.4976
0.0932 | 8.5337 | 52.0501 | -221.2432 | 227.5037 | -1.0807 | -1.5745 | -0.73599 | 2.0496 | 221.2432 | 1.5745 | 8.6649
0.095 | 8.6176 | 53.1494 | -223.6281 | 230.075 | -1.09 | -1.5558 | -0.7453 | 2.0403 | 223.6281 | 1.5558 | 8.8323
0.0968 | 8.8412 | 54.3512 | -226.7211 | 233.373 | -1.118 | -1.5372 | -0.75462 | 2.0496 | 226.7211 | 1.5372 | 8.9996
0.0986 | 9.2511 | 55.6369 | -230.9414 | 237.7889 | -1.1459 | -1.5465 | -0.76394 | 2.0682 | 230.9414 | 1.5465 | 9.167
0.1004 | 9.7821 | 57.4256 | -236.4939 | 243.6209 | -1.1552 | -1.5465 | -0.77325 | 2.0775 | 236.4939 | 1.5465 | 9.3343
0.1022 | 10.3318 | 59.41 | -242.5122 | 249.9466 | -1.1645 | -1.5465 | -0.77325 | 2.0868 | 242.5122 | 1.5465 | 9.5017
0.104 | 10.7696 | 61.2173 | -248.1672 | 255.8811 | -1.1739 | -1.5372 | -0.77325 | 2.0868 | 248.1672 | 1.5372 | 9.669
0.1058 | 11.0118 | 62.8011 | -252.8999 | 260.8653 | -1.1832 | -1.4999 | -0.78257 | 2.0682 | 252.8999 | 1.4999 | 9.8364
0.1076 | 11.2168 | 64.2078 | -256.7848 | 264.9738 | -1.2018 | -1.4627 | -0.78257 | 2.0589 | 256.7848 | 1.4627 | 10.0037
0.1094 | 11.2261 | 65.1395 | -259.6728 | 268.0016 | -1.2204 | -1.4161 | -0.78257 | 2.0403 | 259.6728 | 1.4161 | 10.1711
0.1111 | 11.0771 | 65.9779 | -262.6075 | 271.048 | -1.2577 | -1.3602 | -0.77325 | 2.031 | 262.6075 | 1.3602 | 10.3291
0.1129 | 10.9746 | 66.9096 | -265.9241 | 274.495 | -1.2856 | -1.3136 | -0.77325 | 2.0123 | 265.9241 | 1.3136 | 10.4965
0.1147 | 10.8535 | 68.1393 | -270.3213 | 279.0507 | -1.295 | -1.2577 | -0.76394 | 1.9844 | 270.3213 | 1.2577 | 10.6638
0.1165 | 10.8721 | 69.5088 | -275.4546 | 284.3517 | -1.3136 | -1.2018 | -0.7453 | 1.9564 | 275.4546 | 1.2018 | 10.8312
0.1183 | 10.9932 | 70.9063 | -280.9885 | 290.0719 | -1.3229 | -1.1552 | -0.73599 | 1.9378 | 280.9885 | 1.1552 | 10.9985
0.1201 | 11.1423 | 72.2757 | -286.1124 | 295.3728 | -1.3322 | -1.1086 | -0.71735 | 1.9192 | 286.1124 | 1.1086 | 11.1659
0.1219 | 11.2634 | 73.4216 | -289.6247 | 299.0621 | -1.3415 | -1.0621 | -0.70804 | 1.9098 | 289.6247 | 1.0621 | 11.3332
0.1237 | 11.2261 | 74.3905 | -292.2798 | 301.8756 | -1.3509 | -1.0062 | -0.69872 | 1.8912 | 292.2798 | 1.0062 | 11.5006
0.1255 | 10.8442 | 74.9309 | -293.4444 | 303.1147 | -1.3695 | -0.94094 | -0.68941 | 1.8633 | 293.4444 | 0.94094 | 11.6679
0.1273 | 10.4622 | 75.3501 | -294.9816 | 304.6984 | -1.3881 | -0.88505 | -0.68009 | 1.8446 | 294.9816 | 0.88505 | 11.8353
0.1291 | 10.2759 | 75.8905 | -298.3634 | 308.0989 | -1.4068 | -0.82915 | -0.66146 | 1.826 | 298.3634 | 0.82915 | 12.0026
0.1309 | 10.2293 | 76.4867 | -301.6613 | 311.4434 | -1.4254 | -0.78257 | -0.64282 | 1.8074 | 301.6613 | 0.78257 | 12.17
0.1327 | 10.3597 | 77.3531 | -304.5494 | 314.4526 | -1.4347 | -0.72667 | -0.62419 | 1.7887 | 304.5494 | 0.72667 | 12.3373
0.1345 | 10.5088 | 78.2009 | -307.2138 | 317.2382 | -1.4347 | -0.68009 | -0.60556 | 1.7701 | 307.2138 | 0.68009 | 12.5046
0.1362 | 10.5926 | 78.8158 | -308.5088 | 318.6449 | -1.444 | -0.64282 | -0.59624 | 1.7608 | 308.5088 | 0.64282 | 12.6627
0.138 | 10.3783 | 79.0487 | -308.9839 | 319.148 | -1.4627 | -0.59624 | -0.57761 | 1.7608 | 308.9839 | 0.59624 | 12.83
0.1398 | 9.8473 | 78.8251 | -308.6858 | 318.7847 | -1.4906 | -0.54034 | -0.56829 | 1.7701 | 308.6858 | 0.54034 | 12.9974
0.1416 | 9.2324 | 78.5177 | -308.5926 | 318.589 | -1.5186 | -0.47513 | -0.55898 | 1.7701 | 308.5926 | 0.47513 | 13.1647
0.1434 | 8.8318 | 78.7878 | -310.8379 | 320.8156 | -1.5372 | -0.41923 | -0.54034 | 1.7794 | 310.8379 | 0.41923 | 13.3321
0.1452 | 8.4499 | 79.2443 | -314.1545 | 324.1322 | -1.5558 | -0.36334 | -0.52171 | 1.7887 | 314.1545 | 0.36334 | 13.4994
0.147 | 8.1424 | 79.8685 | -318.794 | 328.7717 | -1.5745 | -0.30744 | -0.49376 | 1.826 | 318.794 | 0.30744 | 13.6668
0.1488 | 8.0027 | 80.6418 | -323.0143 | 333.0479 | -1.5931 | -0.27017 | -0.48445 | 1.8633 | 323.0143 | 0.27017 | 13.8341
0.1506 | 7.8164 | 81.1355 | -326.275 | 336.3272 | -1.6024 | -0.23291 | -0.46581 | 1.8819 | 326.275 | 0.23291 | 14.0015
0.1524 | 7.6953 | 81.4989 | -328.8556 | 338.9078 | -1.6024 | -0.19564 | -0.4565 | 1.8912 | 328.8556 | 0.19564 | 14.1688
0.1542 | 7.7511 | 81.7504 | -330.5604 | 340.622 | -1.6117 | -0.18633 | -0.44718 | 1.9098 | 330.5604 | 0.18633 | 14.3362
0.156 | 7.8536 | 81.6479 | -331.3523 | 341.3673 | -1.6117 | -0.16769 | -0.44718 | 1.9192 | 331.3523 | 0.16769 | 14.5035
0.1578 | 7.7698 | 81.1635 | -331.2126 | 341.1158 | -1.6117 | -0.13043 | -0.43787 | 1.9192 | 331.2126 | 0.13043 | 14.6709
0.1595 | 7.5276 | 80.3343 | -330.4673 | 340.1935 | -1.6117 | -0.07453 | -0.41923 | 1.9471 | 330.4673 | 0.07453 | 14.8289
0.1613 | 7.2015 | 79.235 | -329.0605 | 338.5631 | -1.621 | -0.027949 | -0.40992 | 1.9751 | 329.0605 | 0.027949 | 14.9963
0.1631 | 6.8288 | 77.8655 | -327.9239 | 337.1471 | -1.6397 | 0.018633 | -0.4006 | 2.003 | 327.9239 | -0.018633 | 15.1636
0.1649 | 6.4096 | 76.4681 | -327.6817 | 336.5974 | -1.6676 | 0.055898 | -0.39128 | 2.0403 | 327.6817 | -0.055898 | 15.331
0.1667 | 6.009 | 75.1917 | -328.0078 | 336.6347 | -1.6862 | 0.093163 | -0.37265 | 2.0682 | 328.0078 | -0.093163 | 15.4983
0.1685 | 5.6643 | 74.0738 | -328.9115 | 337.2682 | -1.7049 | 0.13043 | -0.36334 | 2.0868 | 328.9115 | -0.13043 | 15.6657
0.1703 | 5.6457 | 73.2353 | -331.7902 | 339.914 | -1.6956 | 0.13974 | -0.36334 | 2.0868 | 331.7902 | -0.13974 | 15.833
0.1721 | 5.5898 | 72.5739 | -335.0975 | 343.0256 | -1.6862 | 0.13974 | -0.35402 | 2.0962 | 335.0975 | -0.13974 | 16.0004
0.1739 | 5.6364 | 71.9124 | -338.3209 | 346.0534 | -1.6769 | 0.13974 | -0.35402 | 2.1055 | 338.3209 | -0.13974 | 16.1677
0.1757 | 5.9065 | 71.4373 | -341.2928 | 348.8763 | -1.6397 | 0.14906 | -0.36334 | 2.0868 | 341.2928 | -0.14906 | 16.3351
0.1775 | 6.2419 | 70.7386 | -343.6032 | 351.019 | -1.6024 | 0.15838 | -0.36334 | 2.0589 | 343.6032 | -0.15838 | 16.5024
0.1793 | 6.54 | 69.9746 | -344.9355 | 352.1929 | -1.5558 | 0.17701 | -0.36334 | 2.031 | 344.9355 | -0.17701 | 16.6698
0.1811 | 6.9872 | 69.1362 | -345.6156 | 352.7053 | -1.4999 | 0.19564 | -0.37265 | 1.9937 | 345.6156 | -0.19564 | 16.8371
0.1829 | 7.2667 | 68.0275 | -345.6435 | 352.5282 | -1.4533 | 0.22359 | -0.37265 | 1.9657 | 345.6435 | -0.22359 | 17.0045
0.1846 | 7.4158 | 66.6487 | -344.9448 | 351.5873 | -1.4161 | 0.27949 | -0.37265 | 1.9378 | 344.9448 | -0.27949 | 17.1625
0.1864 | 7.2015 | 65.009 | -343.4728 | 349.8079 | -1.3881 | 0.35402 | -0.36334 | 1.9098 | 343.4728 | -0.35402 | 17.3299
0.1882 | 6.9593 | 63.1458 | -342.3735 | 348.3546 | -1.3695 | 0.43787 | -0.35402 | 1.9005 | 342.3735 | -0.43787 | 17.4972
0.19 | 7.0524 | 61.2453 | -342.5412 | 348.1682 | -1.3322 | 0.5124 | -0.35402 | 1.8819 | 342.5412 | -0.5124 | 17.6646
0.1918 | 7.2108 | 59.2795 | -343.538 | 348.8017 | -1.295 | 0.58693 | -0.3447 | 1.8726 | 343.538 | -0.58693 | 17.8319
0.1936 | 7.4437 | 57.3231 | -345.951 | 350.87 | -1.2577 | 0.67077 | -0.33539 | 1.8633 | 345.951 | -0.67077 | 17.9993
0.1954 | 7.7139 | 55.4412 | -349.3607 | 353.9443 | -1.2018 | 0.77325 | -0.32607 | 1.8446 | 349.3607 | -0.77325 | 18.1666
0.1972 | 7.9841 | 53.6618 | -353.7021 | 357.9596 | -1.1366 | 0.86641 | -0.31675 | 1.826 | 353.7021 | -0.86641 | 18.334
0.199 | 8.5337 | 52.0315 | -359.1428 | 363.1209 | -1.0714 | 0.94094 | -0.31675 | 1.826 | 359.1428 | -0.94094 | 18.5013
0.2008 | 9.4653 | 50.7644 | -365.1425 | 368.8597 | -1.0155 | 0.97821 | -0.31675 | 1.7701 | 365.1425 | -0.97821 | 18.6687
0.2026 | 10.4342 | 48.6403 | -369.7447 | 373.1545 | -0.95026 | 1.0434 | -0.31675 | 1.7142 | 369.7447 | -1.0434 | 18.836
0.2044 | 11.3566 | 46.4696 | -374.7755 | 377.9151 | -0.87573 | 1.1273 | -0.31675 | 1.6956 | 374.7755 | -1.1273 | 19.0033
0.2062 | 12.1391 | 44.5691 | -381.3063 | 384.2036 | -0.78257 | 1.2297 | -0.30744 | 1.6769 | 381.3063 | -1.2297 | 19.1707
0.208 | 12.6329 | 42.7431 | -389.6071 | 392.2715 | -0.69872 | 1.3509 | -0.30744 | 1.6956 | 389.6071 | -1.3509 | 19.338
0.2097 | 13.0707 | 40.8333 | -398.9513 | 401.3642 | -0.62419 | 1.4999 | -0.30744 | 1.7515 | 398.9513 | -1.4999 | 19.4961
0.2115 | 13.7602 | 39.0259 | -410.2613 | 412.4599 | -0.53103 | 1.6117 | -0.30744 | 1.8074 | 410.2613 | -1.6117 | 19.6634
0.2133 | 14.4682 | 37.144 | -423.2575 | 425.2605 | -0.43787 | 1.7515 | -0.31675 | 1.9005 | 423.2575 | -1.7515 | 19.8308
0.2151 | 15.2321 | 35.3274 | -437.2692 | 439.1045 | -0.3447 | 1.8726 | -0.32607 | 1.9844 | 437.2692 | -1.8726 | 19.9981
0.2169 | 15.9215 | 33.492 | -452.8833 | 454.5416 | -0.25154 | 1.9937 | -0.33539 | 2.0775 | 452.8833 | -1.9937 | 20.1655
0.2187 | 16.5271 | 31.3773 | -470.4072 | 471.8978 | -0.14906 | 2.1427 | -0.35402 | 2.208 | 470.4072 | -2.1427 | 20.3328
0.2205 | 17.291 | 29.2345 | -489.3845 | 490.7074 | -0.037265 | 2.2732 | -0.37265 | 2.3384 | 489.3845 | -2.2732 | 20.5002
0.2223 | 18.232 | 27.0452 | -509.2748 | 510.4859 | 0.083847 | 2.4036 | -0.4006 | 2.4781 | 509.2748 | -2.4036 | 20.6675
0.2241 | 19.2288 | 24.7534 | -529.2861 | 530.3948 | 0.19564 | 2.5247 | -0.41923 | 2.5992 | 529.2861 | -2.5247 | 20.8349
0.2259 | 20.3747 | 22.3218 | -549.7726 | 550.7788 | 0.31675 | 2.6272 | -0.44718 | 2.7204 | 549.7726 | -2.6272 | 21.0022
0.2277 | 21.9492 | 19.8157 | -572.6628 | 573.613 | 0.4565 | 2.6831 | -0.48445 | 2.7949 | 572.6628 | -2.6831 | 21.1696
0.2295 | 23.5982 | 17.2351 | -596.3541 | 597.2484 | 0.59624 | 2.7297 | -0.5124 | 2.8787 | 596.3541 | -2.7297 | 21.3369
0.2313 | 25.4055 | 14.552 | -621.1633 | 622.0111 | 0.7453 | 2.7576 | -0.54966 | 2.9439 | 621.1633 | -2.7576 | 21.5043
0.2331 | 27.5296 | 11.8224 | -647.603 | 648.4601 | 0.913 | 2.7576 | -0.58693 | 3.0092 | 647.603 | -2.7576 | 21.6716
0.2348 | 29.8494 | 9.0461 | -674.1078 | 674.9928 | 1.09 | 2.7297 | -0.62419 | 3.0744 | 674.1078 | -2.7297 | 21.8297
0.2366 | 32.1878 | 6.2699 | -700.7617 | 701.6933 | 1.2577 | 2.6924 | -0.67077 | 3.1303 | 700.7617 | -2.6924 | 21.997
0.2384 | 34.5634 | 3.5402 | -728.3099 | 729.2882 | 1.4347 | 2.6458 | -0.71735 | 3.1955 | 728.3099 | -2.6458 | 22.1644
0.2402 | 37.1533 | 0.93163 | -757.4047 | 758.4668 | 1.621 | 2.5527 | -0.76394 | 3.2886 | 757.4047 | -2.5527 | 22.3317
0.242 | 39.8551 | -1.4068 | -787.9528 | 789.0987 | 1.826 | 2.4316 | -0.81052 | 3.3911 | 787.9528 | -2.4316 | 22.4991
0.2438 | 42.4916 | -3.6054 | -818.976 | 820.2337 | 2.031 | 2.3198 | -0.84778 | 3.5029 | 818.976 | -2.3198 | 22.6664
0.2456 | 45.156 | -5.627 | -849.2726 | 850.6328 | 2.2452 | 2.1893 | -0.89436 | 3.624 | 849.2726 | -2.1893 | 22.8338
0.2474 | 47.923 | -7.7791 | -880.6964 | 882.1777 | 2.4502 | 2.0496 | -0.94094 | 3.7545 | 880.6964 | -2.0496 | 23.0011
0.2492 | 50.6526 | -10.0616 | -912.6886 | 914.2723 | 2.6645 | 1.9005 | -0.98753 | 3.8942 | 912.6886 | -1.9005 | 23.1685
0.251 | 53.2333 | -12.1578 | -944.0192 | 945.7241 | 2.8787 | 1.7328 | -1.0341 | 4.0526 | 944.0192 | -1.7328 | 23.3358
0.2528 | 55.5996 | -14.4589 | -975.5735 | 977.3808 | 3.0744 | 1.5745 | -1.0807 | 4.2296 | 975.5735 | -1.5745 | 23.5032
0.2546 | 57.9473 | -16.8159 | -1006.4942 | 1008.4134 | 3.2607 | 1.3974 | -1.1273 | 4.4159 | 1006.4942 | -1.3974 | 23.6705
0.2564 | 60.6211 | -19.0425 | -1041.4489 | 1043.4985 | 3.4843 | 1.2204 | -1.1832 | 4.6116 | 1041.4489 | -1.2204 | 23.8379
0.2581 | 63.1924 | -21.2132 | -1077.4843 | 1079.6457 | 3.7172 | 1.0714 | -1.2391 | 4.7793 | 1077.4843 | -1.0714 | 23.9959
0.2599 | 66.2295 | -23.7099 | -1114.2557 | 1116.5755 | 3.9594 | 0.913 | -1.295 | 4.9656 | 1114.2557 | -0.913 | 24.1633
0.2617 | 69.015 | -26.1135 | -1149.5831 | 1152.0426 | 4.1923 | 0.73599 | -1.3509 | 5.1799 | 1149.5831 | -0.73599 | 24.3306
0.2635 | 71.8006 | -28.573 | -1184.5005 | 1187.0997 | 4.4252 | 0.54966 | -1.3974 | 5.3941 | 1184.5005 | -0.54966 | 24.498
0.2653 | 74.5303 | -30.9953 | -1218.7099 | 1221.4675 | 4.6581 | 0.36334 | -1.4533 | 5.627 | 1218.7099 | -0.36334 | 24.6653
0.2671 | 76.6264 | -33.2964 | -1249.1089 | 1251.9784 | 4.8631 | 0.21427 | -1.4999 | 5.832 | 1249.1089 | -0.21427 | 24.8327
0.2689 | 78.499 | -35.5603 | -1277.7938 | 1280.7843 | 5.0494 | 0.083847 | -1.5558 | 6.0276 | 1277.7938 | -0.083847 | 25
0.2707 | 80.6045 | -38.3924 | -1312.0498 | 1315.1707 | 5.273 | -0.083847 | -1.6024 | 6.2605 | 1312.0498 | 0.083847 | 25.1673
0.2725 | 82.71 | -41.3084 | -1347.8057 | 1351.0571 | 5.5059 | -0.27017 | -1.6583 | 6.5121 | 1347.8057 | 0.27017 | 25.3347
0.2743 | 84.9366 | -44.3176 | -1385.1919 | 1388.5831 | 5.7481 | -0.4565 | -1.7142 | 6.7636 | 1385.1919 | 0.4565 | 25.502
0.2761 | 87.0793 | -47.1031 | -1422.4664 | 1425.9879 | 5.9997 | -0.66146 | -1.7701 | 7.0245 | 1422.4664 | 0.66146 | 25.6694
0.2779 | 89.1662 | -49.6744 | -1459.9551 | 1463.5978 | 6.2512 | -0.86641 | -1.826 | 7.304 | 1459.9551 | 0.86641 | 25.8367
0.2797 | 91.2623 | -52.1991 | -1497.8165 | 1501.5896 | 6.5214 | -1.0807 | -1.8912 | 7.5835 | 1497.8165 | 1.0807 | 26.0041
0.2815 | 93.4051 | -54.7332 | -1535.9201 | 1539.8143 | 6.7916 | -1.3043 | -1.9471 | 7.8723 | 1535.9201 | 1.3043 | 26.1714
0.2832 | 95.5758 | -57.1927 | -1574.0796 | 1578.1043 | 7.0617 | -1.5372 | -2.0123 | 8.1704 | 1574.0796 | 1.5372 | 26.3295
0.285 | 97.8117 | -59.7267 | -1613.2639 | 1617.4097 | 7.3505 | -1.7701 | -2.0775 | 8.4871 | 1613.2639 | 1.7701 | 26.4968
0.2868 | 99.9544 | -62.0278 | -1651.2744 | 1655.5412 | 7.6394 | -2.0216 | -2.1427 | 8.8039 | 1651.2744 | 2.0216 | 26.6642
0.2886 | 101.9574 | -64.3383 | -1688.3439 | 1692.7225 | 7.9282 | -2.2545 | -2.208 | 9.1206 | 1688.3439 | 2.2545 | 26.8315
0.2904 | 103.8114 | -66.7232 | -1723.8203 | 1728.3014 | 8.189 | -2.5061 | -2.2545 | 9.4374 | 1723.8203 | 2.5061 | 26.9989
0.2922 | 105.5908 | -69.0896 | -1758.8681 | 1763.4611 | 8.4592 | -2.7576 | -2.3104 | 9.7542 | 1758.8681 | 2.7576 | 27.1662
0.294 | 107.0907 | -71.4186 | -1793.7856 | 1798.4624 | 8.72 | -2.9998 | -2.3663 | 10.0709 | 1793.7856 | 2.9998 | 27.3336
0.2958 | 108.3205 | -73.6173 | -1827.0261 | 1831.7774 | 8.9809 | -3.2421 | -2.4129 | 10.3783 | 1827.0261 | 3.2421 | 27.5009
0.2976 | 109.4664 | -75.76 | -1855.5432 | 1860.3784 | 9.1952 | -3.4936 | -2.4502 | 10.6485 | 1855.5432 | 3.4936 | 27.6683
0.2994 | 110.3794 | -77.8189 | -1882.8213 | 1887.7217 | 9.3908 | -3.7358 | -2.4781 | 10.9094 | 1882.8213 | 3.7358 | 27.8356
0.3012 | 111.5625 | -79.5145 | -1908.9535 | 1913.9191 | 9.6051 | -4.0153 | -2.5154 | 11.1982 | 1908.9535 | 4.0153 | 28.003
0.303 | 112.6712 | -80.7256 | -1933.828 | 1938.8495 | 9.8287 | -4.2855 | -2.5527 | 11.4963 | 1933.828 | 4.2855 | 28.1703
0.3048 | 113.7891 | -81.9274 | -1958.218 | 1963.2861 | 10.0523 | -4.565 | -2.5992 | 11.8037 | 1958.218 | 4.565 | 28.3377
0.3066 | 114.9257 | -83.3808 | -1981.7603 | 1986.9029 | 10.2665 | -4.8817 | -2.6272 | 12.1019 | 1981.7603 | 4.8817 | 28.505
0.3083 | 115.7548 | -84.9738 | -2003.7374 | 2008.9359 | 10.4622 | -5.1892 | -2.6645 | 12.372 | 2003.7374 | 5.1892 | 28.6631
0.3101 | 116.1461 | -86.6042 | -2023.9631 | 2029.1988 | 10.6206 | -5.4873 | -2.6924 | 12.6236 | 2023.9631 | 5.4873 | 28.8304
0.3119 | 116.3884 | -88.2345 | -2043.4155 | 2048.6978 | 10.7789 | -5.7761 | -2.7204 | 12.8937 | 2043.4155 | 5.7761 | 28.9978
0.3137 | 116.7144 | -90.0698 | -2061.7126 | 2067.0416 | 10.928 | -6.0742 | -2.7483 | 13.1453 | 2061.7126 | 6.0742 | 29.1651
0.3155 | 116.8169 | -91.8399 | -2078.2304 | 2083.5966 | 11.0677 | -6.3817 | -2.7763 | 13.4061 | 2078.2304 | 6.3817 | 29.3325
0.3173 | 116.556 | -93.3585 | -2093.3042 | 2098.689 | 11.2168 | -6.6798 | -2.8042 | 13.6949 | 2093.3042 | 6.6798 | 29.4998
0.3191 | 116.3977 | -95.0913 | -2109.0301 | 2114.4428 | 11.3472 | -6.9872 | -2.8415 | 13.9558 | 2109.0301 | 6.9872 | 29.6672
0.3209 | 116.2207 | -96.7217 | -2123.9175 | 2129.3582 | 11.4683 | -7.2947 | -2.8694 | 14.1887 | 2123.9175 | 7.2947 | 29.8345
0.3227 | 115.6803 | -97.8862 | -2137.3423 | 2142.783 | 11.5895 | -7.5928 | -2.888 | 14.4309 | 2137.3423 | 7.5928 | 30.0019
0.3245 | 114.7301 | -98.8644 | -2149.4907 | 2154.8941 | 11.7106 | -7.8723 | -2.916 | 14.6731 | 2149.4907 | 7.8723 | 30.1692
0.3263 | 113.873 | -99.8613 | -2161.4249 | 2166.8097 | 11.8224 | -8.1424 | -2.9439 | 14.9247 | 2161.4249 | 8.1424 | 30.3366
0.3281 | 113.1183 | -100.532 | -2172.9864 | 2178.3339 | 11.9435 | -8.4312 | -2.9626 | 15.1576 | 2172.9864 | 8.4312 | 30.5039
0.3299 | 111.8793 | -101.2028 | -2184.2032 | 2189.5041 | 12.0553 | -8.7107 | -2.9812 | 15.3905 | 2184.2032 | 8.7107 | 30.6713
0.3316 | 110.3235 | -101.8363 | -2194.3766 | 2199.603 | 12.1671 | -8.9995 | -3.0092 | 15.642 | 2194.3766 | 8.9995 | 30.8293
0.3334 | 108.7863 | -102.5909 | -2203.8792 | 2209.0404 | 12.2695 | -9.2697 | -3.0278 | 15.8843 | 2203.8792 | 9.2697 | 30.9967
0.3352 | 107.2677 | -103.3921 | -2212.8601 | 2217.9747 | 12.372 | -9.5492 | -3.0557 | 16.1358 | 2212.8601 | 9.5492 | 31.164
0.337 | 105.7398 | -104.2027 | -2221.4217 | 2226.4805 | 12.4652 | -9.81 | -3.0744 | 16.3594 | 2221.4217 | 9.81 | 31.3313
0.3388 | 104.3238 | -105.2181 | -2229.3313 | 2234.3621 | 12.5397 | -10.0523 | -3.093 | 16.5644 | 2229.3313 | 10.0523 | 31.4987
0.3406 | 102.9822 | -106.4665 | -2236.3092 | 2241.3213 | 12.6049 | -10.3131 | -3.1116 | 16.7786 | 2236.3092 | 10.3131 | 31.666
0.3424 | 101.7432 | -107.9385 | -2242.3368 | 2247.3583 | 12.6515 | -10.574 | -3.1303 | 16.9743 | 2242.3368 | 10.574 | 31.8334
0.3442 | 100.4668 | -109.3732 | -2247.8893 | 2252.9294 | 12.6888 | -10.8069 | -3.1396 | 17.1606 | 2247.8893 | 10.8069 | 32.0007
0.346 | 98.8458 | -110.7893 | -2253.0133 | 2258.0441 | 12.726 | -10.9932 | -3.1396 | 17.319 | 2253.0133 | 10.9932 | 32.1681
0.3478 | 97.0943 | -111.9911 | -2258.0254 | 2263.0376 | 12.754 | -11.1982 | -3.1489 | 17.4774 | 2258.0254 | 11.1982 | 32.3354
0.3496 | 95.3242 | -113.1556 | -2263.0469 | 2268.0311 | 12.7819 | -11.4124 | -3.1582 | 17.645 | 2263.0469 | 11.4124 | 32.5028
0.3514 | 93.284 | -114.199 | -2267.8262 | 2272.7638 | 12.8192 | -11.6081 | -3.1675 | 17.7941 | 2267.8262 | 11.6081 | 32.6701
0.3532 | 90.8245 | -115.0095 | -2271.4875 | 2276.3692 | 12.8658 | -11.7572 | -3.1862 | 17.9338 | 2271.4875 | 11.7572 | 32.8375
0.355 | 88.6724 | -116.053 | -2274.45 | 2279.2945 | 12.9031 | -11.9248 | -3.1955 | 18.055 | 2274.45 | 11.9248 | 33.0048
0.3567 | 86.5669 | -117.2734 | -2276.453 | 2281.2789 | 12.931 | -12.046 | -3.2141 | 18.1574 | 2276.453 | 12.046 | 33.1629
0.3585 | 84.2006 | -118.0467 | -2277.8691 | 2282.6484 | 12.9776 | -12.1764 | -3.2328 | 18.2786 | 2277.8691 | 12.1764 | 33.3302
0.3603 | 81.7225 | -118.6243 | -2279.0895 | 2283.8036 | 13.0335 | -12.2882 | -3.2607 | 18.3997 | 2279.0895 | 12.2882 | 33.4976
0.3621 | 79.1698 | -119.7888 | -2279.8535 | 2284.5489 | 13.0707 | -12.3068 | -3.298 | 18.4462 | 2279.8535 | 12.3068 | 33.6649
0.3639 | 76.5985 | -120.9627 | -2280.5149 | 2285.2103 | 13.1173 | -12.3348 | -3.3352 | 18.5021 | 2280.5149 | 12.3348 | 33.8323
0.3657 | 74.0552 | -121.9874 | -2280.9342 | 2285.5923 | 13.1453 | -12.4186 | -3.3539 | 18.5953 | 2280.9342 | 12.4186 | 33.9996
0.3675 | 71.5491 | -123.0588 | -2281.1205 | 2285.76 | 13.1453 | -12.5025 | -3.3725 | 18.6512 | 2281.1205 | 12.5025 | 34.167
0.3693 | 69.0803 | -123.683 | -2280.9248 | 2285.5178 | 13.1732 | -12.6236 | -3.3911 | 18.7444 | 2280.9248 | 12.6236 | 34.3343
0.3711 | 66.5928 | -124.0277 | -2279.8628 | 2284.4091 | 13.2012 | -12.754 | -3.4098 | 18.8562 | 2279.8628 | 12.754 | 34.5017
0.3729 | 64.2172 | -124.7451 | -2277.9623 | 2282.4807 | 13.2105 | -12.9031 | -3.4284 | 18.968 | 2277.9623 | 12.9031 | 34.669
0.3747 | 61.7111 | -125.0059 | -2275.9127 | 2280.3566 | 13.2012 | -13.0614 | -3.4377 | 19.0704 | 2275.9127 | 13.0614 | 34.8364
0.3765 | 59.028 | -125.1457 | -2273.2855 | 2277.6735 | 13.2012 | -13.108 | -3.4657 | 19.1077 | 2273.2855 | 13.108 | 35.0037
0.3783 | 56.6896 | -125.7419 | -2270.267 | 2274.6364 | 13.2012 | -13.2291 | -3.475 | 19.1915 | 2270.267 | 13.2291 | 35.1711
0.3801 | 54.1835 | -126.7015 | -2266.7641 | 2271.1334 | 13.1919 | -13.3037 | -3.4843 | 19.2381 | 2266.7641 | 13.3037 | 35.3384
0.3818 | 51.8824 | -127.6052 | -2262.7301 | 2267.0995 | 13.1639 | -13.4527 | -3.4657 | 19.322 | 2262.7301 | 13.4527 | 35.4965
0.3836 | 49.0782 | -128.4716 | -2258.4726 | 2262.8419 | 13.1453 | -13.5645 | -3.4657 | 19.3872 | 2258.4726 | 13.5645 | 35.6638
0.3854 | 46.3858 | -129.2169 | -2253.9728 | 2258.3329 | 13.1173 | -13.667 | -3.475 | 19.4431 | 2253.9728 | 13.667 | 35.8312
0.3872 | 44.1033 | -130.2324 | -2248.7743 | 2253.1809 | 13.108 | -13.7229 | -3.4843 | 19.471 | 2248.7743 | 13.7229 | 35.9985
0.389 | 41.9233 | -130.8379 | -2243.9578 | 2248.3644 | 13.0801 | -13.8626 | -3.4936 | 19.5269 | 2243.9578 | 13.8626 | 36.1659
0.3908 | 39.706 | -131.2106 | -2240.1381 | 2244.5354 | 13.0614 | -13.9092 | -3.5216 | 19.5456 | 2240.1381 | 13.9092 | 36.3332
0.3926 | 37.5539 | -131.4807 | -2236.4768 | 2240.8555 | 13.0521 | -13.9558 | -3.5495 | 19.5642 | 2236.4768 | 13.9558 | 36.5006
0.3944 | 35.495 | -131.6391 | -2232.2938 | 2236.6539 | 13.0428 | -14.0117 | -3.5681 | 19.5921 | 2232.2938 | 14.0117 | 36.6679
0.3962 | 33.343 | -131.3876 | -2227.0115 | 2231.3436 | 13.0149 | -14.0769 | -3.5775 | 19.6294 | 2227.0115 | 14.0769 | 36.8353
0.398 | 30.7344 | -130.81 | -2221.3565 | 2225.6234 | 12.9683 | -14.1608 | -3.5961 | 19.6667 | 2221.3565 | 14.1608 | 37.0026
0.3998 | 28.5637 | -130.4094 | -2215.3009 | 2219.5305 | 12.9683 | -14.198 | -3.6054 | 19.6946 | 2215.3009 | 14.198 | 37.17
0.4016 | 26.3558 | -129.8877 | -2208.2671 | 2212.4501 | 12.9496 | -14.2167 | -3.624 | 19.7039 | 2208.2671 | 14.2167 | 37.3373
0.4034 | 24.2037 | -129.5616 | -2200.5439 | 2204.699 | 12.9217 | -14.2167 | -3.624 | 19.6946 | 2200.5439 | 14.2167 | 37.5046
0.4051 | 21.884 | -129.0678 | -2192.5971 | 2196.7056 | 12.8937 | -14.226 | -3.6334 | 19.6946 | 2192.5971 | 14.226 | 37.6627
0.4069 | 19.9275 | -128.1735 | -2184.1845 | 2188.2371 | 12.8658 | -14.198 | -3.652 | 19.6667 | 2184.1845 | 14.198 | 37.83
0.4087 | 18.1761 | -127.3816 | -2175.5111 | 2179.5171 | 12.8285 | -14.1701 | -3.6799 | 19.6294 | 2175.5111 | 14.1701 | 37.9974
0.4105 | 16.3594 | -126.7108 | -2166.437 | 2170.3964 | 12.7913 | -14.1142 | -3.6986 | 19.5735 | 2166.437 | 14.1142 | 38.1647
0.4123 | 14.4402 | -126.0959 | -2157.6518 | 2161.5739 | 12.7354 | -14.0676 | -3.7265 | 19.5176 | 2157.6518 | 14.0676 | 38.3321
0.4141 | 12.2975 | -125.5276 | -2148.9317 | 2152.8166 | 12.6701 | -14.021 | -3.7638 | 19.4524 | 2148.9317 | 14.021 | 38.4994
0.4159 | 10.397 | -124.4935 | -2140.4073 | 2144.2456 | 12.6143 | -13.9372 | -3.801 | 19.3686 | 2140.4073 | 13.9372 | 38.6668
0.4177 | 8.571 | -123.1892 | -2132.2649 | 2136.038 | 12.5584 | -13.8626 | -3.8476 | 19.2847 | 2132.2649 | 13.8626 | 38.8341
0.4195 | 7.3599 | -121.7359 | -2122.6132 | 2126.3118 | 12.5211 | -13.7788 | -3.8942 | 19.2195 | 2122.6132 | 13.7788 | 39.0015
0.4213 | 6.1581 | -120.031 | -2112.6261 | 2116.2409 | 12.4838 | -13.7043 | -3.9408 | 19.1636 | 2112.6261 | 13.7043 | 39.1688
0.4231 | 5.6736 | -118.7733 | -2101.9683 | 2105.5271 | 12.4186 | -13.6111 | -3.9967 | 19.0611 | 2101.9683 | 13.6111 | 39.3362
0.4249 | 4.8165 | -117.8137 | -2092.5682 | 2096.0897 | 12.3534 | -13.5179 | -4.034 | 18.9586 | 2092.5682 | 13.5179 | 39.5035
0.4267 | 3.9874 | -117.0498 | -2083.9786 | 2087.4722 | 12.2789 | -13.4248 | -4.0619 | 18.8562 | 2083.9786 | 13.4248 | 39.6709
0.4285 | 3.3259 | -116.0437 | -2074.7275 | 2078.1838 | 12.195 | -13.3409 | -4.0992 | 18.7444 | 2074.7275 | 13.3409 | 39.8382
0.4302 | 2.9533 | -114.6276 | -2064.3585 | 2067.7589 | 12.1112 | -13.2571 | -4.1364 | 18.6419 | 2064.3585 | 13.2571 | 39.9963
0.432 | 2.6924 | -113.3233 | -2053.3466 | 2056.6912 | 12.0273 | -13.1546 | -4.1644 | 18.5208 | 2053.3466 | 13.1546 | 40.1636
0.4338 | 2.2266 | -111.4507 | -2042.5397 | 2045.8097 | 11.9621 | -13.0149 | -4.211 | 18.3903 | 2042.5397 | 13.0149 | 40.331
0.4356 | 1.7142 | -109.839 | -2031.2763 | 2034.4718 | 11.8783 | -12.8658 | -4.2575 | 18.2413 | 2031.2763 | 12.8658 | 40.4983
0.4374 | 1.2577 | -108.4695 | -2020.1527 | 2023.2923 | 11.7758 | -12.7167 | -4.2948 | 18.0829 | 2020.1527 | 12.7167 | 40.6657
0.4392 | 0.99684 | -107.128 | -2008.9545 | 2012.0289 | 11.6733 | -12.5677 | -4.3321 | 17.9152 | 2008.9545 | 12.5677 | 40.833
0.441 | 0.88505 | -105.5535 | -1998.5296 | 2001.5388 | 11.5801 | -12.4466 | -4.3787 | 17.7848 | 1998.5296 | 12.4466 | 41.0004
0.4428 | 0.90368 | -103.8486 | -1988.7289 | 1991.6635 | 11.5056 | -12.3441 | -4.4159 | 17.673 | 1988.7289 | 12.3441 | 41.1677
0.4446 | 1.0527 | -101.9667 | -1976.6643 | 1979.5057 | 11.4124 | -12.223 | -4.4625 | 17.5332 | 1976.6643 | 12.223 | 41.3351
0.4464 | 1.3136 | -100.355 | -1963.6774 | 1966.4536 | 11.2913 | -12.1019 | -4.4998 | 17.3842 | 1963.6774 | 12.1019 | 41.5024
0.4482 | 1.5186 | -99.0414 | -1950.5507 | 1953.2804 | 11.1702 | -11.9528 | -4.5277 | 17.2165 | 1950.5507 | 11.9528 | 41.6698
0.45 | 1.9192 | -97.6533 | -1937.2843 | 1939.9581 | 11.0491 | -11.8317 | -4.565 | 17.0674 | 1937.2843 | 11.8317 | 41.8371
0.4518 | 2.8135 | -96.1254 | -1923.0211 | 1925.6483 | 10.9187 | -11.7292 | -4.5929 | 16.9277 | 1923.0211 | 11.7292 | 42.0045
0.4536 | 3.7917 | -94.5696 | -1908.6274 | 1911.1894 | 10.7976 | -11.636 | -4.6116 | 16.7879 | 1908.6274 | 11.636 | 42.1718
0.4553 | 4.6395 | -92.9113 | -1894.0009 | 1896.4976 | 10.6671 | -11.5429 | -4.6395 | 16.6482 | 1894.0009 | 11.5429 | 42.3299
0.4571 | 5.1519 | -91.1505 | -1879.4395 | 1881.8618 | 10.546 | -11.4311 | -4.6675 | 16.5085 | 1879.4395 | 11.4311 | 42.4972
0.4589 | 5.4407 | -89.3059 | -1865.0272 | 1867.3749 | 10.4249 | -11.3193 | -4.6861 | 16.3687 | 1865.0272 | 11.3193 | 42.6646
0.4607 | 5.6736 | -87.5358 | -1850.736 | 1853.0185 | 10.3131 | -11.1889 | -4.714 | 16.2197 | 1850.736 | 11.1889 | 42.8319
0.4625 | 5.9811 | -85.9521 | -1835.7461 | 1837.9727 | 10.1734 | -11.1236 | -4.7327 | 16.0892 | 1835.7461 | 11.1236 | 42.9993
0.4643 | 6.3164 | -84.3031 | -1820.5047 | 1822.6661 | 10.043 | -11.0212 | -4.742 | 15.9495 | 1820.5047 | 11.0212 | 43.1666
0.4661 | 6.9499 | -82.7379 | -1804.9838 | 1807.0892 | 9.9218 | -10.928 | -4.7513 | 15.8097 | 1804.9838 | 10.928 | 43.334
0.4679 | 7.807 | -81.2194 | -1789.2392 | 1791.2981 | 9.7914 | -10.8535 | -4.7606 | 15.6793 | 1789.2392 | 10.8535 | 43.5013
0.4697 | 8.9809 | -79.7288 | -1772.6656 | 1774.6686 | 9.661 | -10.7603 | -4.7699 | 15.5396 | 1772.6656 | 10.7603 | 43.6687
0.4715 | 10.1827 | -78.2568 | -1756.8372 | 1758.8029 | 9.5306 | -10.7044 | -4.7699 | 15.4185 | 1756.8372 | 10.7044 | 43.836
0.4733 | 11.1516 | -76.7289 | -1741.484 | 1743.3938 | 9.4001 | -10.6485 | -4.7793 | 15.3067 | 1741.484 | 10.6485 | 44.0033
0.4751 | 11.9248 | -75.2011 | -1726.3636 | 1728.2362 | 9.279 | -10.6019 | -4.7793 | 15.1855 | 1726.3636 | 10.6019 | 44.1707
0.4769 | 12.8006 | -73.8595 | -1711.5321 | 1713.3674 | 9.1486 | -10.5554 | -4.7793 | 15.0831 | 1711.5321 | 10.5554 | 44.338
0.4786 | 13.8906 | -72.5739 | -1697.5856 | 1699.393 | 9.0368 | -10.5181 | -4.7886 | 14.9806 | 1697.5856 | 10.5181 | 44.4961
0.4804 | 14.6359 | -71.3348 | -1685.0552 | 1686.8346 | 8.9343 | -10.4995 | -4.7886 | 14.9061 | 1685.0552 | 10.4995 | 44.6634
0.4822 | 14.8967 | -69.9094 | -1673.8943 | 1675.6365 | 8.8784 | -10.4436 | -4.7979 | 14.8315 | 1673.8943 | 10.4436 | 44.8308
0.484 | 14.9061 | -68.307 | -1662.2955 | 1663.9818 | 8.8039 | -10.4063 | -4.7979 | 14.757 | 1662.2955 | 10.4063 | 44.9981
0.4858 | 15.1017 | -66.863 | -1649.3738 | 1651.0135 | 8.7107 | -10.3504 | -4.7979 | 14.6638 | 1649.3738 | 10.3504 | 45.1655
0.4876 | 15.2694 | -65.6612 | -1636.4987 | 1638.1011 | 8.6082 | -10.3038 | -4.8072 | 14.5614 | 1636.4987 | 10.3038 | 45.3328
0.4894 | 15.465 | -64.6084 | -1623.9124 | 1625.4869 | 8.5151 | -10.2479 | -4.8072 | 14.4682 | 1623.9124 | 10.2479 | 45.5002
0.4912 | 15.9681 | -63.77 | -1611.7547 | 1613.3198 | 8.4126 | -10.2106 | -4.8165 | 14.375 | 1611.7547 | 10.2106 | 45.6675
0.493 | 16.7879 | -63.1551 | -1599.8671 | 1601.4322 | 8.3101 | -10.1734 | -4.8165 | 14.2912 | 1599.8671 | 10.1734 | 45.8349
0.4948 | 17.6264 | -62.5961 | -1587.2901 | 1588.8553 | 8.2076 | -10.1548 | -4.8165 | 14.2073 | 1587.2901 | 10.1548 | 46.0022
0.4966 | 18.4462 | -62.0744 | -1574.2101 | 1575.7845 | 8.0959 | -10.1268 | -4.8165 | 14.1142 | 1574.2101 | 10.1268 | 46.1696
0.4984 | 19.0798 | -61.3664 | -1561.1393 | 1562.7044 | 7.9934 | -10.1175 | -4.8072 | 14.0303 | 1561.1393 | 10.1175 | 46.3369
0.5002 | 19.7319 | -60.7888 | -1548.1524 | 1549.6989 | 7.9002 | -10.0895 | -4.8072 | 13.9465 | 1548.1524 | 10.0895 | 46.5043
0.502 | 20.4213 | -60.1273 | -1534.9885 | 1536.535 | 7.7977 | -10.0709 | -4.7979 | 13.8719 | 1534.9885 | 10.0709 | 46.6716
0.5037 | 21.1387 | -59.5404 | -1521.8618 | 1523.399 | 7.7046 | -10.0616 | -4.7886 | 13.8067 | 1521.8618 | 10.0616 | 46.8297
0.5055 | 21.9026 | -58.823 | -1508.1669 | 1509.6948 | 7.6114 | -10.0709 | -4.7793 | 13.7322 | 1508.1669 | 10.0709 | 46.997
0.5073 | 22.4243 | -58.0777 | -1495.18 | 1496.6892 | 7.5276 | -10.0616 | -4.7699 | 13.667 | 1495.18 | 10.0616 | 47.1644
0.5091 | 22.8715 | -57.2393 | -1482.8639 | 1484.3545 | 7.4437 | -10.0616 | -4.7606 | 13.6018 | 1482.8639 | 10.0616 | 47.3317
0.5109 | 23.3093 | -56.7734 | -1471.0042 | 1472.4762 | 7.3692 | -10.0616 | -4.7606 | 13.5459 | 1471.0042 | 10.0616 | 47.4991
0.5127 | 23.8497 | -56.3635 | -1459.48 | 1460.952 | 7.2947 | -10.0616 | -4.7513 | 13.49 | 1459.48 | 10.0616 | 47.6664
0.5145 | 24.3062 | -55.6834 | -1448.0862 | 1449.5488 | 7.2294 | -10.0616 | -4.742 | 13.4341 | 1448.0862 | 10.0616 | 47.8338
0.5163 | 24.6882 | -54.9288 | -1436.6924 | 1438.1364 | 7.1456 | -10.0709 | -4.7327 | 13.3875 | 1436.6924 | 10.0709 | 48.0011
0.5181 | 24.9304 | -54.3046 | -1426.1836 | 1427.609 | 7.0711 | -10.0895 | -4.7234 | 13.3409 | 1426.1836 | 10.0895 | 48.1685
0.5199 | 24.9304 | -53.6898 | -1416.6717 | 1418.0784 | 7.0058 | -10.0895 | -4.714 | 13.3037 | 1416.6717 | 10.0895 | 48.3358
0.5217 | 24.772 | -53.1028 | -1407.8398 | 1409.2279 | 6.9499 | -10.0895 | -4.7047 | 13.2571 | 1407.8398 | 10.0895 | 48.5032
0.5235 | 24.6043 | -52.7022 | -1400.0235 | 1401.393 | 6.8941 | -10.1175 | -4.6954 | 13.2384 | 1400.0235 | 10.1175 | 48.6705
0.5253 | 24.5111 | -52.3389 | -1392.3562 | 1393.7163 | 6.8475 | -10.1548 | -4.6768 | 13.2384 | 1392.3562 | 10.1548 | 48.8379
0.5271 | 24.3528 | -51.8079 | -1384.4187 | 1385.7602 | 6.7823 | -10.22 | -4.6488 | 13.2384 | 1384.4187 | 10.22 | 49.0052
0.5288 | 24.2689 | -51.5936 | -1377.2079 | 1378.5401 | 6.7264 | -10.2572 | -4.6395 | 13.2198 | 1377.2079 | 10.2572 | 49.1633
0.5306 | 24.418 | -51.4911 | -1369.8853 | 1371.2175 | 6.6425 | -10.3131 | -4.6302 | 13.2105 | 1369.8853 | 10.3131 | 49.3306
0.5324 | 24.4459 | -51.4445 | -1363.0471 | 1364.3887 | 6.5866 | -10.3411 | -4.6209 | 13.2012 | 1363.0471 | 10.3411 | 49.498
0.5342 | 24.39 | -51.3793 | -1357.0288 | 1358.3703 | 6.568 | -10.3877 | -4.6116 | 13.2198 | 1357.0288 | 10.3877 | 49.6653
0.536 | 24.3341 | -51.2582 | -1351.1689 | 1352.5011 | 6.54 | -10.4436 | -4.6022 | 13.2478 | 1351.1689 | 10.4436 | 49.8327
0.5378 | 24.0453 | -51.0905 | -1345.9331 | 1347.2653 | 6.5214 | -10.4901 | -4.5836 | 13.2757 | 1345.9331 | 10.4901 | 50
0.5396 | 23.3466 | -50.839 | -1342.0948 | 1343.3898 | 6.5121 | -10.546 | -4.5743 | 13.3037 | 1342.0948 | 10.546 | 50.1673
0.5414 | 22.6386 | -50.6992 | -1338.508 | 1339.775 | 6.5028 | -10.6019 | -4.5557 | 13.3316 | 1338.508 | 10.6019 | 50.3347
0.5432 | 22.2566 | -50.8856 | -1334.7442 | 1336.0019 | 6.4841 | -10.6765 | -4.5463 | 13.3782 | 1334.7442 | 10.6765 | 50.502
0.545 | 22.0796 | -51.1091 | -1330.1327 | 1331.409 | 6.4655 | -10.7603 | -4.5184 | 13.4341 | 1330.1327 | 10.7603 | 50.6694
0.5468 | 21.8374 | -51.3979 | -1326.0894 | 1327.3751 | 6.4655 | -10.8348 | -4.4998 | 13.4807 | 1326.0894 | 10.8348 | 50.8367
0.5486 | 21.502 | -51.4352 | -1321.8691 | 1323.1548 | 6.4655 | -10.9187 | -4.4718 | 13.5459 | 1321.8691 | 10.9187 | 51.0041
0.5504 | 20.7287 | -51.4352 | -1319.2233 | 1320.4996 | 6.4841 | -10.9932 | -4.4532 | 13.6111 | 1319.2233 | 10.9932 | 51.1714
0.5521 | 20.0859 | -51.7613 | -1317.5836 | 1318.86 | 6.5121 | -11.0677 | -4.4346 | 13.6763 | 1317.5836 | 11.0677 | 51.3295
0.5539 | 19.6201 | -52.1339 | -1315.6086 | 1316.8942 | 6.54 | -11.1702 | -4.4159 | 13.7695 | 1315.6086 | 11.1702 | 51.4968
0.5557 | 19.3313 | -52.7209 | -1313.5776 | 1314.8912 | 6.5587 | -11.2634 | -4.3973 | 13.8533 | 1313.5776 | 11.2634 | 51.6642
0.5575 | 18.9959 | -53.3357 | -1311.7796 | 1313.1118 | 6.5773 | -11.3659 | -4.3787 | 13.9372 | 1311.7796 | 11.3659 | 51.8315
0.5593 | 18.3531 | -54.1183 | -1311.733 | 1313.0932 | 6.6146 | -11.4497 | -4.36 | 14.0117 | 1311.733 | 11.4497 | 51.9989
0.5611 | 17.6637 | -55.0499 | -1312.3013 | 1313.6894 | 6.6425 | -11.5336 | -4.3321 | 14.0955 | 1312.3013 | 11.5336 | 52.1662
0.5629 | 17.3656 | -56.1306 | -1312.5994 | 1314.0435 | 6.6705 | -11.6733 | -4.3134 | 14.2167 | 1312.5994 | 11.6733 | 52.3336
0.5647 | 17.0954 | -57.2206 | -1312.8882 | 1314.3788 | 6.6891 | -11.8224 | -4.2855 | 14.3378 | 1312.8882 | 11.8224 | 52.5009
0.5665 | 16.3501 | -58.0964 | -1313.7453 | 1315.2639 | 6.7264 | -11.9342 | -4.2575 | 14.4496 | 1313.7453 | 11.9342 | 52.6683
0.5683 | 15.2135 | -58.7485 | -1315.208 | 1316.7452 | 6.7823 | -12.0366 | -4.2389 | 14.5614 | 1315.208 | 12.0366 | 52.8356
0.5701 | 14.6825 | -59.5404 | -1317.1924 | 1318.7575 | 6.8382 | -12.1857 | -4.211 | 14.7011 | 1317.1924 | 12.1857 | 53.003
0.5719 | 14.2353 | -60.3509 | -1319.1954 | 1320.7884 | 6.8847 | -12.3441 | -4.183 | 14.8502 | 1319.1954 | 12.3441 | 53.1703
0.5737 | 14.0955 | -61.1987 | -1320.4344 | 1322.0741 | 6.9406 | -12.4838 | -4.1551 | 14.9806 | 1320.4344 | 12.4838 | 53.3377
0.5755 | 13.7415 | -62.1676 | -1320.8071 | 1322.4933 | 6.9872 | -12.5956 | -4.1364 | 15.1017 | 1320.8071 | 12.5956 | 53.505
0.5772 | 13.1453 | -63.1272 | -1321.6921 | 1323.4156 | 7.0431 | -12.6981 | -4.1085 | 15.2135 | 1321.6921 | 12.6981 | 53.6631
0.579 | 11.9342 | -63.8259 | -1324.5802 | 1326.3316 | 7.1176 | -12.7819 | -4.0805 | 15.316 | 1324.5802 | 12.7819 | 53.8304
0.5808 | 11.1143 | -64.6364 | -1328.0924 | 1329.8811 | 7.1922 | -12.8844 | -4.034 | 15.4278 | 1328.0924 | 12.8844 | 53.9978
0.5826 | 10.4995 | -65.5587 | -1331.6326 | 1333.4679 | 7.2667 | -12.9962 | -3.9874 | 15.5489 | 1331.6326 | 12.9962 | 54.1651
0.5844 | 10.1641 | -66.6301 | -1335.2939 | 1337.1851 | 7.3319 | -13.108 | -3.9781 | 15.67 | 1335.2939 | 13.108 | 54.3325
0.5862 | 9.4933 | -67.6362 | -1338.4335 | 1340.3713 | 7.3878 | -13.1825 | -3.9501 | 15.7632 | 1338.4335 | 13.1825 | 54.4998
0.588 | 8.6828 | -68.5213 | -1341.275 | 1343.25 | 7.4623 | -13.2478 | -3.9315 | 15.847 | 1341.275 | 13.2478 | 54.6672
0.5898 | 8.1145 | -69.3132 | -1344.6568 | 1346.6691 | 7.5276 | -13.3409 | -3.9128 | 15.9588 | 1344.6568 | 13.3409 | 54.8345
0.5916 | 7.6487 | -70.1144 | -1348.8305 | 1350.8707 | 7.6021 | -13.4434 | -3.9035 | 16.0799 | 1348.8305 | 13.4434 | 55.0019
0.5934 | 8.012 | -71.3255 | -1353.0414 | 1355.1469 | 7.658 | -13.5925 | -3.8942 | 16.2383 | 1353.0414 | 13.5925 | 55.1692
0.5952 | 8.3194 | -72.6577 | -1355.6034 | 1357.7834 | 7.7046 | -13.7229 | -3.8756 | 16.3687 | 1355.6034 | 13.7229 | 55.3366
0.597 | 8.5617 | -74.0738 | -1356.8984 | 1359.1529 | 7.7511 | -13.8533 | -3.8569 | 16.4898 | 1356.8984 | 13.8533 | 55.5039
0.5988 | 8.5524 | -75.583 | -1358.2679 | 1360.6063 | 7.7884 | -13.9931 | -3.829 | 16.6296 | 1358.2679 | 13.9931 | 55.6713
0.6006 | 8.5058 | -77.26 | -1359.7771 | 1362.2087 | 7.835 | -14.1514 | -3.801 | 16.7786 | 1359.7771 | 14.1514 | 55.8386
0.6023 | 8.3101 | -78.6108 | -1362.4043 | 1364.9011 | 7.8909 | -14.3098 | -3.7731 | 16.9277 | 1362.4043 | 14.3098 | 55.9967
0.6041 | 7.9934 | -79.589 | -1366.6059 | 1369.1493 | 7.9561 | -14.4589 | -3.7545 | 17.0861 | 1366.6059 | 14.4589 | 56.164
0.6059 | 7.6766 | -80.3064 | -1371.2827 | 1373.854 | 8.0213 | -14.5986 | -3.7358 | 17.2351 | 1371.2827 | 14.5986 | 56.3313
0.6077 | 7.5928 | -81.1821 | -1376.1551 | 1378.773 | 8.1052 | -14.7756 | -3.7172 | 17.4121 | 1376.1551 | 14.7756 | 56.4987
0.6095 | 7.6766 | -82.3094 | -1380.9623 | 1383.6361 | 8.1797 | -14.9526 | -3.6986 | 17.6078 | 1380.9623 | 14.9526 | 56.666
0.6113 | 7.2853 | -83.4273 | -1385.9559 | 1388.6855 | 8.2542 | -15.1296 | -3.6706 | 17.7941 | 1385.9559 | 15.1296 | 56.8334
0.6131 | 6.568 | -84.4242 | -1391.3593 | 1394.1356 | 8.3381 | -15.3067 | -3.6334 | 17.9804 | 1391.3593 | 15.3067 | 57.0007
0.6149 | 5.9065 | -85.2254 | -1397.6944 | 1400.5079 | 8.4126 | -15.4837 | -3.6147 | 18.1668 | 1397.6944 | 15.4837 | 57.1681
0.6167 | 5.4873 | -86.2222 | -1404.5512 | 1407.4206 | 8.4685 | -15.67 | -3.5868 | 18.3531 | 1404.5512 | 15.67 | 57.3354
0.6185 | 5.3475 | -87.4054 | -1412.5539 | 1415.4885 | 8.5337 | -15.8843 | -3.5775 | 18.5674 | 1412.5539 | 15.8843 | 57.5028
0.6203 | 5.4221 | -88.8028 | -1420.8081 | 1423.8266 | 8.5896 | -16.0799 | -3.5681 | 18.7723 | 1420.8081 | 16.0799 | 57.6701
0.6221 | 5.1985 | -90.0698 | -1428.1307 | 1431.2144 | 8.6455 | -16.2383 | -3.5681 | 18.9307 | 1428.1307 | 16.2383 | 57.8375
0.6239 | 5.3103 | -91.1971 | -1434.4192 | 1437.5588 | 8.72 | -16.3967 | -3.5588 | 19.1077 | 1434.4192 | 16.3967 | 58.0048
0.6256 | 5.478 | -92.5573 | -1440.3909 | 1443.5957 | 8.7759 | -16.5644 | -3.5495 | 19.2754 | 1440.3909 | 16.5644 | 58.1629
0.6274 | 5.655 | -93.9175 | -1446.3533 | 1449.642 | 8.8412 | -16.7227 | -3.5309 | 19.4431 | 1446.3533 | 16.7227 | 58.3302
0.6292 | 5.2917 | -95.1193 | -1453.6759 | 1457.0298 | 8.8971 | -16.8904 | -3.5216 | 19.6201 | 1453.6759 | 16.8904 | 58.4976
0.631 | 4.9097 | -96.1627 | -1460.4582 | 1463.868 | 8.953 | -17.0674 | -3.5029 | 19.7971 | 1460.4582 | 17.0674 | 58.6649
0.6328 | 4.6209 | -97.0664 | -1468.2094 | 1471.6657 | 9.0182 | -17.2631 | -3.4843 | 19.9927 | 1468.2094 | 17.2631 | 58.8323
0.6346 | 4.2669 | -97.9887 | -1475.5226 | 1479.0256 | 9.0834 | -17.4401 | -3.475 | 20.1791 | 1475.5226 | 17.4401 | 58.9996
0.6364 | 3.9128 | -98.8644 | -1483.1154 | 1486.6556 | 9.1486 | -17.6171 | -3.4657 | 20.3654 | 1483.1154 | 17.6171 | 59.167
0.6382 | 3.801 | -99.3768 | -1492.1615 | 1495.711 | 9.2231 | -17.8034 | -3.4657 | 20.561 | 1492.1615 | 17.8034 | 59.3343
0.64 | 3.6892 | -100.1128 | -1500.9934 | 1504.5615 | 9.307 | -17.9618 | -3.475 | 20.7287 | 1500.9934 | 17.9618 | 59.5017
0.6418 | 3.7079 | -100.6718 | -1510.6637 | 1514.2411 | 9.3815 | -18.1481 | -3.475 | 20.9337 | 1510.6637 | 18.1481 | 59.669
0.6436 | 3.6054 | -101.5382 | -1519.2533 | 1522.8494 | 9.4653 | -18.3065 | -3.4843 | 21.1014 | 1519.2533 | 18.3065 | 59.8364
0.6454 | 3.6613 | -102.6096 | -1527.6752 | 1531.3179 | 9.5399 | -18.4556 | -3.4936 | 21.2691 | 1527.6752 | 18.4556 | 60.0037
0.6472 | 3.801 | -103.6157 | -1536.0412 | 1539.7212 | 9.6237 | -18.6139 | -3.5122 | 21.4368 | 1536.0412 | 18.6139 | 60.1711
0.649 | 4.0619 | -104.4076 | -1544.4352 | 1548.1431 | 9.7076 | -18.7816 | -3.5216 | 21.6138 | 1544.4352 | 18.7816 | 60.3384
0.6507 | 4.36 | -105.274 | -1552.6522 | 1556.388 | 9.7914 | -18.9493 | -3.5309 | 21.7908 | 1552.6522 | 18.9493 | 60.4965
0.6525 | 4.6954 | -106.2336 | -1560.3474 | 1564.1205 | 9.8659 | -19.1077 | -3.5402 | 21.9678 | 1560.3474 | 19.1077 | 60.6638
0.6543 | 5.2264 | -107.5565 | -1566.6732 | 1570.5022 | 9.9312 | -19.294 | -3.5402 | 22.1448 | 1566.6732 | 19.294 | 60.8312
0.6561 | 5.4128 | -108.8142 | -1573.0083 | 1576.9025 | 10.0057 | -19.4524 | -3.5402 | 22.3125 | 1573.0083 | 19.4524 | 60.9985
0.6579 | 5.5339 | -109.6434 | -1579.8557 | 1583.7686 | 10.0895 | -19.6015 | -3.5495 | 22.4802 | 1579.8557 | 19.6015 | 61.1659
0.6597 | 5.655 | -110.193 | -1587.0199 | 1590.9514 | 10.1827 | -19.7598 | -3.5588 | 22.6572 | 1587.0199 | 19.7598 | 61.3332
0.6615 | 5.7016 | -110.6495 | -1594.3705 | 1598.3113 | 10.2852 | -19.9089 | -3.5588 | 22.8342 | 1594.3705 | 19.9089 | 61.5006
0.6633 | 5.7481 | -111.0874 | -1601.6465 | 1605.5966 | 10.3783 | -20.0486 | -3.5681 | 23.0019 | 1601.6465 | 20.0486 | 61.6679
0.6651 | 5.7295 | -111.5998 | -1608.3263 | 1612.295 | 10.4808 | -20.1791 | -3.5775 | 23.1603 | 1608.3263 | 20.1791 | 61.8353
0.6669 | 5.627 | -112.2706 | -1613.1055 | 1617.1022 | 10.574 | -20.3375 | -3.5681 | 23.3187 | 1613.1055 | 20.3375 | 62.0026
0.6687 | 5.4593 | -112.9786 | -1617.0184 | 1621.043 | 10.6578 | -20.5051 | -3.5402 | 23.4957 | 1617.0184 | 20.5051 | 62.17
0.6705 | 5.0587 | -113.3512 | -1621.2946 | 1625.3285 | 10.7603 | -20.6635 | -3.5122 | 23.6634 | 1621.2946 | 20.6635 | 62.3373
0.6723 | 4.6954 | -113.6214 | -1625.8689 | 1629.9028 | 10.8535 | -20.8033 | -3.5029 | 23.8217 | 1625.8689 | 20.8033 | 62.5046
0.6741 | 4.3507 | -113.8543 | -1630.7692 | 1634.8125 | 10.9373 | -20.9337 | -3.4936 | 23.9708 | 1630.7692 | 20.9337 | 62.672
0.6758 | 3.9501 | -114.1711 | -1636.7503 | 1640.7935 | 11.0398 | -21.0548 | -3.5029 | 24.1199 | 1636.7503 | 21.0548 | 62.83
0.6776 | 3.4657 | -114.5624 | -1643.7095 | 1647.7621 | 11.1609 | -21.1666 | -3.5122 | 24.2782 | 1643.7095 | 21.1666 | 62.9974
0.6794 | 3.1023 | -115.1307 | -1650.4638 | 1654.5444 | 11.2913 | -21.2598 | -3.5402 | 24.4273 | 1650.4638 | 21.2598 | 63.1647
0.6812 | 2.8508 | -115.7828 | -1654.3767 | 1658.4852 | 11.3845 | -21.3343 | -3.5402 | 24.5391 | 1654.3767 | 21.3343 | 63.3321
0.683 | 2.5992 | -116.2766 | -1658.2243 | 1662.3514 | 11.5149 | -21.4181 | -3.5495 | 24.6788 | 1658.2243 | 21.4181 | 63.4994
0.6848 | 2.534 | -116.5374 | -1660.6279 | 1664.7737 | 11.636 | -21.4834 | -3.5588 | 24.7813 | 1660.6279 | 21.4834 | 63.6668
0.6866 | 2.4968 | -116.8542 | -1663.1899 | 1667.3543 | 11.7385 | -21.5206 | -3.5681 | 24.8745 | 1663.1899 | 21.5206 | 63.8341
0.6884 | 2.2825 | -117.3479 | -1665.5562 | 1669.7392 | 11.7665 | -21.5393 | -3.5868 | 24.9117 | 1665.5562 | 21.5393 | 64.0015
0.6902 | 1.9378 | -117.9628 | -1668.1275 | 1672.3478 | 11.7758 | -21.5486 | -3.5961 | 24.9211 | 1668.1275 | 21.5486 | 64.1688
0.692 | 1.5745 | -118.5591 | -1671.6863 | 1675.9346 | 11.8131 | -21.5672 | -3.6147 | 24.9583 | 1671.6863 | 21.5672 | 64.3362
0.6938 | 1.3136 | -118.9969 | -1676.0836 | 1680.3598 | 11.9155 | -21.5858 | -3.6427 | 25.0329 | 1676.0836 | 21.5858 | 64.5035
0.6956 | 1.0621 | -119.2764 | -1680.8536 | 1685.1297 | 12.0646 | -21.6231 | -3.6613 | 25.1447 | 1680.8536 | 21.6231 | 64.6709
0.6974 | 0.86641 | -119.4534 | -1685.521 | 1689.8065 | 12.2043 | -21.651 | -3.6799 | 25.2378 | 1685.521 | 21.651 | 64.8382
0.6992 | 0.63351 | -119.5279 | -1689.4991 | 1693.7753 | 12.3441 | -21.6604 | -3.7079 | 25.3123 | 1689.4991 | 21.6604 | 65.0056
0.7009 | -0.16769 | -119.1926 | -1692.0517 | 1696.3 | 12.549 | -21.5858 | -3.7172 | 25.3496 | 1692.0517 | 21.5858 | 65.1636
0.7027 | -0.93163 | -118.6988 | -1694.0175 | 1698.2284 | 12.726 | -21.4927 | -3.7451 | 25.3682 | 1694.0175 | 21.4927 | 65.331
0.7045 | -1.5651 | -118.1398 | -1695.8155 | 1699.9799 | 12.8658 | -21.3995 | -3.7731 | 25.3682 | 1695.8155 | 21.3995 | 65.4983
0.7063 | -1.1273 | -118.0467 | -1696.4304 | 1700.5761 | 13.0149 | -21.3343 | -3.801 | 25.4055 | 1696.4304 | 21.3343 | 65.6657
0.7081 | -0.50308 | -118.2516 | -1696.6633 | 1700.8277 | 13.1266 | -21.2784 | -3.829 | 25.4335 | 1696.6633 | 21.2784 | 65.833
0.7099 | -1.3602 | -118.5497 | -1698.0049 | 1702.1879 | 13.1825 | -21.1852 | -3.8663 | 25.3962 | 1698.0049 | 21.1852 | 66.0004
0.7117 | -2.4222 | -118.8013 | -1700.0451 | 1704.2468 | 13.2478 | -21.0828 | -3.9035 | 25.3589 | 1700.0451 | 21.0828 | 66.1677
0.7135 | -3.8197 | -118.9783 | -1703.6692 | 1707.8801 | 13.3409 | -20.943 | -3.9501 | 25.3123 | 1703.6692 | 20.943 | 66.3351
0.7153 | -4.9749 | -119.3789 | -1708.5602 | 1712.7898 | 13.4807 | -20.8126 | -3.9874 | 25.2937 | 1708.5602 | 20.8126 | 66.5024
0.7171 | -6.1674 | -120.1242 | -1713.4699 | 1717.7461 | 13.6018 | -20.6728 | -4.0246 | 25.2658 | 1713.4699 | 20.6728 | 66.6698
0.7189 | -6.5307 | -120.3385 | -1719.8795 | 1724.1464 | 13.7229 | -20.6542 | -4.0619 | 25.331 | 1719.8795 | 20.6542 | 66.8371
0.7207 | -6.7077 | -120.7856 | -1726.9692 | 1731.264 | 13.8254 | -20.6822 | -4.0898 | 25.4055 | 1726.9692 | 20.6822 | 67.0045
0.7225 | -6.9406 | -120.9254 | -1734.534 | 1738.8102 | 14.0024 | -20.7287 | -4.1085 | 25.5453 | 1734.534 | 20.7287 | 67.1718
0.7242 | -7.2294 | -120.9813 | -1742.2106 | 1746.4775 | 14.1887 | -20.766 | -4.1178 | 25.685 | 1742.2106 | 20.766 | 67.3299
0.726 | -7.8629 | -121.708 | -1750.1202 | 1754.4243 | 14.3191 | -20.7567 | -4.1457 | 25.7502 | 1750.1202 | 20.7567 | 67.4972
0.7278 | -8.7573 | -122.8725 | -1757.9272 | 1762.2965 | 14.4402 | -20.6728 | -4.1737 | 25.7688 | 1757.9272 | 20.6728 | 67.6646
0.7296 | -9.4933 | -123.329 | -1767.0572 | 1771.4545 | 14.5614 | -20.5983 | -4.2203 | 25.7875 | 1767.0572 | 20.5983 | 67.8319
0.7314 | -9.7169 | -124.1022 | -1775.2369 | 1779.6714 | 14.7104 | -20.5704 | -4.2669 | 25.8527 | 1775.2369 | 20.5704 | 67.9993
0.7332 | -9.9684 | -125.1084 | -1782.4663 | 1786.9568 | 14.8781 | -20.5331 | -4.3041 | 25.9179 | 1782.4663 | 20.5331 | 68.1666
0.735 | -10.6765 | -126.4127 | -1788.2517 | 1792.8167 | 15.0272 | -20.4679 | -4.3321 | 25.9645 | 1788.2517 | 20.4679 | 68.334
0.7368 | -11.4218 | -127.5772 | -1794.4284 | 1799.0679 | 15.1855 | -20.4213 | -4.3507 | 26.0111 | 1794.4284 | 20.4213 | 68.5013
0.7386 | -12.3813 | -128.22 | -1802.0957 | 1806.7725 | 15.3626 | -20.3468 | -4.3693 | 26.067 | 1802.0957 | 20.3468 | 68.6687
0.7404 | -13.8067 | -129.2541 | -1808.9525 | 1813.7038 | 15.5302 | -20.1884 | -4.4159 | 26.0577 | 1808.9525 | 20.1884 | 68.836
0.7422 | -14.4123 | -130.074 | -1815.7907 | 1820.5885 | 15.6793 | -20.0859 | -4.4532 | 26.0763 | 1815.7907 | 20.0859 | 69.0033
0.744 | -14.6918 | -130.4932 | -1822.7592 | 1827.5851 | 15.8191 | -20.03 | -4.4811 | 26.1135 | 1822.7592 | 20.03 | 69.1707
0.7458 | -15.5396 | -131.0242 | -1829.262 | 1834.1158 | 16.0333 | -19.9275 | -4.537 | 26.1508 | 1829.262 | 19.9275 | 69.338
0.7476 | -16.0613 | -131.8534 | -1835.5971 | 1840.5068 | 16.2103 | -19.8251 | -4.565 | 26.1788 | 1835.5971 | 19.8251 | 69.5054
0.7493 | -16.5178 | -132.5894 | -1840.8049 | 1845.7611 | 16.2849 | -19.7505 | -4.5743 | 26.1694 | 1840.8049 | 19.7505 | 69.6634
0.7511 | -17.1792 | -133.3347 | -1846.7114 | 1851.7049 | 16.3594 | -19.676 | -4.5929 | 26.1508 | 1846.7114 | 19.676 | 69.8308
0.7529 | -18.9959 | -133.875 | -1854.8352 | 1859.8567 | 16.5178 | -19.4897 | -4.6675 | 26.1042 | 1854.8352 | 19.4897 | 69.9981
0.7547 | -21.1387 | -134.3129 | -1863.3503 | 1868.3904 | 16.7041 | -19.2661 | -4.7606 | 26.039 | 1863.3503 | 19.2661 | 70.1655
0.7565 | -22.6945 | -134.7694 | -1870.617 | 1875.6851 | 16.7973 | -19.0611 | -4.8258 | 25.9552 | 1870.617 | 19.0611 | 70.3328
0.7583 | -23.859 | -135.3097 | -1876.9893 | 1882.0853 | 16.8159 | -18.8748 | -4.891 | 25.8341 | 1876.9893 | 18.8748 | 70.5002
0.7601 | -25.1167 | -136.5674 | -1882.9424 | 1888.113 | 16.76 | -18.6978 | -4.9376 | 25.6757 | 1882.9424 | 18.6978 | 70.6675
0.7619 | -26.8123 | -138.2723 | -1889.1937 | 1894.476 | 16.7507 | -18.4649 | -5.0028 | 25.5173 | 1889.1937 | 18.4649 | 70.8349
0.7637 | -28.5358 | -139.8468 | -1894.5505 | 1899.9447 | 16.7879 | -18.204 | -5.0774 | 25.3682 | 1894.5505 | 18.204 | 71.0022
0.7655 | -30.3152 | -141.2722 | -1898.9665 | 1904.4724 | 16.8718 | -17.9059 | -5.1612 | 25.2192 | 1898.9665 | 17.9059 | 71.1696
0.7673 | -32.1785 | -142.2131 | -1902.4694 | 1908.0685 | 16.9929 | -17.5519 | -5.2451 | 25.0701 | 1902.4694 | 17.5519 | 71.3369
0.7691 | -34.1442 | -142.6789 | -1905.0873 | 1910.7422 | 17.142 | -17.1513 | -5.3475 | 24.8931 | 1905.0873 | 17.1513 | 71.5043
0.7709 | -35.9609 | -143.0609 | -1907.202 | 1912.9129 | 17.2817 | -16.76 | -5.4407 | 24.7347 | 1907.202 | 16.76 | 71.6716
0.7727 | -37.2186 | -143.1447 | -1907.2859 | 1913.0247 | 17.4028 | -16.3873 | -5.4966 | 24.5857 | 1907.2859 | 16.3873 | 71.839
0.7744 | -38.2806 | -143.3497 | -1905.935 | 1911.7111 | 17.5332 | -16.024 | -5.5432 | 24.4459 | 1905.935 | 16.024 | 71.997
0.7762 | -39.4172 | -143.4242 | -1903.4383 | 1909.2516 | 17.645 | -15.6048 | -5.5991 | 24.2689 | 1903.4383 | 15.6048 | 72.1644
0.778 | -40.3954 | -142.8187 | -1899.3298 | 1905.1245 | 17.7848 | -15.2228 | -5.627 | 24.1292 | 1899.3298 | 15.2228 | 72.3317
0.7798 | -41.4109 | -141.8777 | -1893.6934 | 1899.4602 | 17.9059 | -14.8315 | -5.6457 | 23.9894 | 1893.6934 | 14.8315 | 72.4991
0.7816 | -42.3984 | -140.8622 | -1886.6969 | 1892.4357 | 18.027 | -14.4309 | -5.655 | 23.8311 | 1886.6969 | 14.4309 | 72.6664
0.7834 | -43.3114 | -139.6139 | -1877.5669 | 1883.2592 | 18.1388 | -14.021 | -5.655 | 23.6727 | 1877.5669 | 14.021 | 72.8338
0.7852 | -44.1685 | -138.1698 | -1867.0768 | 1872.7132 | 18.2506 | -13.6297 | -5.6457 | 23.5236 | 1867.0768 | 13.6297 | 73.0011
0.787 | -44.8579 | -136.5115 | -1854.9563 | 1860.5275 | 18.3438 | -13.2384 | -5.6177 | 23.3746 | 1854.9563 | 13.2384 | 73.1685
0.7888 | -45.5753 | -134.9371 | -1841.1216 | 1846.6369 | 18.4276 | -12.8658 | -5.5805 | 23.2162 | 1841.1216 | 12.8658 | 73.3358
0.7906 | -46.4976 | -133.6701 | -1825.7032 | 1831.1905 | 18.4556 | -12.5025 | -5.5246 | 23.0299 | 1825.7032 | 12.5025 | 73.5032
0.7924 | -47.4572 | -132.3378 | -1809.7537 | 1815.2224 | 18.4742 | -12.1484 | -5.4687 | 22.8435 | 1809.7537 | 12.1484 | 73.6705
0.7942 | -48.4447 | -130.6702 | -1792.7049 | 1798.127 | 18.5021 | -11.8037 | -5.4034 | 22.6758 | 1792.7049 | 11.8037 | 73.8379
0.796 | -49.1993 | -128.4064 | -1773.532 | 1778.8702 | 18.5208 | -11.487 | -5.3289 | 22.5175 | 1773.532 | 11.487 | 74.0052
0.7977 | -49.7396 | -125.7699 | -1752.943 | 1758.1694 | 18.5394 | -11.1795 | -5.2637 | 22.3684 | 1752.943 | 11.1795 | 74.1633
0.7995 | -50.1868 | -122.9377 | -1730.6398 | 1735.7451 | 18.5487 | -10.8721 | -5.1892 | 22.21 | 1730.6398 | 10.8721 | 74.3306
0.8013 | -50.5036 | -119.9192 | -1705.7374 | 1710.7123 | 18.5301 | -10.5926 | -5.1053 | 22.061 | 1705.7374 | 10.5926 | 74.498
0.8031 | -51.0532 | -116.8821 | -1678.6177 | 1683.4808 | 18.4742 | -10.2945 | -5.0308 | 21.8653 | 1678.6177 | 10.2945 | 74.6653
0.8049 | -51.6402 | -113.7705 | -1650.4825 | 1655.2245 | 18.409 | -10.015 | -4.9469 | 21.651 | 1650.4825 | 10.015 | 74.8327
0.8067 | -52.3296 | -110.4073 | -1620.8846 | 1625.5148 | 18.3158 | -9.7355 | -4.8631 | 21.4461 | 1620.8846 | 9.7355 | 75
0.8085 | -52.9631 | -106.8857 | -1589.8614 | 1594.3612 | 18.2133 | -9.4747 | -4.7699 | 21.2225 | 1589.8614 | 9.4747 | 75.1673
0.8103 | -53.6245 | -103.1499 | -1557.1426 | 1561.512 | 18.0829 | -9.1952 | -4.6861 | 20.9803 | 1557.1426 | 9.1952 | 75.3347
0.8121 | -54.1463 | -99.2278 | -1523.3152 | 1527.5448 | 17.9432 | -8.925 | -4.6022 | 20.7194 | 1523.3152 | 8.925 | 75.502
0.8139 | -54.2953 | -94.7932 | -1487.5965 | 1491.6398 | 17.7662 | -8.6548 | -4.5091 | 20.4399 | 1487.5965 | 8.6548 | 75.6694
0.8157 | -53.9786 | -89.7158 | -1449.4929 | 1453.3126 | 17.5612 | -8.394 | -4.4066 | 20.1418 | 1449.4929 | 8.394 | 75.8367
0.8175 | -53.5034 | -84.4242 | -1410.7558 | 1414.3426 | 17.3376 | -8.1424 | -4.3134 | 19.8251 | 1410.7558 | 8.1424 | 76.0041
0.8193 | -52.9351 | -79.3002 | -1371.9162 | 1375.2887 | 17.0674 | -7.8816 | -4.211 | 19.471 | 1371.9162 | 7.8816 | 76.1714
0.8211 | -52.3389 | -74.2228 | -1333.2909 | 1336.4305 | 16.76 | -7.6207 | -4.1085 | 19.0704 | 1333.2909 | 7.6207 | 76.3388
0.8228 | -51.5191 | -68.7169 | -1293.1377 | 1296.0444 | 16.4339 | -7.3319 | -4.0153 | 18.6419 | 1293.1377 | 7.3319 | 76.4968
0.8246 | -50.6061 | -62.9315 | -1251.3542 | 1254.0093 | 16.0706 | -7.0338 | -3.9035 | 18.1854 | 1251.3542 | 7.0338 | 76.6642
0.8264 | -49.6558 | -57.0157 | -1208.1173 | 1210.5488 | 15.67 | -6.7264 | -3.7917 | 17.6916 | 1208.1173 | 6.7264 | 76.8315
0.8282 | -48.5099 | -50.8576 | -1164.0326 | 1166.2313 | 15.2321 | -6.391 | -3.6706 | 17.1606 | 1164.0326 | 6.391 | 76.9989
0.83 | -47.3826 | -44.7834 | -1120.7305 | 1122.6776 | 14.7943 | -6.0463 | -3.5588 | 16.6109 | 1120.7305 | 6.0463 | 77.1662
0.8318 | -46.1715 | -38.97 | -1076.4502 | 1078.2203 | 14.3005 | -5.7016 | -3.4377 | 16.0147 | 1076.4502 | 5.7016 | 77.3336
0.8336 | -44.9138 | -33.138 | -1031.76 | 1033.3531 | 13.7788 | -5.3289 | -3.3073 | 15.3905 | 1031.76 | 5.3289 | 77.5009
0.8354 | -43.6468 | -27.3992 | -987.7778 | 989.1939 | 13.2478 | -4.9749 | -3.1769 | 14.757 | 987.7778 | 4.9749 | 77.6683
0.8372 | -42.4077 | -21.8746 | -944.8018 | 946.0874 | 12.7074 | -4.6022 | -3.0557 | 14.1142 | 944.8018 | 4.6022 | 77.8356
0.839 | -41.0662 | -16.6109 | -901.863 | 903.0462 | 12.1391 | -4.2389 | -2.9253 | 13.4434 | 901.863 | 4.2389 | 78.003
0.8408 | -39.6501 | -11.487 | -858.9522 | 860.0422 | 11.5615 | -3.8756 | -2.7856 | 12.7726 | 858.9522 | 3.8756 | 78.1703
0.8426 | -38.2154 | -6.3258 | -815.2495 | 816.2557 | 10.9746 | -3.4657 | -2.6551 | 12.0739 | 815.2495 | 3.4657 | 78.3377
0.8444 | -36.6223 | -1.6583 | -772.348 | 773.3169 | 10.369 | -3.0744 | -2.534 | 11.3659 | 772.348 | 3.0744 | 78.505
0.8462 | -35.0851 | 2.5713 | -730.9744 | 731.9247 | 9.7542 | -2.711 | -2.4036 | 10.6671 | 730.9744 | 2.711 | 78.6724
0.8479 | -33.8088 | 6.2605 | -691.7342 | 692.6752 | 9.1486 | -2.3384 | -2.2918 | 9.9684 | 691.7342 | 2.3384 | 78.8304
0.8497 | -32.7095 | 9.4374 | -654.4504 | 655.4007 | 8.543 | -1.9564 | -2.1707 | 9.279 | 654.4504 | 1.9564 | 78.9978
0.8515 | -31.6847 | 11.841 | -620.4274 | 621.4149 | 7.9561 | -1.6303 | -2.0682 | 8.6269 | 620.4274 | 1.6303 | 79.1651
0.8533 | -30.7344 | 13.8813 | -588.4446 | 589.4787 | 7.3785 | -1.3229 | -1.9657 | 8.0027 | 588.4446 | 1.3229 | 79.3325
0.8551 | -29.9053 | 15.5209 | -559.0889 | 560.1696 | 6.8382 | -1.0341 | -1.8726 | 7.4251 | 559.0889 | 1.0341 | 79.4998
0.8569 | -29.1693 | 16.583 | -531.55 | 532.6773 | 6.3164 | -0.76394 | -1.7887 | 6.8754 | 531.55 | 0.76394 | 79.6672
0.8587 | -28.4892 | 17.5612 | -505.0358 | 506.2097 | 5.8134 | -0.5124 | -1.7049 | 6.3444 | 505.0358 | 0.5124 | 79.8345
0.8605 | -28.0141 | 18.4649 | -480.9812 | 482.2296 | 5.3382 | -0.27949 | -1.6303 | 5.8599 | 480.9812 | 0.27949 | 80.0019
0.8623 | -27.7812 | 19.2847 | -459.9264 | 461.2493 | 4.919 | -0.07453 | -1.5558 | 5.4221 | 459.9264 | 0.07453 | 80.1692
0.8641 | -27.6787 | 19.7598 | -441.1634 | 442.5608 | 4.5463 | 0.07453 | -1.4999 | 5.0494 | 441.1634 | -0.07453 | 80.3366
0.8659 | -27.6787 | 19.9648 | -423.9935 | 425.4468 | 4.1923 | 0.22359 | -1.444 | 4.7047 | 423.9935 | -0.22359 | 80.5039
0.8677 | -27.7625 | 20.058 | -409.0688 | 410.578 | 3.8476 | 0.4006 | -1.3974 | 4.36 | 409.0688 | -0.4006 | 80.6713
0.8695 | -27.7625 | 20.1698 | -395.4111 | 396.9669 | 3.5309 | 0.56829 | -1.3509 | 4.0526 | 395.4111 | -0.56829 | 80.8386
0.8712 | -27.8277 | 20.1418 | -383.3838 | 384.9769 | 3.27 | 0.68941 | -1.3043 | 3.8197 | 383.3838 | -0.68941 | 80.9967
0.873 | -28.0141 | 20.1418 | -373.2011 | 374.8501 | 3.0557 | 0.8012 | -1.2763 | 3.6334 | 373.2011 | -0.8012 | 81.164
0.8748 | -28.3215 | 20.5424 | -365.3847 | 367.1082 | 2.8974 | 0.86641 | -1.2577 | 3.5122 | 365.3847 | -0.86641 | 81.3313
0.8766 | -28.3588 | 21.2877 | -357.5684 | 359.3757 | 2.8228 | 0.89436 | -1.2484 | 3.475 | 357.5684 | -0.89436 | 81.4987
0.8784 | -28.0141 | 21.5579 | -348.7365 | 350.5998 | 2.7669 | 0.92231 | -1.2111 | 3.447 | 348.7365 | -0.92231 | 81.666
0.8802 | -27.8277 | 21.3622 | -339.8302 | 341.7214 | 2.6738 | 0.98753 | -1.1552 | 3.3818 | 339.8302 | -0.98753 | 81.8334
0.882 | -27.8836 | 20.9616 | -331.902 | 333.8118 | 2.534 | 1.09 | -1.1086 | 3.2793 | 331.902 | -1.09 | 82.0007
0.8838 | -27.8464 | 20.5331 | -325.8744 | 327.7842 | 2.4036 | 1.1832 | -1.0621 | 3.2048 | 325.8744 | -1.1832 | 82.1681
0.8856 | -27.6973 | 20.2443 | -320.9181 | 322.8093 | 2.3104 | 1.2763 | -1.0155 | 3.1303 | 320.9181 | -1.2763 | 82.3354
0.8874 | -27.6601 | 20.384 | -316.9307 | 318.8406 | 2.2266 | 1.3881 | -0.98753 | 3.0557 | 316.9307 | -1.3881 | 82.5028
0.8892 | -27.7625 | 20.2443 | -313.5862 | 315.5053 | 2.1427 | 1.5092 | -0.95026 | 3.0185 | 313.5862 | -1.5092 | 82.6701
0.891 | -27.9302 | 19.9275 | -310.7633 | 312.6918 | 2.0589 | 1.6303 | -0.913 | 2.9905 | 310.7633 | -1.6303 | 82.8375
0.8928 | -28.1445 | 19.294 | -308.6672 | 310.5863 | 1.9657 | 1.7421 | -0.87573 | 2.9626 | 308.6672 | -1.7421 | 83.0048
0.8946 | -28.3215 | 19.1636 | -308.0057 | 309.9342 | 1.9005 | 1.8353 | -0.8571 | 2.9626 | 308.0057 | -1.8353 | 83.1722
0.8963 | -28.3495 | 19.4897 | -307.6424 | 309.5988 | 1.8726 | 1.9378 | -0.82915 | 3.0092 | 307.6424 | -1.9378 | 83.3302
0.8981 | -27.9116 | 19.3965 | -305.0431 | 306.953 | 1.8167 | 2.0496 | -0.8012 | 3.0464 | 305.0431 | -2.0496 | 83.4976
0.8999 | -27.9209 | 19.2288 | -304.5028 | 306.4126 | 1.7608 | 2.1614 | -0.77325 | 3.0557 | 304.5028 | -2.1614 | 83.6649
0.9017 | -27.7905 | 19.0891 | -304.9686 | 306.8505 | 1.6956 | 2.3104 | -0.73599 | 3.1116 | 304.9686 | -2.3104 | 83.8323
0.9035 | -27.8184 | 19.2009 | -307.3163 | 309.2075 | 1.6583 | 2.4781 | -0.68009 | 3.2048 | 307.3163 | -2.4781 | 83.9996
0.9053 | -27.8184 | 19.4338 | -309.2727 | 311.1546 | 1.621 | 2.6365 | -0.63351 | 3.298 | 309.2727 | -2.6365 | 84.167
0.9071 | -27.5576 | 19.7971 | -308.928 | 310.8192 | 1.5838 | 2.739 | -0.59624 | 3.3445 | 308.928 | -2.739 | 84.3343
0.9089 | -27.3992 | 19.9462 | -308.2852 | 310.1764 | 1.5372 | 2.7949 | -0.54966 | 3.3725 | 308.2852 | -2.7949 | 84.5017
0.9107 | -27.3433 | 20.0114 | -307.4467 | 309.3473 | 1.472 | 2.8322 | -0.5124 | 3.3725 | 307.4467 | -2.8322 | 84.669
0.9125 | -27.1663 | 20.4586 | -305.5183 | 307.4467 | 1.4347 | 2.8601 | -0.47513 | 3.3725 | 305.5183 | -2.8601 | 84.8364
0.9143 | -26.9334 | 21.0734 | -302.8631 | 304.8289 | 1.3881 | 2.916 | -0.43787 | 3.3911 | 302.8631 | -2.916 | 85.0037
0.9161 | -26.5235 | 21.6231 | -300.0776 | 302.0619 | 1.3509 | 2.9626 | -0.4006 | 3.4004 | 300.0776 | -2.9626 | 85.1711
0.9179 | -26.0483 | 22.3964 | -296.8262 | 298.8571 | 1.3229 | 2.9719 | -0.36334 | 3.3911 | 296.8262 | -2.9719 | 85.3384
0.9197 | -25.4894 | 23.2721 | -293.1556 | 295.2331 | 1.295 | 2.9719 | -0.3447 | 3.3632 | 293.1556 | -2.9719 | 85.5058
0.9214 | -24.8559 | 23.7472 | -289.4477 | 291.5345 | 1.267 | 2.9812 | -0.29812 | 3.3539 | 289.4477 | -2.9812 | 85.6638
0.9232 | -24.2037 | 24.2503 | -285.423 | 287.5192 | 1.2297 | 2.9719 | -0.26086 | 3.3352 | 285.423 | -2.9719 | 85.8312
0.925 | -23.5329 | 24.8093 | -280.5693 | 282.6934 | 1.2018 | 2.9533 | -0.24222 | 3.2886 | 280.5693 | -2.9533 | 85.9985
0.9268 | -22.7876 | 25.0888 | -274.9795 | 277.1129 | 1.1552 | 2.9346 | -0.19564 | 3.2514 | 274.9795 | -2.9346 | 86.1659
0.9286 | -22.0051 | 25.303 | -268.8307 | 270.9828 | 1.1086 | 2.9253 | -0.15838 | 3.2141 | 268.8307 | -2.9253 | 86.3332
0.9304 | -21.2039 | 25.4707 | -261.592 | 263.772 | 1.0714 | 2.9346 | -0.10248 | 3.1955 | 261.592 | -2.9346 | 86.5006
0.9322 | -20.4586 | 25.4521 | -254.2973 | 256.496 | 1.0155 | 2.9719 | -0.037265 | 3.1955 | 254.2973 | -2.9719 | 86.6679
0.934 | -19.8157 | 25.0142 | -248.5399 | 250.6733 | 0.95958 | 2.9533 | 0.018633 | 3.1396 | 248.5399 | -2.9533 | 86.8353
0.9358 | -19.2102 | 24.6788 | -242.5868 | 244.6922 | 0.90368 | 2.916 | 0.055898 | 3.0837 | 242.5868 | -2.916 | 87.0026
0.9376 | -18.6605 | 24.1571 | -236.056 | 238.1243 | 0.8571 | 2.8415 | 0.065214 | 2.9905 | 236.056 | -2.8415 | 87.17
0.9394 | -18.0922 | 23.4305 | -230.429 | 232.4227 | 0.8012 | 2.7669 | 0.083847 | 2.9067 | 230.429 | -2.7669 | 87.3373
0.9412 | -17.4028 | 22.7317 | -225.5193 | 227.4198 | 0.73599 | 2.6924 | 0.10248 | 2.8135 | 225.5193 | -2.6924 | 87.5046
0.943 | -16.6575 | 21.9771 | -220.8891 | 222.6872 | 0.66146 | 2.6086 | 0.12111 | 2.711 | 220.8891 | -2.6086 | 87.672
0.9447 | -16.0426 | 20.9616 | -217.0042 | 218.6905 | 0.58693 | 2.534 | 0.13043 | 2.6179 | 217.0042 | -2.534 | 87.83
0.9465 | -15.5861 | 20.0114 | -213.8553 | 215.4298 | 0.52171 | 2.4409 | 0.13043 | 2.5247 | 213.8553 | -2.4409 | 87.9974
0.9483 | -15.2414 | 18.8655 | -211.4797 | 212.933 | 0.43787 | 2.3757 | 0.13974 | 2.4409 | 211.4797 | -2.3757 | 88.1647
0.9501 | -14.9154 | 18.1015 | -209.3463 | 210.7064 | 0.38197 | 2.2825 | 0.14906 | 2.3384 | 209.3463 | -2.2825 | 88.3321
0.9519 | -14.7197 | 17.0954 | -208.1817 | 209.4487 | 0.30744 | 2.2173 | 0.15838 | 2.2639 | 208.1817 | -2.2173 | 88.4994
0.9537 | -14.6266 | 16.3501 | -207.7159 | 208.9177 | 0.25154 | 2.1521 | 0.17701 | 2.1893 | 207.7159 | -2.1521 | 88.6668
0.9555 | -14.5986 | 15.9308 | -207.5109 | 208.6755 | 0.19564 | 2.0775 | 0.18633 | 2.1148 | 207.5109 | -2.0775 | 88.8341
0.9573 | -14.4309 | 15.5209 | -207.8184 | 208.9363 | 0.15838 | 1.9937 | 0.19564 | 2.0216 | 207.8184 | -1.9937 | 89.0015
0.9591 | -14.4216 | 15.2135 | -208.4892 | 209.5885 | 0.1118 | 1.9098 | 0.19564 | 1.9378 | 208.4892 | -1.9098 | 89.1688
0.9609 | -14.4589 | 14.9061 | -209.2345 | 210.3058 | 0.07453 | 1.8167 | 0.19564 | 1.8446 | 209.2345 | -1.8167 | 89.3362
0.9627 | -14.6173 | 14.934 | -209.2065 | 210.2872 | 0.046581 | 1.7142 | 0.18633 | 1.7328 | 209.2065 | -1.7142 | 89.5035
0.9645 | -14.7197 | 15.1669 | -208.6103 | 209.7189 | 0.037265 | 1.6117 | 0.16769 | 1.6397 | 208.6103 | -1.6117 | 89.6709
0.9663 | -14.7384 | 15.5489 | -207.1756 | 208.3121 | 0.037265 | 1.5092 | 0.14906 | 1.5279 | 207.1756 | -1.5092 | 89.8382
0.9681 | -14.4682 | 15.9774 | -204.39 | 205.5638 | 0.055898 | 1.4068 | 0.13043 | 1.4254 | 204.39 | -1.4068 | 90.0056
0.9698 | -14.0862 | 16.3314 | -201.269 | 202.4615 | 0.065214 | 1.295 | 0.1118 | 1.3136 | 201.269 | -1.295 | 90.1636
0.9716 | -13.7136 | 16.6575 | -197.822 | 199.0331 | 0.07453 | 1.1925 | 0.093163 | 1.2111 | 197.822 | -1.1925 | 90.331
0.9734 | -13.3875 | 16.9277 | -193.844 | 195.083 | 0.083847 | 1.09 | 0.065214 | 1.118 | 193.844 | -1.09 | 90.4983
0.9752 | -13.1173 | 16.965 | -188.7014 | 189.9591 | 0.083847 | 0.99684 | 0.037265 | 1.0248 | 188.7014 | -0.99684 | 90.6657
0.977 | -12.9217 | 16.8252 | -182.7017 | 183.9594 | 0.083847 | 0.89436 | 0.0093163 | 0.92231 | 182.7017 | -0.89436 | 90.833
0.9788 | -12.6143 | 16.5178 | -176.7299 | 177.9876 | 0.083847 | 0.8012 | -0.018633 | 0.81983 | 176.7299 | -0.8012 | 91.0004
0.9806 | -12.3627 | 15.8936 | -171.8855 | 173.1059 | 0.07453 | 0.71735 | -0.037265 | 0.7453 | 171.8855 | -0.71735 | 91.1677
0.9824 | -12.1391 | 14.7011 | -168.308 | 169.4446 | 0.037265 | 0.65214 | -0.065214 | 0.68009 | 168.308 | -0.65214 | 91.3351
0.9842 | -11.9062 | 13.4993 | -164.8144 | 165.8485 | 0.0093163 | 0.58693 | -0.083847 | 0.61487 | 164.8144 | -0.58693 | 91.5024
0.986 | -11.6267 | 12.3627 | -162.327 | 163.2679 | -0.037265 | 0.52171 | -0.10248 | 0.56829 | 162.327 | -0.52171 | 91.6698
0.9878 | -11.3379 | 11.4124 | -161.0227 | 161.8891 | -0.07453 | 0.4565 | -0.1118 | 0.53103 | 161.0227 | -0.4565 | 91.8371
0.9896 | -11.0305 | 10.6858 | -160.2308 | 161.0227 | -0.10248 | 0.38197 | -0.13043 | 0.49376 | 160.2308 | -0.38197 | 92.0045
0.9914 | -10.6765 | 10.3877 | -160.2028 | 160.9575 | -0.1118 | 0.29812 | -0.14906 | 0.44718 | 160.2028 | -0.29812 | 92.1718
0.9932 | -10.2386 | 10.4156 | -161.3581 | 162.0754 | -0.13043 | 0.20496 | -0.16769 | 0.43787 | 161.3581 | -0.20496 | 92.3392
0.9949 | -9.7355 | 10.546 | -162.6996 | 163.4077 | -0.13974 | 0.10248 | -0.19564 | 0.46581 | 162.6996 | -0.10248 | 92.4972
0.9967 | -9.3536 | 10.9653 | -162.7276 | 163.4635 | -0.13043 | 0.027949 | -0.21427 | 0.49376 | 162.7276 | -0.027949 | 92.6646
0.9985 | -8.9623 | 11.3752 | -162.4108 | 163.1654 | -0.1118 | -0.046581 | -0.24222 | 0.52171 | 162.4108 | 0.046581 | 92.8319
1.0003 | -8.6176 | 11.5801 | -161.8705 | 162.6158 | -0.10248 | -0.10248 | -0.26086 | 0.54966 | 161.8705 | 0.10248 | 92.9993
1.0021 | -8.366 | 11.7758 | -160.7059 | 161.4605 | -0.093163 | -0.15838 | -0.2888 | 0.54034 | 160.7059 | 0.15838 | 93.1666
1.0039 | -8.2076 | 11.8969 | -158.7122 | 159.4762 | -0.083847 | -0.21427 | -0.30744 | 0.54966 | 158.7122 | 0.21427 | 93.334
1.0057 | -8.04 | 11.5708 | -155.9733 | 156.7186 | -0.093163 | -0.24222 | -0.32607 | 0.55898 | 155.9733 | 0.24222 | 93.5013
1.0075 | -7.9468 | 11.2261 | -152.8709 | 153.5976 | -0.093163 | -0.27949 | -0.3447 | 0.55898 | 152.8709 | 0.27949 | 93.6687
1.0093 | -7.8723 | 10.574 | -150.5791 | 151.2592 | -0.12111 | -0.29812 | -0.36334 | 0.58693 | 150.5791 | 0.29812 | 93.836
1.0111 | -7.6859 | 9.9312 | -149.6102 | 150.2624 | -0.13974 | -0.32607 | -0.37265 | 0.60556 | 149.6102 | 0.32607 | 94.0033
1.0129 | -7.4717 | 9.4933 | -149.3587 | 149.9736 | -0.15838 | -0.35402 | -0.38197 | 0.61487 | 149.3587 | 0.35402 | 94.1707
1.0147 | -7.304 | 9.0647 | -149.6382 | 150.2158 | -0.17701 | -0.38197 | -0.39128 | 0.62419 | 149.6382 | 0.38197 | 94.338
1.0165 | -7.1456 | 8.72 | -150.8493 | 151.3896 | -0.19564 | -0.4006 | -0.39128 | 0.64282 | 150.8493 | 0.4006 | 94.5054
1.0182 | -6.8847 | 8.4406 | -152.5542 | 153.0479 | -0.21427 | -0.42855 | -0.4006 | 0.67077 | 152.5542 | 0.42855 | 94.6634
1.02 | -6.54 | 8.7107 | -154.7621 | 155.2652 | -0.20496 | -0.4565 | -0.4006 | 0.68009 | 154.7621 | 0.4565 | 94.8308
1.0218 | -6.1953 | 9.1859 | -156.7931 | 157.3241 | -0.17701 | -0.48445 | -0.40992 | 0.69872 | 156.7931 | 0.48445 | 94.9981
1.0236 | -5.804 | 9.4001 | -157.911 | 158.4421 | -0.15838 | -0.50308 | -0.41923 | 0.70804 | 157.911 | 0.50308 | 95.1655
1.0254 | -5.6364 | 9.633 | -158.0042 | 158.5445 | -0.14906 | -0.5124 | -0.42855 | 0.71735 | 158.0042 | 0.5124 | 95.3328
1.0272 | -5.5059 | 9.633 | -157.8831 | 158.4234 | -0.13043 | -0.52171 | -0.43787 | 0.72667 | 157.8831 | 0.52171 | 95.5002
1.029 | -5.6084 | 9.456 | -156.7558 | 157.2869 | -0.13043 | -0.5124 | -0.44718 | 0.72667 | 156.7558 | 0.5124 | 95.6675
1.0308 | -5.7947 | 8.9064 | -155.3118 | 155.8149 | -0.13974 | -0.48445 | -0.4565 | 0.71735 | 155.3118 | 0.48445 | 95.8349
1.0326 | -5.9345 | 8.2542 | -154.9578 | 155.4236 | -0.14906 | -0.46581 | -0.4565 | 0.70804 | 154.9578 | 0.46581 | 96.0022
1.0344 | -6.0556 | 7.4996 | -156.1409 | 156.5602 | -0.17701 | -0.43787 | -0.4565 | 0.68941 | 156.1409 | 0.43787 | 96.1696
1.0362 | -6.214 | 6.8009 | -157.8924 | 158.2744 | -0.19564 | -0.40992 | -0.4565 | 0.68009 | 157.8924 | 0.40992 | 96.3369
1.038 | -6.3164 | 6.6052 | -160.5289 | 160.8923 | -0.19564 | -0.4006 | -0.4565 | 0.67077 | 160.5289 | 0.4006 | 96.5043
1.0398 | -6.3817 | 6.717 | -164.0318 | 164.3859 | -0.18633 | -0.4006 | -0.4565 | 0.66146 | 164.0318 | 0.4006 | 96.6716
1.0416 | -6.3817 | 6.7077 | -167.0224 | 167.3764 | -0.17701 | -0.39128 | -0.4565 | 0.65214 | 167.0224 | 0.39128 | 96.839
1.0433 | -6.3444 | 6.717 | -169.249 | 169.5937 | -0.16769 | -0.38197 | -0.4565 | 0.65214 | 169.249 | 0.38197 | 96.997
1.0451 | -6.1767 | 6.7357 | -170.9073 | 171.2426 | -0.15838 | -0.38197 | -0.46581 | 0.64282 | 170.9073 | 0.38197 | 97.1644
1.0469 | -5.9345 | 6.6239 | -171.6991 | 172.0252 | -0.16769 | -0.36334 | -0.46581 | 0.64282 | 171.6991 | 0.36334 | 97.3317
1.0487 | -5.7761 | 6.6984 | -172.6121 | 172.9196 | -0.15838 | -0.36334 | -0.46581 | 0.63351 | 172.6121 | 0.36334 | 97.4991
1.0505 | -5.6084 | 6.8102 | -173.4413 | 173.7487 | -0.15838 | -0.35402 | -0.46581 | 0.63351 | 173.4413 | 0.35402 | 97.6664
1.0523 | -5.4128 | 6.9965 | -174.0841 | 174.4009 | -0.15838 | -0.33539 | -0.4565 | 0.61487 | 174.0841 | 0.33539 | 97.8338
1.0541 | -5.2451 | 7.304 | -173.9257 | 174.2518 | -0.14906 | -0.33539 | -0.4565 | 0.60556 | 173.9257 | 0.33539 | 98.0011
1.0559 | -5.0774 | 7.2947 | -173.2084 | 173.5251 | -0.15838 | -0.32607 | -0.44718 | 0.60556 | 173.2084 | 0.32607 | 98.1685
1.0577 | -4.8352 | 7.7977 | -172.3699 | 172.7146 | -0.15838 | -0.32607 | -0.43787 | 0.59624 | 172.3699 | 0.32607 | 98.3358
1.0595 | -4.6209 | 8.2542 | -171.0191 | 171.3824 | -0.14906 | -0.32607 | -0.43787 | 0.58693 | 171.0191 | 0.32607 | 98.5032
1.0613 | -4.4718 | 8.3381 | -169.6682 | 170.0408 | -0.16769 | -0.31675 | -0.42855 | 0.58693 | 169.6682 | 0.31675 | 98.6705
1.0631 | -4.3414 | 8.2542 | -169.6309 | 170.0036 | -0.17701 | -0.31675 | -0.41923 | 0.58693 | 169.6309 | 0.31675 | 98.8379
1.0649 | -4.1457 | 8.2356 | -170.0502 | 170.4321 | -0.19564 | -0.31675 | -0.40992 | 0.58693 | 170.0502 | 0.31675 | 99.0052
1.0667 | -3.7917 | 8.7107 | -171.5128 | 171.9414 | -0.19564 | -0.33539 | -0.4006 | 0.58693 | 171.5128 | 0.33539 | 99.1726
1.0684 | -3.4843 | 9.1113 | -173.2177 | 173.6742 | -0.20496 | -0.3447 | -0.39128 | 0.58693 | 173.2177 | 0.3447 | 99.3306
1.0702 | -3.2141 | 9.633 | -173.5904 | 174.0841 | -0.20496 | -0.35402 | -0.38197 | 0.58693 | 173.5904 | 0.35402 | 99.498
1.072 | -3.0092 | 10.3318 | -172.5656 | 173.0966 | -0.19564 | -0.35402 | -0.38197 | 0.57761 | 172.5656 | 0.35402 | 99.6653
1.0738 | -2.8694 | 10.8814 | -171.5128 | 172.0718 | -0.18633 | -0.3447 | -0.36334 | 0.56829 | 171.5128 | 0.3447 | 99.8327
1.0756 | -2.5899 | 11.5336 | -171.075 | 171.7085 | -0.18633 | -0.35402 | -0.36334 | 0.56829 | 171.075 | 0.35402 | 100
| Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My
Minimum | -54.2953 | -143.4242 | -2281.1205 | 149.9736 | -1.7049 | -21.6604 | -5.655 | 0.43787 | 149.3587 | -2.9812
Maximum | 116.8169 | 81.7504 | -149.3587 | 2285.76 | 18.5487 | 2.9812 | 0.19564 | 26.1788 | 2281.1205 | 21.6604

## Diagr_AVER75
AVER75 = Loads at Average Level in Subject with 75 kg Body Weight.
Contact forces and moments, measured telemetrically in knee implants of 8 subjects.
Averages of one individual subject.
Fx  lateral, Fy  anterior, -Fz  distal, Fres = Resultant Force
Mx around x, My around y, Mz around z, Mres = Resultant Moment
z-axis = axis of implant stem!
From Bergmann et al, PlosOne (2014) "Standardized Loads Acting in Knee Implants"
Subject K8L
Subject K8L
G.Bergmann et al., PlosOne, 2014

## Dat_HIGH100
Data File: | k8l_2266_0_6_i1.akf
Curve Alignment Program: | 1.5.7
------------------------MEASURING INFOS------------------------
Diagram Title #1 | Forces and Moments at Knee Joint
Diagram Title #2 | Measured Relative to Implant
Comment #1 | Knee Joint; Gaitanalysis; Stairs up;
Comment #2 | K8L, 25 Months PO *
BodyWeight [N]: | 775
Patient(en): | K8L | [7 empty] | Subject | K8L
Implant Type: | Knee Joint
Varus Angle Leg [°] | 4 | [7 empty] | HIGH100 Loads in N (forces) and Nm (moments)
Axes Directions X, Y, Z: | Lateral, Anterior, Superior
Axes Origin: | Tibial Plateau at Stem Axis | [7 empty] | = AVER75 Loads * 1/Caver
Transformation X Y Z (°): | 11.5 | -1 | 12
Displacement X, Y, Z (mm): | 0 | 0 | 0 | [5 empty] | Caver = | 0.53
Activity Code: | KneeActivities | 1P_1_15_4_1
---------------------------DATA--------------------------------
Number of Data Sets = | 601
Number of analog channels = | 0
Time-Offset(sec) = | 0
Time | Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My | Time
[s] | [N] | [N] | [N] | [N] | [Nm] | [Nm] | [Nm] | [Nm] | [N] | (Nm] | [%Cycle]
0 | -4.6933 | 40.9917 | -303.3242 | 306.2773 | -0.68554 | -1.9687 | -1.2832 | 2.4609 | 303.3242 | 1.9687 | 0
0.0018 | -5.1855 | 40.3764 | -303.1308 | 306.0136 | -0.68554 | -1.8984 | -1.2832 | 2.4082 | 303.1308 | 1.8984 | 0.16735
0.0036 | -5.3261 | 39.9897 | -302.6211 | 305.4687 | -0.70312 | -1.8633 | -1.2656 | 2.373 | 302.6211 | 1.8633 | 0.3347
0.0054 | -5.2909 | 39.7436 | -302.0762 | 304.8886 | -0.70312 | -1.8457 | -1.248 | 2.3554 | 302.0762 | 1.8457 | 0.50205
0.0072 | -5.1327 | 40.0776 | -305.996 | 308.8085 | -0.68554 | -1.8457 | -1.2305 | 2.3379 | 305.996 | 1.8457 | 0.66939
0.009 | -5.2206 | 39.7788 | -309.5468 | 312.3241 | -0.70312 | -1.8105 | -1.2129 | 2.3027 | 309.5468 | 1.8105 | 0.83674
0.0108 | -5.2734 | 39.7964 | -312.2889 | 315.0838 | -0.68554 | -1.7929 | -1.1953 | 2.2851 | 312.2889 | 1.7929 | 1.0041
0.0125 | -5.1503 | 40.3764 | -313.8709 | 316.7186 | -0.65038 | -1.7754 | -1.1777 | 2.25 | 313.8709 | 1.7754 | 1.1621
0.0143 | -5.08 | 40.8335 | -313.9588 | 316.8592 | -0.6328 | -1.7578 | -1.1777 | 2.2148 | 313.9588 | 1.7578 | 1.3295
0.0161 | -5.3085 | 40.9565 | -312.1659 | 315.0838 | -0.61523 | -1.7051 | -1.1601 | 2.1797 | 312.1659 | 1.7051 | 1.4968
0.0179 | -5.9413 | 40.6049 | -309.4589 | 312.3592 | -0.61523 | -1.6347 | -1.1601 | 2.1093 | 309.4589 | 1.6347 | 1.6642
0.0197 | -6.6444 | 40.0249 | -306.6113 | 309.4765 | -0.6328 | -1.5644 | -1.1426 | 2.0566 | 306.6113 | 1.5644 | 1.8315
0.0215 | -6.7499 | 40.0073 | -306.2773 | 309.1601 | -0.65038 | -1.5293 | -1.1074 | 2.0215 | 306.2773 | 1.5293 | 1.9989
0.0233 | -6.5917 | 40.9213 | -307.9648 | 310.9706 | -0.65038 | -1.5117 | -1.0898 | 2.0039 | 307.9648 | 1.5117 | 2.1662
0.0251 | -6.4511 | 41.9057 | -309.7753 | 312.9217 | -0.66796 | -1.5117 | -1.0723 | 1.9863 | 309.7753 | 1.5117 | 2.3336
0.0269 | -5.9765 | 42.9428 | -311.0057 | 314.2928 | -0.66796 | -1.5117 | -1.0547 | 1.9863 | 311.0057 | 1.5117 | 2.5009
0.0287 | -5.6425 | 43.9799 | -310.1796 | 313.6073 | -0.65038 | -1.5293 | -1.0371 | 1.9687 | 310.1796 | 1.5293 | 2.6683
0.0305 | -5.1152 | 45.2807 | -309.037 | 312.6581 | -0.61523 | -1.5469 | -1.0195 | 1.9687 | 309.037 | 1.5469 | 2.8356
0.0323 | -4.5351 | 46.7221 | -307.2089 | 310.9706 | -0.6328 | -1.582 | -1.0195 | 2.0039 | 307.2089 | 1.582 | 3.003
0.0341 | -3.6738 | 48.638 | -305.9257 | 309.9511 | -0.65038 | -1.6347 | -1.0195 | 2.039 | 305.9257 | 1.6347 | 3.1703
0.0359 | -3.5859 | 49.0423 | -302.9375 | 307.0859 | -0.66796 | -1.6172 | -1.0019 | 2.0215 | 302.9375 | 1.6172 | 3.3377
0.0376 | -3.7617 | 49.1302 | -299.2285 | 303.4121 | -0.72069 | -1.5996 | -0.98436 | 2.0215 | 299.2285 | 1.5996 | 3.4957
0.0394 | -3.7968 | 49.4994 | -296.8204 | 301.0918 | -0.75585 | -1.582 | -0.96678 | 2.0215 | 296.8204 | 1.582 | 3.6631
0.0412 | -3.3222 | 50.431 | -298.7891 | 303.2012 | -0.79101 | -1.6172 | -0.94921 | 2.0566 | 298.7891 | 1.6172 | 3.8304
0.043 | -2.8476 | 51.4329 | -302.0762 | 306.6113 | -0.84374 | -1.6523 | -0.93163 | 2.0918 | 302.0762 | 1.6523 | 3.9978
0.0448 | -2.3379 | 52.5579 | -305.0644 | 309.7401 | -0.87889 | -1.7051 | -0.93163 | 2.1269 | 305.0644 | 1.7051 | 4.1651
0.0466 | -1.7754 | 54.0696 | -307.2616 | 312.1483 | -0.89647 | -1.7402 | -0.91405 | 2.1621 | 307.2616 | 1.7402 | 4.3325
0.0484 | -1.2129 | 55.5989 | -308.3163 | 313.4491 | -0.91405 | -1.7754 | -0.91405 | 2.2148 | 308.3163 | 1.7754 | 4.4998
0.0502 | -0.79101 | 57.6028 | -307.6132 | 313.1327 | -0.93163 | -1.8281 | -0.93163 | 2.25 | 307.6132 | 1.8281 | 4.6672
0.052 | -0.38671 | 58.9035 | -306.3652 | 312.1835 | -0.96678 | -1.8457 | -0.94921 | 2.3027 | 306.3652 | 1.8457 | 4.8345
0.0538 | 0.017578 | 59.8176 | -305.4335 | 311.4628 | -1.0019 | -1.8808 | -0.94921 | 2.3379 | 305.4335 | 1.8808 | 5.0019
0.0556 | 0.42187 | 60.8371 | -306.5058 | 312.6932 | -1.0547 | -1.916 | -0.94921 | 2.3906 | 306.5058 | 1.916 | 5.1692
0.0574 | 0.82616 | 61.2766 | -308.4394 | 314.6795 | -1.125 | -1.9511 | -0.94921 | 2.4433 | 308.4394 | 1.9511 | 5.3366
0.0592 | 1.248 | 61.9972 | -311.744 | 318.0545 | -1.2129 | -1.9863 | -0.94921 | 2.5136 | 311.744 | 1.9863 | 5.5039
0.061 | 2.0742 | 63.6847 | -316.49 | 323.0466 | -1.2305 | -2.0566 | -0.94921 | 2.6015 | 316.49 | 2.0566 | 5.6713
0.0627 | 3.0586 | 65.5128 | -321.6052 | 328.4079 | -1.2656 | -2.1269 | -0.96678 | 2.6543 | 321.6052 | 2.1269 | 5.8293
0.0645 | 3.7441 | 66.8487 | -324.7517 | 331.7477 | -1.3183 | -2.1797 | -1.0019 | 2.7246 | 324.7517 | 2.1797 | 5.9967
0.0663 | 4.3593 | 68.1495 | -327.529 | 334.7007 | -1.3711 | -2.2148 | -1.0195 | 2.7949 | 327.529 | 2.2148 | 6.164
0.0681 | 5.0097 | 69.3975 | -331.9586 | 339.2886 | -1.4414 | -2.2675 | -1.0371 | 2.8828 | 331.9586 | 2.2675 | 6.3313
0.0699 | 6.1698 | 70.7159 | -339.5171 | 346.9701 | -1.5293 | -2.3379 | -1.0547 | 2.9882 | 339.5171 | 2.3379 | 6.4987
0.0717 | 7.4003 | 72.5088 | -347.6908 | 355.3372 | -1.582 | -2.4257 | -1.0723 | 3.0937 | 347.6908 | 2.4257 | 6.666
0.0735 | 8.6307 | 74.3018 | -355.636 | 363.5285 | -1.6347 | -2.5136 | -1.0898 | 3.1992 | 355.636 | 2.5136 | 6.8334
0.0753 | 9.8085 | 76.499 | -362.8078 | 371.0167 | -1.6699 | -2.6015 | -1.125 | 3.2871 | 362.8078 | 2.6015 | 7.0007
0.0771 | 10.6698 | 78.6259 | -368.046 | 376.6065 | -1.7226 | -2.6718 | -1.1601 | 3.3925 | 368.046 | 2.6718 | 7.1681
0.0789 | 11.1268 | 80.2958 | -371.2979 | 380.1572 | -1.7578 | -2.707 | -1.1953 | 3.4453 | 371.2979 | 2.707 | 7.3354
0.0807 | 11.4081 | 81.4384 | -373.7061 | 382.7411 | -1.8105 | -2.7422 | -1.2305 | 3.498 | 373.7061 | 2.7422 | 7.5028
0.0825 | 11.6893 | 82.6161 | -376.4483 | 385.6942 | -1.8633 | -2.7597 | -1.248 | 3.5507 | 376.4483 | 2.7597 | 7.6701
0.0843 | 12.3045 | 84.4794 | -381.5634 | 391.1082 | -1.916 | -2.7949 | -1.2656 | 3.621 | 381.5634 | 2.7949 | 7.8375
0.086 | 13.078 | 86.87 | -388.454 | 398.403 | -1.9687 | -2.8476 | -1.2832 | 3.6914 | 388.454 | 2.8476 | 7.9955
0.0878 | 13.8162 | 89.6297 | -396.698 | 407.0689 | -1.9863 | -2.9004 | -1.3183 | 3.7617 | 396.698 | 2.9004 | 8.1629
0.0896 | 15.0994 | 93.075 | -405.9615 | 416.8774 | -1.9863 | -2.9882 | -1.3535 | 3.832 | 405.9615 | 2.9882 | 8.3302
0.0914 | 15.8904 | 95.9753 | -412.6587 | 424.0668 | -2.0039 | -3.0058 | -1.3711 | 3.8671 | 412.6587 | 3.0058 | 8.4976
0.0932 | 16.1014 | 98.2077 | -417.4399 | 429.2523 | -2.039 | -2.9707 | -1.3887 | 3.8671 | 417.4399 | 2.9707 | 8.6649
0.095 | 16.2596 | 100.2819 | -421.9399 | 434.1038 | -2.0566 | -2.9355 | -1.4062 | 3.8496 | 421.9399 | 2.9355 | 8.8323
0.0968 | 16.6814 | 102.5495 | -427.7757 | 440.3263 | -2.1093 | -2.9004 | -1.4238 | 3.8671 | 427.7757 | 2.9004 | 8.9996
0.0986 | 17.4549 | 104.9752 | -435.7385 | 448.6583 | -2.1621 | -2.9179 | -1.4414 | 3.9023 | 435.7385 | 2.9179 | 9.167
0.1004 | 18.4568 | 108.3502 | -446.2149 | 459.662 | -2.1797 | -2.9179 | -1.459 | 3.9199 | 446.2149 | 2.9179 | 9.3343
0.1022 | 19.4939 | 112.0943 | -457.5702 | 471.5974 | -2.1972 | -2.9179 | -1.459 | 3.9374 | 457.5702 | 2.9179 | 9.5017
0.104 | 20.32 | 115.5044 | -468.24 | 482.7945 | -2.2148 | -2.9004 | -1.459 | 3.9374 | 468.24 | 2.9004 | 9.669
0.1058 | 20.7771 | 118.4926 | -477.1696 | 492.1987 | -2.2324 | -2.83 | -1.4765 | 3.9023 | 477.1696 | 2.83 | 9.8364
0.1076 | 21.1638 | 121.1469 | -484.4996 | 499.9506 | -2.2675 | -2.7597 | -1.4765 | 3.8847 | 484.4996 | 2.7597 | 10.0037
0.1094 | 21.1814 | 122.9047 | -489.9487 | 505.6634 | -2.3027 | -2.6718 | -1.4765 | 3.8496 | 489.9487 | 2.6718 | 10.1711
0.1111 | 20.9001 | 124.4867 | -495.4858 | 511.4113 | -2.373 | -2.5664 | -1.459 | 3.832 | 495.4858 | 2.5664 | 10.3291
0.1129 | 20.7068 | 126.2445 | -501.7435 | 517.9152 | -2.4257 | -2.4785 | -1.459 | 3.7968 | 501.7435 | 2.4785 | 10.4965
0.1147 | 20.4783 | 128.5647 | -510.0403 | 526.5108 | -2.4433 | -2.373 | -1.4414 | 3.7441 | 510.0403 | 2.373 | 10.6638
0.1165 | 20.5134 | 131.1487 | -519.7257 | 536.5126 | -2.4785 | -2.2675 | -1.4062 | 3.6914 | 519.7257 | 2.2675 | 10.8312
0.1183 | 20.7419 | 133.7854 | -530.167 | 547.3054 | -2.4961 | -2.1797 | -1.3887 | 3.6562 | 530.167 | 2.1797 | 10.9985
0.1201 | 21.0232 | 136.3693 | -539.8348 | 557.3072 | -2.5136 | -2.0918 | -1.3535 | 3.621 | 539.8348 | 2.0918 | 11.1659
0.1219 | 21.2517 | 138.5314 | -546.4617 | 564.2681 | -2.5312 | -2.0039 | -1.3359 | 3.6035 | 546.4617 | 2.0039 | 11.3332
0.1237 | 21.1814 | 140.3595 | -551.4714 | 569.5766 | -2.5488 | -1.8984 | -1.3183 | 3.5683 | 551.4714 | 1.8984 | 11.5006
0.1255 | 20.4607 | 141.379 | -553.6686 | 571.9145 | -2.584 | -1.7754 | -1.3008 | 3.5156 | 553.6686 | 1.7754 | 11.6679
0.1273 | 19.74 | 142.17 | -556.569 | 574.9027 | -2.6191 | -1.6699 | -1.2832 | 3.4804 | 556.569 | 1.6699 | 11.8353
0.1291 | 19.3884 | 143.1896 | -562.9497 | 581.3186 | -2.6543 | -1.5644 | -1.248 | 3.4453 | 562.9497 | 1.5644 | 12.0026
0.1309 | 19.3005 | 144.3145 | -569.1723 | 587.6291 | -2.6894 | -1.4765 | -1.2129 | 3.4101 | 569.1723 | 1.4765 | 12.17
0.1327 | 19.5466 | 145.9493 | -574.6215 | 593.3068 | -2.707 | -1.3711 | -1.1777 | 3.375 | 574.6215 | 1.3711 | 12.3373
0.1345 | 19.8279 | 147.5489 | -579.6487 | 598.5626 | -2.707 | -1.2832 | -1.1426 | 3.3398 | 579.6487 | 1.2832 | 12.5046
0.1362 | 19.9861 | 148.709 | -582.0921 | 601.2168 | -2.7246 | -1.2129 | -1.125 | 3.3222 | 582.0921 | 1.2129 | 12.6627
0.138 | 19.5818 | 149.1485 | -582.9885 | 602.166 | -2.7597 | -1.125 | -1.0898 | 3.3222 | 582.9885 | 1.125 | 12.83
0.1398 | 18.5798 | 148.7266 | -582.4261 | 601.4805 | -2.8125 | -1.0195 | -1.0723 | 3.3398 | 582.4261 | 1.0195 | 12.9974
0.1416 | 17.4197 | 148.1465 | -582.2503 | 601.1114 | -2.8652 | -0.89647 | -1.0547 | 3.3398 | 582.2503 | 0.89647 | 13.1647
0.1434 | 16.6638 | 148.6563 | -586.4865 | 605.3125 | -2.9004 | -0.79101 | -1.0195 | 3.3574 | 586.4865 | 0.79101 | 13.3321
0.1452 | 15.9432 | 149.5176 | -592.7443 | 611.5702 | -2.9355 | -0.68554 | -0.98436 | 3.375 | 592.7443 | 0.68554 | 13.4994
0.147 | 15.3631 | 150.6953 | -601.4981 | 620.324 | -2.9707 | -0.58007 | -0.93163 | 3.4453 | 601.4981 | 0.58007 | 13.6668
0.1488 | 15.0994 | 152.1543 | -609.4609 | 628.3923 | -3.0058 | -0.50976 | -0.91405 | 3.5156 | 609.4609 | 0.50976 | 13.8341
0.1506 | 14.7479 | 153.0859 | -615.6131 | 634.5797 | -3.0234 | -0.43945 | -0.87889 | 3.5507 | 615.6131 | 0.43945 | 14.0015
0.1524 | 14.5193 | 153.7714 | -620.4822 | 639.4488 | -3.0234 | -0.36914 | -0.86132 | 3.5683 | 620.4822 | 0.36914 | 14.1688
0.1542 | 14.6248 | 154.2461 | -623.699 | 642.6831 | -3.041 | -0.35156 | -0.84374 | 3.6035 | 623.699 | 0.35156 | 14.3362
0.156 | 14.8182 | 154.0527 | -625.1931 | 644.0893 | -3.041 | -0.3164 | -0.84374 | 3.621 | 625.1931 | 0.3164 | 14.5035
0.1578 | 14.66 | 153.1386 | -624.9294 | 643.6147 | -3.041 | -0.24609 | -0.82616 | 3.621 | 624.9294 | 0.24609 | 14.6709
0.1595 | 14.2029 | 151.5742 | -623.5232 | 641.8745 | -3.041 | -0.14062 | -0.79101 | 3.6738 | 623.5232 | 0.14062 | 14.8289
0.1613 | 13.5877 | 149.5 | -620.8689 | 638.7984 | -3.0586 | -0.052734 | -0.77343 | 3.7265 | 620.8689 | 0.052734 | 14.9963
0.1631 | 12.8846 | 146.9161 | -618.7244 | 636.1265 | -3.0937 | 0.035156 | -0.75585 | 3.7792 | 618.7244 | -0.035156 | 15.1636
0.1649 | 12.0936 | 144.2794 | -618.2674 | 635.0894 | -3.1464 | 0.10547 | -0.73827 | 3.8496 | 618.2674 | -0.10547 | 15.331
0.1667 | 11.3377 | 141.8712 | -618.8826 | 635.1597 | -3.1816 | 0.17578 | -0.70312 | 3.9023 | 618.8826 | -0.17578 | 15.4983
0.1685 | 10.6874 | 139.7619 | -620.5877 | 636.355 | -3.2168 | 0.24609 | -0.68554 | 3.9374 | 620.5877 | -0.24609 | 15.6657
0.1703 | 10.6522 | 138.1799 | -626.0192 | 641.3472 | -3.1992 | 0.26367 | -0.68554 | 3.9374 | 626.0192 | -0.26367 | 15.833
0.1721 | 10.5467 | 136.9318 | -632.2594 | 647.2182 | -3.1816 | 0.26367 | -0.66796 | 3.955 | 632.2594 | -0.26367 | 16.0004
0.1739 | 10.6346 | 135.6838 | -638.3413 | 652.931 | -3.164 | 0.26367 | -0.66796 | 3.9726 | 638.3413 | -0.26367 | 16.1677
0.1757 | 11.1444 | 134.7873 | -643.9487 | 658.2571 | -3.0937 | 0.28125 | -0.68554 | 3.9374 | 643.9487 | -0.28125 | 16.3351
0.1775 | 11.7772 | 133.469 | -648.308 | 662.3 | -3.0234 | 0.29882 | -0.68554 | 3.8847 | 648.308 | -0.29882 | 16.5024
0.1793 | 12.3397 | 132.0276 | -650.8217 | 664.5148 | -2.9355 | 0.33398 | -0.68554 | 3.832 | 650.8217 | -0.33398 | 16.6698
0.1811 | 13.1834 | 130.4456 | -652.1048 | 665.4816 | -2.83 | 0.36914 | -0.70312 | 3.7617 | 652.1048 | -0.36914 | 16.8371
0.1829 | 13.7108 | 128.3538 | -652.1576 | 665.1476 | -2.7422 | 0.42187 | -0.70312 | 3.7089 | 652.1576 | -0.42187 | 17.0045
0.1846 | 13.992 | 125.7523 | -650.8392 | 663.3723 | -2.6718 | 0.52734 | -0.70312 | 3.6562 | 650.8392 | -0.52734 | 17.1625
0.1864 | 13.5877 | 122.6586 | -648.0619 | 660.0149 | -2.6191 | 0.66796 | -0.68554 | 3.6035 | 648.0619 | -0.66796 | 17.3299
0.1882 | 13.1307 | 119.143 | -645.9877 | 657.2727 | -2.584 | 0.82616 | -0.66796 | 3.5859 | 645.9877 | -0.82616 | 17.4972
0.19 | 13.3065 | 115.5571 | -646.3041 | 656.9212 | -2.5136 | 0.96678 | -0.66796 | 3.5507 | 646.3041 | -0.96678 | 17.6646
0.1918 | 13.6053 | 111.8482 | -648.185 | 658.1165 | -2.4433 | 1.1074 | -0.65038 | 3.5332 | 648.185 | -1.1074 | 17.8319
0.1936 | 14.0447 | 108.1568 | -652.7376 | 662.0188 | -2.373 | 1.2656 | -0.6328 | 3.5156 | 652.7376 | -1.2656 | 17.9993
0.1954 | 14.5545 | 104.6061 | -659.1712 | 667.8195 | -2.2675 | 1.459 | -0.61523 | 3.4804 | 659.1712 | -1.459 | 18.1666
0.1972 | 15.0643 | 101.2487 | -667.3625 | 675.3956 | -2.1445 | 1.6347 | -0.59765 | 3.4453 | 667.3625 | -1.6347 | 18.334
0.199 | 16.1014 | 98.1726 | -677.6279 | 685.1337 | -2.0215 | 1.7754 | -0.59765 | 3.4453 | 677.6279 | -1.7754 | 18.5013
0.2008 | 17.8591 | 95.782 | -688.9481 | 695.9617 | -1.916 | 1.8457 | -0.59765 | 3.3398 | 688.9481 | -1.8457 | 18.6687
0.2026 | 19.6872 | 91.7742 | -697.6316 | 704.0651 | -1.7929 | 1.9687 | -0.59765 | 3.2343 | 697.6316 | -1.9687 | 18.836
0.2044 | 21.4275 | 87.6786 | -707.1237 | 713.0474 | -1.6523 | 2.1269 | -0.59765 | 3.1992 | 707.1237 | -2.1269 | 19.0033
0.2062 | 22.904 | 84.0927 | -719.4458 | 724.9125 | -1.4765 | 2.3203 | -0.58007 | 3.164 | 719.4458 | -2.3203 | 19.1707
0.208 | 23.8356 | 80.6474 | -735.1077 | 740.1349 | -1.3183 | 2.5488 | -0.58007 | 3.1992 | 735.1077 | -2.5488 | 19.338
0.2097 | 24.6618 | 77.0439 | -752.7383 | 757.291 | -1.1777 | 2.83 | -0.58007 | 3.3046 | 752.7383 | -2.83 | 19.4961
0.2115 | 25.9626 | 73.6338 | -774.0779 | 778.2263 | -1.0019 | 3.041 | -0.58007 | 3.4101 | 774.0779 | -3.041 | 19.6634
0.2133 | 27.2985 | 70.0831 | -798.599 | 802.3783 | -0.82616 | 3.3046 | -0.59765 | 3.5859 | 798.599 | -3.3046 | 19.8308
0.2151 | 28.7399 | 66.6554 | -825.0362 | 828.499 | -0.65038 | 3.5332 | -0.61523 | 3.7441 | 825.0362 | -3.5332 | 19.9981
0.2169 | 30.0406 | 63.1925 | -854.4968 | 857.6256 | -0.4746 | 3.7617 | -0.6328 | 3.9199 | 854.4968 | -3.7617 | 20.1655
0.2187 | 31.1832 | 59.2024 | -887.5608 | 890.3732 | -0.28125 | 4.0429 | -0.66796 | 4.166 | 887.5608 | -4.0429 | 20.3328
0.2205 | 32.6246 | 55.1594 | -923.367 | 925.863 | -0.070312 | 4.289 | -0.70312 | 4.4121 | 923.367 | -4.289 | 20.5002
0.2223 | 34.3999 | 51.0286 | -960.8958 | 963.1809 | 0.1582 | 4.5351 | -0.75585 | 4.6757 | 960.8958 | -4.5351 | 20.6675
0.2241 | 36.2808 | 46.7045 | -998.6531 | 1000.7449 | 0.36914 | 4.7636 | -0.79101 | 4.9042 | 998.6531 | -4.7636 | 20.8349
0.2259 | 38.4429 | 42.1166 | -1037.3069 | 1039.2053 | 0.59765 | 4.957 | -0.84374 | 5.1327 | 1037.3069 | -4.957 | 21.0022
0.2277 | 41.4135 | 37.3882 | -1080.4958 | 1082.2887 | 0.86132 | 5.0624 | -0.91405 | 5.2734 | 1080.4958 | -5.0624 | 21.1696
0.2295 | 44.5248 | 32.5191 | -1125.1964 | 1126.8838 | 1.125 | 5.1503 | -0.96678 | 5.4316 | 1125.1964 | -5.1503 | 21.3369
0.2313 | 47.9349 | 27.4567 | -1172.0063 | 1173.6059 | 1.4062 | 5.2031 | -1.0371 | 5.5546 | 1172.0063 | -5.2031 | 21.5043
0.2331 | 51.9427 | 22.3064 | -1221.8924 | 1223.5096 | 1.7226 | 5.2031 | -1.1074 | 5.6777 | 1221.8924 | -5.2031 | 21.6716
0.2348 | 56.3196 | 17.0681 | -1271.9015 | 1273.5714 | 2.0566 | 5.1503 | -1.1777 | 5.8007 | 1271.9015 | -5.1503 | 21.8297
0.2366 | 60.7316 | 11.8299 | -1322.1919 | 1323.9497 | 2.373 | 5.08 | -1.2656 | 5.9062 | 1322.1919 | -5.08 | 21.997
0.2384 | 65.214 | 6.6796 | -1374.1697 | 1376.0154 | 2.707 | 4.9921 | -1.3535 | 6.0292 | 1374.1697 | -4.9921 | 22.1644
0.2402 | 70.1007 | 1.7578 | -1429.0655 | 1431.0694 | 3.0586 | 4.8163 | -1.4414 | 6.205 | 1429.0655 | -4.8163 | 22.3317
0.242 | 75.1982 | -2.6543 | -1486.7034 | 1488.8655 | 3.4453 | 4.5878 | -1.5293 | 6.3984 | 1486.7034 | -4.5878 | 22.4991
0.2438 | 80.1728 | -6.8026 | -1545.2378 | 1547.6108 | 3.832 | 4.3769 | -1.5996 | 6.6093 | 1545.2378 | -4.3769 | 22.6664
0.2456 | 85.2001 | -10.6171 | -1602.4011 | 1604.9675 | 4.2363 | 4.1308 | -1.6875 | 6.8378 | 1602.4011 | -4.1308 | 22.8338
0.2474 | 90.4207 | -14.6775 | -1661.6914 | 1664.4863 | 4.623 | 3.8671 | -1.7754 | 7.0839 | 1661.6914 | -3.8671 | 23.0011
0.2492 | 95.571 | -18.9841 | -1722.0539 | 1725.0421 | 5.0273 | 3.5859 | -1.8633 | 7.3476 | 1722.0539 | -3.5859 | 23.1685
0.251 | 100.4401 | -22.9392 | -1781.1683 | 1784.3851 | 5.4316 | 3.2695 | -1.9511 | 7.6464 | 1781.1683 | -3.2695 | 23.3358
0.2528 | 104.9049 | -27.2809 | -1840.7047 | 1844.1148 | 5.8007 | 2.9707 | -2.039 | 7.9804 | 1840.7047 | -2.9707 | 23.5032
0.2546 | 109.3345 | -31.7281 | -1899.0457 | 1902.6668 | 6.1523 | 2.6367 | -2.1269 | 8.3319 | 1899.0457 | -2.6367 | 23.6705
0.2564 | 114.3794 | -35.9292 | -1964.998 | 1968.8651 | 6.5741 | 2.3027 | -2.2324 | 8.7011 | 1964.998 | -2.3027 | 23.8379
0.2581 | 119.2309 | -40.0249 | -2032.9893 | 2037.0674 | 7.0136 | 2.0215 | -2.3379 | 9.0175 | 2032.9893 | -2.0215 | 23.9959
0.2599 | 124.9613 | -44.7357 | -2102.3693 | 2106.7462 | 7.4706 | 1.7226 | -2.4433 | 9.369 | 2102.3693 | -1.7226 | 24.1633
0.2617 | 130.2171 | -49.2708 | -2169.0247 | 2173.6652 | 7.9101 | 1.3887 | -2.5488 | 9.7733 | 2169.0247 | -1.3887 | 24.3306
0.2635 | 135.4729 | -53.9114 | -2234.9066 | 2239.8108 | 8.3495 | 1.0371 | -2.6367 | 10.1776 | 2234.9066 | -1.0371 | 24.498
0.2653 | 140.6232 | -58.4817 | -2299.4527 | 2304.6557 | 8.7889 | 0.68554 | -2.7422 | 10.6171 | 2299.4527 | -0.68554 | 24.6653
0.2671 | 144.5782 | -62.8234 | -2356.8093 | 2362.2233 | 9.1757 | 0.40429 | -2.83 | 11.0038 | 2356.8093 | -0.40429 | 24.8327
0.2689 | 148.1114 | -67.0948 | -2410.9317 | 2416.5742 | 9.5272 | 0.1582 | -2.9355 | 11.3729 | 2410.9317 | -0.1582 | 25
0.2707 | 152.084 | -72.4385 | -2475.5656 | 2481.4542 | 9.9491 | -0.1582 | -3.0234 | 11.8123 | 2475.5656 | 0.1582 | 25.1673
0.2725 | 156.0566 | -77.9404 | -2543.0296 | 2549.1643 | 10.3885 | -0.50976 | -3.1289 | 12.287 | 2543.0296 | 0.50976 | 25.3347
0.2743 | 160.2577 | -83.6181 | -2613.5697 | 2619.968 | 10.8456 | -0.86132 | -3.2343 | 12.7616 | 2613.5697 | 0.86132 | 25.502
0.2761 | 164.3006 | -88.8738 | -2683.8988 | 2690.5433 | 11.3202 | -1.248 | -3.3398 | 13.2537 | 2683.8988 | 1.248 | 25.6694
0.2779 | 168.2381 | -93.7253 | -2754.6323 | 2761.5053 | 11.7948 | -1.6347 | -3.4453 | 13.7811 | 2754.6323 | 1.6347 | 25.8367
0.2797 | 172.1931 | -98.489 | -2826.0689 | 2833.1879 | 12.3045 | -2.039 | -3.5683 | 14.3084 | 2826.0689 | 2.039 | 26.0041
0.2815 | 176.236 | -103.2701 | -2897.9625 | 2905.31 | 12.8143 | -2.4609 | -3.6738 | 14.8533 | 2897.9625 | 2.4609 | 26.1714
0.2832 | 180.3316 | -107.9107 | -2969.9616 | 2977.5552 | 13.324 | -2.9004 | -3.7968 | 15.4158 | 2969.9616 | 2.9004 | 26.3295
0.285 | 184.5503 | -112.6919 | -3043.8942 | 3051.7164 | 13.869 | -3.3398 | -3.9199 | 16.0135 | 3043.8942 | 3.3398 | 26.4968
0.2868 | 188.5933 | -117.0336 | -3115.612 | 3123.6627 | 14.4139 | -3.8144 | -4.0429 | 16.6111 | 3115.612 | 3.8144 | 26.6642
0.2886 | 192.3725 | -121.393 | -3185.5545 | 3193.8161 | 14.9588 | -4.2539 | -4.166 | 17.2088 | 3185.5545 | 4.2539 | 26.8315
0.2904 | 195.8705 | -125.8929 | -3252.4911 | 3260.9461 | 15.451 | -4.7285 | -4.2539 | 17.8064 | 3252.4911 | 4.7285 | 26.9989
0.2922 | 199.2279 | -130.3577 | -3318.6191 | 3327.2851 | 15.9607 | -5.2031 | -4.3593 | 18.4041 | 3318.6191 | 5.2031 | 27.1662
0.294 | 202.0579 | -134.7522 | -3384.5011 | 3393.3252 | 16.4529 | -5.6601 | -4.4648 | 19.0017 | 3384.5011 | 5.6601 | 27.3336
0.2958 | 204.3782 | -138.9005 | -3447.219 | 3456.1838 | 16.9451 | -6.1171 | -4.5527 | 19.5818 | 3447.219 | 6.1171 | 27.5009
0.2976 | 206.5403 | -142.9435 | -3501.025 | 3510.1479 | 17.3494 | -6.5917 | -4.623 | 20.0915 | 3501.025 | 6.5917 | 27.6683
0.2994 | 208.2629 | -146.8282 | -3552.4931 | 3561.739 | 17.7185 | -7.0487 | -4.6757 | 20.5837 | 3552.4931 | 7.0487 | 27.8356
0.3012 | 210.4953 | -150.0274 | -3601.7991 | 3611.1681 | 18.1228 | -7.5761 | -4.746 | 21.1286 | 3601.7991 | 7.5761 | 28.003
0.303 | 212.5871 | -152.3125 | -3648.7321 | 3658.2065 | 18.5447 | -8.0858 | -4.8163 | 21.6911 | 3648.7321 | 8.0858 | 28.1703
0.3048 | 214.6964 | -154.58 | -3694.751 | 3704.3134 | 18.9666 | -8.6132 | -4.9042 | 22.2712 | 3694.751 | 8.6132 | 28.3377
0.3066 | 216.8409 | -157.3222 | -3739.1703 | 3748.8733 | 19.3708 | -9.2108 | -4.957 | 22.8337 | 3739.1703 | 9.2108 | 28.505
0.3083 | 218.4054 | -160.328 | -3780.6366 | 3790.4451 | 19.74 | -9.7909 | -5.0273 | 23.3434 | 3780.6366 | 9.7909 | 28.6631
0.3101 | 219.1436 | -163.4041 | -3818.7982 | 3828.677 | 20.0388 | -10.3534 | -5.08 | 23.8181 | 3818.7982 | 10.3534 | 28.8304
0.3119 | 219.6007 | -166.4803 | -3855.5009 | 3865.4675 | 20.3376 | -10.8983 | -5.1327 | 24.3278 | 3855.5009 | 10.8983 | 28.9978
0.3137 | 220.2159 | -169.9431 | -3890.0239 | 3900.0784 | 20.6189 | -11.4608 | -5.1855 | 24.8024 | 3890.0239 | 11.4608 | 29.1651
0.3155 | 220.4093 | -173.2829 | -3921.1895 | 3931.3143 | 20.8825 | -12.0409 | -5.2382 | 25.2946 | 3921.1895 | 12.0409 | 29.3325
0.3173 | 219.9171 | -176.1481 | -3949.6305 | 3959.7905 | 21.1638 | -12.6034 | -5.2909 | 25.8395 | 3949.6305 | 12.6034 | 29.4998
0.3191 | 219.6183 | -179.4176 | -3979.302 | 3989.5148 | 21.4099 | -13.1834 | -5.3613 | 26.3317 | 3979.302 | 13.1834 | 29.6672
0.3209 | 219.2843 | -182.4937 | -4007.3915 | 4017.657 | 21.6384 | -13.7635 | -5.414 | 26.7711 | 4007.3915 | 13.7635 | 29.8345
0.3227 | 218.2648 | -184.691 | -4032.7212 | 4042.9867 | 21.8669 | -14.326 | -5.4491 | 27.2282 | 4032.7212 | 14.326 | 30.0019
0.3245 | 216.4718 | -186.5366 | -4055.6428 | 4065.838 | 22.0954 | -14.8533 | -5.5019 | 27.6852 | 4055.6428 | 14.8533 | 30.1692
0.3263 | 214.8546 | -188.4175 | -4078.1601 | 4088.3201 | 22.3064 | -15.3631 | -5.5546 | 28.1598 | 4078.1601 | 15.3631 | 30.3366
0.3281 | 213.4308 | -189.6831 | -4099.9743 | 4110.064 | 22.5349 | -15.908 | -5.5898 | 28.5992 | 4099.9743 | 15.908 | 30.5039
0.3299 | 211.093 | -190.9487 | -4121.1381 | 4131.1399 | 22.7458 | -16.4353 | -5.6249 | 29.0387 | 4121.1381 | 16.4353 | 30.6713
0.3316 | 208.1575 | -192.144 | -4140.3331 | 4150.1943 | 22.9567 | -16.9802 | -5.6777 | 29.5133 | 4140.3331 | 16.9802 | 30.8293
0.3334 | 205.2571 | -193.5678 | -4158.2626 | 4168.0007 | 23.1501 | -17.49 | -5.7128 | 29.9703 | 4158.2626 | 17.49 | 30.9967
0.3352 | 202.3919 | -195.0795 | -4175.2077 | 4184.8579 | 23.3434 | -18.0173 | -5.7656 | 30.4449 | 4175.2077 | 18.0173 | 31.164
0.337 | 199.5091 | -196.6088 | -4191.3618 | 4200.9066 | 23.5192 | -18.5095 | -5.8007 | 30.8668 | 4191.3618 | 18.5095 | 31.3313
0.3388 | 196.8373 | -198.5248 | -4206.2854 | 4215.7775 | 23.6598 | -18.9666 | -5.8359 | 31.2535 | 4206.2854 | 18.9666 | 31.4987
0.3406 | 194.3061 | -200.8802 | -4219.4512 | 4228.9081 | 23.7829 | -19.4587 | -5.871 | 31.6578 | 4219.4512 | 19.4587 | 31.666
0.3424 | 191.9682 | -203.6575 | -4230.8241 | 4240.2986 | 23.8708 | -19.9509 | -5.9062 | 32.0269 | 4230.8241 | 19.9509 | 31.8334
0.3442 | 189.56 | -206.3645 | -4241.3006 | 4250.8102 | 23.9411 | -20.3904 | -5.9238 | 32.3785 | 4241.3006 | 20.3904 | 32.0007
0.346 | 186.5015 | -209.0364 | -4250.9684 | 4260.4605 | 24.0114 | -20.7419 | -5.9238 | 32.6773 | 4250.9684 | 20.7419 | 32.1681
0.3478 | 183.1968 | -211.3039 | -4260.4253 | 4269.8822 | 24.0641 | -21.1286 | -5.9413 | 32.9761 | 4260.4253 | 21.1286 | 32.3354
0.3496 | 179.857 | -213.5011 | -4269.8998 | 4279.304 | 24.1169 | -21.5329 | -5.9589 | 33.2925 | 4269.8998 | 21.5329 | 32.5028
0.3514 | 176.0075 | -215.4699 | -4278.9173 | 4288.2336 | 24.1872 | -21.9021 | -5.9765 | 33.5738 | 4278.9173 | 21.9021 | 32.6701
0.3532 | 171.3669 | -216.9991 | -4285.8254 | 4295.0362 | 24.2751 | -22.1833 | -6.0116 | 33.8375 | 4285.8254 | 22.1833 | 32.8375
0.355 | 167.3064 | -218.9679 | -4291.4151 | 4300.5557 | 24.3454 | -22.4997 | -6.0292 | 34.066 | 4291.4151 | 22.4997 | 33.0048
0.3567 | 163.3338 | -221.2706 | -4295.1944 | 4304.2997 | 24.3981 | -22.7282 | -6.0644 | 34.2593 | 4295.1944 | 22.7282 | 33.1629
0.3585 | 158.869 | -222.7295 | -4297.8662 | 4306.8837 | 24.486 | -22.9743 | -6.0995 | 34.4878 | 4297.8662 | 22.9743 | 33.3302
0.3603 | 154.1933 | -223.8194 | -4300.1689 | 4309.0634 | 24.5915 | -23.1852 | -6.1523 | 34.7163 | 4300.1689 | 23.1852 | 33.4976
0.3621 | 149.377 | -226.0166 | -4301.6103 | 4310.4696 | 24.6618 | -23.2204 | -6.2226 | 34.8042 | 4301.6103 | 23.2204 | 33.6649
0.3639 | 144.5255 | -228.2314 | -4302.8584 | 4311.7176 | 24.7497 | -23.2731 | -6.2929 | 34.9097 | 4302.8584 | 23.2731 | 33.8323
0.3657 | 139.7267 | -230.165 | -4303.6494 | 4312.4383 | 24.8024 | -23.4313 | -6.328 | 35.0855 | 4303.6494 | 23.4313 | 33.9996
0.3675 | 134.9983 | -232.1864 | -4304.0009 | 4312.7547 | 24.8024 | -23.5895 | -6.3632 | 35.191 | 4304.0009 | 23.5895 | 34.167
0.3693 | 130.3401 | -233.3642 | -4303.6318 | 4312.2977 | 24.8551 | -23.8181 | -6.3984 | 35.3667 | 4303.6318 | 23.8181 | 34.3343
0.3711 | 125.6468 | -234.0145 | -4301.6279 | 4310.2059 | 24.9079 | -24.0641 | -6.4335 | 35.5777 | 4301.6279 | 24.0641 | 34.5017
0.3729 | 121.1644 | -235.368 | -4298.042 | 4306.5673 | 24.9255 | -24.3454 | -6.4687 | 35.7886 | 4298.042 | 24.3454 | 34.669
0.3747 | 116.436 | -235.8602 | -4294.1749 | 4302.5595 | 24.9079 | -24.6442 | -6.4862 | 35.982 | 4294.1749 | 24.6442 | 34.8364
0.3765 | 111.3736 | -236.1239 | -4289.2179 | 4297.4971 | 24.9079 | -24.7321 | -6.539 | 36.0523 | 4289.2179 | 24.7321 | 35.0037
0.3783 | 106.9615 | -237.2489 | -4283.5227 | 4291.7667 | 24.9079 | -24.9606 | -6.5566 | 36.2105 | 4283.5227 | 24.9606 | 35.1711
0.3801 | 102.2331 | -239.0594 | -4276.9134 | 4285.1574 | 24.8903 | -25.1012 | -6.5741 | 36.2984 | 4276.9134 | 25.1012 | 35.3384
0.3818 | 97.8913 | -240.7645 | -4269.3022 | 4277.5462 | 24.8376 | -25.3825 | -6.539 | 36.4566 | 4269.3022 | 25.3825 | 35.4965
0.3836 | 92.6004 | -242.3992 | -4261.2691 | 4269.5131 | 24.8024 | -25.5934 | -6.539 | 36.5796 | 4261.2691 | 25.5934 | 35.6638
0.3854 | 87.5204 | -243.8054 | -4252.7789 | 4261.0054 | 24.7497 | -25.7868 | -6.5566 | 36.6851 | 4252.7789 | 25.7868 | 35.8312
0.3872 | 83.2138 | -245.7214 | -4242.9705 | 4251.2848 | 24.7321 | -25.8922 | -6.5741 | 36.7378 | 4242.9705 | 25.8922 | 35.9985
0.389 | 79.1005 | -246.864 | -4233.8827 | 4242.197 | 24.6794 | -26.1559 | -6.5917 | 36.8433 | 4233.8827 | 26.1559 | 36.1659
0.3908 | 74.917 | -247.5671 | -4226.6758 | 4234.9725 | 24.6442 | -26.2438 | -6.6444 | 36.8784 | 4226.6758 | 26.2438 | 36.3332
0.3926 | 70.8565 | -248.0769 | -4219.7676 | 4228.0293 | 24.6266 | -26.3317 | -6.6972 | 36.9136 | 4219.7676 | 26.3317 | 36.5006
0.3944 | 66.9718 | -248.3757 | -4211.8752 | 4220.1016 | 24.6091 | -26.4372 | -6.7323 | 36.9663 | 4211.8752 | 26.4372 | 36.6679
0.3962 | 62.9113 | -247.9011 | -4201.9085 | 4210.0822 | 24.5563 | -26.5602 | -6.7499 | 37.0366 | 4201.9085 | 26.5602 | 36.8353
0.398 | 57.9895 | -246.8113 | -4191.2387 | 4199.2894 | 24.4684 | -26.7184 | -6.7851 | 37.1069 | 4191.2387 | 26.7184 | 37.0026
0.3998 | 53.8938 | -246.0554 | -4179.8131 | 4187.7934 | 24.4684 | -26.7887 | -6.8026 | 37.1597 | 4179.8131 | 26.7887 | 37.17
0.4016 | 49.7279 | -245.071 | -4166.5418 | 4174.4342 | 24.4333 | -26.8239 | -6.8378 | 37.1773 | 4166.5418 | 26.8239 | 37.3373
0.4034 | 45.6674 | -244.4558 | -4151.9697 | 4159.8094 | 24.3805 | -26.8239 | -6.8378 | 37.1597 | 4151.9697 | 26.8239 | 37.5046
0.4051 | 41.2905 | -243.5242 | -4136.9757 | 4144.7276 | 24.3278 | -26.8414 | -6.8554 | 37.1597 | 4136.9757 | 26.8414 | 37.6627
0.4069 | 37.5991 | -241.8367 | -4121.1029 | 4128.7493 | 24.2751 | -26.7887 | -6.8905 | 37.1069 | 4121.1029 | 26.7887 | 37.83
0.4087 | 34.2945 | -240.3426 | -4104.7379 | 4112.2964 | 24.2048 | -26.736 | -6.9433 | 37.0366 | 4104.7379 | 26.736 | 37.9974
0.4105 | 30.8668 | -239.077 | -4087.617 | 4095.0876 | 24.1345 | -26.6305 | -6.9784 | 36.9312 | 4087.617 | 26.6305 | 38.1647
0.4123 | 27.2457 | -237.9168 | -4071.041 | 4078.4413 | 24.029 | -26.5426 | -7.0312 | 36.8257 | 4071.041 | 26.5426 | 38.3321
0.4141 | 23.2028 | -236.8446 | -4054.5881 | 4061.9181 | 23.9059 | -26.4547 | -7.1015 | 36.7026 | 4054.5881 | 26.4547 | 38.4994
0.4159 | 19.6169 | -234.8934 | -4038.5044 | 4045.7465 | 23.8005 | -26.2965 | -7.1718 | 36.5444 | 4038.5044 | 26.2965 | 38.6668
0.4177 | 16.1717 | -232.4325 | -4023.1413 | 4030.2603 | 23.695 | -26.1559 | -7.2597 | 36.3862 | 4023.1413 | 26.1559 | 38.8341
0.4195 | 13.8865 | -229.6904 | -4004.9306 | 4011.909 | 23.6247 | -25.9977 | -7.3476 | 36.2632 | 4004.9306 | 25.9977 | 39.0015
0.4213 | 11.619 | -226.4736 | -3986.0871 | 3992.9073 | 23.5544 | -25.8571 | -7.4355 | 36.1577 | 3986.0871 | 25.8571 | 39.1688
0.4231 | 10.7049 | -224.1006 | -3965.978 | 3972.6927 | 23.4313 | -25.6813 | -7.5409 | 35.9644 | 3965.978 | 25.6813 | 39.3362
0.4249 | 9.0878 | -222.2901 | -3948.2419 | 3954.8863 | 23.3083 | -25.5055 | -7.6112 | 35.771 | 3948.2419 | 25.5055 | 39.5035
0.4267 | 7.5233 | -220.8487 | -3932.035 | 3938.6267 | 23.1677 | -25.3298 | -7.664 | 35.5777 | 3932.035 | 25.3298 | 39.6709
0.4285 | 6.2753 | -218.9503 | -3914.5802 | 3921.1016 | 23.0095 | -25.1715 | -7.7343 | 35.3667 | 3914.5802 | 25.1715 | 39.8382
0.4302 | 5.5722 | -216.2785 | -3895.016 | 3901.4319 | 22.8513 | -25.0133 | -7.8046 | 35.1734 | 3895.016 | 25.0133 | 39.9963
0.432 | 5.08 | -213.8175 | -3874.2389 | 3880.5494 | 22.6931 | -24.82 | -7.8573 | 34.9449 | 3874.2389 | 24.82 | 40.1636
0.4338 | 4.2011 | -210.2844 | -3853.8485 | 3860.0184 | 22.57 | -24.5563 | -7.9452 | 34.6988 | 3853.8485 | 24.5563 | 40.331
0.4356 | 3.2343 | -207.2434 | -3832.5969 | 3838.6261 | 22.4118 | -24.2751 | -8.0331 | 34.4175 | 3832.5969 | 24.2751 | 40.4983
0.4374 | 2.373 | -204.6595 | -3811.6089 | 3817.5326 | 22.2185 | -23.9938 | -8.1034 | 34.1187 | 3811.6089 | 23.9938 | 40.6657
0.4392 | 1.8808 | -202.1282 | -3790.4802 | 3796.2809 | 22.0251 | -23.7126 | -8.1737 | 33.8023 | 3790.4802 | 23.7126 | 40.833
0.441 | 1.6699 | -199.1576 | -3770.8106 | 3776.4882 | 21.8493 | -23.4841 | -8.2616 | 33.5562 | 3770.8106 | 23.4841 | 41.0004
0.4428 | 1.7051 | -195.9408 | -3752.3186 | 3757.8557 | 21.7087 | -23.2907 | -8.3319 | 33.3453 | 3752.3186 | 23.2907 | 41.1677
0.4446 | 1.9863 | -192.3901 | -3729.5552 | 3734.9165 | 21.5329 | -23.0622 | -8.4198 | 33.0816 | 3729.5552 | 23.0622 | 41.3351
0.4464 | 2.4785 | -189.3491 | -3705.0516 | 3710.2899 | 21.3044 | -22.8337 | -8.4901 | 32.8004 | 3705.0516 | 22.8337 | 41.5024
0.4482 | 2.8652 | -186.8706 | -3680.2844 | 3685.4347 | 21.0759 | -22.5524 | -8.5429 | 32.484 | 3680.2844 | 22.5524 | 41.6698
0.45 | 3.621 | -184.2515 | -3655.2535 | 3660.2983 | 20.8474 | -22.3239 | -8.6132 | 32.2027 | 3655.2535 | 22.3239 | 41.8371
0.4518 | 5.3085 | -181.3687 | -3628.3417 | 3633.2987 | 20.6013 | -22.1306 | -8.6659 | 31.939 | 3628.3417 | 22.1306 | 42.0045
0.4536 | 7.1542 | -178.4332 | -3601.1839 | 3606.0178 | 20.3728 | -21.9548 | -8.7011 | 31.6754 | 3601.1839 | 21.9548 | 42.1718
0.4553 | 8.7538 | -175.3044 | -3573.5866 | 3578.2974 | 20.1267 | -21.779 | -8.7538 | 31.4117 | 3573.5866 | 21.779 | 42.3299
0.4571 | 9.7206 | -171.9821 | -3546.1123 | 3550.6826 | 19.8982 | -21.5681 | -8.8065 | 31.148 | 3546.1123 | 21.5681 | 42.4972
0.4589 | 10.2655 | -168.5017 | -3518.9193 | 3523.3489 | 19.6697 | -21.3571 | -8.8417 | 30.8844 | 3518.9193 | 21.3571 | 42.6646
0.4607 | 10.7049 | -165.1619 | -3491.9548 | 3496.2614 | 19.4587 | -21.1111 | -8.8944 | 30.6031 | 3491.9548 | 21.1111 | 42.8319
0.4625 | 11.285 | -162.1737 | -3463.672 | 3467.8731 | 19.1951 | -20.988 | -8.9296 | 30.357 | 3463.672 | 20.988 | 42.9993
0.4643 | 11.9178 | -159.0624 | -3434.9145 | 3438.9926 | 18.949 | -20.7947 | -8.9471 | 30.0934 | 3434.9145 | 20.7947 | 43.1666
0.4661 | 13.1131 | -156.1093 | -3405.6297 | 3409.6023 | 18.7205 | -20.6189 | -8.9647 | 29.8297 | 3405.6297 | 20.6189 | 43.334
0.4679 | 14.7303 | -153.2441 | -3375.9231 | 3379.8078 | 18.4744 | -20.4783 | -8.9823 | 29.5836 | 3375.9231 | 20.4783 | 43.5013
0.4697 | 16.9451 | -150.4316 | -3344.652 | 3348.4313 | 18.2283 | -20.3025 | -8.9999 | 29.3199 | 3344.652 | 20.3025 | 43.6687
0.4715 | 19.2126 | -147.6543 | -3314.7872 | 3318.4961 | 17.9822 | -20.197 | -8.9999 | 29.0914 | 3314.7872 | 20.197 | 43.836
0.4733 | 21.0407 | -144.7716 | -3285.8188 | 3289.4223 | 17.7361 | -20.0915 | -9.0175 | 28.8805 | 3285.8188 | 20.0915 | 44.0033
0.4751 | 22.4997 | -141.8888 | -3257.2899 | 3260.823 | 17.5076 | -20.0036 | -9.0175 | 28.652 | 3257.2899 | 20.0036 | 44.1707
0.4769 | 24.152 | -139.3576 | -3229.3059 | 3232.7687 | 17.2615 | -19.9158 | -9.0175 | 28.4586 | 3229.3059 | 19.9158 | 44.338
0.4786 | 26.2086 | -136.9318 | -3202.9917 | 3206.4019 | 17.0506 | -19.8454 | -9.035 | 28.2653 | 3202.9917 | 19.8454 | 44.4961
0.4804 | 27.6149 | -134.594 | -3179.3495 | 3182.7068 | 16.8572 | -19.8103 | -9.035 | 28.1246 | 3179.3495 | 19.8103 | 44.6634
0.4822 | 28.1071 | -131.9045 | -3158.2911 | 3161.5782 | 16.7517 | -19.7048 | -9.0526 | 27.984 | 3158.2911 | 19.7048 | 44.8308
0.484 | 28.1246 | -128.8811 | -3136.4067 | 3139.5883 | 16.6111 | -19.6345 | -9.0526 | 27.8434 | 3136.4067 | 19.6345 | 44.9981
0.4858 | 28.4938 | -126.1566 | -3112.0261 | 3115.1198 | 16.4353 | -19.529 | -9.0526 | 27.6676 | 3112.0261 | 19.529 | 45.1655
0.4876 | 28.8102 | -123.889 | -3087.7335 | 3090.7569 | 16.242 | -19.4412 | -9.0702 | 27.4743 | 3087.7335 | 19.4412 | 45.3328
0.4894 | 29.1793 | -121.9027 | -3063.9857 | 3066.9564 | 16.0662 | -19.3357 | -9.0702 | 27.2985 | 3063.9857 | 19.3357 | 45.5002
0.4912 | 30.1285 | -120.3207 | -3041.0466 | 3043.9997 | 15.8728 | -19.2654 | -9.0878 | 27.1227 | 3041.0466 | 19.2654 | 45.6675
0.493 | 31.6754 | -119.1606 | -3018.6172 | 3021.5703 | 15.6795 | -19.1951 | -9.0878 | 26.9645 | 3018.6172 | 19.1951 | 45.8349
0.4948 | 33.2574 | -118.1059 | -2994.887 | 2997.8401 | 15.4861 | -19.1599 | -9.0878 | 26.8063 | 2994.887 | 19.1599 | 46.0022
0.4966 | 34.8042 | -117.1215 | -2970.2076 | 2973.1783 | 15.2752 | -19.1072 | -9.0878 | 26.6305 | 2970.2076 | 19.1072 | 46.1696
0.4984 | 35.9995 | -115.7856 | -2945.5459 | 2948.4989 | 15.0818 | -19.0896 | -9.0702 | 26.4723 | 2945.5459 | 19.0896 | 46.3369
0.5002 | 37.23 | -114.6958 | -2921.0423 | 2923.9602 | 14.9061 | -19.0369 | -9.0702 | 26.3141 | 2921.0423 | 19.0369 | 46.5043
0.502 | 38.5308 | -113.4478 | -2896.2047 | 2899.1226 | 14.7127 | -19.0017 | -9.0526 | 26.1735 | 2896.2047 | 19.0017 | 46.6716
0.5037 | 39.8842 | -112.3403 | -2871.4374 | 2874.3378 | 14.5369 | -18.9841 | -9.035 | 26.0504 | 2871.4374 | 18.9841 | 46.8297
0.5055 | 41.3256 | -110.9868 | -2845.5979 | 2848.4807 | 14.3611 | -19.0017 | -9.0175 | 25.9098 | 2845.5979 | 19.0017 | 46.997
0.5073 | 42.31 | -109.5806 | -2821.0943 | 2823.942 | 14.2029 | -18.9841 | -8.9999 | 25.7868 | 2821.0943 | 18.9841 | 47.1644
0.5091 | 43.1537 | -107.9986 | -2797.8564 | 2800.6688 | 14.0447 | -18.9841 | -8.9823 | 25.6637 | 2797.8564 | 18.9841 | 47.3317
0.5109 | 43.9799 | -107.1197 | -2775.4797 | 2778.257 | 13.9041 | -18.9841 | -8.9823 | 25.5583 | 2775.4797 | 18.9841 | 47.4991
0.5127 | 44.9994 | -106.3463 | -2753.7358 | 2756.5131 | 13.7635 | -18.9841 | -8.9647 | 25.4528 | 2753.7358 | 18.9841 | 47.6664
0.5145 | 45.8607 | -105.0631 | -2732.2381 | 2734.9978 | 13.6404 | -18.9841 | -8.9471 | 25.3473 | 2732.2381 | 18.9841 | 47.8338
0.5163 | 46.5814 | -103.6393 | -2710.7403 | 2713.4649 | 13.4822 | -19.0017 | -8.9296 | 25.2594 | 2710.7403 | 19.0017 | 48.0011
0.5181 | 47.0385 | -102.4616 | -2690.9124 | 2693.6018 | 13.3416 | -19.0369 | -8.912 | 25.1715 | 2690.9124 | 19.0369 | 48.1685
0.5199 | 47.0385 | -101.3014 | -2672.9654 | 2675.6197 | 13.2186 | -19.0369 | -8.8944 | 25.1012 | 2672.9654 | 19.0369 | 48.3358
0.5217 | 46.7396 | -100.194 | -2656.3016 | 2658.9207 | 13.1131 | -19.0369 | -8.8768 | 25.0133 | 2656.3016 | 19.0369 | 48.5032
0.5235 | 46.4232 | -99.4382 | -2641.5537 | 2644.1376 | 13.0076 | -19.0896 | -8.8593 | 24.9782 | 2641.5537 | 19.0896 | 48.6705
0.5253 | 46.2474 | -98.7526 | -2627.0871 | 2629.6535 | 12.9198 | -19.1599 | -8.8241 | 24.9782 | 2627.0871 | 19.1599 | 48.8379
0.5271 | 45.9486 | -97.7507 | -2612.1107 | 2614.6419 | 12.7967 | -19.283 | -8.7714 | 24.9782 | 2612.1107 | 19.283 | 49.0052
0.5288 | 45.7904 | -97.3464 | -2598.5054 | 2601.0191 | 12.6912 | -19.3533 | -8.7538 | 24.943 | 2598.5054 | 19.3533 | 49.1633
0.5306 | 46.0717 | -97.153 | -2584.6892 | 2587.2028 | 12.533 | -19.4587 | -8.7362 | 24.9255 | 2584.6892 | 19.4587 | 49.3306
0.5324 | 46.1244 | -97.0651 | -2571.787 | 2574.3182 | 12.4276 | -19.5115 | -8.7186 | 24.9079 | 2571.787 | 19.5115 | 49.498
0.5342 | 46.0189 | -96.9421 | -2560.4317 | 2562.9629 | 12.3924 | -19.5994 | -8.7011 | 24.943 | 2560.4317 | 19.5994 | 49.6653
0.536 | 45.9135 | -96.7136 | -2549.3752 | 2551.8888 | 12.3397 | -19.7048 | -8.6835 | 24.9958 | 2549.3752 | 19.7048 | 49.8327
0.5378 | 45.3686 | -96.3972 | -2539.4964 | 2542.0101 | 12.3045 | -19.7927 | -8.6483 | 25.0485 | 2539.4964 | 19.7927 | 50
0.5396 | 44.0502 | -95.9226 | -2532.2543 | 2534.6977 | 12.287 | -19.8982 | -8.6307 | 25.1012 | 2532.2543 | 19.8982 | 50.1673
0.5414 | 42.7143 | -95.6589 | -2525.4868 | 2527.8774 | 12.2694 | -20.0036 | -8.5956 | 25.154 | 2525.4868 | 20.0036 | 50.3347
0.5432 | 41.9936 | -96.0105 | -2518.3854 | 2520.7584 | 12.2342 | -20.1443 | -8.578 | 25.2419 | 2518.3854 | 20.1443 | 50.502
0.545 | 41.6596 | -96.4323 | -2509.6843 | 2512.0925 | 12.1991 | -20.3025 | -8.5253 | 25.3473 | 2509.6843 | 20.3025 | 50.6694
0.5468 | 41.2026 | -96.9773 | -2502.0555 | 2504.4813 | 12.1991 | -20.4431 | -8.4901 | 25.4352 | 2502.0555 | 20.4431 | 50.8367
0.5486 | 40.5698 | -97.0476 | -2494.0927 | 2496.5185 | 12.1991 | -20.6013 | -8.4374 | 25.5583 | 2494.0927 | 20.6013 | 51.0041
0.5504 | 39.1108 | -97.0476 | -2489.1006 | 2491.5088 | 12.2342 | -20.7419 | -8.4022 | 25.6813 | 2489.1006 | 20.7419 | 51.1714
0.5521 | 37.8979 | -97.6628 | -2486.0069 | 2488.4151 | 12.287 | -20.8825 | -8.3671 | 25.8044 | 2486.0069 | 20.8825 | 51.3295
0.5539 | 37.0191 | -98.3659 | -2482.2804 | 2484.7061 | 12.3397 | -21.0759 | -8.3319 | 25.9801 | 2482.2804 | 21.0759 | 51.4968
0.5557 | 36.4741 | -99.4733 | -2478.4484 | 2480.9269 | 12.3748 | -21.2517 | -8.2968 | 26.1383 | 2478.4484 | 21.2517 | 51.6642
0.5575 | 35.8413 | -100.6335 | -2475.0559 | 2477.5695 | 12.41 | -21.445 | -8.2616 | 26.2965 | 2475.0559 | 21.445 | 51.8315
0.5593 | 34.6285 | -102.11 | -2474.968 | 2477.5343 | 12.4803 | -21.6032 | -8.2265 | 26.4372 | 2474.968 | 21.6032 | 51.9989
0.5611 | 33.3277 | -103.8678 | -2476.0402 | 2478.6593 | 12.533 | -21.7614 | -8.1737 | 26.5954 | 2476.0402 | 21.7614 | 52.1662
0.5629 | 32.7652 | -105.9068 | -2476.6027 | 2479.3273 | 12.5858 | -22.0251 | -8.1386 | 26.8239 | 2476.6027 | 22.0251 | 52.3336
0.5647 | 32.2554 | -107.9634 | -2477.1476 | 2479.9601 | 12.6209 | -22.3064 | -8.0858 | 27.0524 | 2477.1476 | 22.3064 | 52.5009
0.5665 | 30.8492 | -109.6158 | -2478.7648 | 2481.63 | 12.6912 | -22.5173 | -8.0331 | 27.2633 | 2478.7648 | 22.5173 | 52.6683
0.5683 | 28.7047 | -110.8462 | -2481.5245 | 2484.4249 | 12.7967 | -22.7106 | -7.9979 | 27.4743 | 2481.5245 | 22.7106 | 52.8356
0.5701 | 27.7028 | -112.3403 | -2485.2686 | 2488.2217 | 12.9022 | -22.9919 | -7.9452 | 27.7379 | 2485.2686 | 22.9919 | 53.003
0.5719 | 26.859 | -113.8696 | -2489.0479 | 2492.0537 | 12.9901 | -23.2907 | -7.8925 | 28.0192 | 2489.0479 | 23.2907 | 53.1703
0.5737 | 26.5954 | -115.4692 | -2491.3857 | 2494.4794 | 13.0955 | -23.5544 | -7.8397 | 28.2653 | 2491.3857 | 23.5544 | 53.3377
0.5755 | 25.9274 | -117.2973 | -2492.0888 | 2495.2704 | 13.1834 | -23.7653 | -7.8046 | 28.4938 | 2492.0888 | 23.7653 | 53.505
0.5772 | 24.8024 | -119.1078 | -2493.7587 | 2497.0106 | 13.2889 | -23.9587 | -7.7519 | 28.7047 | 2493.7587 | 23.9587 | 53.6631
0.579 | 22.5173 | -120.4262 | -2499.2079 | 2502.5125 | 13.4295 | -24.1169 | -7.6991 | 28.8981 | 2499.2079 | 24.1169 | 53.8304
0.5808 | 20.9704 | -121.9555 | -2505.8347 | 2509.2097 | 13.5701 | -24.3102 | -7.6112 | 29.109 | 2505.8347 | 24.3102 | 53.9978
0.5826 | 19.8103 | -123.6957 | -2512.5143 | 2515.9772 | 13.7108 | -24.5212 | -7.5233 | 29.3375 | 2512.5143 | 24.5212 | 54.1651
0.5844 | 19.1775 | -125.7171 | -2519.4225 | 2522.9908 | 13.8338 | -24.7321 | -7.5058 | 29.566 | 2519.4225 | 24.7321 | 54.3325
0.5862 | 17.9119 | -127.6155 | -2525.3462 | 2529.0024 | 13.9393 | -24.8727 | -7.453 | 29.7418 | 2525.3462 | 24.8727 | 54.4998
0.588 | 16.3826 | -129.2854 | -2530.7075 | 2534.434 | 14.0799 | -24.9958 | -7.4179 | 29.9 | 2530.7075 | 24.9958 | 54.6672
0.5898 | 15.3103 | -130.7796 | -2537.0883 | 2540.8851 | 14.2029 | -25.1715 | -7.3827 | 30.1109 | 2537.0883 | 25.1715 | 54.8345
0.5916 | 14.4315 | -132.2913 | -2544.9631 | 2548.8127 | 14.3436 | -25.3649 | -7.3651 | 30.3395 | 2544.9631 | 25.3649 | 55.0019
0.5934 | 15.117 | -134.5764 | -2552.9084 | 2556.881 | 14.449 | -25.6462 | -7.3476 | 30.6383 | 2552.9084 | 25.6462 | 55.1692
0.5952 | 15.6971 | -137.09 | -2557.7423 | 2561.8555 | 14.5369 | -25.8922 | -7.3124 | 30.8844 | 2557.7423 | 25.8922 | 55.3366
0.597 | 16.1541 | -139.7619 | -2560.1856 | 2564.4395 | 14.6248 | -26.1383 | -7.2772 | 31.1129 | 2560.1856 | 26.1383 | 55.5039
0.5988 | 16.1365 | -142.6095 | -2562.7696 | 2567.1816 | 14.6951 | -26.402 | -7.2245 | 31.3765 | 2562.7696 | 26.402 | 55.6713
0.6006 | 16.0486 | -145.7735 | -2565.6172 | 2570.205 | 14.783 | -26.7008 | -7.1718 | 31.6578 | 2565.6172 | 26.7008 | 55.8386
0.6023 | 15.6795 | -148.3223 | -2570.5741 | 2575.285 | 14.8885 | -26.9997 | -7.119 | 31.939 | 2570.5741 | 26.9997 | 55.9967
0.6041 | 15.0818 | -150.168 | -2578.5018 | 2583.3005 | 15.0115 | -27.2809 | -7.0839 | 32.2379 | 2578.5018 | 27.2809 | 56.164
0.6059 | 14.4842 | -151.5215 | -2587.3259 | 2592.1774 | 15.1346 | -27.5446 | -7.0487 | 32.5191 | 2587.3259 | 27.5446 | 56.3313
0.6077 | 14.326 | -153.1738 | -2596.5191 | 2601.4585 | 15.2928 | -27.8785 | -7.0136 | 32.8531 | 2596.5191 | 27.8785 | 56.4987
0.6095 | 14.4842 | -155.3007 | -2605.5893 | 2610.6342 | 15.4334 | -28.2125 | -6.9784 | 33.2222 | 2605.5893 | 28.2125 | 56.666
0.6113 | 13.7459 | -157.4101 | -2615.0111 | 2620.1614 | 15.574 | -28.5465 | -6.9257 | 33.5738 | 2615.0111 | 28.5465 | 56.8334
0.6131 | 12.3924 | -159.2909 | -2625.2063 | 2630.4445 | 15.7322 | -28.8805 | -6.8554 | 33.9253 | 2625.2063 | 28.8805 | 57.0007
0.6149 | 11.1444 | -160.8026 | -2637.1592 | 2642.4677 | 15.8728 | -29.2145 | -6.8202 | 34.2769 | 2637.1592 | 29.2145 | 57.1681
0.6167 | 10.3534 | -162.6834 | -2650.0966 | 2655.5105 | 15.9783 | -29.566 | -6.7675 | 34.6285 | 2650.0966 | 29.566 | 57.3354
0.6185 | 10.0897 | -164.9158 | -2665.196 | 2670.733 | 16.1014 | -29.9703 | -6.7499 | 35.0327 | 2665.196 | 29.9703 | 57.5028
0.6203 | 10.2303 | -167.5525 | -2680.77 | 2686.4652 | 16.2068 | -30.3395 | -6.7323 | 35.4195 | 2680.77 | 30.3395 | 57.6701
0.6221 | 9.8085 | -169.9431 | -2694.5862 | 2700.4045 | 16.3123 | -30.6383 | -6.7323 | 35.7183 | 2694.5862 | 30.6383 | 57.8375
0.6239 | 10.0194 | -172.07 | -2706.4513 | 2712.375 | 16.4529 | -30.9371 | -6.7148 | 36.0523 | 2706.4513 | 30.9371 | 58.0048
0.6256 | 10.3358 | -174.6364 | -2717.7187 | 2723.7655 | 16.5584 | -31.2535 | -6.6972 | 36.3687 | 2717.7187 | 31.2535 | 58.1629
0.6274 | 10.6698 | -177.2028 | -2728.9686 | 2735.1736 | 16.6814 | -31.5523 | -6.662 | 36.6851 | 2728.9686 | 31.5523 | 58.3302
0.6292 | 9.9842 | -179.4703 | -2742.7848 | 2749.1128 | 16.7869 | -31.8687 | -6.6444 | 37.0191 | 2742.7848 | 31.8687 | 58.4976
0.631 | 9.2636 | -181.4391 | -2755.5815 | 2762.015 | 16.8924 | -32.2027 | -6.6093 | 37.353 | 2755.5815 | 32.2027 | 58.6649
0.6328 | 8.7186 | -183.1441 | -2770.2063 | 2776.7277 | 17.0154 | -32.5718 | -6.5741 | 37.7222 | 2770.2063 | 32.5718 | 58.8323
0.6346 | 8.0507 | -184.8843 | -2784.005 | 2790.6143 | 17.1385 | -32.9058 | -6.5566 | 38.0737 | 2784.005 | 32.9058 | 58.9996
0.6364 | 7.3827 | -186.5366 | -2798.331 | 2805.0106 | 17.2615 | -33.2398 | -6.539 | 38.4253 | 2798.331 | 33.2398 | 59.167
0.6382 | 7.1718 | -187.5034 | -2815.3991 | 2822.0963 | 17.4021 | -33.5914 | -6.539 | 38.7944 | 2815.3991 | 33.5914 | 59.3343
0.64 | 6.9608 | -188.8921 | -2832.0629 | 2838.7953 | 17.5603 | -33.8902 | -6.5566 | 39.1108 | 2832.0629 | 33.8902 | 59.5017
0.6418 | 6.996 | -189.9468 | -2850.3088 | 2857.0587 | 17.7009 | -34.2417 | -6.5566 | 39.4975 | 2850.3088 | 34.2417 | 59.669
0.6436 | 6.8026 | -191.5815 | -2866.5156 | 2873.3007 | 17.8591 | -34.5406 | -6.5741 | 39.8139 | 2866.5156 | 34.5406 | 59.8364
0.6454 | 6.9081 | -193.603 | -2882.406 | 2889.279 | 17.9998 | -34.8218 | -6.5917 | 40.1303 | 2882.406 | 34.8218 | 60.0037
0.6472 | 7.1718 | -195.5014 | -2898.191 | 2905.1343 | 18.158 | -35.1206 | -6.6269 | 40.4467 | 2898.191 | 35.1206 | 60.1711
0.649 | 7.664 | -196.9955 | -2914.0287 | 2921.0247 | 18.3162 | -35.437 | -6.6444 | 40.7807 | 2914.0287 | 35.437 | 60.3384
0.6507 | 8.2265 | -198.6302 | -2929.5324 | 2936.5811 | 18.4744 | -35.7534 | -6.662 | 41.1147 | 2929.5324 | 35.7534 | 60.4965
0.6525 | 8.8593 | -200.4408 | -2944.0517 | 2951.1708 | 18.615 | -36.0523 | -6.6796 | 41.4487 | 2944.0517 | 36.0523 | 60.6638
0.6543 | 9.8612 | -202.9368 | -2955.9871 | 2963.2116 | 18.738 | -36.4038 | -6.6796 | 41.7827 | 2955.9871 | 36.4038 | 60.8312
0.6561 | 10.2128 | -205.3098 | -2967.9401 | 2975.2877 | 18.8787 | -36.7026 | -6.6796 | 42.0991 | 2967.9401 | 36.7026 | 60.9985
0.6579 | 10.4413 | -206.8743 | -2980.8598 | 2988.2426 | 19.0369 | -36.9839 | -6.6972 | 42.4155 | 2980.8598 | 36.9839 | 61.1659
0.6597 | 10.6698 | -207.9114 | -2994.3773 | 3001.7951 | 19.2126 | -37.2827 | -6.7148 | 42.7494 | 2994.3773 | 37.2827 | 61.3332
0.6615 | 10.7577 | -208.7727 | -3008.2462 | 3015.6817 | 19.406 | -37.564 | -6.7148 | 43.0834 | 3008.2462 | 37.564 | 61.5006
0.6633 | 10.8456 | -209.5988 | -3021.9746 | 3029.4276 | 19.5818 | -37.8276 | -6.7323 | 43.3998 | 3021.9746 | 37.8276 | 61.6679
0.6651 | 10.8104 | -210.5656 | -3034.5779 | 3042.0661 | 19.7751 | -38.0737 | -6.7499 | 43.6987 | 3034.5779 | 38.0737 | 61.8353
0.6669 | 10.6171 | -211.8312 | -3043.5954 | 3051.1363 | 19.9509 | -38.3726 | -6.7323 | 43.9975 | 3043.5954 | 38.3726 | 62.0026
0.6687 | 10.3006 | -213.1672 | -3050.9781 | 3058.5717 | 20.1091 | -38.689 | -6.6796 | 44.3315 | 3050.9781 | 38.689 | 62.17
0.6705 | 9.5448 | -213.8703 | -3059.0463 | 3066.6576 | 20.3025 | -38.9878 | -6.6269 | 44.6479 | 3059.0463 | 38.9878 | 62.3373
0.6723 | 8.8593 | -214.38 | -3067.6771 | 3075.2883 | 20.4783 | -39.2514 | -6.6093 | 44.9467 | 3067.6771 | 39.2514 | 62.5046
0.6741 | 8.2089 | -214.8195 | -3076.9231 | 3084.5519 | 20.6365 | -39.4975 | -6.5917 | 45.2279 | 3076.9231 | 39.4975 | 62.672
0.6758 | 7.453 | -215.4171 | -3088.2081 | 3095.8369 | 20.8298 | -39.726 | -6.6093 | 45.5092 | 3088.2081 | 39.726 | 62.83
0.6776 | 6.539 | -216.1554 | -3101.3388 | 3108.9851 | 21.0583 | -39.937 | -6.6269 | 45.808 | 3101.3388 | 39.937 | 62.9974
0.6794 | 5.8534 | -217.2277 | -3114.0827 | 3121.7819 | 21.3044 | -40.1128 | -6.6796 | 46.0892 | 3114.0827 | 40.1128 | 63.1647
0.6812 | 5.3788 | -218.4581 | -3121.4655 | 3129.2173 | 21.4802 | -40.2534 | -6.6796 | 46.3002 | 3121.4655 | 40.2534 | 63.3321
0.683 | 4.9042 | -219.3897 | -3128.7251 | 3136.5121 | 21.7263 | -40.4116 | -6.6972 | 46.5639 | 3128.7251 | 40.4116 | 63.4994
0.6848 | 4.7812 | -219.8819 | -3133.2602 | 3141.0824 | 21.9548 | -40.5346 | -6.7148 | 46.7572 | 3133.2602 | 40.5346 | 63.6668
0.6866 | 4.7109 | -220.4796 | -3138.0941 | 3145.9515 | 22.1482 | -40.6049 | -6.7323 | 46.933 | 3138.0941 | 40.6049 | 63.8341
0.6884 | 4.3066 | -221.4112 | -3142.5589 | 3150.4514 | 22.2009 | -40.6401 | -6.7675 | 47.0033 | 3142.5589 | 40.6401 | 64.0015
0.6902 | 3.6562 | -222.5713 | -3147.4104 | 3155.3732 | 22.2185 | -40.6577 | -6.7851 | 47.0209 | 3147.4104 | 40.6577 | 64.1688
0.692 | 2.9707 | -223.6963 | -3154.1252 | 3162.1407 | 22.2888 | -40.6928 | -6.8202 | 47.0912 | 3154.1252 | 40.6928 | 64.3362
0.6938 | 2.4785 | -224.5225 | -3162.422 | 3170.4902 | 22.4821 | -40.728 | -6.873 | 47.2318 | 3162.422 | 40.728 | 64.5035
0.6956 | 2.0039 | -225.0498 | -3171.4218 | 3179.4901 | 22.7634 | -40.7983 | -6.9081 | 47.4427 | 3171.4218 | 40.7983 | 64.6709
0.6974 | 1.6347 | -225.3838 | -3180.2284 | 3188.3142 | 23.027 | -40.851 | -6.9433 | 47.6185 | 3180.2284 | 40.851 | 64.8382
0.6992 | 1.1953 | -225.5244 | -3187.7341 | 3195.8024 | 23.2907 | -40.8686 | -6.996 | 47.7591 | 3187.7341 | 40.8686 | 65.0056
0.7009 | -0.3164 | -224.8916 | -3192.5505 | 3200.566 | 23.6774 | -40.728 | -7.0136 | 47.8295 | 3192.5505 | 40.728 | 65.1636
0.7027 | -1.7578 | -223.96 | -3196.2594 | 3204.2046 | 24.0114 | -40.5522 | -7.0663 | 47.8646 | 3196.2594 | 40.5522 | 65.331
0.7045 | -2.9531 | -222.9053 | -3199.6519 | 3207.5093 | 24.2751 | -40.3764 | -7.119 | 47.8646 | 3199.6519 | 40.3764 | 65.4983
0.7063 | -2.1269 | -222.7295 | -3200.8121 | 3208.6342 | 24.5563 | -40.2534 | -7.1718 | 47.9349 | 3200.8121 | 40.2534 | 65.6657
0.7081 | -0.94921 | -223.1163 | -3201.2515 | 3209.1088 | 24.7673 | -40.1479 | -7.2245 | 47.9877 | 3201.2515 | 40.1479 | 65.833
0.7099 | -2.5664 | -223.6787 | -3203.7827 | 3211.6752 | 24.8727 | -39.9721 | -7.2948 | 47.9173 | 3203.7827 | 39.9721 | 66.0004
0.7117 | -4.5703 | -224.1533 | -3207.6323 | 3215.5599 | 24.9958 | -39.7788 | -7.3651 | 47.847 | 3207.6323 | 39.7788 | 66.1677
0.7135 | -7.2069 | -224.4873 | -3214.4701 | 3222.4153 | 25.1715 | -39.5151 | -7.453 | 47.7591 | 3214.4701 | 39.5151 | 66.3351
0.7153 | -9.3866 | -225.2432 | -3223.6985 | 3231.6789 | 25.4352 | -39.269 | -7.5233 | 47.724 | 3223.6985 | 39.269 | 66.5024
0.7171 | -11.6366 | -226.6494 | -3232.9621 | 3241.0303 | 25.6637 | -39.0054 | -7.5937 | 47.6713 | 3232.9621 | 39.0054 | 66.6698
0.7189 | -12.3221 | -227.0537 | -3245.0556 | 3253.1063 | 25.8922 | -38.9702 | -7.664 | 47.7943 | 3245.0556 | 38.9702 | 66.8371
0.7207 | -12.6561 | -227.8974 | -3258.4324 | 3266.5358 | 26.0856 | -39.0229 | -7.7167 | 47.9349 | 3258.4324 | 39.0229 | 67.0045
0.7225 | -13.0955 | -228.1611 | -3272.7057 | 3280.7739 | 26.4196 | -39.1108 | -7.7519 | 48.1986 | 3272.7057 | 39.1108 | 67.1718
0.7242 | -13.6404 | -228.2666 | -3287.1899 | 3295.2405 | 26.7711 | -39.1811 | -7.7694 | 48.4623 | 3287.1899 | 39.1811 | 67.3299
0.726 | -14.8357 | -229.6377 | -3302.1135 | 3310.2345 | 27.0172 | -39.1636 | -7.8222 | 48.5853 | 3302.1135 | 39.1636 | 67.4972
0.7278 | -16.5232 | -231.8349 | -3316.8438 | 3325.0878 | 27.2457 | -39.0054 | -7.8749 | 48.6205 | 3316.8438 | 39.0054 | 67.6646
0.7296 | -17.9119 | -232.6962 | -3334.0701 | 3342.3669 | 27.4743 | -38.8647 | -7.9628 | 48.6556 | 3334.0701 | 38.8647 | 67.8319
0.7314 | -18.3337 | -234.1552 | -3349.5035 | 3357.8706 | 27.7555 | -38.812 | -8.0507 | 48.7787 | 3349.5035 | 38.812 | 67.9993
0.7332 | -18.8084 | -236.0536 | -3363.144 | 3371.6165 | 28.0719 | -38.7417 | -8.121 | 48.9017 | 3363.144 | 38.7417 | 68.1666
0.735 | -20.1443 | -238.5145 | -3374.0598 | 3382.673 | 28.3531 | -38.6186 | -8.1737 | 48.9896 | 3374.0598 | 38.6186 | 68.334
0.7368 | -21.5505 | -240.7117 | -3385.714 | 3394.4678 | 28.652 | -38.5308 | -8.2089 | 49.0775 | 3385.714 | 38.5308 | 68.5013
0.7386 | -23.361 | -241.9246 | -3400.1806 | 3409.0047 | 28.986 | -38.3901 | -8.244 | 49.183 | 3400.1806 | 38.3901 | 68.6687
0.7404 | -26.0504 | -243.8757 | -3413.1179 | 3422.0827 | 29.3024 | -38.0913 | -8.3319 | 49.1654 | 3413.1179 | 38.0913 | 68.836
0.7422 | -27.193 | -245.4226 | -3426.0201 | 3435.0727 | 29.5836 | -37.8979 | -8.4022 | 49.2005 | 3426.0201 | 37.8979 | 69.0033
0.744 | -27.7203 | -246.2136 | -3439.1684 | 3448.2737 | 29.8473 | -37.7925 | -8.455 | 49.2708 | 3439.1684 | 37.7925 | 69.1707
0.7458 | -29.3199 | -247.2155 | -3451.4377 | 3460.5958 | 30.2516 | -37.5991 | -8.5604 | 49.3412 | 3451.4377 | 37.5991 | 69.338
0.7476 | -30.3043 | -248.78 | -3463.3907 | 3472.6543 | 30.5855 | -37.4058 | -8.6132 | 49.3939 | 3463.3907 | 37.4058 | 69.5054
0.7493 | -31.1656 | -250.1686 | -3473.2168 | 3482.5682 | 30.7262 | -37.2651 | -8.6307 | 49.3763 | 3473.2168 | 37.2651 | 69.6634
0.7511 | -32.4136 | -251.5749 | -3484.3611 | 3493.7829 | 30.8668 | -37.1245 | -8.6659 | 49.3412 | 3484.3611 | 37.1245 | 69.8308
0.7529 | -35.8413 | -252.5944 | -3499.6891 | 3509.1636 | 31.1656 | -36.773 | -8.8065 | 49.2533 | 3499.6891 | 36.773 | 69.9981
0.7547 | -39.8842 | -253.4205 | -3515.7553 | 3525.2649 | 31.5172 | -36.3511 | -8.9823 | 49.1302 | 3515.7553 | 36.3511 | 70.1655
0.7565 | -42.8198 | -254.2819 | -3529.466 | 3539.0284 | 31.6929 | -35.9644 | -9.1054 | 48.972 | 3529.466 | 35.9644 | 70.3328
0.7583 | -45.017 | -255.3014 | -3541.4893 | 3551.1044 | 31.7281 | -35.6128 | -9.2284 | 48.7435 | 3541.4893 | 35.6128 | 70.5002
0.7601 | -47.39 | -257.6744 | -3552.7216 | 3562.4773 | 31.6226 | -35.2788 | -9.3163 | 48.4447 | 3552.7216 | 35.2788 | 70.6675
0.7619 | -50.5892 | -260.8912 | -3564.5164 | 3574.483 | 31.6051 | -34.8394 | -9.4393 | 48.1459 | 3564.5164 | 34.8394 | 70.8349
0.7637 | -53.8411 | -263.8618 | -3574.6236 | 3584.8012 | 31.6754 | -34.3472 | -9.58 | 47.8646 | 3574.6236 | 34.3472 | 71.0022
0.7655 | -57.1985 | -266.5512 | -3582.9556 | 3593.3441 | 31.8336 | -33.7847 | -9.7382 | 47.5834 | 3582.9556 | 33.7847 | 71.1696
0.7673 | -60.7141 | -268.3266 | -3589.5649 | 3600.1292 | 32.0621 | -33.1168 | -9.8964 | 47.3021 | 3589.5649 | 33.1168 | 71.3369
0.7691 | -64.423 | -269.2055 | -3594.5042 | 3605.174 | 32.3433 | -32.3609 | -10.0897 | 46.9681 | 3594.5042 | 32.3609 | 71.5043
0.7709 | -67.8507 | -269.9262 | -3598.4944 | 3609.2697 | 32.607 | -31.6226 | -10.2655 | 46.6693 | 3598.4944 | 31.6226 | 71.6716
0.7727 | -70.2237 | -270.0844 | -3598.6526 | 3609.4806 | 32.8355 | -30.9195 | -10.371 | 46.3881 | 3598.6526 | 30.9195 | 71.839
0.7744 | -72.2276 | -270.4711 | -3596.1038 | 3607.0021 | 33.0816 | -30.234 | -10.4588 | 46.1244 | 3596.1038 | 30.234 | 71.997
0.7762 | -74.3721 | -270.6117 | -3591.393 | 3602.3616 | 33.2925 | -29.443 | -10.5643 | 45.7904 | 3591.393 | 29.443 | 72.1644
0.778 | -76.2178 | -269.4692 | -3583.6411 | 3594.5746 | 33.5562 | -28.7223 | -10.6171 | 45.5268 | 3583.6411 | 28.7223 | 72.3317
0.7798 | -78.1338 | -267.6938 | -3573.0065 | 3583.8872 | 33.7847 | -27.984 | -10.6522 | 45.2631 | 3573.0065 | 27.984 | 72.4991
0.7816 | -79.997 | -265.7778 | -3559.8055 | 3570.6335 | 34.0132 | -27.2282 | -10.6698 | 44.9643 | 3559.8055 | 27.2282 | 72.6664
0.7834 | -81.7196 | -263.4224 | -3542.5791 | 3553.3192 | 34.2242 | -26.4547 | -10.6698 | 44.6654 | 3542.5791 | 26.4547 | 72.8338
0.7852 | -83.3368 | -260.6978 | -3522.7864 | 3533.4211 | 34.4351 | -25.7165 | -10.6522 | 44.3842 | 3522.7864 | 25.7165 | 73.0011
0.787 | -84.6376 | -257.5689 | -3499.9176 | 3510.4292 | 34.6109 | -24.9782 | -10.5995 | 44.1029 | 3499.9176 | 24.9782 | 73.1685
0.7888 | -85.9911 | -254.5983 | -3473.8144 | 3484.2205 | 34.7691 | -24.2751 | -10.5292 | 43.8041 | 3473.8144 | 24.2751 | 73.3358
0.7906 | -87.7313 | -252.2077 | -3444.723 | 3455.0764 | 34.8218 | -23.5895 | -10.4237 | 43.4526 | 3444.723 | 23.5895 | 73.5032
0.7924 | -89.5418 | -249.694 | -3414.6296 | 3424.9479 | 34.857 | -22.9216 | -10.3182 | 43.101 | 3414.6296 | 22.9216 | 73.6705
0.7942 | -91.4051 | -246.5476 | -3382.4621 | 3392.6924 | 34.9097 | -22.2712 | -10.1952 | 42.7846 | 3382.4621 | 22.2712 | 73.8379
0.796 | -92.8289 | -242.2762 | -3346.2868 | 3356.3589 | 34.9449 | -21.6735 | -10.0546 | 42.4858 | 3346.2868 | 21.6735 | 74.0052
0.7977 | -93.8484 | -237.3016 | -3307.4396 | 3317.3008 | 34.98 | -21.0935 | -9.9315 | 42.2045 | 3307.4396 | 21.0935 | 74.1633
0.7995 | -94.6921 | -231.9579 | -3265.3581 | 3274.9908 | 34.9976 | -20.5134 | -9.7909 | 41.9057 | 3265.3581 | 20.5134 | 74.3306
0.8013 | -95.2898 | -226.2627 | -3218.3724 | 3227.759 | 34.9624 | -19.9861 | -9.6327 | 41.6245 | 3218.3724 | 19.9861 | 74.498
0.8031 | -96.3269 | -220.5323 | -3167.2031 | 3176.3788 | 34.857 | -19.4236 | -9.4921 | 41.2553 | 3167.2031 | 19.4236 | 74.6653
0.8049 | -97.4343 | -214.6613 | -3114.1179 | 3123.065 | 34.7339 | -18.8962 | -9.3339 | 40.851 | 3114.1179 | 18.8962 | 74.8327
0.8067 | -98.735 | -208.3157 | -3058.2729 | 3067.0091 | 34.5581 | -18.3689 | -9.1757 | 40.4643 | 3058.2729 | 18.3689 | 75
0.8085 | -99.9303 | -201.6712 | -2999.7385 | 3008.2286 | 34.3648 | -17.8767 | -8.9999 | 40.0425 | 2999.7385 | 17.8767 | 75.1673
0.8103 | -101.1784 | -194.6225 | -2938.0049 | 2946.249 | 34.1187 | -17.3494 | -8.8417 | 39.5854 | 2938.0049 | 17.3494 | 75.3347
0.8121 | -102.1627 | -187.2222 | -2874.1796 | 2882.16 | 33.855 | -16.8396 | -8.6835 | 39.0932 | 2874.1796 | 16.8396 | 75.502
0.8139 | -102.444 | -178.8551 | -2806.7859 | 2814.4147 | 33.5211 | -16.3299 | -8.5077 | 38.5659 | 2806.7859 | 16.3299 | 75.6694
0.8157 | -101.8463 | -169.2752 | -2734.8923 | 2742.0993 | 33.1343 | -15.8377 | -8.3143 | 38.0034 | 2734.8923 | 15.8377 | 75.8367
0.8175 | -100.9499 | -159.2909 | -2661.8034 | 2668.5709 | 32.7125 | -15.3631 | -8.1386 | 37.4058 | 2661.8034 | 15.3631 | 76.0041
0.8193 | -99.8776 | -149.6231 | -2588.5212 | 2594.8844 | 32.2027 | -14.8709 | -7.9452 | 36.7378 | 2588.5212 | 14.8709 | 76.1714
0.8211 | -98.7526 | -140.0431 | -2515.6432 | 2521.567 | 31.6226 | -14.3787 | -7.7519 | 35.982 | 2515.6432 | 14.3787 | 76.3388
0.8228 | -97.2058 | -129.6546 | -2439.8825 | 2445.3668 | 31.0074 | -13.8338 | -7.5761 | 35.1734 | 2439.8825 | 13.8338 | 76.4968
0.8246 | -95.4831 | -118.7387 | -2361.0456 | 2366.0553 | 30.3219 | -13.2713 | -7.3651 | 34.3121 | 2361.0456 | 13.2713 | 76.6642
0.8264 | -93.6902 | -107.5767 | -2279.4666 | 2284.0544 | 29.566 | -12.6912 | -7.1542 | 33.3804 | 2279.4666 | 12.6912 | 76.8315
0.8282 | -91.5281 | -95.9577 | -2196.288 | 2200.4364 | 28.7399 | -12.0584 | -6.9257 | 32.3785 | 2196.288 | 12.0584 | 76.9989
0.83 | -89.4012 | -84.497 | -2114.5859 | 2118.2597 | 27.9137 | -11.4081 | -6.7148 | 31.3414 | 2114.5859 | 11.4081 | 77.1662
0.8318 | -87.1161 | -73.5283 | -2031.0382 | 2034.378 | 26.9821 | -10.7577 | -6.4862 | 30.2164 | 2031.0382 | 10.7577 | 77.3336
0.8336 | -84.743 | -62.5246 | -1946.717 | 1949.7228 | 25.9977 | -10.0546 | -6.2402 | 29.0387 | 1946.717 | 10.0546 | 77.5009
0.8354 | -82.3524 | -51.6966 | -1863.7317 | 1866.4036 | 24.9958 | -9.3866 | -5.9941 | 27.8434 | 1863.7317 | 9.3866 | 77.6683
0.8372 | -80.0146 | -41.2729 | -1782.6449 | 1785.0706 | 23.9763 | -8.6835 | -5.7656 | 26.6305 | 1782.6449 | 8.6835 | 77.8356
0.839 | -77.4834 | -31.3414 | -1701.6284 | 1703.8608 | 22.904 | -7.9979 | -5.5195 | 25.3649 | 1701.6284 | 7.9979 | 78.003
0.8408 | -74.8115 | -21.6735 | -1620.6646 | 1622.7212 | 21.8142 | -7.3124 | -5.2558 | 24.0993 | 1620.6646 | 7.3124 | 78.1703
0.8426 | -72.1045 | -11.9354 | -1538.2066 | 1540.1051 | 20.7068 | -6.539 | -5.0097 | 22.781 | 1538.2066 | 6.539 | 78.3377
0.8444 | -69.0987 | -3.1289 | -1457.2604 | 1459.0885 | 19.5642 | -5.8007 | -4.7812 | 21.445 | 1457.2604 | 5.8007 | 78.505
0.8462 | -66.1984 | 4.8515 | -1379.197 | 1380.9899 | 18.4041 | -5.1152 | -4.5351 | 20.1267 | 1379.197 | 5.1152 | 78.6724
0.8479 | -63.7902 | 11.8123 | -1305.1589 | 1306.9343 | 17.2615 | -4.4121 | -4.3242 | 18.8084 | 1305.1589 | 4.4121 | 78.8304
0.8497 | -61.716 | 17.8064 | -1234.8121 | 1236.6051 | 16.1189 | -3.6914 | -4.0957 | 17.5076 | 1234.8121 | 3.6914 | 78.9978
0.8515 | -59.7824 | 22.3415 | -1170.6177 | 1172.4809 | 15.0115 | -3.0761 | -3.9023 | 16.2771 | 1170.6177 | 3.0761 | 79.1651
0.8533 | -57.9895 | 26.1911 | -1110.2727 | 1112.2239 | 13.9217 | -2.4961 | -3.7089 | 15.0994 | 1110.2727 | 2.4961 | 79.3325
0.8551 | -56.4251 | 29.2848 | -1054.8848 | 1056.9238 | 12.9022 | -1.9511 | -3.5332 | 14.0096 | 1054.8848 | 1.9511 | 79.4998
0.8569 | -55.0364 | 31.2887 | -1002.9245 | 1005.0514 | 11.9178 | -1.4414 | -3.375 | 12.9725 | 1002.9245 | 1.4414 | 79.6672
0.8587 | -53.7532 | 33.1343 | -952.8978 | 955.1126 | 10.9686 | -0.96678 | -3.2168 | 11.9705 | 952.8978 | 0.96678 | 79.8345
0.8605 | -52.8567 | 34.8394 | -907.5117 | 909.8671 | 10.0721 | -0.52734 | -3.0761 | 11.0565 | 907.5117 | 0.52734 | 80.0019
0.8623 | -52.4173 | 36.3862 | -867.7856 | 870.2817 | 9.2811 | -0.14062 | -2.9355 | 10.2303 | 867.7856 | 0.14062 | 80.1692
0.8641 | -52.2239 | 37.2827 | -832.3838 | 835.0204 | 8.578 | 0.14062 | -2.83 | 9.5272 | 832.3838 | -0.14062 | 80.3366
0.8659 | -52.2239 | 37.6694 | -799.9877 | 802.7298 | 7.9101 | 0.42187 | -2.7246 | 8.8768 | 799.9877 | -0.42187 | 80.5039
0.8677 | -52.3821 | 37.8452 | -771.8279 | 774.6755 | 7.2597 | 0.75585 | -2.6367 | 8.2265 | 771.8279 | -0.75585 | 80.6713
0.8695 | -52.3821 | 38.0561 | -746.0587 | 748.9942 | 6.662 | 1.0723 | -2.5488 | 7.6464 | 746.0587 | -1.0723 | 80.8386
0.8712 | -52.5052 | 38.0034 | -723.3656 | 726.3715 | 6.1698 | 1.3008 | -2.4609 | 7.2069 | 723.3656 | -1.3008 | 80.9967
0.873 | -52.8567 | 38.0034 | -704.153 | 707.2643 | 5.7656 | 1.5117 | -2.4082 | 6.8554 | 704.153 | -1.5117 | 81.164
0.8748 | -53.4368 | 38.7593 | -689.4051 | 692.657 | 5.4667 | 1.6347 | -2.373 | 6.6269 | 689.4051 | -1.6347 | 81.3313
0.8766 | -53.5071 | 40.1655 | -674.6573 | 678.0674 | 5.3261 | 1.6875 | -2.3554 | 6.5566 | 674.6573 | -1.6875 | 81.4987
0.8784 | -52.8567 | 40.6753 | -657.9934 | 661.509 | 5.2206 | 1.7402 | -2.2851 | 6.5038 | 657.9934 | -1.7402 | 81.666
0.8802 | -52.5052 | 40.3061 | -641.189 | 644.7573 | 5.0449 | 1.8633 | -2.1797 | 6.3808 | 641.189 | -1.8633 | 81.8334
0.882 | -52.6106 | 39.5503 | -626.2302 | 629.8336 | 4.7812 | 2.0566 | -2.0918 | 6.1874 | 626.2302 | -2.0566 | 82.0007
0.8838 | -52.5403 | 38.7417 | -614.8573 | 618.4607 | 4.5351 | 2.2324 | -2.0039 | 6.0468 | 614.8573 | -2.2324 | 82.1681
0.8856 | -52.2591 | 38.1968 | -605.5058 | 609.0741 | 4.3593 | 2.4082 | -1.916 | 5.9062 | 605.5058 | -2.4082 | 82.3354
0.8874 | -52.1888 | 38.4604 | -597.9825 | 601.586 | 4.2011 | 2.6191 | -1.8633 | 5.7656 | 597.9825 | -2.6191 | 82.5028
0.8892 | -52.3821 | 38.1968 | -591.672 | 595.2931 | 4.0429 | 2.8476 | -1.7929 | 5.6952 | 591.672 | -2.8476 | 82.6701
0.891 | -52.6985 | 37.5991 | -586.3459 | 589.9845 | 3.8847 | 3.0761 | -1.7226 | 5.6425 | 586.3459 | -3.0761 | 82.8375
0.8928 | -53.1028 | 36.4038 | -582.3909 | 586.0119 | 3.7089 | 3.2871 | -1.6523 | 5.5898 | 582.3909 | -3.2871 | 83.0048
0.8946 | -53.4368 | 36.1577 | -581.1429 | 584.7815 | 3.5859 | 3.4628 | -1.6172 | 5.5898 | 581.1429 | -3.4628 | 83.1722
0.8963 | -53.4895 | 36.773 | -580.4573 | 584.1487 | 3.5332 | 3.6562 | -1.5644 | 5.6777 | 580.4573 | -3.6562 | 83.3302
0.8981 | -52.6634 | 36.5972 | -575.5531 | 579.1566 | 3.4277 | 3.8671 | -1.5117 | 5.748 | 575.5531 | -3.8671 | 83.4976
0.8999 | -52.681 | 36.2808 | -574.5336 | 578.137 | 3.3222 | 4.0781 | -1.459 | 5.7656 | 574.5336 | -4.0781 | 83.6649
0.9017 | -52.4349 | 36.0171 | -575.4125 | 578.9632 | 3.1992 | 4.3593 | -1.3887 | 5.871 | 575.4125 | -4.3593 | 83.8323
0.9035 | -52.4876 | 36.228 | -579.8421 | 583.4104 | 3.1289 | 4.6757 | -1.2832 | 6.0468 | 579.8421 | -4.6757 | 83.9996
0.9053 | -52.4876 | 36.6675 | -583.5335 | 587.0842 | 3.0586 | 4.9745 | -1.1953 | 6.2226 | 583.5335 | -4.9745 | 84.167
0.9071 | -51.9954 | 37.353 | -582.8831 | 586.4514 | 2.9882 | 5.1679 | -1.125 | 6.3105 | 582.8831 | -5.1679 | 84.3343
0.9089 | -51.6966 | 37.6343 | -581.6702 | 585.2385 | 2.9004 | 5.2734 | -1.0371 | 6.3632 | 581.6702 | -5.2734 | 84.5017
0.9107 | -51.5911 | 37.7573 | -580.0882 | 583.6741 | 2.7773 | 5.3437 | -0.96678 | 6.3632 | 580.0882 | -5.3437 | 84.669
0.9125 | -51.2571 | 38.6011 | -576.4496 | 580.0882 | 2.707 | 5.3964 | -0.89647 | 6.3632 | 576.4496 | -5.3964 | 84.8364
0.9143 | -50.8177 | 39.7612 | -571.4399 | 575.1488 | 2.6191 | 5.5019 | -0.82616 | 6.3984 | 571.4399 | -5.5019 | 85.0037
0.9161 | -50.0443 | 40.7983 | -566.1841 | 569.9282 | 2.5488 | 5.5898 | -0.75585 | 6.4159 | 566.1841 | -5.5898 | 85.1711
0.9179 | -49.1478 | 42.2573 | -560.0494 | 563.8814 | 2.4961 | 5.6073 | -0.68554 | 6.3984 | 560.0494 | -5.6073 | 85.3384
0.9197 | -48.0931 | 43.9096 | -553.1237 | 557.0436 | 2.4433 | 5.6073 | -0.65038 | 6.3456 | 553.1237 | -5.6073 | 85.5058
0.9214 | -46.8978 | 44.8061 | -546.1277 | 550.0651 | 2.3906 | 5.6249 | -0.56249 | 6.328 | 546.1277 | -5.6249 | 85.6638
0.9232 | -45.6674 | 45.7553 | -538.534 | 542.4891 | 2.3203 | 5.6073 | -0.49218 | 6.2929 | 538.534 | -5.6073 | 85.8312
0.925 | -44.4018 | 46.8099 | -529.376 | 533.3837 | 2.2675 | 5.5722 | -0.45703 | 6.205 | 529.376 | -5.5722 | 85.9985
0.9268 | -42.9955 | 47.3373 | -518.8292 | 522.8546 | 2.1797 | 5.537 | -0.36914 | 6.1347 | 518.8292 | -5.537 | 86.1659
0.9286 | -41.519 | 47.7416 | -507.2278 | 511.2883 | 2.0918 | 5.5195 | -0.29882 | 6.0644 | 507.2278 | -5.5195 | 86.3332
0.9304 | -40.0073 | 48.058 | -493.5698 | 497.683 | 2.0215 | 5.537 | -0.19336 | 6.0292 | 493.5698 | -5.537 | 86.5006
0.9322 | -38.6011 | 48.0228 | -479.8063 | 483.9547 | 1.916 | 5.6073 | -0.070312 | 6.0292 | 479.8063 | -5.6073 | 86.6679
0.934 | -37.3882 | 47.1967 | -468.9431 | 472.9685 | 1.8105 | 5.5722 | 0.035156 | 5.9238 | 468.9431 | -5.5722 | 86.8353
0.9358 | -36.2456 | 46.5639 | -457.7109 | 461.6835 | 1.7051 | 5.5019 | 0.10547 | 5.8183 | 457.7109 | -5.5019 | 87.0026
0.9376 | -35.2085 | 45.5795 | -445.3888 | 449.2911 | 1.6172 | 5.3613 | 0.12305 | 5.6425 | 445.3888 | -5.3613 | 87.17
0.9394 | -34.1363 | 44.2084 | -434.7717 | 438.5334 | 1.5117 | 5.2206 | 0.1582 | 5.4843 | 434.7717 | -5.2206 | 87.3373
0.9412 | -32.8355 | 42.8901 | -425.5082 | 429.0941 | 1.3887 | 5.08 | 0.19336 | 5.3085 | 425.5082 | -5.08 | 87.5046
0.943 | -31.4293 | 41.4663 | -416.7719 | 420.1645 | 1.248 | 4.9218 | 0.22851 | 5.1152 | 416.7719 | -4.9218 | 87.672
0.9447 | -30.2691 | 39.5503 | -409.442 | 412.6236 | 1.1074 | 4.7812 | 0.24609 | 4.9394 | 409.442 | -4.7812 | 87.83
0.9465 | -29.4078 | 37.7573 | -403.5006 | 406.4713 | 0.98436 | 4.6054 | 0.24609 | 4.7636 | 403.5006 | -4.6054 | 87.9974
0.9483 | -28.7574 | 35.5952 | -399.0183 | 401.7604 | 0.82616 | 4.4824 | 0.26367 | 4.6054 | 399.0183 | -4.4824 | 88.1647
0.9501 | -28.1422 | 34.1539 | -394.9929 | 397.5593 | 0.72069 | 4.3066 | 0.28125 | 4.4121 | 394.9929 | -4.3066 | 88.3321
0.9519 | -27.7731 | 32.2554 | -392.7957 | 395.1863 | 0.58007 | 4.1835 | 0.29882 | 4.2714 | 392.7957 | -4.1835 | 88.4994
0.9537 | -27.5973 | 30.8492 | -391.9168 | 394.1844 | 0.4746 | 4.0605 | 0.33398 | 4.1308 | 391.9168 | -4.0605 | 88.6668
0.9555 | -27.5446 | 30.0582 | -391.5301 | 393.7273 | 0.36914 | 3.9199 | 0.35156 | 3.9902 | 391.5301 | -3.9199 | 88.8341
0.9573 | -27.2282 | 29.2848 | -392.1102 | 394.2195 | 0.29882 | 3.7617 | 0.36914 | 3.8144 | 392.1102 | -3.7617 | 89.0015
0.9591 | -27.2106 | 28.7047 | -393.3758 | 395.45 | 0.21093 | 3.6035 | 0.36914 | 3.6562 | 393.3758 | -3.6035 | 89.1688
0.9609 | -27.2809 | 28.1246 | -394.782 | 396.8035 | 0.14062 | 3.4277 | 0.36914 | 3.4804 | 394.782 | -3.4277 | 89.3362
0.9627 | -27.5797 | 28.1774 | -394.7293 | 396.7683 | 0.087889 | 3.2343 | 0.35156 | 3.2695 | 394.7293 | -3.2343 | 89.5035
0.9645 | -27.7731 | 28.6168 | -393.6043 | 395.696 | 0.070312 | 3.041 | 0.3164 | 3.0937 | 393.6043 | -3.041 | 89.6709
0.9663 | -27.8082 | 29.3375 | -390.8973 | 393.0418 | 0.070312 | 2.8476 | 0.28125 | 2.8828 | 390.8973 | -2.8476 | 89.8382
0.9681 | -27.2985 | 30.1461 | -385.6415 | 387.8563 | 0.10547 | 2.6543 | 0.24609 | 2.6894 | 385.6415 | -2.6543 | 90.0056
0.9698 | -26.5778 | 30.8141 | -379.7529 | 382.0029 | 0.12305 | 2.4433 | 0.21093 | 2.4785 | 379.7529 | -2.4433 | 90.1636
0.9716 | -25.8747 | 31.4293 | -373.2491 | 375.5342 | 0.14062 | 2.25 | 0.17578 | 2.2851 | 373.2491 | -2.25 | 90.331
0.9734 | -25.2594 | 31.939 | -365.7433 | 368.0812 | 0.1582 | 2.0566 | 0.12305 | 2.1093 | 365.7433 | -2.0566 | 90.4983
0.9752 | -24.7497 | 32.0094 | -356.0403 | 358.4133 | 0.1582 | 1.8808 | 0.070312 | 1.9336 | 356.0403 | -1.8808 | 90.6657
0.977 | -24.3805 | 31.7457 | -344.7201 | 347.0932 | 0.1582 | 1.6875 | 0.017578 | 1.7402 | 344.7201 | -1.6875 | 90.833
0.9788 | -23.8005 | 31.1656 | -333.4527 | 335.8257 | 0.1582 | 1.5117 | -0.035156 | 1.5469 | 333.4527 | -1.5117 | 91.0004
0.9806 | -23.3259 | 29.9879 | -324.3122 | 326.6149 | 0.14062 | 1.3535 | -0.070312 | 1.4062 | 324.3122 | -1.3535 | 91.1677
0.9824 | -22.904 | 27.7379 | -317.5623 | 319.7068 | 0.070312 | 1.2305 | -0.12305 | 1.2832 | 317.5623 | -1.2305 | 91.3351
0.9842 | -22.4646 | 25.4704 | -310.9706 | 312.9217 | 0.017578 | 1.1074 | -0.1582 | 1.1601 | 310.9706 | -1.1074 | 91.5024
0.986 | -21.9372 | 23.3259 | -306.2773 | 308.0527 | -0.070312 | 0.98436 | -0.19336 | 1.0723 | 306.2773 | -0.98436 | 91.6698
0.9878 | -21.3923 | 21.5329 | -303.8164 | 305.4511 | -0.14062 | 0.86132 | -0.21093 | 1.0019 | 303.8164 | -0.86132 | 91.8371
0.9896 | -20.8122 | 20.1618 | -302.3223 | 303.8164 | -0.19336 | 0.72069 | -0.24609 | 0.93163 | 302.3223 | -0.72069 | 92.0045
0.9914 | -20.1443 | 19.5994 | -302.2695 | 303.6933 | -0.21093 | 0.56249 | -0.28125 | 0.84374 | 302.2695 | -0.56249 | 92.1718
0.9932 | -19.3181 | 19.6521 | -304.4492 | 305.8027 | -0.24609 | 0.38671 | -0.3164 | 0.82616 | 304.4492 | -0.38671 | 92.3392
0.9949 | -18.3689 | 19.8982 | -306.9804 | 308.3163 | -0.26367 | 0.19336 | -0.36914 | 0.87889 | 306.9804 | -0.19336 | 92.4972
0.9967 | -17.6482 | 20.6892 | -307.0331 | 308.4218 | -0.24609 | 0.052734 | -0.40429 | 0.93163 | 307.0331 | -0.052734 | 92.6646
0.9985 | -16.9099 | 21.4626 | -306.4355 | 307.8593 | -0.21093 | -0.087889 | -0.45703 | 0.98436 | 306.4355 | 0.087889 | 92.8319
1.0003 | -16.2596 | 21.8493 | -305.416 | 306.8222 | -0.19336 | -0.19336 | -0.49218 | 1.0371 | 305.416 | 0.19336 | 92.9993
1.0021 | -15.785 | 22.2185 | -303.2187 | 304.6425 | -0.17578 | -0.29882 | -0.54491 | 1.0195 | 303.2187 | 0.29882 | 93.1666
1.0039 | -15.4861 | 22.447 | -299.4571 | 300.8984 | -0.1582 | -0.40429 | -0.58007 | 1.0371 | 299.4571 | 0.40429 | 93.334
1.0057 | -15.1697 | 21.8317 | -294.2892 | 295.6954 | -0.17578 | -0.45703 | -0.61523 | 1.0547 | 294.2892 | 0.45703 | 93.5013
1.0075 | -14.9939 | 21.1814 | -288.4357 | 289.8068 | -0.17578 | -0.52734 | -0.65038 | 1.0547 | 288.4357 | 0.52734 | 93.6687
1.0093 | -14.8533 | 19.9509 | -284.1116 | 285.3947 | -0.22851 | -0.56249 | -0.68554 | 1.1074 | 284.1116 | 0.56249 | 93.836
1.0111 | -14.5018 | 18.738 | -282.2835 | 283.5139 | -0.26367 | -0.61523 | -0.70312 | 1.1426 | 282.2835 | 0.61523 | 94.0033
1.0129 | -14.0975 | 17.9119 | -281.8089 | 282.969 | -0.29882 | -0.66796 | -0.72069 | 1.1601 | 281.8089 | 0.66796 | 94.1707
1.0147 | -13.7811 | 17.1033 | -282.3362 | 283.426 | -0.33398 | -0.72069 | -0.73827 | 1.1777 | 282.3362 | 0.72069 | 94.338
1.0165 | -13.4822 | 16.4529 | -284.6213 | 285.6408 | -0.36914 | -0.75585 | -0.73827 | 1.2129 | 284.6213 | 0.75585 | 94.5054
1.0182 | -12.9901 | 15.9256 | -287.8381 | 288.7697 | -0.40429 | -0.80858 | -0.75585 | 1.2656 | 287.8381 | 0.80858 | 94.6634
1.02 | -12.3397 | 16.4353 | -292.004 | 292.9532 | -0.38671 | -0.86132 | -0.75585 | 1.2832 | 292.004 | 0.86132 | 94.8308
1.0218 | -11.6893 | 17.3318 | -295.836 | 296.838 | -0.33398 | -0.91405 | -0.77343 | 1.3183 | 295.836 | 0.91405 | 94.9981
1.0236 | -10.951 | 17.7361 | -297.9454 | 298.9473 | -0.29882 | -0.94921 | -0.79101 | 1.3359 | 297.9454 | 0.94921 | 95.1655
1.0254 | -10.6346 | 18.1755 | -298.1211 | 299.1407 | -0.28125 | -0.96678 | -0.80858 | 1.3535 | 298.1211 | 0.96678 | 95.3328
1.0272 | -10.3885 | 18.1755 | -297.8926 | 298.9121 | -0.24609 | -0.98436 | -0.82616 | 1.3711 | 297.8926 | 0.98436 | 95.5002
1.029 | -10.5819 | 17.8416 | -295.7657 | 296.7676 | -0.24609 | -0.96678 | -0.84374 | 1.3711 | 295.7657 | 0.96678 | 95.6675
1.0308 | -10.9335 | 16.8045 | -293.0411 | 293.9903 | -0.26367 | -0.91405 | -0.86132 | 1.3535 | 293.0411 | 0.91405 | 95.8349
1.0326 | -11.1971 | 15.574 | -292.3732 | 293.2521 | -0.28125 | -0.87889 | -0.86132 | 1.3359 | 292.3732 | 0.87889 | 96.0022
1.0344 | -11.4256 | 14.1502 | -294.6056 | 295.3966 | -0.33398 | -0.82616 | -0.86132 | 1.3008 | 294.6056 | 0.82616 | 96.1696
1.0362 | -11.7245 | 12.8319 | -297.9102 | 298.6309 | -0.36914 | -0.77343 | -0.86132 | 1.2832 | 297.9102 | 0.77343 | 96.3369
1.038 | -11.9178 | 12.4627 | -302.8848 | 303.5703 | -0.36914 | -0.75585 | -0.86132 | 1.2656 | 302.8848 | 0.75585 | 96.5043
1.0398 | -12.0409 | 12.6737 | -309.494 | 310.162 | -0.35156 | -0.75585 | -0.86132 | 1.248 | 309.494 | 0.75585 | 96.6716
1.0416 | -12.0409 | 12.6561 | -315.1365 | 315.8045 | -0.33398 | -0.73827 | -0.86132 | 1.2305 | 315.1365 | 0.73827 | 96.839
1.0433 | -11.9705 | 12.6737 | -319.3377 | 319.988 | -0.3164 | -0.72069 | -0.86132 | 1.2305 | 319.3377 | 0.72069 | 96.997
1.0451 | -11.6541 | 12.7088 | -322.4665 | 323.0993 | -0.29882 | -0.72069 | -0.87889 | 1.2129 | 322.4665 | 0.72069 | 97.1644
1.0469 | -11.1971 | 12.4979 | -323.9606 | 324.5759 | -0.3164 | -0.68554 | -0.87889 | 1.2129 | 323.9606 | 0.68554 | 97.3317
1.0487 | -10.8983 | 12.6385 | -325.6833 | 326.2634 | -0.29882 | -0.68554 | -0.87889 | 1.1953 | 325.6833 | 0.68554 | 97.4991
1.0505 | -10.5819 | 12.8494 | -327.2477 | 327.8278 | -0.29882 | -0.66796 | -0.87889 | 1.1953 | 327.2477 | 0.66796 | 97.6664
1.0523 | -10.2128 | 13.201 | -328.4606 | 329.0582 | -0.29882 | -0.6328 | -0.86132 | 1.1601 | 328.4606 | 0.6328 | 97.8338
1.0541 | -9.8964 | 13.7811 | -328.1618 | 328.777 | -0.28125 | -0.6328 | -0.86132 | 1.1426 | 328.1618 | 0.6328 | 98.0011
1.0559 | -9.58 | 13.7635 | -326.8083 | 327.4059 | -0.29882 | -0.61523 | -0.84374 | 1.1426 | 326.8083 | 0.61523 | 98.1685
1.0577 | -9.1229 | 14.7127 | -325.2263 | 325.8766 | -0.29882 | -0.61523 | -0.82616 | 1.125 | 325.2263 | 0.61523 | 98.3358
1.0595 | -8.7186 | 15.574 | -322.6775 | 323.363 | -0.28125 | -0.61523 | -0.82616 | 1.1074 | 322.6775 | 0.61523 | 98.5032
1.0613 | -8.4374 | 15.7322 | -320.1287 | 320.8318 | -0.3164 | -0.59765 | -0.80858 | 1.1074 | 320.1287 | 0.59765 | 98.6705
1.0631 | -8.1913 | 15.574 | -320.0584 | 320.7615 | -0.33398 | -0.59765 | -0.79101 | 1.1074 | 320.0584 | 0.59765 | 98.8379
1.0649 | -7.8222 | 15.5389 | -320.8494 | 321.5701 | -0.36914 | -0.59765 | -0.77343 | 1.1074 | 320.8494 | 0.59765 | 99.0052
1.0667 | -7.1542 | 16.4353 | -323.6091 | 324.4177 | -0.36914 | -0.6328 | -0.75585 | 1.1074 | 323.6091 | 0.6328 | 99.1726
1.0684 | -6.5741 | 17.1912 | -326.8258 | 327.6872 | -0.38671 | -0.65038 | -0.73827 | 1.1074 | 326.8258 | 0.65038 | 99.3306
1.0702 | -6.0644 | 18.1755 | -327.529 | 328.4606 | -0.38671 | -0.66796 | -0.72069 | 1.1074 | 327.529 | 0.66796 | 99.498
1.072 | -5.6777 | 19.4939 | -325.5954 | 326.5973 | -0.36914 | -0.66796 | -0.72069 | 1.0898 | 325.5954 | 0.66796 | 99.6653
1.0738 | -5.414 | 20.531 | -323.6091 | 324.6638 | -0.35156 | -0.65038 | -0.68554 | 1.0723 | 323.6091 | 0.65038 | 99.8327
1.0756 | -4.8867 | 21.7614 | -322.7829 | 323.9782 | -0.35156 | -0.66796 | -0.68554 | 1.0723 | 322.7829 | 0.66796 | 100
| Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My
Minimum | -102.444 | -270.6117 | -4304.0009 | 282.969 | -3.2168 | -40.8686 | -10.6698 | 0.82616 | 281.8089 | -5.6249
Maximum | 220.4093 | 154.2461 | -281.8089 | 4312.7547 | 34.9976 | 5.6249 | 0.36914 | 49.3939 | 4304.0009 | 40.8686

## Diagr_HIGH100
HIGH100 = Loads at High Level in Subject with 100kg Body Weight.
Contact forces and moments, measured telemetrically in knee implants of 8 subjects.
Averages of one individual subject.
Fx  lateral, Fy  anterior, -Fz  distal, Fres = Resultant Force
Mx around x, My around y, Mz around z, Mres = Resultant Moment
z-axis = axis of implant stem!
From Bergmann et al, PlosOne (2014) "Standardized Loads Acting in Knee Implants"
Subject K8L
Subject K8L
G.Bergmann et al., PlosOne, 2014
G.Bergmann et al., PlosOne, 2014