## Dat_Individ
Data File: | k2l_2242_0_21_i1.akf
Curve Alignment Program: | 1.5.7
------------------------MEASURING INFOS------------------------
Diagram Title #1 | Forces and Moments at Knee Joint
Diagram Title #2 | Measured Relative to Implant
Comment #1 | Knee Joint; Gaitanalysis; Level Walking;
Comment #2 | K2L, 45,5 Months PO *
BodyWeight [N]: | 902.5
Patient(en): | k2l | [7 empty] | Subject | k2l
Implant Type: | Knee Joint
Varus Angle Leg [°] | 5 | [7 empty] | Individual loads in N (forces) and Nm (moments)
Axes Directions X, Y, Z: | Lateral, Anterior, Superior
Axes Origin: | Tibial Plateau at Stem Axis | [7 empty] | BW [kg] = | 91.998
Transformation X Y Z (°): | 11 | -0.5 | 0.5
Displacement X, Y, Z (mm): | 0 | 0 | 0
Activity Code: | KneeActivities | 1P_1_15_2
---------------------------DATA--------------------------------
Number of Data Sets = | 601
Number of analog channels = | 0
Time-Offset(sec) = | 0
Time | Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My | Time
[s] | [N] | [N] | [N] | [N] | [Nm] | [Nm] | [Nm] | [Nm] | [N] | (Nm] | [%Cycle]
0 | 32.03 | 10.6 | -722.88 | 724.21 | 0.32 | -6.86 | 0.33 | 6.96 | 722.88 | 6.86 | 0
0.0019 | 30.51 | 7.74 | -725.34 | 726.62 | 0.43 | -6.86 | 0.4 | 6.98 | 725.34 | 6.86 | 0.16678
0.0038 | 28.64 | 4.51 | -728.31 | 729.51 | 0.56 | -6.86 | 0.47 | 6.98 | 728.31 | 6.86 | 0.33357
0.0057 | 26.55 | 1.23 | -731.76 | 732.89 | 0.7 | -6.84 | 0.55 | 6.99 | 731.76 | 6.84 | 0.50035
0.0076 | 24.46 | -2.55 | -736.57 | 737.68 | 0.85 | -6.84 | 0.64 | 7.01 | 736.57 | 6.84 | 0.66713
0.0095 | 22.57 | -6.46 | -742 | 743.08 | 0.99 | -6.84 | 0.72 | 7.03 | 742 | 6.84 | 0.83392
0.0114 | 20.55 | -10.52 | -748.37 | 749.45 | 1.14 | -6.86 | 0.82 | 7.09 | 748.37 | 6.86 | 1.0007
0.0133 | 18.36 | -15.02 | -755.77 | 756.87 | 1.31 | -6.87 | 0.93 | 7.14 | 755.77 | 6.87 | 1.1675
0.0152 | 15.82 | -19.49 | -763.48 | 764.63 | 1.49 | -6.9 | 1.04 | 7.22 | 763.48 | 6.9 | 1.3343
0.0171 | 13.18 | -23.67 | -770.94 | 772.14 | 1.7 | -6.95 | 1.17 | 7.34 | 770.94 | 6.95 | 1.5011
0.019 | 10.66 | -28.23 | -780.05 | 781.35 | 1.91 | -7.02 | 1.3 | 7.48 | 780.05 | 7.02 | 1.6678
0.0209 | 8.43 | -33.3 | -791.21 | 792.63 | 2.1 | -7.08 | 1.43 | 7.61 | 791.21 | 7.08 | 1.8346
0.0228 | 6.39 | -38.62 | -804.17 | 805.76 | 2.3 | -7.16 | 1.55 | 7.76 | 804.17 | 7.16 | 2.0014
0.0247 | 5.07 | -43.82 | -820.1 | 821.86 | 2.52 | -7.28 | 1.66 | 7.95 | 820.1 | 7.28 | 2.1682
0.0266 | 4.48 | -48.79 | -839.11 | 841.09 | 2.75 | -7.44 | 1.75 | 8.18 | 839.11 | 7.44 | 2.335
0.0285 | 4.66 | -53.93 | -862.07 | 864.3 | 2.95 | -7.62 | 1.81 | 8.43 | 862.07 | 7.62 | 2.5018
0.0304 | 6.26 | -59.59 | -893.17 | 895.68 | 3.1 | -7.83 | 1.82 | 8.67 | 893.17 | 7.83 | 2.6685
0.0323 | 8.65 | -65.28 | -929.09 | 931.9 | 3.24 | -8.06 | 1.8 | 8.93 | 929.09 | 8.06 | 2.8353
0.0342 | 10.95 | -70.56 | -966.28 | 969.36 | 3.38 | -8.29 | 1.77 | 9.17 | 966.28 | 8.29 | 3.0021
0.0361 | 13.73 | -76.86 | -1004.62 | 1008.07 | 3.49 | -8.5 | 1.73 | 9.4 | 1004.62 | 8.5 | 3.1689
0.038 | 16.87 | -84.01 | -1044.06 | 1048.01 | 3.61 | -8.7 | 1.71 | 9.62 | 1044.06 | 8.7 | 3.3357
0.0399 | 20.65 | -92.09 | -1082.64 | 1087.18 | 3.73 | -8.92 | 1.7 | 9.85 | 1082.64 | 8.92 | 3.5025
0.0418 | 24.6 | -100.44 | -1119.18 | 1124.39 | 3.87 | -9.13 | 1.71 | 10.09 | 1119.18 | 9.13 | 3.6692
0.0437 | 28.28 | -108.52 | -1151.46 | 1157.33 | 4.03 | -9.32 | 1.74 | 10.34 | 1151.46 | 9.32 | 3.836
0.0456 | 31.42 | -116.31 | -1178.86 | 1185.38 | 4.24 | -9.5 | 1.8 | 10.6 | 1178.86 | 9.5 | 4.0028
0.0475 | 33.56 | -123.42 | -1200.24 | 1207.42 | 4.49 | -9.65 | 1.88 | 10.86 | 1200.24 | 9.65 | 4.1696
0.0494 | 34.64 | -129.66 | -1215.84 | 1223.65 | 4.8 | -9.77 | 2 | 11.13 | 1215.84 | 9.77 | 4.3364
0.0513 | 34.82 | -135.15 | -1227.62 | 1235.97 | 5.13 | -9.88 | 2.11 | 11.41 | 1227.62 | 9.88 | 4.5032
0.0532 | 34.69 | -140.3 | -1236.25 | 1245.12 | 5.48 | -9.99 | 2.25 | 11.7 | 1236.25 | 9.99 | 4.6699
0.0551 | 34.33 | -145.3 | -1243.42 | 1252.79 | 5.82 | -10.11 | 2.39 | 11.99 | 1243.42 | 10.11 | 4.8367
0.057 | 33.56 | -150.15 | -1249.35 | 1259.23 | 6.18 | -10.23 | 2.53 | 12.31 | 1249.35 | 10.23 | 5.0035
0.0589 | 32.4 | -154.39 | -1254.14 | 1264.46 | 6.54 | -10.39 | 2.68 | 12.67 | 1254.14 | 10.39 | 5.1703
0.0608 | 30.93 | -158.31 | -1257.25 | 1268 | 6.91 | -10.57 | 2.84 | 13.05 | 1257.25 | 10.57 | 5.3371
0.0627 | 29.46 | -162.34 | -1258.56 | 1269.79 | 7.29 | -10.77 | 3.02 | 13.46 | 1258.56 | 10.77 | 5.5039
0.0646 | 28.1 | -166.46 | -1259.28 | 1271.02 | 7.66 | -10.97 | 3.21 | 13.87 | 1259.28 | 10.97 | 5.6706
0.0665 | 26.51 | -170.46 | -1259.37 | 1271.61 | 8.04 | -11.17 | 3.4 | 14.29 | 1259.37 | 11.17 | 5.8374
0.0684 | 24.45 | -174.07 | -1257.98 | 1270.68 | 8.41 | -11.39 | 3.59 | 14.71 | 1257.98 | 11.39 | 6.0042
0.0703 | 22.35 | -177.74 | -1257.08 | 1270.25 | 8.76 | -11.63 | 3.78 | 15.15 | 1257.08 | 11.63 | 6.171
0.0721 | 20.37 | -181.76 | -1257.98 | 1271.67 | 9.11 | -11.92 | 3.96 | 15.61 | 1257.98 | 11.92 | 6.329
0.074 | 18.49 | -186.11 | -1260.7 | 1274.95 | 9.46 | -12.22 | 4.14 | 16.1 | 1260.7 | 12.22 | 6.4958
0.0759 | 16.7 | -190.61 | -1264.65 | 1279.51 | 9.81 | -12.54 | 4.32 | 16.6 | 1264.65 | 12.54 | 6.6626
0.0778 | 15.12 | -195.12 | -1269.57 | 1285.01 | 10.14 | -12.91 | 4.48 | 17.11 | 1269.57 | 12.91 | 6.8294
0.0797 | 13.86 | -199.8 | -1275.16 | 1291.22 | 10.44 | -13.32 | 4.65 | 17.65 | 1275.16 | 13.32 | 6.9961
0.0816 | 12.94 | -204.68 | -1281.61 | 1298.32 | 10.74 | -13.77 | 4.81 | 18.21 | 1281.61 | 13.77 | 7.1629
0.0835 | 12.3 | -209.65 | -1289.3 | 1306.69 | 11.05 | -14.21 | 4.97 | 18.77 | 1289.3 | 14.21 | 7.3297
0.0854 | 11.9 | -214.63 | -1297.93 | 1315.99 | 11.36 | -14.67 | 5.13 | 19.35 | 1297.93 | 14.67 | 7.4965
0.0873 | 11.68 | -219.52 | -1308.15 | 1326.86 | 11.66 | -15.16 | 5.27 | 19.94 | 1308.15 | 15.16 | 7.6633
0.0892 | 11.79 | -224.33 | -1319.93 | 1339.25 | 11.95 | -15.68 | 5.41 | 20.55 | 1319.93 | 15.68 | 7.8301
0.0911 | 12.27 | -229.13 | -1332.71 | 1352.65 | 12.24 | -16.2 | 5.54 | 21.15 | 1332.71 | 16.2 | 7.9968
0.093 | 13.02 | -233.9 | -1346.1 | 1366.66 | 12.52 | -16.73 | 5.66 | 21.76 | 1346.1 | 16.73 | 8.1636
0.0949 | 14.15 | -238.66 | -1359.9 | 1381.08 | 12.8 | -17.28 | 5.78 | 22.38 | 1359.9 | 17.28 | 8.3304
0.0968 | 15.74 | -243.41 | -1374.48 | 1396.27 | 13.06 | -17.87 | 5.9 | 23.01 | 1374.48 | 17.87 | 8.4972
0.0987 | 17.79 | -248.12 | -1390.8 | 1413.18 | 13.31 | -18.46 | 6 | 23.64 | 1390.8 | 18.46 | 8.664
0.1006 | 20.14 | -252.65 | -1408.67 | 1431.59 | 13.55 | -19.05 | 6.08 | 24.26 | 1408.67 | 19.05 | 8.8308
0.1025 | 23.08 | -257.02 | -1428.15 | 1451.57 | 13.77 | -19.67 | 6.15 | 24.9 | 1428.15 | 19.67 | 8.9975
0.1044 | 26.43 | -261.02 | -1448.56 | 1472.41 | 13.99 | -20.28 | 6.21 | 25.53 | 1448.56 | 20.28 | 9.1643
0.1063 | 29.88 | -264.72 | -1470.01 | 1494.24 | 14.22 | -20.89 | 6.26 | 26.16 | 1470.01 | 20.89 | 9.3311
0.1082 | 33.58 | -267.98 | -1492.54 | 1517.07 | 14.44 | -21.48 | 6.29 | 26.77 | 1492.54 | 21.48 | 9.4979
0.1101 | 37.56 | -270.85 | -1516.23 | 1541 | 14.65 | -22.07 | 6.3 | 27.36 | 1516.23 | 22.07 | 9.6647
0.112 | 41.74 | -273.32 | -1541.58 | 1566.51 | 14.85 | -22.64 | 6.29 | 27.93 | 1541.58 | 22.64 | 9.8315
0.1139 | 46.1 | -275.21 | -1567.93 | 1592.93 | 15.04 | -23.19 | 6.26 | 28.48 | 1567.93 | 23.19 | 9.9982
0.1158 | 50.33 | -276.96 | -1593.26 | 1618.33 | 15.24 | -23.74 | 6.24 | 29.03 | 1593.26 | 23.74 | 10.165
0.1177 | 54.32 | -278.76 | -1616.67 | 1641.83 | 15.43 | -24.26 | 6.22 | 29.56 | 1616.67 | 24.26 | 10.3318
0.1196 | 58.56 | -280.3 | -1639.43 | 1664.66 | 15.62 | -24.78 | 6.21 | 30.08 | 1639.43 | 24.78 | 10.4986
0.1215 | 62.85 | -281.56 | -1661.49 | 1686.75 | 15.78 | -25.27 | 6.18 | 30.57 | 1661.49 | 25.27 | 10.6654
0.1234 | 67.06 | -282.71 | -1682.98 | 1708.28 | 15.94 | -25.75 | 6.16 | 31.04 | 1682.98 | 25.75 | 10.8322
0.1253 | 71.11 | -283.69 | -1703.25 | 1728.58 | 16.09 | -26.19 | 6.13 | 31.47 | 1703.25 | 26.19 | 10.9989
0.1272 | 74.77 | -284.61 | -1722.82 | 1748.16 | 16.24 | -26.6 | 6.1 | 31.89 | 1722.82 | 26.6 | 11.1657
0.1291 | 78.25 | -285.24 | -1741.72 | 1767.05 | 16.38 | -26.99 | 6.07 | 32.28 | 1741.72 | 26.99 | 11.3325
0.131 | 81.76 | -285.71 | -1759.85 | 1785.16 | 16.5 | -27.37 | 6.03 | 32.64 | 1759.85 | 27.37 | 11.4993
0.1329 | 85.18 | -286.03 | -1777.03 | 1802.3 | 16.6 | -27.72 | 5.99 | 32.98 | 1777.03 | 27.72 | 11.6661
0.1348 | 88.56 | -286.29 | -1794.15 | 1819.36 | 16.71 | -28.07 | 5.95 | 33.31 | 1794.15 | 28.07 | 11.8329
0.1367 | 91.81 | -286.31 | -1812.31 | 1837.41 | 16.82 | -28.41 | 5.89 | 33.64 | 1812.31 | 28.41 | 11.9996
0.1386 | 94.99 | -286.1 | -1831.15 | 1856.09 | 16.95 | -28.74 | 5.84 | 33.97 | 1831.15 | 28.74 | 12.1664
0.1405 | 97.91 | -285.47 | -1848.38 | 1873.12 | 17.05 | -29.02 | 5.78 | 34.25 | 1848.38 | 29.02 | 12.3332
0.1424 | 100.32 | -284.99 | -1862.44 | 1887.03 | 17.12 | -29.25 | 5.73 | 34.47 | 1862.44 | 29.25 | 12.5
0.1443 | 102.16 | -284.79 | -1874.88 | 1899.37 | 17.17 | -29.41 | 5.68 | 34.62 | 1874.88 | 29.41 | 12.6668
0.1462 | 103.8 | -284.68 | -1886.03 | 1910.44 | 17.2 | -29.55 | 5.63 | 34.75 | 1886.03 | 29.55 | 12.8336
0.1481 | 105.25 | -284.62 | -1895.81 | 1920.16 | 17.22 | -29.67 | 5.59 | 34.85 | 1895.81 | 29.67 | 13.0004
0.15 | 106.48 | -284.55 | -1904.11 | 1928.41 | 17.23 | -29.78 | 5.56 | 34.95 | 1904.11 | 29.78 | 13.1671
0.1519 | 107.4 | -284.36 | -1911.21 | 1935.44 | 17.24 | -29.87 | 5.53 | 35.02 | 1911.21 | 29.87 | 13.3339
0.1538 | 108.09 | -284 | -1917.76 | 1941.89 | 17.24 | -29.95 | 5.5 | 35.09 | 1917.76 | 29.95 | 13.5007
0.1557 | 108.44 | -283.57 | -1923.79 | 1947.82 | 17.24 | -30.01 | 5.46 | 35.13 | 1923.79 | 30.01 | 13.6675
0.1576 | 108.52 | -283.25 | -1929.46 | 1953.38 | 17.24 | -30.05 | 5.43 | 35.16 | 1929.46 | 30.05 | 13.8343
0.1595 | 108.38 | -283.12 | -1934.81 | 1958.63 | 17.23 | -30.07 | 5.39 | 35.17 | 1934.81 | 30.07 | 14.0011
0.1614 | 108.11 | -282.74 | -1939.11 | 1962.82 | 17.22 | -30.09 | 5.35 | 35.17 | 1939.11 | 30.09 | 14.1678
0.1633 | 107.55 | -282.14 | -1942.69 | 1966.24 | 17.21 | -30.1 | 5.31 | 35.17 | 1942.69 | 30.1 | 14.3346
0.1652 | 106.84 | -281.22 | -1945.73 | 1969.08 | 17.21 | -30.1 | 5.27 | 35.17 | 1945.73 | 30.1 | 14.5014
0.1671 | 106.22 | -280.06 | -1947.56 | 1970.69 | 17.21 | -30.12 | 5.24 | 35.18 | 1947.56 | 30.12 | 14.6682
0.169 | 105.23 | -278.63 | -1948.36 | 1971.21 | 17.2 | -30.13 | 5.21 | 35.18 | 1948.36 | 30.13 | 14.835
0.1709 | 103.91 | -277.12 | -1948.34 | 1970.92 | 17.17 | -30.14 | 5.18 | 35.17 | 1948.34 | 30.14 | 15.0018
0.1728 | 102.43 | -275.5 | -1947.78 | 1970.06 | 17.12 | -30.14 | 5.14 | 35.14 | 1947.78 | 30.14 | 15.1685
0.1747 | 100.73 | -273.59 | -1946.16 | 1968.12 | 17.07 | -30.13 | 5.1 | 35.11 | 1946.16 | 30.13 | 15.3353
0.1766 | 98.88 | -271.72 | -1943.94 | 1965.58 | 17.01 | -30.12 | 5.06 | 35.06 | 1943.94 | 30.12 | 15.5021
0.1785 | 96.92 | -269.83 | -1940.9 | 1962.22 | 16.95 | -30.1 | 5.02 | 35.01 | 1940.9 | 30.1 | 15.6689
0.1804 | 94.92 | -267.83 | -1937.18 | 1958.18 | 16.87 | -30.08 | 4.98 | 34.94 | 1937.18 | 30.08 | 15.8357
0.1823 | 93.03 | -265.71 | -1933.49 | 1954.16 | 16.78 | -30.07 | 4.92 | 34.88 | 1933.49 | 30.07 | 16.0025
0.1842 | 91.13 | -263.48 | -1929.71 | 1950.04 | 16.68 | -30.06 | 4.87 | 34.82 | 1929.71 | 30.06 | 16.1692
0.1861 | 89.16 | -261.15 | -1925.48 | 1945.47 | 16.58 | -30.05 | 4.81 | 34.75 | 1925.48 | 30.05 | 16.336
0.188 | 87.37 | -258.89 | -1919.52 | 1939.22 | 16.47 | -30.03 | 4.77 | 34.68 | 1919.52 | 30.03 | 16.5028
0.1899 | 85.41 | -256.55 | -1911.74 | 1931.14 | 16.37 | -30 | 4.73 | 34.6 | 1911.74 | 30 | 16.6696
0.1918 | 83.59 | -254.23 | -1902.51 | 1921.63 | 16.26 | -29.97 | 4.7 | 34.51 | 1902.51 | 29.97 | 16.8364
0.1937 | 81.76 | -251.74 | -1891.12 | 1909.97 | 16.13 | -29.94 | 4.67 | 34.42 | 1891.12 | 29.94 | 17.0032
0.1956 | 79.9 | -249.03 | -1878.26 | 1896.83 | 16 | -29.91 | 4.65 | 34.32 | 1878.26 | 29.91 | 17.1699
0.1975 | 78.1 | -246.05 | -1866.49 | 1884.73 | 15.87 | -29.89 | 4.61 | 34.23 | 1866.49 | 29.89 | 17.3367
0.1994 | 76.32 | -243.04 | -1856.36 | 1874.26 | 15.75 | -29.87 | 4.58 | 34.14 | 1856.36 | 29.87 | 17.5035
0.2013 | 74.53 | -240.14 | -1848.1 | 1865.64 | 15.63 | -29.83 | 4.53 | 34.05 | 1848.1 | 29.83 | 17.6703
0.2032 | 72.72 | -237.43 | -1841.16 | 1858.36 | 15.53 | -29.79 | 4.48 | 33.96 | 1841.16 | 29.79 | 17.8371
0.2051 | 70.86 | -235.02 | -1835.86 | 1852.75 | 15.45 | -29.76 | 4.43 | 33.89 | 1835.86 | 29.76 | 18.0039
0.207 | 69.01 | -232.88 | -1832.64 | 1849.23 | 15.38 | -29.73 | 4.38 | 33.83 | 1832.64 | 29.73 | 18.1706
0.2089 | 67.18 | -230.81 | -1831.7 | 1847.98 | 15.31 | -29.71 | 4.32 | 33.77 | 1831.7 | 29.71 | 18.3374
0.2108 | 65.4 | -228.77 | -1831.46 | 1847.44 | 15.23 | -29.69 | 4.25 | 33.71 | 1831.46 | 29.69 | 18.5042
0.2126 | 63.6 | -226.92 | -1832.68 | 1848.35 | 15.14 | -29.67 | 4.17 | 33.64 | 1832.68 | 29.67 | 18.6622
0.2145 | 61.89 | -225.27 | -1834.49 | 1849.89 | 15.04 | -29.64 | 4.09 | 33.57 | 1834.49 | 29.64 | 18.829
0.2164 | 60.4 | -223.33 | -1834.83 | 1849.95 | 14.94 | -29.61 | 4.01 | 33.48 | 1834.83 | 29.61 | 18.9958
0.2183 | 59.13 | -221.45 | -1833.71 | 1848.58 | 14.85 | -29.57 | 3.95 | 33.41 | 1833.71 | 29.57 | 19.1626
0.2202 | 57.97 | -219.72 | -1831.82 | 1846.48 | 14.77 | -29.54 | 3.89 | 33.33 | 1831.82 | 29.54 | 19.3294
0.2221 | 56.58 | -217.91 | -1828.63 | 1843.06 | 14.69 | -29.5 | 3.85 | 33.26 | 1828.63 | 29.5 | 19.4961
0.224 | 55.16 | -215.87 | -1824.22 | 1838.41 | 14.6 | -29.47 | 3.8 | 33.19 | 1824.22 | 29.47 | 19.6629
0.2259 | 53.59 | -213.73 | -1819.17 | 1833.1 | 14.51 | -29.42 | 3.75 | 33.1 | 1819.17 | 29.42 | 19.8297
0.2278 | 51.95 | -211.53 | -1814.02 | 1827.7 | 14.43 | -29.38 | 3.71 | 33.02 | 1814.02 | 29.38 | 19.9965
0.2297 | 50.28 | -209.32 | -1808.71 | 1822.13 | 14.34 | -29.33 | 3.67 | 32.94 | 1808.71 | 29.33 | 20.1633
0.2316 | 48.52 | -207.28 | -1803.89 | 1817.07 | 14.25 | -29.3 | 3.62 | 32.86 | 1803.89 | 29.3 | 20.3301
0.2335 | 46.97 | -205.53 | -1800.27 | 1813.25 | 14.16 | -29.27 | 3.57 | 32.79 | 1800.27 | 29.27 | 20.4968
0.2354 | 45.53 | -203.81 | -1797.73 | 1810.5 | 14.08 | -29.24 | 3.52 | 32.73 | 1797.73 | 29.24 | 20.6636
0.2373 | 44.31 | -201.93 | -1795.68 | 1808.22 | 13.99 | -29.22 | 3.46 | 32.66 | 1795.68 | 29.22 | 20.8304
0.2392 | 43.25 | -200.23 | -1792.85 | 1805.18 | 13.91 | -29.19 | 3.41 | 32.6 | 1792.85 | 29.19 | 20.9972
0.2411 | 42.11 | -198.76 | -1788.58 | 1800.73 | 13.81 | -29.16 | 3.37 | 32.52 | 1788.58 | 29.16 | 21.164
0.243 | 40.82 | -197.09 | -1782.82 | 1794.78 | 13.7 | -29.11 | 3.33 | 32.43 | 1782.82 | 29.11 | 21.3308
0.2449 | 39.47 | -194.91 | -1775.79 | 1787.5 | 13.59 | -29.06 | 3.29 | 32.33 | 1775.79 | 29.06 | 21.4975
0.2468 | 38.14 | -192.73 | -1768.12 | 1779.59 | 13.47 | -29.01 | 3.25 | 32.23 | 1768.12 | 29.01 | 21.6643
0.2487 | 36.7 | -190.47 | -1758.83 | 1770.06 | 13.34 | -28.96 | 3.22 | 32.12 | 1758.83 | 28.96 | 21.8311
0.2506 | 35.26 | -188.18 | -1749.77 | 1760.75 | 13.22 | -28.9 | 3.19 | 32.02 | 1749.77 | 28.9 | 21.9979
0.2525 | 33.96 | -186.27 | -1742.11 | 1752.89 | 13.1 | -28.85 | 3.16 | 31.92 | 1742.11 | 28.85 | 22.1647
0.2544 | 32.52 | -184.26 | -1736.09 | 1746.63 | 12.99 | -28.81 | 3.12 | 31.82 | 1736.09 | 28.81 | 22.3315
0.2563 | 31.03 | -182.09 | -1730.47 | 1740.76 | 12.88 | -28.75 | 3.07 | 31.72 | 1730.47 | 28.75 | 22.4982
0.2582 | 29.58 | -180.02 | -1723.76 | 1733.83 | 12.77 | -28.69 | 3.03 | 31.61 | 1723.76 | 28.69 | 22.665
0.2601 | 28.16 | -178.05 | -1716.25 | 1726.12 | 12.66 | -28.62 | 3 | 31.5 | 1716.25 | 28.62 | 22.8318
0.262 | 26.7 | -176.1 | -1708.37 | 1718.05 | 12.55 | -28.55 | 2.96 | 31.38 | 1708.37 | 28.55 | 22.9986
0.2639 | 25.3 | -174.21 | -1701.11 | 1710.6 | 12.44 | -28.47 | 2.93 | 31.27 | 1701.11 | 28.47 | 23.1654
0.2658 | 24.03 | -172.54 | -1694.91 | 1704.23 | 12.34 | -28.42 | 2.9 | 31.17 | 1694.91 | 28.42 | 23.3322
0.2677 | 22.87 | -171.05 | -1689.5 | 1698.68 | 12.24 | -28.38 | 2.87 | 31.09 | 1689.5 | 28.38 | 23.4989
0.2696 | 21.73 | -169.65 | -1684.81 | 1693.84 | 12.16 | -28.35 | 2.84 | 31.03 | 1684.81 | 28.35 | 23.6657
0.2715 | 20.52 | -168.19 | -1680.27 | 1689.15 | 12.08 | -28.32 | 2.81 | 30.96 | 1680.27 | 28.32 | 23.8325
0.2734 | 19.36 | -166.7 | -1675.16 | 1683.9 | 11.99 | -28.29 | 2.78 | 30.89 | 1675.16 | 28.29 | 23.9993
0.2753 | 18.33 | -165.27 | -1670.46 | 1679.06 | 11.9 | -28.26 | 2.75 | 30.82 | 1670.46 | 28.26 | 24.1661
0.2772 | 17.48 | -163.96 | -1666.65 | 1675.12 | 11.82 | -28.22 | 2.71 | 30.75 | 1666.65 | 28.22 | 24.3329
0.2791 | 16.63 | -162.72 | -1663.58 | 1671.92 | 11.74 | -28.19 | 2.67 | 30.68 | 1663.58 | 28.19 | 24.4996
0.281 | 15.79 | -161.45 | -1661.12 | 1669.32 | 11.66 | -28.15 | 2.63 | 30.62 | 1661.12 | 28.15 | 24.6664
0.2829 | 15 | -160.22 | -1658.85 | 1666.92 | 11.59 | -28.12 | 2.59 | 30.56 | 1658.85 | 28.12 | 24.8332
0.2848 | 14.27 | -159.23 | -1656.56 | 1664.52 | 11.52 | -28.09 | 2.55 | 30.5 | 1656.56 | 28.09 | 25
0.2867 | 13.47 | -158.43 | -1654.39 | 1662.27 | 11.44 | -28.06 | 2.51 | 30.43 | 1654.39 | 28.06 | 25.1668
0.2886 | 12.76 | -157.74 | -1652.52 | 1660.34 | 11.37 | -28.02 | 2.48 | 30.37 | 1652.52 | 28.02 | 25.3336
0.2905 | 12.24 | -157.18 | -1651.2 | 1658.96 | 11.31 | -28 | 2.44 | 30.33 | 1651.2 | 28 | 25.5004
0.2924 | 11.78 | -156.64 | -1650.34 | 1658.05 | 11.24 | -27.99 | 2.4 | 30.28 | 1650.34 | 27.99 | 25.6671
0.2943 | 11.33 | -156.01 | -1650.1 | 1657.74 | 11.18 | -27.98 | 2.36 | 30.25 | 1650.1 | 27.98 | 25.8339
0.2962 | 10.9 | -155.27 | -1650.47 | 1658.03 | 11.12 | -27.98 | 2.31 | 30.22 | 1650.47 | 27.98 | 26.0007
0.2981 | 10.49 | -154.38 | -1651.83 | 1659.3 | 11.07 | -27.98 | 2.26 | 30.19 | 1651.83 | 27.98 | 26.1675
0.3 | 10.16 | -153.57 | -1654.08 | 1661.46 | 11.03 | -27.98 | 2.2 | 30.17 | 1654.08 | 27.98 | 26.3343
0.3019 | 9.88 | -152.86 | -1656.92 | 1664.22 | 10.99 | -27.98 | 2.14 | 30.16 | 1656.92 | 27.98 | 26.5011
0.3038 | 9.54 | -151.97 | -1659.98 | 1667.18 | 10.94 | -27.98 | 2.08 | 30.13 | 1659.98 | 27.98 | 26.6678
0.3057 | 9.16 | -150.97 | -1662.94 | 1670.02 | 10.89 | -27.96 | 2 | 30.09 | 1662.94 | 27.96 | 26.8346
0.3076 | 8.77 | -149.94 | -1665.35 | 1672.31 | 10.84 | -27.94 | 1.93 | 30.05 | 1665.35 | 27.94 | 27.0014
0.3095 | 8.53 | -149.12 | -1666.22 | 1673.09 | 10.81 | -27.93 | 1.88 | 30.02 | 1666.22 | 27.93 | 27.1682
0.3114 | 8.47 | -148.58 | -1665.57 | 1672.39 | 10.76 | -27.93 | 1.84 | 30 | 1665.57 | 27.93 | 27.335
0.3133 | 8.49 | -148.19 | -1663.86 | 1670.64 | 10.72 | -27.93 | 1.81 | 29.98 | 1663.86 | 27.93 | 27.5018
0.3152 | 8.51 | -148 | -1661.76 | 1668.53 | 10.67 | -27.93 | 1.79 | 29.97 | 1661.76 | 27.93 | 27.6685
0.3171 | 8.52 | -147.88 | -1660.42 | 1667.17 | 10.62 | -27.94 | 1.77 | 29.96 | 1660.42 | 27.94 | 27.8353
0.319 | 8.48 | -147.71 | -1660.59 | 1667.32 | 10.58 | -27.95 | 1.73 | 29.95 | 1660.59 | 27.95 | 28.0021
0.3209 | 8.32 | -147.37 | -1662.33 | 1669.02 | 10.55 | -27.96 | 1.69 | 29.95 | 1662.33 | 27.96 | 28.1689
0.3228 | 8.25 | -147.01 | -1664.81 | 1671.46 | 10.52 | -27.96 | 1.64 | 29.94 | 1664.81 | 27.96 | 28.3357
0.3247 | 8.27 | -146.62 | -1668.24 | 1674.83 | 10.5 | -27.99 | 1.59 | 29.95 | 1668.24 | 27.99 | 28.5025
0.3266 | 8.35 | -146.3 | -1672.73 | 1679.28 | 10.48 | -28.02 | 1.54 | 29.96 | 1672.73 | 28.02 | 28.6692
0.3285 | 8.47 | -146.09 | -1677.62 | 1684.13 | 10.46 | -28.05 | 1.48 | 29.99 | 1677.62 | 28.05 | 28.836
0.3304 | 8.78 | -146.15 | -1680.93 | 1687.43 | 10.44 | -28.1 | 1.44 | 30.02 | 1680.93 | 28.1 | 29.0028
0.3323 | 9.3 | -146.5 | -1682.72 | 1689.24 | 10.43 | -28.16 | 1.42 | 30.08 | 1682.72 | 28.16 | 29.1696
0.3342 | 9.86 | -146.89 | -1684.61 | 1691.16 | 10.42 | -28.23 | 1.4 | 30.14 | 1684.61 | 28.23 | 29.3364
0.3361 | 10.33 | -147.09 | -1687.83 | 1694.38 | 10.41 | -28.29 | 1.37 | 30.19 | 1687.83 | 28.29 | 29.5032
0.338 | 10.69 | -146.97 | -1692.6 | 1699.12 | 10.4 | -28.35 | 1.32 | 30.24 | 1692.6 | 28.35 | 29.6699
0.3399 | 11.12 | -146.72 | -1698.84 | 1705.32 | 10.42 | -28.41 | 1.26 | 30.3 | 1698.84 | 28.41 | 29.8367
0.3418 | 11.74 | -146.64 | -1706.26 | 1712.71 | 10.43 | -28.47 | 1.2 | 30.36 | 1706.26 | 28.47 | 30.0035
0.3437 | 12.44 | -146.73 | -1713.55 | 1719.99 | 10.45 | -28.53 | 1.14 | 30.42 | 1713.55 | 28.53 | 30.1703
0.3456 | 13.17 | -146.83 | -1720.16 | 1726.58 | 10.47 | -28.59 | 1.08 | 30.48 | 1720.16 | 28.59 | 30.3371
0.3475 | 13.95 | -146.7 | -1725.07 | 1731.48 | 10.5 | -28.65 | 1.04 | 30.55 | 1725.07 | 28.65 | 30.5039
0.3494 | 14.66 | -146.62 | -1729.19 | 1735.58 | 10.51 | -28.71 | 1 | 30.61 | 1729.19 | 28.71 | 30.6706
0.3513 | 15.21 | -146.36 | -1733.32 | 1739.69 | 10.53 | -28.76 | 0.95 | 30.66 | 1733.32 | 28.76 | 30.8374
0.3531 | 15.56 | -145.86 | -1738.8 | 1745.11 | 10.57 | -28.81 | 0.9 | 30.72 | 1738.8 | 28.81 | 30.9954
0.355 | 15.74 | -145.22 | -1745.59 | 1751.82 | 10.62 | -28.86 | 0.84 | 30.78 | 1745.59 | 28.86 | 31.1622
0.3569 | 15.97 | -144.6 | -1752.64 | 1758.81 | 10.66 | -28.91 | 0.78 | 30.85 | 1752.64 | 28.91 | 31.329
0.3588 | 16.29 | -144.17 | -1758.85 | 1764.97 | 10.7 | -28.97 | 0.73 | 30.91 | 1758.85 | 28.97 | 31.4958
0.3607 | 16.74 | -144 | -1763.72 | 1769.81 | 10.72 | -29.02 | 0.69 | 30.96 | 1763.72 | 29.02 | 31.6626
0.3626 | 17.15 | -143.81 | -1767.33 | 1773.41 | 10.74 | -29.06 | 0.65 | 31.01 | 1767.33 | 29.06 | 31.8294
0.3645 | 17.51 | -143.68 | -1769.79 | 1775.85 | 10.76 | -29.09 | 0.62 | 31.05 | 1769.79 | 29.09 | 31.9961
0.3664 | 17.82 | -143.58 | -1771.83 | 1777.89 | 10.78 | -29.13 | 0.59 | 31.1 | 1771.83 | 29.13 | 32.1629
0.3683 | 18.06 | -143.35 | -1774.51 | 1780.54 | 10.8 | -29.18 | 0.56 | 31.15 | 1774.51 | 29.18 | 32.3297
0.3702 | 18.39 | -143.09 | -1778.64 | 1784.65 | 10.83 | -29.24 | 0.52 | 31.21 | 1778.64 | 29.24 | 32.4965
0.3721 | 18.92 | -142.91 | -1784.43 | 1790.41 | 10.85 | -29.3 | 0.48 | 31.28 | 1784.43 | 29.3 | 32.6633
0.374 | 19.44 | -142.64 | -1791.47 | 1797.42 | 10.89 | -29.37 | 0.42 | 31.36 | 1791.47 | 29.37 | 32.8301
0.3759 | 19.61 | -141.88 | -1799.34 | 1805.22 | 10.95 | -29.44 | 0.37 | 31.44 | 1799.34 | 29.44 | 32.9968
0.3778 | 19.83 | -141.22 | -1807.2 | 1813.02 | 11 | -29.5 | 0.31 | 31.52 | 1807.2 | 29.5 | 33.1636
0.3797 | 20.15 | -140.71 | -1814.14 | 1819.91 | 11.04 | -29.56 | 0.26 | 31.59 | 1814.14 | 29.56 | 33.3304
0.3816 | 20.51 | -140.28 | -1819.76 | 1825.5 | 11.08 | -29.62 | 0.21 | 31.66 | 1819.76 | 29.62 | 33.4972
0.3835 | 21.07 | -140.06 | -1824.89 | 1830.61 | 11.1 | -29.68 | 0.17 | 31.72 | 1824.89 | 29.68 | 33.664
0.3854 | 21.82 | -139.75 | -1830.81 | 1836.49 | 11.13 | -29.75 | 0.12 | 31.79 | 1830.81 | 29.75 | 33.8308
0.3873 | 22.65 | -139.18 | -1837.67 | 1843.3 | 11.18 | -29.84 | 0.07 | 31.89 | 1837.67 | 29.84 | 33.9975
0.3892 | 23.43 | -138.7 | -1843.5 | 1849.09 | 11.22 | -29.92 | 0.04 | 31.98 | 1843.5 | 29.92 | 34.1643
0.3911 | 24.19 | -138.24 | -1848.54 | 1854.09 | 11.25 | -29.99 | 0 | 32.06 | 1848.54 | 29.99 | 34.3311
0.393 | 24.95 | -137.74 | -1852.64 | 1858.15 | 11.29 | -30.06 | -0.04 | 32.14 | 1852.64 | 30.06 | 34.4979
0.3949 | 25.57 | -137.32 | -1855.58 | 1861.05 | 11.31 | -30.1 | -0.07 | 32.19 | 1855.58 | 30.1 | 34.6647
0.3968 | 26.18 | -137.01 | -1857.79 | 1863.24 | 11.34 | -30.15 | -0.1 | 32.24 | 1857.79 | 30.15 | 34.8315
0.3987 | 26.76 | -136.93 | -1859.18 | 1864.63 | 11.36 | -30.18 | -0.12 | 32.28 | 1859.18 | 30.18 | 34.9982
0.4006 | 27.34 | -136.87 | -1859.9 | 1865.35 | 11.38 | -30.22 | -0.13 | 32.32 | 1859.9 | 30.22 | 35.165
0.4025 | 27.96 | -136.54 | -1860.03 | 1865.47 | 11.42 | -30.27 | -0.13 | 32.39 | 1860.03 | 30.27 | 35.3318
0.4044 | 28.56 | -136.08 | -1860.14 | 1865.56 | 11.45 | -30.33 | -0.14 | 32.45 | 1860.14 | 30.33 | 35.4986
0.4063 | 29 | -135.48 | -1861.2 | 1866.58 | 11.48 | -30.38 | -0.16 | 32.51 | 1861.2 | 30.38 | 35.6654
0.4082 | 29.39 | -134.99 | -1863.35 | 1868.71 | 11.51 | -30.43 | -0.18 | 32.57 | 1863.35 | 30.43 | 35.8322
0.4101 | 29.67 | -134.4 | -1866.1 | 1871.41 | 11.54 | -30.49 | -0.21 | 32.63 | 1866.1 | 30.49 | 35.9989
0.412 | 29.94 | -133.81 | -1869.04 | 1874.31 | 11.59 | -30.55 | -0.24 | 32.7 | 1869.04 | 30.55 | 36.1657
0.4139 | 30.39 | -133.17 | -1872.04 | 1877.26 | 11.63 | -30.61 | -0.28 | 32.77 | 1872.04 | 30.61 | 36.3325
0.4158 | 31.01 | -132.59 | -1875.28 | 1880.46 | 11.66 | -30.67 | -0.31 | 32.85 | 1875.28 | 30.67 | 36.4993
0.4177 | 31.7 | -132.1 | -1879.18 | 1884.33 | 11.69 | -30.75 | -0.33 | 32.93 | 1879.18 | 30.75 | 36.6661
0.4196 | 32.31 | -131.54 | -1883.64 | 1888.75 | 11.72 | -30.83 | -0.36 | 33.02 | 1883.64 | 30.83 | 36.8329
0.4215 | 32.77 | -130.82 | -1888.59 | 1893.64 | 11.76 | -30.9 | -0.4 | 33.1 | 1888.59 | 30.9 | 36.9996
0.4234 | 33.25 | -130.06 | -1893.22 | 1898.23 | 11.81 | -30.99 | -0.43 | 33.2 | 1893.22 | 30.99 | 37.1664
0.4253 | 33.71 | -129.21 | -1898.09 | 1903.04 | 11.85 | -31.09 | -0.46 | 33.31 | 1898.09 | 31.09 | 37.3332
0.4272 | 34.16 | -128.37 | -1904.61 | 1909.49 | 11.89 | -31.19 | -0.5 | 33.42 | 1904.61 | 31.19 | 37.5
0.4291 | 34.68 | -127.4 | -1912.2 | 1917.02 | 11.93 | -31.3 | -0.54 | 33.54 | 1912.2 | 31.3 | 37.6668
0.431 | 35.16 | -126.46 | -1920.15 | 1924.9 | 11.97 | -31.39 | -0.59 | 33.64 | 1920.15 | 31.39 | 37.8336
0.4329 | 35.68 | -125.46 | -1928.79 | 1933.48 | 12.01 | -31.5 | -0.64 | 33.76 | 1928.79 | 31.5 | 38.0004
0.4348 | 36.3 | -124.27 | -1937.96 | 1942.58 | 12.06 | -31.62 | -0.7 | 33.89 | 1937.96 | 31.62 | 38.1671
0.4367 | 36.95 | -123.27 | -1946.35 | 1950.89 | 12.1 | -31.73 | -0.75 | 34.01 | 1946.35 | 31.73 | 38.3339
0.4386 | 37.68 | -122.54 | -1953.91 | 1958.4 | 12.14 | -31.84 | -0.79 | 34.12 | 1953.91 | 31.84 | 38.5007
0.4405 | 38.53 | -121.9 | -1961.54 | 1966 | 12.16 | -31.95 | -0.84 | 34.24 | 1961.54 | 31.95 | 38.6675
0.4424 | 39.38 | -121.22 | -1969.73 | 1974.15 | 12.19 | -32.06 | -0.88 | 34.36 | 1969.73 | 32.06 | 38.8343
0.4443 | 40.11 | -120.47 | -1977.57 | 1981.94 | 12.23 | -32.17 | -0.93 | 34.47 | 1977.57 | 32.17 | 39.0011
0.4462 | 40.82 | -119.79 | -1984.45 | 1988.78 | 12.28 | -32.26 | -0.97 | 34.57 | 1984.45 | 32.26 | 39.1678
0.4481 | 41.52 | -119.01 | -1990.31 | 1994.59 | 12.33 | -32.34 | -1.01 | 34.66 | 1990.31 | 32.34 | 39.3346
0.45 | 42.08 | -118.07 | -1995.3 | 1999.52 | 12.38 | -32.39 | -1.05 | 34.73 | 1995.3 | 32.39 | 39.5014
0.4519 | 42.64 | -117.28 | -1998.17 | 2002.35 | 12.44 | -32.45 | -1.08 | 34.81 | 1998.17 | 32.45 | 39.6682
0.4538 | 43.17 | -116.39 | -2000.46 | 2004.59 | 12.51 | -32.52 | -1.1 | 34.89 | 2000.46 | 32.52 | 39.835
0.4557 | 43.63 | -115.15 | -2003.13 | 2007.21 | 12.58 | -32.6 | -1.12 | 34.99 | 2003.13 | 32.6 | 40.0018
0.4576 | 44.21 | -113.88 | -2006.6 | 2010.6 | 12.63 | -32.69 | -1.14 | 35.1 | 2006.6 | 32.69 | 40.1685
0.4595 | 44.89 | -112.62 | -2010.83 | 2014.77 | 12.68 | -32.79 | -1.16 | 35.21 | 2010.83 | 32.79 | 40.3353
0.4614 | 45.59 | -111.36 | -2016.89 | 2020.77 | 12.73 | -32.91 | -1.19 | 35.34 | 2016.89 | 32.91 | 40.5021
0.4633 | 46.26 | -110.07 | -2024.24 | 2028.05 | 12.77 | -33.03 | -1.23 | 35.47 | 2024.24 | 33.03 | 40.6689
0.4652 | 46.94 | -108.7 | -2032.72 | 2036.46 | 12.81 | -33.16 | -1.27 | 35.61 | 2032.72 | 33.16 | 40.8357
0.4671 | 47.74 | -107.35 | -2041.59 | 2045.27 | 12.85 | -33.29 | -1.31 | 35.75 | 2041.59 | 33.29 | 41.0025
0.469 | 48.44 | -105.87 | -2051.08 | 2054.68 | 12.89 | -33.43 | -1.35 | 35.89 | 2051.08 | 33.43 | 41.1692
0.4709 | 49.08 | -104.37 | -2060.93 | 2064.47 | 12.92 | -33.57 | -1.4 | 36.03 | 2060.93 | 33.57 | 41.336
0.4728 | 49.72 | -102.89 | -2070.88 | 2074.35 | 12.96 | -33.7 | -1.44 | 36.17 | 2070.88 | 33.7 | 41.5028
0.4747 | 50.36 | -101.3 | -2080.69 | 2084.09 | 12.99 | -33.83 | -1.49 | 36.31 | 2080.69 | 33.83 | 41.6696
0.4766 | 50.99 | -99.69 | -2090.28 | 2093.61 | 13.03 | -33.97 | -1.53 | 36.45 | 2090.28 | 33.97 | 41.8364
0.4785 | 51.67 | -98.06 | -2100.25 | 2103.52 | 13.06 | -34.1 | -1.57 | 36.58 | 2100.25 | 34.1 | 42.0032
0.4804 | 52.52 | -96.53 | -2110.66 | 2113.87 | 13.09 | -34.22 | -1.61 | 36.71 | 2110.66 | 34.22 | 42.1699
0.4823 | 53.52 | -95.14 | -2121.04 | 2124.2 | 13.12 | -34.35 | -1.65 | 36.84 | 2121.04 | 34.35 | 42.3367
0.4842 | 54.52 | -93.79 | -2130.94 | 2134.06 | 13.14 | -34.48 | -1.69 | 36.97 | 2130.94 | 34.48 | 42.5035
0.4861 | 55.56 | -92.44 | -2140.19 | 2143.27 | 13.16 | -34.6 | -1.72 | 37.1 | 2140.19 | 34.6 | 42.6703
0.488 | 56.59 | -91.12 | -2148.33 | 2151.36 | 13.18 | -34.73 | -1.74 | 37.23 | 2148.33 | 34.73 | 42.8371
0.4899 | 57.59 | -89.81 | -2155.37 | 2158.36 | 13.2 | -34.86 | -1.76 | 37.35 | 2155.37 | 34.86 | 43.0039
0.4918 | 58.63 | -88.49 | -2162.33 | 2165.3 | 13.22 | -34.99 | -1.78 | 37.48 | 2162.33 | 34.99 | 43.1706
0.4936 | 59.66 | -87.15 | -2169.84 | 2172.78 | 13.24 | -35.1 | -1.8 | 37.59 | 2169.84 | 35.1 | 43.3287
0.4955 | 60.54 | -85.76 | -2178.55 | 2181.45 | 13.26 | -35.21 | -1.83 | 37.7 | 2178.55 | 35.21 | 43.4954
0.4974 | 61.41 | -84.37 | -2187.68 | 2190.55 | 13.28 | -35.31 | -1.87 | 37.81 | 2187.68 | 35.31 | 43.6622
0.4993 | 62.24 | -83 | -2196.97 | 2199.79 | 13.3 | -35.41 | -1.9 | 37.91 | 2196.97 | 35.41 | 43.829
0.5012 | 63.08 | -81.68 | -2206.2 | 2208.99 | 13.31 | -35.51 | -1.94 | 38.01 | 2206.2 | 35.51 | 43.9958
0.5031 | 63.85 | -80.32 | -2215.35 | 2218.1 | 13.33 | -35.6 | -1.98 | 38.1 | 2215.35 | 35.6 | 44.1626
0.505 | 64.56 | -78.93 | -2224.59 | 2227.31 | 13.34 | -35.69 | -2.02 | 38.19 | 2224.59 | 35.69 | 44.3294
0.5069 | 65.26 | -77.52 | -2233.68 | 2236.37 | 13.36 | -35.78 | -2.07 | 38.28 | 2233.68 | 35.78 | 44.4961
0.5088 | 65.95 | -76.06 | -2242.15 | 2244.79 | 13.37 | -35.86 | -2.11 | 38.37 | 2242.15 | 35.86 | 44.6629
0.5107 | 66.65 | -74.66 | -2250.07 | 2252.68 | 13.38 | -35.94 | -2.14 | 38.45 | 2250.07 | 35.94 | 44.8297
0.5126 | 67.36 | -73.32 | -2257.55 | 2260.12 | 13.39 | -36.02 | -2.18 | 38.52 | 2257.55 | 36.02 | 44.9965
0.5145 | 68.06 | -72.04 | -2264.61 | 2267.16 | 13.4 | -36.09 | -2.21 | 38.6 | 2264.61 | 36.09 | 45.1633
0.5164 | 68.74 | -70.82 | -2271.25 | 2273.77 | 13.41 | -36.16 | -2.24 | 38.67 | 2271.25 | 36.16 | 45.3301
0.5183 | 69.36 | -69.63 | -2277.22 | 2279.72 | 13.41 | -36.23 | -2.26 | 38.73 | 2277.22 | 36.23 | 45.4968
0.5202 | 69.94 | -68.39 | -2282.31 | 2284.77 | 13.4 | -36.3 | -2.29 | 38.79 | 2282.31 | 36.3 | 45.6636
0.5221 | 70.45 | -67.1 | -2287.22 | 2289.66 | 13.4 | -36.36 | -2.31 | 38.85 | 2287.22 | 36.36 | 45.8304
0.524 | 70.85 | -65.75 | -2292.19 | 2294.6 | 13.39 | -36.42 | -2.34 | 38.91 | 2292.19 | 36.42 | 45.9972
0.5259 | 71.19 | -64.39 | -2297 | 2299.36 | 13.39 | -36.47 | -2.37 | 38.96 | 2297 | 36.47 | 46.164
0.5278 | 71.68 | -62.98 | -2300.81 | 2303.14 | 13.4 | -36.53 | -2.4 | 39.01 | 2300.81 | 36.53 | 46.3308
0.5297 | 72.6 | -61.54 | -2304.06 | 2306.38 | 13.43 | -36.59 | -2.41 | 39.08 | 2304.06 | 36.59 | 46.4975
0.5316 | 73.17 | -59.73 | -2307.26 | 2309.53 | 13.48 | -36.63 | -2.44 | 39.14 | 2307.26 | 36.63 | 46.6643
0.5335 | 73.35 | -57.95 | -2310.92 | 2313.13 | 13.51 | -36.66 | -2.47 | 39.18 | 2310.92 | 36.66 | 46.8311
0.5354 | 73.35 | -56.27 | -2314.47 | 2316.64 | 13.52 | -36.69 | -2.51 | 39.21 | 2314.47 | 36.69 | 46.9979
0.5373 | 73.12 | -54.58 | -2317.31 | 2319.42 | 13.53 | -36.71 | -2.55 | 39.24 | 2317.31 | 36.71 | 47.1647
0.5392 | 72.78 | -52.73 | -2319.67 | 2321.73 | 13.53 | -36.73 | -2.59 | 39.26 | 2319.67 | 36.73 | 47.3315
0.5411 | 72.4 | -50.9 | -2321.6 | 2323.61 | 13.54 | -36.75 | -2.64 | 39.28 | 2321.6 | 36.75 | 47.4982
0.543 | 71.94 | -49 | -2323.01 | 2324.96 | 13.54 | -36.76 | -2.69 | 39.3 | 2323.01 | 36.76 | 47.665
0.5449 | 71.29 | -46.95 | -2323.88 | 2325.76 | 13.55 | -36.76 | -2.74 | 39.31 | 2323.88 | 36.76 | 47.8318
0.5468 | 70.51 | -44.77 | -2324.22 | 2326.02 | 13.55 | -36.76 | -2.8 | 39.31 | 2324.22 | 36.76 | 47.9986
0.5487 | 69.75 | -42.81 | -2323.93 | 2325.68 | 13.54 | -36.76 | -2.85 | 39.3 | 2323.93 | 36.76 | 48.1654
0.5506 | 68.99 | -40.82 | -2323.11 | 2324.81 | 13.54 | -36.76 | -2.89 | 39.31 | 2323.11 | 36.76 | 48.3322
0.5525 | 68.07 | -38.51 | -2321.66 | 2323.29 | 13.53 | -36.76 | -2.93 | 39.3 | 2321.66 | 36.76 | 48.4989
0.5544 | 66.95 | -36.04 | -2319.7 | 2321.29 | 13.52 | -36.74 | -2.99 | 39.29 | 2319.7 | 36.74 | 48.6657
0.5563 | 65.82 | -33.65 | -2317.24 | 2318.79 | 13.5 | -36.71 | -3.04 | 39.26 | 2317.24 | 36.71 | 48.8325
0.5582 | 64.69 | -31.32 | -2314.29 | 2315.79 | 13.47 | -36.69 | -3.09 | 39.23 | 2314.29 | 36.69 | 48.9993
0.5601 | 63.61 | -28.93 | -2310.97 | 2312.43 | 13.45 | -36.66 | -3.14 | 39.2 | 2310.97 | 36.66 | 49.1661
0.562 | 62.26 | -25.95 | -2307.3 | 2308.72 | 13.43 | -36.6 | -3.21 | 39.14 | 2307.3 | 36.6 | 49.3329
0.5639 | 60.79 | -22.71 | -2303.23 | 2304.6 | 13.42 | -36.54 | -3.27 | 39.09 | 2303.23 | 36.54 | 49.4996
0.5658 | 59.3 | -19.5 | -2298.43 | 2299.76 | 13.4 | -36.49 | -3.33 | 39.03 | 2298.43 | 36.49 | 49.6664
0.5677 | 57.7 | -16.25 | -2292.97 | 2294.23 | 13.39 | -36.44 | -3.38 | 38.99 | 2292.97 | 36.44 | 49.8332
0.5696 | 56.03 | -12.92 | -2287.23 | 2288.41 | 13.37 | -36.39 | -3.43 | 38.95 | 2287.23 | 36.39 | 50
0.5715 | 54.15 | -9.29 | -2281.12 | 2282.23 | 13.35 | -36.33 | -3.48 | 38.89 | 2281.12 | 36.33 | 50.1668
0.5734 | 52.1 | -5.46 | -2274.63 | 2275.67 | 13.33 | -36.26 | -3.54 | 38.81 | 2274.63 | 36.26 | 50.3336
0.5753 | 49.84 | -1.49 | -2267.88 | 2268.85 | 13.31 | -36.18 | -3.59 | 38.74 | 2267.88 | 36.18 | 50.5004
0.5772 | 47.43 | 2.79 | -2260.45 | 2261.36 | 13.29 | -36.08 | -3.65 | 38.65 | 2260.45 | 36.08 | 50.6671
0.5791 | 45.01 | 7.22 | -2252.49 | 2253.36 | 13.27 | -35.97 | -3.72 | 38.54 | 2252.49 | 35.97 | 50.8339
0.581 | 42.58 | 11.73 | -2244.27 | 2245.1 | 13.24 | -35.86 | -3.78 | 38.43 | 2244.27 | 35.86 | 51.0007
0.5829 | 40.16 | 16.3 | -2235.83 | 2236.63 | 13.21 | -35.73 | -3.85 | 38.31 | 2235.83 | 35.73 | 51.1675
0.5848 | 37.81 | 20.97 | -2227.05 | 2227.83 | 13.16 | -35.59 | -3.91 | 38.17 | 2227.05 | 35.59 | 51.3343
0.5867 | 35.58 | 25.66 | -2218.01 | 2218.8 | 13.11 | -35.46 | -3.98 | 38.03 | 2218.01 | 35.46 | 51.5011
0.5886 | 33.47 | 30.37 | -2208.83 | 2209.61 | 13.04 | -35.32 | -4.04 | 37.89 | 2208.83 | 35.32 | 51.6678
0.5905 | 31.39 | 35.27 | -2199.34 | 2200.14 | 12.97 | -35.17 | -4.11 | 37.73 | 2199.34 | 35.17 | 51.8346
0.5924 | 29.15 | 40.34 | -2189.36 | 2190.19 | 12.88 | -35 | -4.17 | 37.55 | 2189.36 | 35 | 52.0014
0.5943 | 26.68 | 45.79 | -2177.73 | 2178.61 | 12.79 | -34.8 | -4.24 | 37.34 | 2177.73 | 34.8 | 52.1682
0.5962 | 24.29 | 51.33 | -2166.36 | 2167.31 | 12.69 | -34.59 | -4.3 | 37.12 | 2166.36 | 34.59 | 52.335
0.5981 | 22.18 | 56.85 | -2156.32 | 2157.38 | 12.58 | -34.36 | -4.38 | 36.87 | 2156.32 | 34.36 | 52.5018
0.6 | 20.71 | 61.85 | -2149.11 | 2150.29 | 12.46 | -34.11 | -4.48 | 36.61 | 2149.11 | 34.11 | 52.6685
0.6019 | 19.58 | 66.64 | -2143.55 | 2144.86 | 12.31 | -33.85 | -4.58 | 36.33 | 2143.55 | 33.85 | 52.8353
0.6038 | 18.93 | 71.23 | -2140.67 | 2142.12 | 12.13 | -33.55 | -4.71 | 36.01 | 2140.67 | 33.55 | 53.0021
0.6057 | 19.3 | 75.66 | -2143.53 | 2145.14 | 11.92 | -33.21 | -4.88 | 35.65 | 2143.53 | 33.21 | 53.1689
0.6076 | 20.16 | 79.75 | -2148.63 | 2150.4 | 11.69 | -32.84 | -5.07 | 35.26 | 2148.63 | 32.84 | 53.3357
0.6095 | 21.15 | 83.68 | -2153.77 | 2155.7 | 11.45 | -32.44 | -5.26 | 34.84 | 2153.77 | 32.44 | 53.5025
0.6114 | 21.84 | 87.61 | -2155.83 | 2157.91 | 11.22 | -32 | -5.44 | 34.39 | 2155.83 | 32 | 53.6692
0.6133 | 22.63 | 91.37 | -2153.79 | 2156.03 | 10.98 | -31.53 | -5.6 | 33.9 | 2153.79 | 31.53 | 53.836
0.6152 | 23.44 | 95.06 | -2149.6 | 2152.01 | 10.74 | -31 | -5.75 | 33.36 | 2149.6 | 31 | 54.0028
0.6171 | 24.35 | 98.22 | -2144.07 | 2146.65 | 10.49 | -30.42 | -5.9 | 32.78 | 2144.07 | 30.42 | 54.1696
0.619 | 25.43 | 101.22 | -2137.08 | 2139.83 | 10.23 | -29.82 | -6.05 | 32.17 | 2137.08 | 29.82 | 54.3364
0.6209 | 26.6 | 103.67 | -2128.75 | 2131.64 | 9.95 | -29.2 | -6.18 | 31.53 | 2128.75 | 29.2 | 54.5032
0.6228 | 27.72 | 105.58 | -2118.99 | 2122 | 9.67 | -28.53 | -6.31 | 30.87 | 2118.99 | 28.53 | 54.6699
0.6247 | 28.76 | 107.04 | -2107.6 | 2110.72 | 9.38 | -27.85 | -6.42 | 30.18 | 2107.6 | 27.85 | 54.8367
0.6266 | 29.81 | 108.2 | -2094.76 | 2097.96 | 9.09 | -27.16 | -6.52 | 29.48 | 2094.76 | 27.16 | 55.0035
0.6285 | 30.76 | 109.42 | -2080.39 | 2083.69 | 8.8 | -26.46 | -6.61 | 28.77 | 2080.39 | 26.46 | 55.1703
0.6304 | 31.57 | 110.75 | -2063.96 | 2067.36 | 8.51 | -25.75 | -6.7 | 28.05 | 2063.96 | 25.75 | 55.3371
0.6323 | 32.4 | 111.87 | -2045.66 | 2049.17 | 8.21 | -25.03 | -6.77 | 27.32 | 2045.66 | 25.03 | 55.5039
0.6341 | 33.41 | 112.49 | -2026.89 | 2030.5 | 7.91 | -24.31 | -6.84 | 26.59 | 2026.89 | 24.31 | 55.6619
0.636 | 34.48 | 112.11 | -2007.7 | 2011.35 | 7.59 | -23.55 | -6.89 | 25.82 | 2007.7 | 23.55 | 55.8287
0.6379 | 35.44 | 111.63 | -1987.73 | 1991.42 | 7.27 | -22.77 | -6.94 | 25.03 | 1987.73 | 22.77 | 55.9954
0.6398 | 36.42 | 111.07 | -1967.31 | 1971.03 | 6.94 | -21.99 | -6.99 | 24.24 | 1967.31 | 21.99 | 56.1622
0.6417 | 37.43 | 110.47 | -1946.07 | 1949.82 | 6.61 | -21.23 | -7.03 | 23.46 | 1946.07 | 21.23 | 56.329
0.6436 | 38.58 | 109.08 | -1924.5 | 1928.24 | 6.25 | -20.44 | -7.06 | 22.65 | 1924.5 | 20.44 | 56.4958
0.6455 | 39.89 | 106.89 | -1903.47 | 1907.15 | 5.89 | -19.62 | -7.09 | 21.82 | 1903.47 | 19.62 | 56.6626
0.6474 | 41.01 | 104.27 | -1880.97 | 1884.57 | 5.53 | -18.77 | -7.1 | 20.96 | 1880.97 | 18.77 | 56.8294
0.6493 | 41.97 | 101.4 | -1856.76 | 1860.27 | 5.17 | -17.9 | -7.11 | 20.09 | 1856.76 | 17.9 | 56.9961
0.6512 | 42.64 | 98.19 | -1829.25 | 1832.65 | 4.83 | -17.01 | -7.08 | 19.19 | 1829.25 | 17.01 | 57.1629
0.6531 | 43.12 | 94.88 | -1800.07 | 1803.38 | 4.5 | -16.12 | -7.05 | 18.3 | 1800.07 | 16.12 | 57.3297
0.655 | 43.7 | 90.89 | -1770.3 | 1773.48 | 4.17 | -15.23 | -6.99 | 17.4 | 1770.3 | 15.23 | 57.4965
0.6569 | 44.22 | 86.53 | -1739.52 | 1742.55 | 3.84 | -14.32 | -6.92 | 16.5 | 1739.52 | 14.32 | 57.6633
0.6588 | 44.62 | 81.37 | -1706.98 | 1709.81 | 3.52 | -13.4 | -6.84 | 15.58 | 1706.98 | 13.4 | 57.8301
0.6607 | 44.63 | 76.05 | -1672.42 | 1675.04 | 3.21 | -12.46 | -6.73 | 14.65 | 1672.42 | 12.46 | 57.9968
0.6626 | 44.32 | 70.53 | -1636.58 | 1638.97 | 2.92 | -11.53 | -6.61 | 13.72 | 1636.58 | 11.53 | 58.1636
0.6645 | 43.72 | 64.44 | -1598.44 | 1600.58 | 2.66 | -10.59 | -6.46 | 12.81 | 1598.44 | 10.59 | 58.3304
0.6664 | 42.56 | 58.33 | -1557.3 | 1559.19 | 2.45 | -9.66 | -6.28 | 11.91 | 1557.3 | 9.66 | 58.4972
0.6683 | 40.93 | 52.19 | -1513.29 | 1514.94 | 2.26 | -8.72 | -6.08 | 11.02 | 1513.29 | 8.72 | 58.664
0.6702 | 38.75 | 46.15 | -1466.1 | 1467.52 | 2.11 | -7.79 | -5.85 | 10.14 | 1466.1 | 7.79 | 58.8308
0.6721 | 35.9 | 40.48 | -1416.22 | 1417.43 | 1.99 | -6.87 | -5.61 | 9.31 | 1416.22 | 6.87 | 58.9975
0.674 | 32.43 | 35.07 | -1363.34 | 1364.36 | 1.91 | -6 | -5.34 | 8.54 | 1363.34 | 6 | 59.1643
0.6759 | 28.7 | 29.79 | -1309 | 1309.83 | 1.85 | -5.14 | -5.06 | 7.8 | 1309 | 5.14 | 59.3311
0.6778 | 24.7 | 24.64 | -1253.85 | 1254.52 | 1.81 | -4.3 | -4.77 | 7.07 | 1253.85 | 4.3 | 59.4979
0.6797 | 20.4 | 19.7 | -1198.39 | 1198.9 | 1.77 | -3.48 | -4.48 | 6.38 | 1198.39 | 3.48 | 59.6647
0.6816 | 15.77 | 15.33 | -1144.65 | 1145.03 | 1.77 | -2.7 | -4.21 | 5.78 | 1144.65 | 2.7 | 59.8315
0.6835 | 10.85 | 11.51 | -1091.53 | 1091.82 | 1.78 | -1.97 | -3.94 | 5.29 | 1091.53 | 1.97 | 59.9982
0.6854 | 5.87 | 7.99 | -1039.27 | 1039.51 | 1.81 | -1.3 | -3.68 | 4.87 | 1039.27 | 1.3 | 60.165
0.6873 | 0.83 | 4.76 | -987.61 | 987.82 | 1.85 | -0.65 | -3.43 | 4.51 | 987.61 | 0.65 | 60.3318
0.6892 | -4.12 | 1.86 | -937.35 | 937.55 | 1.89 | -0.04 | -3.18 | 4.22 | 937.35 | 0.04 | 60.4986
0.6911 | -9 | -0.54 | -889.97 | 890.2 | 1.94 | 0.51 | -2.96 | 4.03 | 889.97 | -0.51 | 60.6654
0.693 | -13.53 | -2.6 | -847.09 | 847.38 | 2 | 1.02 | -2.76 | 3.92 | 847.09 | -1.02 | 60.8322
0.6949 | -17.33 | -4.5 | -809.94 | 810.3 | 2.05 | 1.46 | -2.59 | 3.88 | 809.94 | -1.46 | 60.9989
0.6968 | -20.42 | -6.44 | -777.62 | 778.08 | 2.08 | 1.84 | -2.45 | 3.91 | 777.62 | -1.84 | 61.1657
0.6987 | -22.8 | -8.38 | -749.2 | 749.76 | 2.1 | 2.17 | -2.32 | 3.98 | 749.2 | -2.17 | 61.3325
0.7006 | -24.67 | -10.18 | -723.4 | 724.04 | 2.12 | 2.46 | -2.21 | 4.07 | 723.4 | -2.46 | 61.4993
0.7025 | -25.94 | -11.77 | -700.86 | 701.57 | 2.13 | 2.72 | -2.11 | 4.16 | 700.86 | -2.72 | 61.6661
0.7044 | -26.81 | -13.01 | -680.94 | 681.72 | 2.16 | 2.94 | -2.02 | 4.27 | 680.94 | -2.94 | 61.8329
0.7063 | -27.5 | -13.94 | -662.43 | 663.27 | 2.19 | 3.14 | -1.93 | 4.37 | 662.43 | -3.14 | 61.9996
0.7082 | -28.1 | -14.67 | -644.49 | 645.39 | 2.22 | 3.33 | -1.84 | 4.48 | 644.49 | -3.33 | 62.1664
0.7101 | -28.6 | -15.18 | -627.36 | 628.31 | 2.26 | 3.5 | -1.76 | 4.57 | 627.36 | -3.5 | 62.3332
0.712 | -28.86 | -15.27 | -612.27 | 613.25 | 2.3 | 3.65 | -1.68 | 4.67 | 612.27 | -3.65 | 62.5
0.7139 | -29 | -15.08 | -599 | 600 | 2.34 | 3.79 | -1.61 | 4.77 | 599 | -3.79 | 62.6668
0.7158 | -29.17 | -14.64 | -586.93 | 587.95 | 2.38 | 3.91 | -1.55 | 4.86 | 586.93 | -3.91 | 62.8336
0.7177 | -29.4 | -14.01 | -575.46 | 576.49 | 2.43 | 4.04 | -1.48 | 4.96 | 575.46 | -4.04 | 63.0004
0.7196 | -29.65 | -13.21 | -565.16 | 566.2 | 2.47 | 4.14 | -1.43 | 5.05 | 565.16 | -4.14 | 63.1671
0.7215 | -29.96 | -12.19 | -555.56 | 556.62 | 2.51 | 4.24 | -1.38 | 5.13 | 555.56 | -4.24 | 63.3339
0.7234 | -30.29 | -11.08 | -546.51 | 547.58 | 2.54 | 4.32 | -1.33 | 5.2 | 546.51 | -4.32 | 63.5007
0.7253 | -30.56 | -9.91 | -538.25 | 539.34 | 2.57 | 4.4 | -1.28 | 5.27 | 538.25 | -4.4 | 63.6675
0.7272 | -30.86 | -8.73 | -530.51 | 531.62 | 2.59 | 4.47 | -1.23 | 5.32 | 530.51 | -4.47 | 63.8343
0.7291 | -31.12 | -7.69 | -522.97 | 524.09 | 2.6 | 4.54 | -1.18 | 5.38 | 522.97 | -4.54 | 64.0011
0.731 | -31.21 | -6.77 | -515.14 | 516.28 | 2.61 | 4.59 | -1.13 | 5.42 | 515.14 | -4.59 | 64.1678
0.7329 | -31.2 | -5.92 | -507.32 | 508.46 | 2.61 | 4.63 | -1.08 | 5.45 | 507.32 | -4.63 | 64.3346
0.7348 | -31.28 | -5.13 | -499.88 | 501.04 | 2.61 | 4.68 | -1.03 | 5.47 | 499.88 | -4.68 | 64.5014
0.7367 | -31.33 | -4.35 | -492.77 | 493.94 | 2.6 | 4.72 | -0.98 | 5.49 | 492.77 | -4.72 | 64.6682
0.7386 | -31.44 | -3.6 | -486.08 | 487.27 | 2.59 | 4.75 | -0.94 | 5.51 | 486.08 | -4.75 | 64.835
0.7405 | -31.6 | -2.88 | -479.89 | 481.09 | 2.57 | 4.79 | -0.9 | 5.53 | 479.89 | -4.79 | 65.0018
0.7424 | -31.64 | -2.21 | -474.02 | 475.23 | 2.55 | 4.82 | -0.86 | 5.53 | 474.02 | -4.82 | 65.1685
0.7443 | -31.59 | -1.58 | -468.29 | 469.5 | 2.53 | 4.84 | -0.82 | 5.53 | 468.29 | -4.84 | 65.3353
0.7462 | -31.48 | -0.88 | -462.41 | 463.61 | 2.52 | 4.85 | -0.78 | 5.54 | 462.41 | -4.85 | 65.5021
0.7481 | -31.24 | -0.2 | -456.3 | 457.5 | 2.5 | 4.86 | -0.74 | 5.52 | 456.3 | -4.86 | 65.6689
0.75 | -30.93 | 0.49 | -450.45 | 451.64 | 2.47 | 4.86 | -0.7 | 5.51 | 450.45 | -4.86 | 65.8357
0.7519 | -30.71 | 1.18 | -445.47 | 446.64 | 2.45 | 4.85 | -0.66 | 5.48 | 445.47 | -4.85 | 66.0025
0.7538 | -30.49 | 1.77 | -441.29 | 442.46 | 2.41 | 4.84 | -0.64 | 5.46 | 441.29 | -4.84 | 66.1692
0.7557 | -30.26 | 2.37 | -437.31 | 438.46 | 2.38 | 4.83 | -0.61 | 5.43 | 437.31 | -4.83 | 66.336
0.7576 | -30.09 | 2.87 | -433.24 | 434.39 | 2.35 | 4.81 | -0.59 | 5.4 | 433.24 | -4.81 | 66.5028
0.7595 | -30 | 3.22 | -428.8 | 429.95 | 2.32 | 4.81 | -0.56 | 5.37 | 428.8 | -4.81 | 66.6696
0.7614 | -29.92 | 3.44 | -423.42 | 424.58 | 2.28 | 4.8 | -0.53 | 5.35 | 423.42 | -4.8 | 66.8364
0.7633 | -29.86 | 3.61 | -416.95 | 418.13 | 2.24 | 4.79 | -0.49 | 5.32 | 416.95 | -4.79 | 67.0032
0.7652 | -29.86 | 3.8 | -410.01 | 411.21 | 2.21 | 4.78 | -0.44 | 5.29 | 410.01 | -4.78 | 67.1699
0.7671 | -29.78 | 4.06 | -402.84 | 404.05 | 2.17 | 4.76 | -0.4 | 5.26 | 402.84 | -4.76 | 67.3367
0.769 | -29.63 | 4.29 | -395.67 | 396.89 | 2.14 | 4.74 | -0.35 | 5.22 | 395.67 | -4.74 | 67.5035
0.7709 | -29.4 | 4.51 | -388.93 | 390.16 | 2.11 | 4.71 | -0.31 | 5.17 | 388.93 | -4.71 | 67.6703
0.7728 | -29.14 | 4.82 | -382.95 | 384.18 | 2.08 | 4.67 | -0.28 | 5.12 | 382.95 | -4.67 | 67.8371
0.7746 | -28.86 | 5.19 | -377.88 | 379.11 | 2.05 | 4.63 | -0.25 | 5.07 | 377.88 | -4.63 | 67.9951
0.7765 | -28.53 | 5.65 | -373.87 | 375.09 | 2.01 | 4.58 | -0.23 | 5.02 | 373.87 | -4.58 | 68.1619
0.7784 | -28.27 | 6.11 | -370.61 | 371.83 | 1.98 | 4.54 | -0.22 | 4.97 | 370.61 | -4.54 | 68.3287
0.7803 | -28.03 | 6.4 | -367.94 | 369.16 | 1.95 | 4.51 | -0.21 | 4.92 | 367.94 | -4.51 | 68.4954
0.7822 | -27.73 | 6.6 | -364.92 | 366.13 | 1.91 | 4.47 | -0.2 | 4.87 | 364.92 | -4.47 | 68.6622
0.7841 | -27.28 | 6.7 | -360.43 | 361.61 | 1.87 | 4.41 | -0.17 | 4.81 | 360.43 | -4.41 | 68.829
0.786 | -26.82 | 6.65 | -355.21 | 356.38 | 1.83 | 4.36 | -0.14 | 4.74 | 355.21 | -4.36 | 68.9958
0.7879 | -26.39 | 6.62 | -349.59 | 350.74 | 1.79 | 4.31 | -0.11 | 4.67 | 349.59 | -4.31 | 69.1626
0.7898 | -26.08 | 6.66 | -344 | 345.15 | 1.75 | 4.25 | -0.08 | 4.61 | 344 | -4.25 | 69.3294
0.7917 | -25.8 | 6.7 | -338.85 | 339.99 | 1.72 | 4.19 | -0.05 | 4.54 | 338.85 | -4.19 | 69.4961
0.7936 | -25.4 | 6.76 | -334.26 | 335.4 | 1.68 | 4.13 | -0.03 | 4.46 | 334.26 | -4.13 | 69.6629
0.7955 | -24.93 | 6.87 | -330.44 | 331.55 | 1.64 | 4.07 | -0.01 | 4.39 | 330.44 | -4.07 | 69.8297
0.7974 | -24.51 | 6.99 | -327.38 | 328.48 | 1.61 | 4.01 | 0 | 4.32 | 327.38 | -4.01 | 69.9965
0.7993 | -24.09 | 7.18 | -324.83 | 325.91 | 1.57 | 3.96 | 0.01 | 4.26 | 324.83 | -3.96 | 70.1633
0.8012 | -23.66 | 7.38 | -322.47 | 323.53 | 1.54 | 3.91 | 0.02 | 4.21 | 322.47 | -3.91 | 70.3301
0.8031 | -23.25 | 7.58 | -320.4 | 321.45 | 1.51 | 3.86 | 0.03 | 4.15 | 320.4 | -3.86 | 70.4968
0.805 | -22.89 | 7.82 | -318.84 | 319.86 | 1.48 | 3.81 | 0.04 | 4.09 | 318.84 | -3.81 | 70.6636
0.8069 | -22.57 | 8.1 | -317.81 | 318.82 | 1.45 | 3.76 | 0.03 | 4.04 | 317.81 | -3.76 | 70.8304
0.8088 | -22.25 | 8.4 | -316.71 | 317.71 | 1.42 | 3.71 | 0.03 | 3.98 | 316.71 | -3.71 | 70.9972
0.8107 | -21.89 | 8.67 | -314.55 | 315.54 | 1.39 | 3.66 | 0.04 | 3.92 | 314.55 | -3.66 | 71.164
0.8126 | -21.49 | 8.88 | -311.48 | 312.45 | 1.36 | 3.6 | 0.05 | 3.86 | 311.48 | -3.6 | 71.3308
0.8145 | -21.1 | 8.98 | -307.68 | 308.63 | 1.33 | 3.55 | 0.08 | 3.79 | 307.68 | -3.55 | 71.4975
0.8164 | -20.75 | 9.02 | -303.27 | 304.22 | 1.3 | 3.49 | 0.1 | 3.73 | 303.27 | -3.49 | 71.6643
0.8183 | -20.44 | 9.08 | -299.12 | 300.06 | 1.27 | 3.43 | 0.13 | 3.66 | 299.12 | -3.43 | 71.8311
0.8202 | -20.23 | 9.24 | -296.29 | 297.23 | 1.24 | 3.38 | 0.14 | 3.6 | 296.29 | -3.38 | 71.9979
0.8221 | -20.1 | 9.5 | -295.28 | 296.22 | 1.21 | 3.33 | 0.14 | 3.55 | 295.28 | -3.33 | 72.1647
0.824 | -20.04 | 9.83 | -295.1 | 296.05 | 1.19 | 3.3 | 0.14 | 3.51 | 295.1 | -3.3 | 72.3315
0.8259 | -20.11 | 10.29 | -295.74 | 296.7 | 1.17 | 3.26 | 0.13 | 3.47 | 295.74 | -3.26 | 72.4982
0.8278 | -20.25 | 10.86 | -297.03 | 298.02 | 1.15 | 3.24 | 0.11 | 3.44 | 297.03 | -3.24 | 72.665
0.8297 | -20.38 | 11.43 | -298.48 | 299.49 | 1.13 | 3.21 | 0.09 | 3.41 | 298.48 | -3.21 | 72.8318
0.8316 | -20.34 | 11.91 | -298.92 | 299.94 | 1.11 | 3.18 | 0.09 | 3.37 | 298.92 | -3.18 | 72.9986
0.8335 | -20.09 | 12.12 | -298.22 | 299.23 | 1.08 | 3.14 | 0.09 | 3.33 | 298.22 | -3.14 | 73.1654
0.8354 | -19.71 | 12.17 | -296.68 | 297.67 | 1.05 | 3.1 | 0.11 | 3.28 | 296.68 | -3.1 | 73.3322
0.8373 | -19.26 | 12.14 | -294.84 | 295.8 | 1.02 | 3.05 | 0.13 | 3.23 | 294.84 | -3.05 | 73.4989
0.8392 | -18.8 | 12.12 | -292.88 | 293.82 | 0.99 | 3.01 | 0.15 | 3.18 | 292.88 | -3.01 | 73.6657
0.8411 | -18.41 | 12.08 | -291.35 | 292.26 | 0.96 | 2.97 | 0.17 | 3.13 | 291.35 | -2.97 | 73.8325
0.843 | -18.17 | 12.16 | -290.84 | 291.74 | 0.93 | 2.93 | 0.18 | 3.09 | 290.84 | -2.93 | 73.9993
0.8449 | -18.13 | 12.41 | -291.43 | 292.33 | 0.91 | 2.91 | 0.18 | 3.06 | 291.43 | -2.91 | 74.1661
0.8468 | -18.19 | 12.74 | -292.52 | 293.44 | 0.89 | 2.89 | 0.17 | 3.03 | 292.52 | -2.89 | 74.3329
0.8487 | -18.36 | 13.19 | -294.26 | 295.2 | 0.87 | 2.88 | 0.16 | 3.02 | 294.26 | -2.88 | 74.4996
0.8506 | -18.51 | 13.65 | -296.36 | 297.32 | 0.86 | 2.88 | 0.15 | 3.01 | 296.36 | -2.88 | 74.6664
0.8525 | -18.76 | 14.19 | -298.69 | 299.68 | 0.84 | 2.87 | 0.13 | 3 | 298.69 | -2.87 | 74.8332
0.8544 | -19.05 | 14.76 | -301.03 | 302.06 | 0.83 | 2.88 | 0.12 | 3 | 301.03 | -2.88 | 75
0.8563 | -19.2 | 15.3 | -302.91 | 303.96 | 0.82 | 2.87 | 0.11 | 2.99 | 302.91 | -2.87 | 75.1668
0.8582 | -19.16 | 15.66 | -303.6 | 304.66 | 0.8 | 2.86 | 0.11 | 2.98 | 303.6 | -2.86 | 75.3336
0.8601 | -18.93 | 15.84 | -303.08 | 304.14 | 0.78 | 2.85 | 0.12 | 2.96 | 303.08 | -2.85 | 75.5004
0.862 | -18.65 | 15.93 | -302.09 | 303.15 | 0.75 | 2.83 | 0.14 | 2.94 | 302.09 | -2.83 | 75.6671
0.8639 | -18.34 | 15.98 | -301.04 | 302.08 | 0.73 | 2.82 | 0.16 | 2.92 | 301.04 | -2.82 | 75.8339
0.8658 | -18.07 | 16.03 | -300.46 | 301.49 | 0.71 | 2.8 | 0.18 | 2.9 | 300.46 | -2.8 | 76.0007
0.8677 | -17.78 | 16.21 | -300.52 | 301.54 | 0.69 | 2.78 | 0.19 | 2.88 | 300.52 | -2.78 | 76.1675
0.8696 | -17.61 | 16.52 | -301.66 | 302.69 | 0.67 | 2.77 | 0.19 | 2.86 | 301.66 | -2.77 | 76.3343
0.8715 | -17.61 | 16.99 | -303.92 | 304.96 | 0.66 | 2.77 | 0.18 | 2.86 | 303.92 | -2.77 | 76.5011
0.8734 | -17.75 | 17.55 | -306.73 | 307.81 | 0.64 | 2.78 | 0.17 | 2.86 | 306.73 | -2.78 | 76.6678
0.8753 | -18 | 18.18 | -309.97 | 311.09 | 0.63 | 2.79 | 0.15 | 2.87 | 309.97 | -2.79 | 76.8346
0.8772 | -18.25 | 18.8 | -313.03 | 314.2 | 0.62 | 2.81 | 0.14 | 2.89 | 313.03 | -2.81 | 77.0014
0.8791 | -18.41 | 19.31 | -315.26 | 316.46 | 0.61 | 2.82 | 0.13 | 2.9 | 315.26 | -2.82 | 77.1682
0.881 | -18.49 | 19.64 | -316.66 | 317.87 | 0.59 | 2.83 | 0.14 | 2.9 | 316.66 | -2.83 | 77.335
0.8829 | -18.47 | 19.96 | -317.42 | 318.66 | 0.58 | 2.84 | 0.14 | 2.91 | 317.42 | -2.84 | 77.5018
0.8848 | -18.52 | 20.24 | -317.86 | 319.12 | 0.56 | 2.85 | 0.15 | 2.92 | 317.86 | -2.85 | 77.6685
0.8867 | -18.63 | 20.48 | -317.82 | 319.1 | 0.55 | 2.86 | 0.16 | 2.92 | 317.82 | -2.86 | 77.8353
0.8886 | -18.76 | 20.64 | -316.96 | 318.26 | 0.53 | 2.88 | 0.18 | 2.94 | 316.96 | -2.88 | 78.0021
0.8905 | -18.76 | 20.75 | -315.66 | 316.98 | 0.51 | 2.89 | 0.21 | 2.94 | 315.66 | -2.89 | 78.1689
0.8924 | -18.76 | 20.83 | -314.27 | 315.59 | 0.49 | 2.9 | 0.23 | 2.95 | 314.27 | -2.9 | 78.3357
0.8943 | -18.71 | 20.84 | -312.48 | 313.81 | 0.47 | 2.9 | 0.26 | 2.96 | 312.48 | -2.9 | 78.5025
0.8962 | -18.66 | 20.89 | -310.88 | 312.21 | 0.46 | 2.91 | 0.28 | 2.96 | 310.88 | -2.91 | 78.6692
0.8981 | -18.71 | 21.11 | -310.31 | 311.66 | 0.44 | 2.92 | 0.3 | 2.97 | 310.31 | -2.92 | 78.836
0.9 | -18.89 | 21.48 | -311.05 | 312.43 | 0.44 | 2.93 | 0.3 | 2.99 | 311.05 | -2.93 | 79.0028
0.9019 | -19.16 | 21.99 | -312.88 | 314.3 | 0.43 | 2.95 | 0.3 | 3.01 | 312.88 | -2.95 | 79.1696
0.9038 | -19.46 | 22.56 | -315.3 | 316.77 | 0.43 | 2.98 | 0.29 | 3.03 | 315.3 | -2.98 | 79.3364
0.9057 | -19.79 | 23.19 | -317.83 | 319.35 | 0.42 | 3 | 0.28 | 3.05 | 317.83 | -3 | 79.5032
0.9076 | -20.14 | 23.76 | -320.16 | 321.73 | 0.42 | 3.03 | 0.27 | 3.08 | 320.16 | -3.03 | 79.6699
0.9095 | -20.42 | 24.1 | -321.7 | 323.3 | 0.41 | 3.05 | 0.27 | 3.1 | 321.7 | -3.05 | 79.8367
0.9114 | -20.6 | 24.25 | -322.31 | 323.94 | 0.39 | 3.08 | 0.27 | 3.12 | 322.31 | -3.08 | 80.0035
0.9133 | -20.8 | 24.31 | -322.03 | 323.68 | 0.38 | 3.1 | 0.29 | 3.14 | 322.03 | -3.1 | 80.1703
0.9151 | -21.21 | 24.23 | -320.82 | 322.5 | 0.36 | 3.12 | 0.3 | 3.17 | 320.82 | -3.12 | 80.3283
0.917 | -21.58 | 24.05 | -318.49 | 320.19 | 0.34 | 3.14 | 0.32 | 3.19 | 318.49 | -3.14 | 80.4951
0.9189 | -21.86 | 23.78 | -315.47 | 317.19 | 0.33 | 3.16 | 0.35 | 3.2 | 315.47 | -3.16 | 80.6619
0.9208 | -22.11 | 23.47 | -312.16 | 313.89 | 0.31 | 3.17 | 0.38 | 3.21 | 312.16 | -3.17 | 80.8287
0.9227 | -22.26 | 23.22 | -308.83 | 310.57 | 0.3 | 3.17 | 0.41 | 3.22 | 308.83 | -3.17 | 80.9954
0.9246 | -22.35 | 22.96 | -306.13 | 307.88 | 0.29 | 3.17 | 0.43 | 3.22 | 306.13 | -3.17 | 81.1622
0.9265 | -22.36 | 22.76 | -304.48 | 306.22 | 0.28 | 3.17 | 0.44 | 3.21 | 304.48 | -3.17 | 81.329
0.9284 | -22.36 | 22.77 | -304.84 | 306.58 | 0.27 | 3.17 | 0.44 | 3.21 | 304.84 | -3.17 | 81.4958
0.9303 | -22.35 | 22.92 | -306.4 | 308.14 | 0.26 | 3.17 | 0.43 | 3.22 | 306.4 | -3.17 | 81.6626
0.9322 | -22.28 | 22.97 | -308.24 | 309.97 | 0.25 | 3.18 | 0.42 | 3.22 | 308.24 | -3.18 | 81.8294
0.9341 | -22.16 | 22.91 | -309.63 | 311.34 | 0.24 | 3.18 | 0.42 | 3.22 | 309.63 | -3.18 | 81.9961
0.936 | -22.13 | 22.88 | -311.17 | 312.87 | 0.23 | 3.19 | 0.41 | 3.23 | 311.17 | -3.19 | 82.1629
0.9379 | -22.2 | 22.89 | -312.87 | 314.57 | 0.22 | 3.2 | 0.4 | 3.24 | 312.87 | -3.2 | 82.3297
0.9398 | -22.17 | 22.76 | -313.87 | 315.55 | 0.21 | 3.2 | 0.4 | 3.24 | 313.87 | -3.2 | 82.4965
0.9417 | -22.15 | 22.47 | -313.52 | 315.18 | 0.2 | 3.2 | 0.4 | 3.24 | 313.52 | -3.2 | 82.6633
0.9436 | -22.13 | 22.1 | -312.23 | 313.87 | 0.18 | 3.21 | 0.41 | 3.24 | 312.23 | -3.21 | 82.8301
0.9455 | -22.05 | 21.62 | -310.28 | 311.9 | 0.17 | 3.2 | 0.43 | 3.24 | 310.28 | -3.2 | 82.9968
0.9474 | -21.92 | 21.15 | -307.94 | 309.53 | 0.16 | 3.19 | 0.44 | 3.23 | 307.94 | -3.19 | 83.1636
0.9493 | -21.75 | 20.76 | -305.57 | 307.13 | 0.14 | 3.18 | 0.46 | 3.23 | 305.57 | -3.18 | 83.3304
0.9512 | -21.6 | 20.46 | -303.86 | 305.4 | 0.14 | 3.17 | 0.47 | 3.21 | 303.86 | -3.17 | 83.4972
0.9531 | -21.4 | 20.2 | -302.83 | 304.35 | 0.13 | 3.15 | 0.48 | 3.2 | 302.83 | -3.15 | 83.664
0.955 | -21.16 | 19.94 | -302.31 | 303.79 | 0.13 | 3.14 | 0.48 | 3.19 | 302.31 | -3.14 | 83.8308
0.9569 | -20.94 | 19.8 | -302.6 | 304.07 | 0.13 | 3.13 | 0.48 | 3.17 | 302.6 | -3.13 | 83.9975
0.9588 | -20.89 | 19.84 | -303.98 | 305.44 | 0.12 | 3.12 | 0.47 | 3.17 | 303.98 | -3.12 | 84.1643
0.9607 | -21 | 19.92 | -305.91 | 307.38 | 0.13 | 3.12 | 0.45 | 3.16 | 305.91 | -3.12 | 84.3311
0.9626 | -21.12 | 19.98 | -308.16 | 309.62 | 0.13 | 3.12 | 0.42 | 3.16 | 308.16 | -3.12 | 84.4979
0.9645 | -21.22 | 19.98 | -309.83 | 311.29 | 0.13 | 3.12 | 0.4 | 3.16 | 309.83 | -3.12 | 84.6647
0.9664 | -21.3 | 19.84 | -309.97 | 311.42 | 0.13 | 3.11 | 0.39 | 3.15 | 309.97 | -3.11 | 84.8315
0.9683 | -21.43 | 19.53 | -308.98 | 310.43 | 0.13 | 3.1 | 0.39 | 3.14 | 308.98 | -3.1 | 84.9982
0.9702 | -21.51 | 19.13 | -306.88 | 308.32 | 0.12 | 3.09 | 0.4 | 3.12 | 306.88 | -3.09 | 85.165
0.9721 | -21.51 | 18.66 | -303.44 | 304.87 | 0.12 | 3.07 | 0.42 | 3.11 | 303.44 | -3.07 | 85.3318
0.974 | -21.42 | 18.2 | -299.38 | 300.79 | 0.12 | 3.04 | 0.44 | 3.08 | 299.38 | -3.04 | 85.4986
0.9759 | -21.29 | 17.73 | -295.28 | 296.68 | 0.11 | 3.01 | 0.46 | 3.05 | 295.28 | -3.01 | 85.6654
0.9778 | -21.22 | 17.31 | -292.33 | 293.71 | 0.11 | 2.98 | 0.48 | 3.03 | 292.33 | -2.98 | 85.8322
0.9797 | -21.14 | 16.99 | -290.26 | 291.63 | 0.11 | 2.95 | 0.48 | 3 | 290.26 | -2.95 | 85.9989
0.9816 | -21.01 | 16.62 | -288.66 | 290.01 | 0.1 | 2.92 | 0.48 | 2.97 | 288.66 | -2.92 | 86.1657
0.9835 | -21.02 | 16.41 | -287.94 | 289.28 | 0.1 | 2.89 | 0.48 | 2.95 | 287.94 | -2.89 | 86.3325
0.9854 | -21.11 | 16.34 | -287.95 | 289.29 | 0.11 | 2.87 | 0.46 | 2.92 | 287.95 | -2.87 | 86.4993
0.9873 | -21.14 | 16.33 | -288.45 | 289.79 | 0.11 | 2.85 | 0.44 | 2.9 | 288.45 | -2.85 | 86.6661
0.9892 | -21.18 | 16.48 | -289.65 | 290.99 | 0.12 | 2.82 | 0.42 | 2.87 | 289.65 | -2.82 | 86.8329
0.9911 | -21.21 | 16.71 | -291.04 | 292.39 | 0.14 | 2.8 | 0.4 | 2.84 | 291.04 | -2.8 | 86.9996
0.993 | -21.25 | 16.86 | -291.99 | 293.34 | 0.15 | 2.77 | 0.38 | 2.82 | 291.99 | -2.77 | 87.1664
0.9949 | -21.3 | 16.81 | -292.13 | 293.48 | 0.15 | 2.75 | 0.37 | 2.79 | 292.13 | -2.75 | 87.3332
0.9968 | -21.17 | 16.66 | -291.31 | 292.65 | 0.15 | 2.72 | 0.36 | 2.76 | 291.31 | -2.72 | 87.5
0.9987 | -20.96 | 16.43 | -290.1 | 291.42 | 0.15 | 2.69 | 0.36 | 2.73 | 290.1 | -2.69 | 87.6668
1.0006 | -20.81 | 16.26 | -288.47 | 289.78 | 0.16 | 2.64 | 0.36 | 2.68 | 288.47 | -2.64 | 87.8336
1.0025 | -20.75 | 16.16 | -286.11 | 287.42 | 0.16 | 2.59 | 0.37 | 2.63 | 286.11 | -2.59 | 88.0004
1.0044 | -20.67 | 16.09 | -282.87 | 284.19 | 0.17 | 2.53 | 0.38 | 2.58 | 282.87 | -2.53 | 88.1671
1.0063 | -20.55 | 16.06 | -279.51 | 280.85 | 0.17 | 2.47 | 0.39 | 2.52 | 279.51 | -2.47 | 88.3339
1.0082 | -20.34 | 16.11 | -277.12 | 278.47 | 0.18 | 2.41 | 0.39 | 2.46 | 277.12 | -2.41 | 88.5007
1.0101 | -19.92 | 16.28 | -276.15 | 277.48 | 0.19 | 2.35 | 0.39 | 2.4 | 276.15 | -2.35 | 88.6675
1.012 | -19.52 | 16.65 | -276 | 277.33 | 0.2 | 2.28 | 0.38 | 2.34 | 276 | -2.28 | 88.8343
1.0139 | -19.25 | 17.25 | -276.6 | 277.95 | 0.21 | 2.22 | 0.36 | 2.27 | 276.6 | -2.22 | 89.0011
1.0158 | -19.05 | 17.91 | -277.63 | 279.01 | 0.23 | 2.16 | 0.34 | 2.21 | 277.63 | -2.16 | 89.1678
1.0177 | -18.92 | 18.48 | -278.58 | 280 | 0.24 | 2.1 | 0.32 | 2.16 | 278.58 | -2.1 | 89.3346
1.0196 | -18.71 | 18.92 | -278.71 | 280.14 | 0.25 | 2.04 | 0.31 | 2.1 | 278.71 | -2.04 | 89.5014
1.0215 | -18.43 | 19.26 | -278.09 | 279.54 | 0.26 | 1.98 | 0.31 | 2.04 | 278.09 | -1.98 | 89.6682
1.0234 | -18.12 | 19.74 | -276.9 | 278.36 | 0.26 | 1.91 | 0.31 | 1.98 | 276.9 | -1.91 | 89.835
1.0253 | -17.86 | 20.39 | -276.01 | 277.51 | 0.27 | 1.84 | 0.31 | 1.9 | 276.01 | -1.84 | 90.0018
1.0272 | -17.74 | 21.34 | -276.23 | 277.81 | 0.28 | 1.76 | 0.3 | 1.83 | 276.23 | -1.76 | 90.1685
1.0291 | -17.71 | 22.49 | -276.59 | 278.27 | 0.3 | 1.67 | 0.29 | 1.75 | 276.59 | -1.67 | 90.3353
1.031 | -17.76 | 23.62 | -275.1 | 276.9 | 0.31 | 1.58 | 0.29 | 1.66 | 275.1 | -1.58 | 90.5021
1.0329 | -17.74 | 24.72 | -271.85 | 273.78 | 0.33 | 1.46 | 0.31 | 1.57 | 271.85 | -1.46 | 90.6689
1.0348 | -17.82 | 25.73 | -267.79 | 269.85 | 0.35 | 1.34 | 0.34 | 1.48 | 267.79 | -1.34 | 90.8357
1.0367 | -18.08 | 26.77 | -263.16 | 265.39 | 0.36 | 1.21 | 0.37 | 1.38 | 263.16 | -1.21 | 91.0025
1.0386 | -18.35 | 27.9 | -257.23 | 259.68 | 0.37 | 1.05 | 0.42 | 1.29 | 257.23 | -1.05 | 91.1692
1.0405 | -18.64 | 29.07 | -251.23 | 253.91 | 0.39 | 0.89 | 0.48 | 1.21 | 251.23 | -0.89 | 91.336
1.0424 | -18.99 | 30.49 | -245.55 | 248.52 | 0.4 | 0.71 | 0.53 | 1.13 | 245.55 | -0.71 | 91.5028
1.0443 | -19.23 | 32.14 | -241.41 | 244.67 | 0.43 | 0.54 | 0.56 | 1.08 | 241.41 | -0.54 | 91.6696
1.0462 | -19.44 | 33.89 | -238.38 | 241.94 | 0.45 | 0.35 | 0.59 | 1.05 | 238.38 | -0.35 | 91.8364
1.0481 | -19.53 | 36.02 | -236.4 | 240.32 | 0.48 | 0.14 | 0.62 | 1.08 | 236.4 | -0.14 | 92.0032
1.05 | -19.64 | 38.37 | -235.09 | 239.41 | 0.5 | -0.08 | 0.65 | 1.12 | 235.09 | 0.08 | 92.1699
1.0519 | -19.71 | 40.89 | -234.91 | 239.67 | 0.53 | -0.31 | 0.67 | 1.19 | 234.91 | 0.31 | 92.3367
1.0538 | -19.63 | 43.39 | -235.97 | 241.13 | 0.54 | -0.54 | 0.69 | 1.27 | 235.97 | 0.54 | 92.5035
1.0556 | -19.39 | 45.67 | -238.56 | 244.04 | 0.54 | -0.78 | 0.7 | 1.38 | 238.56 | 0.78 | 92.6615
1.0575 | -18.63 | 47.99 | -244.53 | 250.27 | 0.52 | -1.06 | 0.7 | 1.53 | 244.53 | 1.06 | 92.8283
1.0594 | -17.17 | 50.31 | -254.97 | 260.87 | 0.48 | -1.35 | 0.68 | 1.74 | 254.97 | 1.35 | 92.9951
1.0613 | -15.26 | 52.76 | -269.04 | 275.07 | 0.42 | -1.66 | 0.65 | 2 | 269.04 | 1.66 | 93.1619
1.0632 | -12.92 | 55 | -285.26 | 291.34 | 0.34 | -1.97 | 0.61 | 2.26 | 285.26 | 1.97 | 93.3287
1.0651 | -10.1 | 56.94 | -303.34 | 309.37 | 0.24 | -2.3 | 0.58 | 2.53 | 303.34 | 2.3 | 93.4954
1.067 | -6.89 | 58.73 | -322.61 | 328.54 | 0.13 | -2.62 | 0.53 | 2.81 | 322.61 | 2.62 | 93.6622
1.0689 | -3.52 | 60.59 | -342.48 | 348.32 | 0.03 | -2.95 | 0.49 | 3.09 | 342.48 | 2.95 | 93.829
1.0708 | -0.16 | 62.52 | -362.39 | 368.16 | -0.08 | -3.27 | 0.44 | 3.37 | 362.39 | 3.27 | 93.9958
1.0727 | 2.99 | 64.33 | -381.44 | 387.19 | -0.19 | -3.57 | 0.4 | 3.65 | 381.44 | 3.57 | 94.1626
1.0746 | 5.92 | 65.97 | -399.57 | 405.35 | -0.3 | -3.85 | 0.37 | 3.92 | 399.57 | 3.85 | 94.3294
1.0765 | 8.77 | 67.37 | -417.4 | 423.21 | -0.42 | -4.11 | 0.34 | 4.19 | 417.4 | 4.11 | 94.4961
1.0784 | 11.48 | 68.7 | -434.8 | 440.65 | -0.53 | -4.37 | 0.31 | 4.45 | 434.8 | 4.37 | 94.6629
1.0803 | 14.06 | 69.86 | -452.49 | 458.37 | -0.64 | -4.61 | 0.28 | 4.7 | 452.49 | 4.61 | 94.8297
1.0822 | 16.6 | 70.87 | -470.41 | 476.3 | -0.76 | -4.85 | 0.25 | 4.95 | 470.41 | 4.85 | 94.9965
1.0841 | 19.18 | 71.58 | -488.69 | 494.55 | -0.88 | -5.08 | 0.22 | 5.2 | 488.69 | 5.08 | 95.1633
1.086 | 21.73 | 71.74 | -507.34 | 513.11 | -1.01 | -5.31 | 0.19 | 5.44 | 507.34 | 5.31 | 95.3301
1.0879 | 24.27 | 71.42 | -526.24 | 531.89 | -1.14 | -5.52 | 0.15 | 5.67 | 526.24 | 5.52 | 95.4968
1.0898 | 26.86 | 70.72 | -545.34 | 550.82 | -1.27 | -5.72 | 0.12 | 5.9 | 545.34 | 5.72 | 95.6636
1.0917 | 29.49 | 69.8 | -564.76 | 570.06 | -1.41 | -5.91 | 0.08 | 6.12 | 564.76 | 5.91 | 95.8304
1.0936 | 32.11 | 68.58 | -583.59 | 588.69 | -1.54 | -6.1 | 0.05 | 6.33 | 583.59 | 6.1 | 95.9972
1.0955 | 34.6 | 66.89 | -601.2 | 606.11 | -1.66 | -6.27 | 0.02 | 6.53 | 601.2 | 6.27 | 96.164
1.0974 | 37 | 64.84 | -618.11 | 622.82 | -1.77 | -6.43 | -0.02 | 6.72 | 618.11 | 6.43 | 96.3308
1.0993 | 39.34 | 62.57 | -633.76 | 638.28 | -1.88 | -6.58 | -0.04 | 6.89 | 633.76 | 6.58 | 96.4975
1.1012 | 41.63 | 60.33 | -647.6 | 651.96 | -1.95 | -6.73 | -0.07 | 7.05 | 647.6 | 6.73 | 96.6643
1.1031 | 43.6 | 58.45 | -659.5 | 663.74 | -1.97 | -6.86 | -0.08 | 7.19 | 659.5 | 6.86 | 96.8311
1.105 | 45.22 | 56.85 | -669.97 | 674.12 | -1.97 | -6.98 | -0.11 | 7.3 | 669.97 | 6.98 | 96.9979
1.1069 | 46.64 | 55.39 | -679.13 | 683.2 | -1.95 | -7.08 | -0.13 | 7.39 | 679.13 | 7.08 | 97.1647
1.1088 | 48.06 | 53.98 | -687.41 | 691.41 | -1.93 | -7.18 | -0.15 | 7.48 | 687.41 | 7.18 | 97.3315
1.1107 | 49.22 | 53.01 | -694.95 | 698.9 | -1.88 | -7.27 | -0.17 | 7.55 | 694.95 | 7.27 | 97.4982
1.1126 | 50.25 | 52.18 | -701.74 | 705.65 | -1.81 | -7.34 | -0.2 | 7.61 | 701.74 | 7.34 | 97.665
1.1145 | 51.19 | 51.22 | -707.59 | 711.45 | -1.75 | -7.41 | -0.21 | 7.66 | 707.59 | 7.41 | 97.8318
1.1164 | 51.92 | 50.24 | -712.81 | 716.62 | -1.67 | -7.46 | -0.23 | 7.69 | 712.81 | 7.46 | 97.9986
1.1183 | 52.44 | 49.21 | -717.24 | 720.99 | -1.6 | -7.5 | -0.25 | 7.71 | 717.24 | 7.5 | 98.1654
1.1202 | 52.61 | 48.24 | -720.37 | 724.05 | -1.51 | -7.52 | -0.26 | 7.71 | 720.37 | 7.52 | 98.3322
1.1221 | 52.28 | 47.24 | -721.64 | 725.22 | -1.4 | -7.53 | -0.26 | 7.7 | 721.64 | 7.53 | 98.4989
1.124 | 51.79 | 46.11 | -721.55 | 725.02 | -1.3 | -7.54 | -0.25 | 7.69 | 721.55 | 7.54 | 98.6657
1.1259 | 51.23 | 44.83 | -720.19 | 723.55 | -1.2 | -7.54 | -0.24 | 7.67 | 720.19 | 7.54 | 98.8325
1.1278 | 50.33 | 43.39 | -717.33 | 720.55 | -1.09 | -7.52 | -0.21 | 7.63 | 717.33 | 7.52 | 98.9993
1.1297 | 49.26 | 41.72 | -714.31 | 717.38 | -0.98 | -7.49 | -0.18 | 7.59 | 714.31 | 7.49 | 99.1661
1.1316 | 48.14 | 39.78 | -711.96 | 714.87 | -0.87 | -7.47 | -0.14 | 7.56 | 711.96 | 7.47 | 99.3329
1.1335 | 46.95 | 37.75 | -710.63 | 713.37 | -0.77 | -7.45 | -0.1 | 7.53 | 710.63 | 7.45 | 99.4996
1.1354 | 45.83 | 35.78 | -710.38 | 712.96 | -0.66 | -7.43 | -0.07 | 7.5 | 710.38 | 7.43 | 99.6664
1.1373 | 44.91 | 33.86 | -711 | 713.42 | -0.56 | -7.41 | -0.04 | 7.48 | 711 | 7.41 | 99.8332
1.1392 | 44.19 | 32.26 | -711.59 | 713.88 | -0.48 | -7.41 | -0.01 | 7.46 | 711.59 | 7.41 | 100
| Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My
Minimum | -31.64 | -286.31 | -2324.22 | 239.41 | -1.97 | -36.76 | -7.11 | 1.05 | 234.91 | -4.86
Maximum | 108.52 | 112.49 | -234.91 | 2326.02 | 17.24 | 4.86 | 6.3 | 39.31 | 2324.22 | 36.76

## Diagr_Individ
Individual Loads
Contact forces and moments, measured telemetrically in knee implants of 8 subjects.
Averages of one individual subject.
Fx  lateral, Fy  anterior, -Fz  distal, Fres = Resultant Force
Mx around x, My around y, Mz around z, Mres = Resultant Moment
z-axis = axis of implant stem!
From Bergmann et al, PlosOne (2014) "Standardized Loads Acting in Knee Implants"
Subject K2L
Subject K2L

## Dat_BW
Data File: | k2l_2242_0_21_i1.akf
Curve Alignment Program: | 1.5.7
------------------------MEASURING INFOS------------------------
Diagram Title #1 | Forces and Moments at Knee Joint
Diagram Title #2 | Measured Relative to Implant
Comment #1 | Knee Joint; Gaitanalysis; Level Walking;
Comment #2 | K2L, 45,5 Months PO *
BodyWeight [N]: | 902.5
Patient(en): | k2l | [7 empty] | Subject | k2l
Implant Type: | Knee Joint
Varus Angle Leg [°] | 5 | [7 empty] | Normalized Loads in %BW (Forces) and %BWm (moments)
Axes Directions X, Y, Z: | Lateral, Anterior, Superior
Axes Origin: | Tibial Plateau at Stem Axis | [7 empty] | BW [kg] = | 91.998
Transformation X Y Z (°): | 11 | -0.5 | 0.5
Displacement X, Y, Z (mm): | 0 | 0 | 0
Activity Code: | KneeActivities | 1P_1_15_2
---------------------------DATA--------------------------------
Number of Data Sets = | 601
Number of analog channels = | 0
Time-Offset(sec) = | 0
Time | Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My | Time
[s] | [%BW] | [%BW] | [%BW] | [%BW] | [%BWm] | [%BWm] | [%BWm] | [%BWm] | [%BW] | [%BWm] | [%Cycle]
0 | 3.549 | 1.1745 | -80.0975 | 80.2449 | 0.035457 | -0.76011 | 0.036565 | 0.77119 | 80.0975 | 0.76011 | 0
0.0018 | 3.3806 | 0.85762 | -80.3701 | 80.5119 | 0.047645 | -0.76011 | 0.044321 | 0.77341 | 80.3701 | 0.76011 | 0.16735
0.0036 | 3.1734 | 0.49972 | -80.6992 | 80.8321 | 0.06205 | -0.76011 | 0.052078 | 0.77341 | 80.6992 | 0.76011 | 0.3347
0.0054 | 2.9418 | 0.13629 | -81.0814 | 81.2066 | 0.077562 | -0.75789 | 0.060942 | 0.77452 | 81.0814 | 0.75789 | 0.50205
0.0072 | 2.7102 | -0.28255 | -81.6144 | 81.7374 | 0.094183 | -0.75789 | 0.070914 | 0.77673 | 81.6144 | 0.75789 | 0.66939
0.009 | 2.5008 | -0.71579 | -82.2161 | 82.3357 | 0.1097 | -0.75789 | 0.079778 | 0.77895 | 82.2161 | 0.75789 | 0.83674
0.0108 | 2.277 | -1.1657 | -82.9219 | 83.0416 | 0.12632 | -0.76011 | 0.090859 | 0.7856 | 82.9219 | 0.76011 | 1.0041
0.0125 | 2.0343 | -1.6643 | -83.7418 | 83.8637 | 0.14515 | -0.76122 | 0.10305 | 0.79114 | 83.7418 | 0.76122 | 1.1621
0.0143 | 1.7529 | -2.1596 | -84.5961 | 84.7235 | 0.1651 | -0.76454 | 0.11524 | 0.8 | 84.5961 | 0.76454 | 1.3295
0.0161 | 1.4604 | -2.6227 | -85.4227 | 85.5557 | 0.18837 | -0.77008 | 0.12964 | 0.8133 | 85.4227 | 0.77008 | 1.4968
0.0179 | 1.1812 | -3.128 | -86.4321 | 86.5762 | 0.21163 | -0.77784 | 0.14404 | 0.82881 | 86.4321 | 0.77784 | 1.6642
0.0197 | 0.93407 | -3.6898 | -87.6687 | 87.826 | 0.23269 | -0.78449 | 0.15845 | 0.84321 | 87.6687 | 0.78449 | 1.8315
0.0215 | 0.70803 | -4.2792 | -89.1047 | 89.2809 | 0.25485 | -0.79335 | 0.17175 | 0.85983 | 89.1047 | 0.79335 | 1.9989
0.0233 | 0.56177 | -4.8554 | -90.8698 | 91.0648 | 0.27922 | -0.80665 | 0.18393 | 0.88089 | 90.8698 | 0.80665 | 2.1662
0.0251 | 0.4964 | -5.4061 | -92.9762 | 93.1956 | 0.30471 | -0.82438 | 0.19391 | 0.90637 | 92.9762 | 0.82438 | 2.3336
0.0269 | 0.51634 | -5.9756 | -95.5202 | 95.7673 | 0.32687 | -0.84432 | 0.20055 | 0.93407 | 95.5202 | 0.84432 | 2.5009
0.0287 | 0.69363 | -6.6028 | -98.9662 | 99.2443 | 0.34349 | -0.86759 | 0.20166 | 0.96066 | 98.9662 | 0.86759 | 2.6683
0.0305 | 0.95845 | -7.2332 | -102.9463 | 103.2576 | 0.359 | -0.89307 | 0.19945 | 0.98947 | 102.9463 | 0.89307 | 2.8356
0.0323 | 1.2133 | -7.8183 | -107.067 | 107.4083 | 0.37452 | -0.91856 | 0.19612 | 1.0161 | 107.067 | 0.91856 | 3.003
0.0341 | 1.5213 | -8.5163 | -111.3152 | 111.6975 | 0.3867 | -0.94183 | 0.19169 | 1.0416 | 111.3152 | 0.94183 | 3.1703
0.0359 | 1.8693 | -9.3086 | -115.6853 | 116.123 | 0.4 | -0.96399 | 0.18947 | 1.0659 | 115.6853 | 0.96399 | 3.3377
0.0376 | 2.2881 | -10.2039 | -119.9601 | 120.4632 | 0.4133 | -0.98837 | 0.18837 | 1.0914 | 119.9601 | 0.98837 | 3.4957
0.0394 | 2.7258 | -11.1291 | -124.0089 | 124.5861 | 0.42881 | -1.0116 | 0.18947 | 1.118 | 124.0089 | 1.0116 | 3.6631
0.0412 | 3.1335 | -12.0244 | -127.5856 | 128.236 | 0.44654 | -1.0327 | 0.1928 | 1.1457 | 127.5856 | 1.0327 | 3.8304
0.043 | 3.4814 | -12.8875 | -130.6216 | 131.344 | 0.46981 | -1.0526 | 0.19945 | 1.1745 | 130.6216 | 1.0526 | 3.9978
0.0448 | 3.7186 | -13.6753 | -132.9906 | 133.7861 | 0.49751 | -1.0693 | 0.20831 | 1.2033 | 132.9906 | 1.0693 | 4.1651
0.0466 | 3.8382 | -14.3668 | -134.7191 | 135.5845 | 0.53186 | -1.0825 | 0.22161 | 1.2332 | 134.7191 | 1.0825 | 4.3325
0.0484 | 3.8582 | -14.9751 | -136.0244 | 136.9496 | 0.56842 | -1.0947 | 0.2338 | 1.2643 | 136.0244 | 1.0947 | 4.4998
0.0502 | 3.8438 | -15.5457 | -136.9806 | 137.9634 | 0.6072 | -1.1069 | 0.24931 | 1.2964 | 136.9806 | 1.1069 | 4.6672
0.052 | 3.8039 | -16.0997 | -137.7751 | 138.8133 | 0.64488 | -1.1202 | 0.26482 | 1.3285 | 137.7751 | 1.1202 | 4.8345
0.0538 | 3.7186 | -16.6371 | -138.4321 | 139.5269 | 0.68476 | -1.1335 | 0.28033 | 1.364 | 138.4321 | 1.1335 | 5.0019
0.0556 | 3.59 | -17.1069 | -138.9629 | 140.1064 | 0.72465 | -1.1512 | 0.29695 | 1.4039 | 138.9629 | 1.1512 | 5.1692
0.0574 | 3.4271 | -17.5413 | -139.3075 | 140.4986 | 0.76565 | -1.1712 | 0.31468 | 1.446 | 139.3075 | 1.1712 | 5.3366
0.0592 | 3.2643 | -17.9878 | -139.4526 | 140.697 | 0.80776 | -1.1934 | 0.33463 | 1.4914 | 139.4526 | 1.1934 | 5.5039
0.061 | 3.1136 | -18.4443 | -139.5324 | 140.8332 | 0.84875 | -1.2155 | 0.35568 | 1.5368 | 139.5324 | 1.2155 | 5.6713
0.0627 | 2.9374 | -18.8875 | -139.5424 | 140.8986 | 0.89086 | -1.2377 | 0.37673 | 1.5834 | 139.5424 | 1.2377 | 5.8293
0.0645 | 2.7091 | -19.2875 | -139.3884 | 140.7956 | 0.93186 | -1.262 | 0.39778 | 1.6299 | 139.3884 | 1.262 | 5.9967
0.0663 | 2.4765 | -19.6942 | -139.2886 | 140.7479 | 0.97064 | -1.2886 | 0.41884 | 1.6787 | 139.2886 | 1.2886 | 6.164
0.0681 | 2.2571 | -20.1396 | -139.3884 | 140.9053 | 1.0094 | -1.3208 | 0.43878 | 1.7296 | 139.3884 | 1.3208 | 6.3313
0.0699 | 2.0488 | -20.6216 | -139.6898 | 141.2687 | 1.0482 | -1.354 | 0.45873 | 1.7839 | 139.6898 | 1.354 | 6.4987
0.0717 | 1.8504 | -21.1202 | -140.1274 | 141.774 | 1.087 | -1.3895 | 0.47867 | 1.8393 | 140.1274 | 1.3895 | 6.666
0.0735 | 1.6753 | -21.6199 | -140.6726 | 142.3834 | 1.1235 | -1.4305 | 0.4964 | 1.8958 | 140.6726 | 1.4305 | 6.8334
0.0753 | 1.5357 | -22.1385 | -141.292 | 143.0715 | 1.1568 | -1.4759 | 0.51524 | 1.9557 | 141.292 | 1.4759 | 7.0007
0.0771 | 1.4338 | -22.6792 | -142.0066 | 143.8582 | 1.19 | -1.5258 | 0.53296 | 2.0177 | 142.0066 | 1.5258 | 7.1681
0.0789 | 1.3629 | -23.2299 | -142.8587 | 144.7856 | 1.2244 | -1.5745 | 0.55069 | 2.0798 | 142.8587 | 1.5745 | 7.3354
0.0807 | 1.3186 | -23.7817 | -143.815 | 145.8161 | 1.2587 | -1.6255 | 0.56842 | 2.144 | 143.815 | 1.6255 | 7.5028
0.0825 | 1.2942 | -24.3235 | -144.9474 | 147.0205 | 1.292 | -1.6798 | 0.58393 | 2.2094 | 144.9474 | 1.6798 | 7.6701
0.0843 | 1.3064 | -24.8565 | -146.2526 | 148.3934 | 1.3241 | -1.7374 | 0.59945 | 2.277 | 146.2526 | 1.7374 | 7.8375
0.086 | 1.3596 | -25.3884 | -147.6687 | 149.8781 | 1.3562 | -1.795 | 0.61385 | 2.3435 | 147.6687 | 1.795 | 7.9955
0.0878 | 1.4427 | -25.9169 | -149.1524 | 151.4305 | 1.3873 | -1.8537 | 0.62715 | 2.4111 | 149.1524 | 1.8537 | 8.1629
0.0896 | 1.5679 | -26.4443 | -150.6814 | 153.0283 | 1.4183 | -1.9147 | 0.64044 | 2.4798 | 150.6814 | 1.9147 | 8.3302
0.0914 | 1.744 | -26.9706 | -152.297 | 154.7114 | 1.4471 | -1.9801 | 0.65374 | 2.5496 | 152.297 | 1.9801 | 8.4976
0.0932 | 1.9712 | -27.4925 | -154.1053 | 156.585 | 1.4748 | -2.0454 | 0.66482 | 2.6194 | 154.1053 | 2.0454 | 8.6649
0.095 | 2.2316 | -27.9945 | -156.0853 | 158.6249 | 1.5014 | -2.1108 | 0.67368 | 2.6881 | 156.0853 | 2.1108 | 8.8323
0.0968 | 2.5573 | -28.4787 | -158.2438 | 160.8388 | 1.5258 | -2.1795 | 0.68144 | 2.759 | 158.2438 | 2.1795 | 8.9996
0.0986 | 2.9285 | -28.9219 | -160.5053 | 163.1479 | 1.5501 | -2.2471 | 0.68809 | 2.8288 | 160.5053 | 2.2471 | 9.167
0.1004 | 3.3108 | -29.3319 | -162.882 | 165.5668 | 1.5756 | -2.3147 | 0.69363 | 2.8986 | 162.882 | 2.3147 | 9.3343
0.1022 | 3.7208 | -29.6931 | -165.3784 | 168.0964 | 1.6 | -2.3801 | 0.69695 | 2.9662 | 165.3784 | 2.3801 | 9.5017
0.104 | 4.1618 | -30.0111 | -168.0033 | 170.7479 | 1.6233 | -2.4454 | 0.69806 | 3.0316 | 168.0033 | 2.4454 | 9.669
0.1058 | 4.6249 | -30.2848 | -170.8122 | 173.5745 | 1.6454 | -2.5086 | 0.69695 | 3.0947 | 170.8122 | 2.5086 | 9.8364
0.1076 | 5.108 | -30.4942 | -173.7319 | 176.5019 | 1.6665 | -2.5695 | 0.69363 | 3.1557 | 173.7319 | 2.5695 | 10.0037
0.1094 | 5.5767 | -30.6881 | -176.5385 | 179.3163 | 1.6886 | -2.6305 | 0.69141 | 3.2166 | 176.5385 | 2.6305 | 10.1711
0.1111 | 6.0188 | -30.8875 | -179.1324 | 181.9202 | 1.7097 | -2.6881 | 0.6892 | 3.2753 | 179.1324 | 2.6881 | 10.3291
0.1129 | 6.4886 | -31.0582 | -181.6543 | 184.4499 | 1.7307 | -2.7457 | 0.68809 | 3.333 | 181.6543 | 2.7457 | 10.4965
0.1147 | 6.964 | -31.1978 | -184.0986 | 186.8975 | 1.7485 | -2.8 | 0.68476 | 3.3873 | 184.0986 | 2.8 | 10.6638
0.1165 | 7.4305 | -31.3252 | -186.4798 | 189.2831 | 1.7662 | -2.8532 | 0.68255 | 3.4393 | 186.4798 | 2.8532 | 10.8312
0.1183 | 7.8792 | -31.4338 | -188.7258 | 191.5324 | 1.7828 | -2.9019 | 0.67922 | 3.487 | 188.7258 | 2.9019 | 10.9985
0.1201 | 8.2848 | -31.5357 | -190.8942 | 193.7019 | 1.7994 | -2.9474 | 0.6759 | 3.5335 | 190.8942 | 2.9474 | 11.1659
0.1219 | 8.6704 | -31.6055 | -192.9884 | 195.795 | 1.815 | -2.9906 | 0.67258 | 3.5767 | 192.9884 | 2.9906 | 11.3332
0.1237 | 9.0593 | -31.6576 | -194.9972 | 197.8017 | 1.8283 | -3.0327 | 0.66814 | 3.6166 | 194.9972 | 3.0327 | 11.5006
0.1255 | 9.4382 | -31.6931 | -196.9008 | 199.7008 | 1.8393 | -3.0715 | 0.66371 | 3.6543 | 196.9008 | 3.0715 | 11.6679
0.1273 | 9.8127 | -31.7219 | -198.7978 | 201.5911 | 1.8515 | -3.1102 | 0.65928 | 3.6909 | 198.7978 | 3.1102 | 11.8353
0.1291 | 10.1729 | -31.7241 | -200.81 | 203.5911 | 1.8637 | -3.1479 | 0.65263 | 3.7274 | 200.81 | 3.1479 | 12.0026
0.1309 | 10.5252 | -31.7008 | -202.8975 | 205.6609 | 1.8781 | -3.1845 | 0.64709 | 3.764 | 202.8975 | 3.1845 | 12.17
0.1327 | 10.8488 | -31.631 | -204.8066 | 207.5479 | 1.8892 | -3.2155 | 0.64044 | 3.795 | 204.8066 | 3.2155 | 12.3373
0.1345 | 11.1158 | -31.5778 | -206.3645 | 209.0892 | 1.897 | -3.241 | 0.6349 | 3.8194 | 206.3645 | 3.241 | 12.5046
0.1362 | 11.3197 | -31.5557 | -207.7429 | 210.4565 | 1.9025 | -3.2587 | 0.62936 | 3.836 | 207.7429 | 3.2587 | 12.6627
0.138 | 11.5014 | -31.5435 | -208.9784 | 211.6831 | 1.9058 | -3.2742 | 0.62382 | 3.8504 | 208.9784 | 3.2742 | 12.83
0.1398 | 11.662 | -31.5368 | -210.062 | 212.7601 | 1.908 | -3.2875 | 0.61939 | 3.8615 | 210.062 | 3.2875 | 12.9974
0.1416 | 11.7983 | -31.5291 | -210.9817 | 213.6742 | 1.9091 | -3.2997 | 0.61607 | 3.8726 | 210.9817 | 3.2997 | 13.1647
0.1434 | 11.9003 | -31.508 | -211.7684 | 214.4532 | 1.9102 | -3.3097 | 0.61274 | 3.8803 | 211.7684 | 3.3097 | 13.3321
0.1452 | 11.9767 | -31.4681 | -212.4942 | 215.1679 | 1.9102 | -3.3186 | 0.60942 | 3.8881 | 212.4942 | 3.3186 | 13.4994
0.147 | 12.0155 | -31.4205 | -213.1623 | 215.8249 | 1.9102 | -3.3252 | 0.60499 | 3.8925 | 213.1623 | 3.3252 | 13.6668
0.1488 | 12.0244 | -31.385 | -213.7906 | 216.441 | 1.9102 | -3.3296 | 0.60166 | 3.8958 | 213.7906 | 3.3296 | 13.8341
0.1506 | 12.0089 | -31.3706 | -214.3834 | 217.0227 | 1.9091 | -3.3319 | 0.59723 | 3.897 | 214.3834 | 3.3319 | 14.0015
0.1524 | 11.9789 | -31.3285 | -214.8598 | 217.487 | 1.908 | -3.3341 | 0.5928 | 3.897 | 214.8598 | 3.3341 | 14.1688
0.1542 | 11.9169 | -31.262 | -215.2565 | 217.8659 | 1.9069 | -3.3352 | 0.58837 | 3.897 | 215.2565 | 3.3352 | 14.3362
0.156 | 11.8382 | -31.1601 | -215.5934 | 218.1806 | 1.9069 | -3.3352 | 0.58393 | 3.897 | 215.5934 | 3.3352 | 14.5035
0.1578 | 11.7695 | -31.0316 | -215.7961 | 218.359 | 1.9069 | -3.3374 | 0.58061 | 3.8981 | 215.7961 | 3.3374 | 14.6709
0.1595 | 11.6598 | -30.8731 | -215.8848 | 218.4166 | 1.9058 | -3.3385 | 0.57729 | 3.8981 | 215.8848 | 3.3385 | 14.8289
0.1613 | 11.5136 | -30.7058 | -215.8825 | 218.3845 | 1.9025 | -3.3396 | 0.57396 | 3.897 | 215.8825 | 3.3396 | 14.9963
0.1631 | 11.3496 | -30.5263 | -215.8205 | 218.2892 | 1.897 | -3.3396 | 0.56953 | 3.8936 | 215.8205 | 3.3396 | 15.1636
0.1649 | 11.1612 | -30.3147 | -215.641 | 218.0742 | 1.8914 | -3.3385 | 0.5651 | 3.8903 | 215.641 | 3.3385 | 15.331
0.1667 | 10.9562 | -30.1075 | -215.395 | 217.7928 | 1.8848 | -3.3374 | 0.56066 | 3.8848 | 215.395 | 3.3374 | 15.4983
0.1685 | 10.7391 | -29.8981 | -215.0582 | 217.4205 | 1.8781 | -3.3352 | 0.55623 | 3.8792 | 215.0582 | 3.3352 | 15.6657
0.1703 | 10.5175 | -29.6765 | -214.646 | 216.9729 | 1.8693 | -3.333 | 0.5518 | 3.8715 | 214.646 | 3.333 | 15.833
0.1721 | 10.308 | -29.4416 | -214.2371 | 216.5274 | 1.8593 | -3.3319 | 0.54515 | 3.8648 | 214.2371 | 3.3319 | 16.0004
0.1739 | 10.0975 | -29.1945 | -213.8183 | 216.0709 | 1.8482 | -3.3307 | 0.53961 | 3.8582 | 213.8183 | 3.3307 | 16.1677
0.1757 | 9.8792 | -28.9363 | -213.3496 | 215.5645 | 1.8371 | -3.3296 | 0.53296 | 3.8504 | 213.3496 | 3.3296 | 16.3351
0.1775 | 9.6809 | -28.6859 | -212.6892 | 214.872 | 1.8249 | -3.3274 | 0.52853 | 3.8427 | 212.6892 | 3.3274 | 16.5024
0.1793 | 9.4637 | -28.4266 | -211.8271 | 213.9767 | 1.8139 | -3.3241 | 0.5241 | 3.8338 | 211.8271 | 3.3241 | 16.6698
0.1811 | 9.262 | -28.1695 | -210.8044 | 212.923 | 1.8017 | -3.3208 | 0.52078 | 3.8238 | 210.8044 | 3.3208 | 16.8371
0.1829 | 9.0593 | -27.8936 | -209.5424 | 211.631 | 1.7873 | -3.3175 | 0.51745 | 3.8139 | 209.5424 | 3.3175 | 17.0045
0.1846 | 8.8532 | -27.5934 | -208.1175 | 210.1751 | 1.7729 | -3.3141 | 0.51524 | 3.8028 | 208.1175 | 3.3141 | 17.1625
0.1864 | 8.6537 | -27.2632 | -206.8133 | 208.8343 | 1.7584 | -3.3119 | 0.5108 | 3.7928 | 206.8133 | 3.3119 | 17.3299
0.1882 | 8.4565 | -26.9296 | -205.6909 | 207.6742 | 1.7452 | -3.3097 | 0.50748 | 3.7828 | 205.6909 | 3.3097 | 17.4972
0.19 | 8.2582 | -26.6083 | -204.7756 | 206.7191 | 1.7319 | -3.3053 | 0.50194 | 3.7729 | 204.7756 | 3.3053 | 17.6646
0.1918 | 8.0576 | -26.308 | -204.0066 | 205.9125 | 1.7208 | -3.3008 | 0.4964 | 3.7629 | 204.0066 | 3.3008 | 17.8319
0.1936 | 7.8515 | -26.041 | -203.4194 | 205.2909 | 1.7119 | -3.2975 | 0.49086 | 3.7551 | 203.4194 | 3.2975 | 17.9993
0.1954 | 7.6465 | -25.8039 | -203.0626 | 204.9008 | 1.7042 | -3.2942 | 0.48532 | 3.7485 | 203.0626 | 3.2942 | 18.1666
0.1972 | 7.4438 | -25.5745 | -202.9584 | 204.7623 | 1.6964 | -3.292 | 0.47867 | 3.7418 | 202.9584 | 3.292 | 18.334
0.199 | 7.2465 | -25.3485 | -202.9319 | 204.7025 | 1.6875 | -3.2898 | 0.47091 | 3.7352 | 202.9319 | 3.2898 | 18.5013
0.2008 | 7.0471 | -25.1435 | -203.067 | 204.8033 | 1.6776 | -3.2875 | 0.46205 | 3.7274 | 203.067 | 3.2875 | 18.6687
0.2026 | 6.8576 | -24.9607 | -203.2676 | 204.974 | 1.6665 | -3.2842 | 0.45319 | 3.7197 | 203.2676 | 3.2842 | 18.836
0.2044 | 6.6925 | -24.7457 | -203.3053 | 204.9806 | 1.6554 | -3.2809 | 0.44432 | 3.7097 | 203.3053 | 3.2809 | 19.0033
0.2062 | 6.5518 | -24.5374 | -203.1812 | 204.8288 | 1.6454 | -3.2765 | 0.43767 | 3.7019 | 203.1812 | 3.2765 | 19.1707
0.208 | 6.4233 | -24.3457 | -202.9717 | 204.5961 | 1.6366 | -3.2731 | 0.43102 | 3.6931 | 202.9717 | 3.2731 | 19.338
0.2097 | 6.2693 | -24.1452 | -202.6183 | 204.2172 | 1.6277 | -3.2687 | 0.42659 | 3.6853 | 202.6183 | 3.2687 | 19.4961
0.2115 | 6.1119 | -23.9191 | -202.1296 | 203.7019 | 1.6177 | -3.2654 | 0.42105 | 3.6776 | 202.1296 | 3.2654 | 19.6634
0.2133 | 5.938 | -23.682 | -201.5701 | 203.1136 | 1.6078 | -3.2598 | 0.41551 | 3.6676 | 201.5701 | 3.2598 | 19.8308
0.2151 | 5.7562 | -23.4382 | -200.9994 | 202.5152 | 1.5989 | -3.2554 | 0.41108 | 3.6587 | 200.9994 | 3.2554 | 19.9981
0.2169 | 5.5712 | -23.1934 | -200.4111 | 201.8981 | 1.5889 | -3.2499 | 0.40665 | 3.6499 | 200.4111 | 3.2499 | 20.1655
0.2187 | 5.3762 | -22.9673 | -199.877 | 201.3374 | 1.5789 | -3.2465 | 0.40111 | 3.641 | 199.877 | 3.2465 | 20.3328
0.2205 | 5.2044 | -22.7734 | -199.4759 | 200.9141 | 1.569 | -3.2432 | 0.39557 | 3.6332 | 199.4759 | 3.2432 | 20.5002
0.2223 | 5.0449 | -22.5828 | -199.1945 | 200.6094 | 1.5601 | -3.2399 | 0.39003 | 3.6266 | 199.1945 | 3.2399 | 20.6675
0.2241 | 4.9097 | -22.3745 | -198.9673 | 200.3568 | 1.5501 | -3.2377 | 0.38338 | 3.6188 | 198.9673 | 3.2377 | 20.8349
0.2259 | 4.7922 | -22.1861 | -198.6537 | 200.0199 | 1.5413 | -3.2343 | 0.37784 | 3.6122 | 198.6537 | 3.2343 | 21.0022
0.2277 | 4.6659 | -22.0233 | -198.1806 | 199.5269 | 1.5302 | -3.231 | 0.37341 | 3.6033 | 198.1806 | 3.231 | 21.1696
0.2295 | 4.523 | -21.8382 | -197.5424 | 198.8676 | 1.518 | -3.2255 | 0.36898 | 3.5934 | 197.5424 | 3.2255 | 21.3369
0.2313 | 4.3734 | -21.5967 | -196.7634 | 198.0609 | 1.5058 | -3.2199 | 0.36454 | 3.5823 | 196.7634 | 3.2199 | 21.5043
0.2331 | 4.226 | -21.3551 | -195.9136 | 197.1845 | 1.4925 | -3.2144 | 0.36011 | 3.5712 | 195.9136 | 3.2144 | 21.6716
0.2348 | 4.0665 | -21.1047 | -194.8842 | 196.1285 | 1.4781 | -3.2089 | 0.35679 | 3.559 | 194.8842 | 3.2089 | 21.8297
0.2366 | 3.9069 | -20.851 | -193.8803 | 195.097 | 1.4648 | -3.2022 | 0.35346 | 3.5479 | 193.8803 | 3.2022 | 21.997
0.2384 | 3.7629 | -20.6393 | -193.0316 | 194.226 | 1.4515 | -3.1967 | 0.35014 | 3.5368 | 193.0316 | 3.1967 | 22.1644
0.2402 | 3.6033 | -20.4166 | -192.3645 | 193.5324 | 1.4393 | -3.1922 | 0.34571 | 3.5258 | 192.3645 | 3.1922 | 22.3317
0.242 | 3.4382 | -20.1762 | -191.7418 | 192.882 | 1.4271 | -3.1856 | 0.34017 | 3.5147 | 191.7418 | 3.1856 | 22.4991
0.2438 | 3.2776 | -19.9468 | -190.9983 | 192.1141 | 1.415 | -3.1789 | 0.33573 | 3.5025 | 190.9983 | 3.1789 | 22.6664
0.2456 | 3.1202 | -19.7285 | -190.1662 | 191.2598 | 1.4028 | -3.1712 | 0.33241 | 3.4903 | 190.1662 | 3.1712 | 22.8338
0.2474 | 2.9584 | -19.5125 | -189.2931 | 190.3657 | 1.3906 | -3.1634 | 0.32798 | 3.477 | 189.2931 | 3.1634 | 23.0011
0.2492 | 2.8033 | -19.303 | -188.4886 | 189.5402 | 1.3784 | -3.1546 | 0.32465 | 3.4648 | 188.4886 | 3.1546 | 23.1685
0.251 | 2.6626 | -19.118 | -187.8017 | 188.8343 | 1.3673 | -3.149 | 0.32133 | 3.4537 | 187.8017 | 3.149 | 23.3358
0.2528 | 2.5341 | -18.9529 | -187.2022 | 188.2194 | 1.3562 | -3.1446 | 0.31801 | 3.4449 | 187.2022 | 3.1446 | 23.5032
0.2546 | 2.4078 | -18.7978 | -186.6825 | 187.6831 | 1.3474 | -3.1413 | 0.31468 | 3.4382 | 186.6825 | 3.1413 | 23.6705
0.2564 | 2.2737 | -18.636 | -186.1795 | 187.1634 | 1.3385 | -3.138 | 0.31136 | 3.4305 | 186.1795 | 3.138 | 23.8379
0.2581 | 2.1452 | -18.4709 | -185.6133 | 186.5817 | 1.3285 | -3.1346 | 0.30803 | 3.4227 | 185.6133 | 3.1346 | 23.9959
0.2599 | 2.031 | -18.3125 | -185.0925 | 186.0454 | 1.3186 | -3.1313 | 0.30471 | 3.415 | 185.0925 | 3.1313 | 24.1633
0.2617 | 1.9368 | -18.1673 | -184.6704 | 185.6089 | 1.3097 | -3.1269 | 0.30028 | 3.4072 | 184.6704 | 3.1269 | 24.3306
0.2635 | 1.8427 | -18.0299 | -184.3302 | 185.2543 | 1.3008 | -3.1235 | 0.29584 | 3.3994 | 184.3302 | 3.1235 | 24.498
0.2653 | 1.7496 | -17.8892 | -184.0576 | 184.9662 | 1.292 | -3.1191 | 0.29141 | 3.3928 | 184.0576 | 3.1191 | 24.6653
0.2671 | 1.662 | -17.7529 | -183.8061 | 184.7003 | 1.2842 | -3.1158 | 0.28698 | 3.3861 | 183.8061 | 3.1158 | 24.8327
0.2689 | 1.5812 | -17.6432 | -183.5524 | 184.4343 | 1.2765 | -3.1125 | 0.28255 | 3.3795 | 183.5524 | 3.1125 | 25
0.2707 | 1.4925 | -17.5546 | -183.3119 | 184.185 | 1.2676 | -3.1091 | 0.27812 | 3.3717 | 183.3119 | 3.1091 | 25.1673
0.2725 | 1.4139 | -17.4781 | -183.1047 | 183.9712 | 1.2598 | -3.1047 | 0.27479 | 3.3651 | 183.1047 | 3.1047 | 25.3347
0.2743 | 1.3562 | -17.4161 | -182.9584 | 183.8183 | 1.2532 | -3.1025 | 0.27036 | 3.3607 | 182.9584 | 3.1025 | 25.502
0.2761 | 1.3053 | -17.3562 | -182.8632 | 183.7175 | 1.2454 | -3.1014 | 0.26593 | 3.3551 | 182.8632 | 3.1014 | 25.6694
0.2779 | 1.2554 | -17.2864 | -182.8366 | 183.6831 | 1.2388 | -3.1003 | 0.2615 | 3.3518 | 182.8366 | 3.1003 | 25.8367
0.2797 | 1.2078 | -17.2044 | -182.8776 | 183.7152 | 1.2321 | -3.1003 | 0.25596 | 3.3485 | 182.8776 | 3.1003 | 26.0041
0.2815 | 1.1623 | -17.1058 | -183.0283 | 183.856 | 1.2266 | -3.1003 | 0.25042 | 3.3452 | 183.0283 | 3.1003 | 26.1714
0.2832 | 1.1258 | -17.0161 | -183.2776 | 184.0953 | 1.2222 | -3.1003 | 0.24377 | 3.3429 | 183.2776 | 3.1003 | 26.3295
0.285 | 1.0947 | -16.9374 | -183.5922 | 184.4011 | 1.2177 | -3.1003 | 0.23712 | 3.3418 | 183.5922 | 3.1003 | 26.4968
0.2868 | 1.0571 | -16.8388 | -183.9313 | 184.7291 | 1.2122 | -3.1003 | 0.23047 | 3.3385 | 183.9313 | 3.1003 | 26.6642
0.2886 | 1.015 | -16.728 | -184.2593 | 185.0438 | 1.2066 | -3.0981 | 0.22161 | 3.3341 | 184.2593 | 3.0981 | 26.8315
0.2904 | 0.97175 | -16.6139 | -184.5263 | 185.2975 | 1.2011 | -3.0958 | 0.21385 | 3.3296 | 184.5263 | 3.0958 | 26.9989
0.2922 | 0.94515 | -16.523 | -184.6227 | 185.3839 | 1.1978 | -3.0947 | 0.20831 | 3.3263 | 184.6227 | 3.0947 | 27.1662
0.294 | 0.9385 | -16.4632 | -184.5507 | 185.3064 | 1.1922 | -3.0947 | 0.20388 | 3.3241 | 184.5507 | 3.0947 | 27.3336
0.2958 | 0.94072 | -16.4199 | -184.3612 | 185.1125 | 1.1878 | -3.0947 | 0.20055 | 3.3219 | 184.3612 | 3.0947 | 27.5009
0.2976 | 0.94294 | -16.3989 | -184.1285 | 184.8787 | 1.1823 | -3.0947 | 0.19834 | 3.3208 | 184.1285 | 3.0947 | 27.6683
0.2994 | 0.94404 | -16.3856 | -183.9801 | 184.728 | 1.1767 | -3.0958 | 0.19612 | 3.3197 | 183.9801 | 3.0958 | 27.8356
0.3012 | 0.93961 | -16.3668 | -183.9989 | 184.7446 | 1.1723 | -3.097 | 0.19169 | 3.3186 | 183.9989 | 3.097 | 28.003
0.303 | 0.92188 | -16.3291 | -184.1917 | 184.933 | 1.169 | -3.0981 | 0.18726 | 3.3186 | 184.1917 | 3.0981 | 28.1703
0.3048 | 0.91413 | -16.2892 | -184.4665 | 185.2033 | 1.1657 | -3.0981 | 0.18172 | 3.3175 | 184.4665 | 3.0981 | 28.3377
0.3066 | 0.91634 | -16.246 | -184.8465 | 185.5767 | 1.1634 | -3.1014 | 0.17618 | 3.3186 | 184.8465 | 3.1014 | 28.505
0.3083 | 0.92521 | -16.2105 | -185.344 | 186.0698 | 1.1612 | -3.1047 | 0.17064 | 3.3197 | 185.344 | 3.1047 | 28.6631
0.3101 | 0.9385 | -16.1873 | -185.8859 | 186.6072 | 1.159 | -3.108 | 0.16399 | 3.323 | 185.8859 | 3.108 | 28.8304
0.3119 | 0.97285 | -16.1939 | -186.2526 | 186.9729 | 1.1568 | -3.1136 | 0.15956 | 3.3263 | 186.2526 | 3.1136 | 28.9978
0.3137 | 1.0305 | -16.2327 | -186.451 | 187.1734 | 1.1557 | -3.1202 | 0.15734 | 3.333 | 186.451 | 3.1202 | 29.1651
0.3155 | 1.0925 | -16.2759 | -186.6604 | 187.3861 | 1.1546 | -3.128 | 0.15512 | 3.3396 | 186.6604 | 3.128 | 29.3325
0.3173 | 1.1446 | -16.2981 | -187.0172 | 187.7429 | 1.1535 | -3.1346 | 0.1518 | 3.3452 | 187.0172 | 3.1346 | 29.4998
0.3191 | 1.1845 | -16.2848 | -187.5457 | 188.2681 | 1.1524 | -3.1413 | 0.14626 | 3.3507 | 187.5457 | 3.1413 | 29.6672
0.3209 | 1.2321 | -16.2571 | -188.2371 | 188.9551 | 1.1546 | -3.1479 | 0.13961 | 3.3573 | 188.2371 | 3.1479 | 29.8345
0.3227 | 1.3008 | -16.2482 | -189.0593 | 189.774 | 1.1557 | -3.1546 | 0.13296 | 3.364 | 189.0593 | 3.1546 | 30.0019
0.3245 | 1.3784 | -16.2582 | -189.867 | 190.5806 | 1.1579 | -3.1612 | 0.12632 | 3.3706 | 189.867 | 3.1612 | 30.1692
0.3263 | 1.4593 | -16.2693 | -190.5994 | 191.3108 | 1.1601 | -3.1679 | 0.11967 | 3.3773 | 190.5994 | 3.1679 | 30.3366
0.3281 | 1.5457 | -16.2548 | -191.1435 | 191.8537 | 1.1634 | -3.1745 | 0.11524 | 3.385 | 191.1435 | 3.1745 | 30.5039
0.3299 | 1.6244 | -16.246 | -191.6 | 192.308 | 1.1645 | -3.1812 | 0.1108 | 3.3917 | 191.6 | 3.1812 | 30.6713
0.3316 | 1.6853 | -16.2172 | -192.0576 | 192.7634 | 1.1668 | -3.1867 | 0.10526 | 3.3972 | 192.0576 | 3.1867 | 30.8293
0.3334 | 1.7241 | -16.1618 | -192.6648 | 193.364 | 1.1712 | -3.1922 | 0.099723 | 3.4039 | 192.6648 | 3.1922 | 30.9967
0.3352 | 1.744 | -16.0909 | -193.4172 | 194.1075 | 1.1767 | -3.1978 | 0.093075 | 3.4105 | 193.4172 | 3.1978 | 31.164
0.337 | 1.7695 | -16.0222 | -194.1983 | 194.882 | 1.1812 | -3.2033 | 0.086427 | 3.4183 | 194.1983 | 3.2033 | 31.3313
0.3388 | 1.805 | -15.9745 | -194.8864 | 195.5645 | 1.1856 | -3.21 | 0.080886 | 3.4249 | 194.8864 | 3.21 | 31.4987
0.3406 | 1.8548 | -15.9557 | -195.426 | 196.1008 | 1.1878 | -3.2155 | 0.076454 | 3.4305 | 195.426 | 3.2155 | 31.666
0.3424 | 1.9003 | -15.9346 | -195.826 | 196.4997 | 1.19 | -3.2199 | 0.072022 | 3.436 | 195.826 | 3.2199 | 31.8334
0.3442 | 1.9402 | -15.9202 | -196.0986 | 196.7701 | 1.1922 | -3.2233 | 0.068698 | 3.4404 | 196.0986 | 3.2233 | 32.0007
0.346 | 1.9745 | -15.9091 | -196.3247 | 196.9961 | 1.1945 | -3.2277 | 0.065374 | 3.446 | 196.3247 | 3.2277 | 32.1681
0.3478 | 2.0011 | -15.8837 | -196.6216 | 197.2898 | 1.1967 | -3.2332 | 0.06205 | 3.4515 | 196.6216 | 3.2332 | 32.3354
0.3496 | 2.0377 | -15.8548 | -197.0792 | 197.7452 | 1.2 | -3.2399 | 0.057618 | 3.4582 | 197.0792 | 3.2399 | 32.5028
0.3514 | 2.0964 | -15.8349 | -197.7208 | 198.3834 | 1.2022 | -3.2465 | 0.053186 | 3.4659 | 197.7208 | 3.2465 | 32.6701
0.3532 | 2.154 | -15.805 | -198.5008 | 199.1601 | 1.2066 | -3.2543 | 0.046537 | 3.4748 | 198.5008 | 3.2543 | 32.8375
0.355 | 2.1729 | -15.7208 | -199.3729 | 200.0244 | 1.2133 | -3.262 | 0.040997 | 3.4837 | 199.3729 | 3.262 | 33.0048
0.3567 | 2.1972 | -15.6476 | -200.2438 | 200.8886 | 1.2188 | -3.2687 | 0.034349 | 3.4925 | 200.2438 | 3.2687 | 33.1629
0.3585 | 2.2327 | -15.5911 | -201.0127 | 201.6521 | 1.2233 | -3.2753 | 0.028809 | 3.5003 | 201.0127 | 3.2753 | 33.3302
0.3603 | 2.2726 | -15.5435 | -201.6355 | 202.2715 | 1.2277 | -3.282 | 0.023269 | 3.508 | 201.6355 | 3.282 | 33.4976
0.3621 | 2.3346 | -15.5191 | -202.2039 | 202.8377 | 1.2299 | -3.2886 | 0.018837 | 3.5147 | 202.2039 | 3.2886 | 33.6649
0.3639 | 2.4177 | -15.4848 | -202.8598 | 203.4892 | 1.2332 | -3.2964 | 0.013296 | 3.5224 | 202.8598 | 3.2964 | 33.8323
0.3657 | 2.5097 | -15.4216 | -203.6199 | 204.2438 | 1.2388 | -3.3064 | 0.0077562 | 3.5335 | 203.6199 | 3.3064 | 33.9996
0.3675 | 2.5961 | -15.3684 | -204.2659 | 204.8853 | 1.2432 | -3.3152 | 0.0044321 | 3.5435 | 204.2659 | 3.3152 | 34.167
0.3693 | 2.6803 | -15.3175 | -204.8244 | 205.4393 | 1.2465 | -3.323 | 0 | 3.5524 | 204.8244 | 3.323 | 34.3343
0.3711 | 2.7645 | -15.262 | -205.2787 | 205.8892 | 1.251 | -3.3307 | -0.0044321 | 3.5612 | 205.2787 | 3.3307 | 34.5017
0.3729 | 2.8332 | -15.2155 | -205.6044 | 206.2105 | 1.2532 | -3.3352 | -0.0077562 | 3.5668 | 205.6044 | 3.3352 | 34.669
0.3747 | 2.9008 | -15.1812 | -205.8493 | 206.4532 | 1.2565 | -3.3407 | -0.01108 | 3.5723 | 205.8493 | 3.3407 | 34.8364
0.3765 | 2.9651 | -15.1723 | -206.0033 | 206.6072 | 1.2587 | -3.344 | -0.013296 | 3.5767 | 206.0033 | 3.344 | 35.0037
0.3783 | 3.0294 | -15.1657 | -206.0831 | 206.687 | 1.2609 | -3.3485 | -0.014404 | 3.5812 | 206.0831 | 3.3485 | 35.1711
0.3801 | 3.0981 | -15.1291 | -206.0975 | 206.7003 | 1.2654 | -3.354 | -0.014404 | 3.5889 | 206.0975 | 3.354 | 35.3384
0.3818 | 3.1645 | -15.0781 | -206.1097 | 206.7102 | 1.2687 | -3.3607 | -0.015512 | 3.5956 | 206.1097 | 3.3607 | 35.4965
0.3836 | 3.2133 | -15.0116 | -206.2271 | 206.8233 | 1.272 | -3.3662 | -0.017729 | 3.6022 | 206.2271 | 3.3662 | 35.6638
0.3854 | 3.2565 | -14.9573 | -206.4654 | 207.0593 | 1.2753 | -3.3717 | -0.019945 | 3.6089 | 206.4654 | 3.3717 | 35.8312
0.3872 | 3.2875 | -14.892 | -206.7701 | 207.3584 | 1.2787 | -3.3784 | -0.023269 | 3.6155 | 206.7701 | 3.3784 | 35.9985
0.389 | 3.3175 | -14.8266 | -207.0958 | 207.6798 | 1.2842 | -3.385 | -0.026593 | 3.6233 | 207.0958 | 3.385 | 36.1659
0.3908 | 3.3673 | -14.7557 | -207.4283 | 208.0066 | 1.2886 | -3.3917 | -0.031025 | 3.631 | 207.4283 | 3.3917 | 36.3332
0.3926 | 3.436 | -14.6914 | -207.7873 | 208.3612 | 1.292 | -3.3983 | -0.034349 | 3.6399 | 207.7873 | 3.3983 | 36.5006
0.3944 | 3.5125 | -14.6371 | -208.2194 | 208.79 | 1.2953 | -3.4072 | -0.036565 | 3.6488 | 208.2194 | 3.4072 | 36.6679
0.3962 | 3.5801 | -14.5751 | -208.7136 | 209.2798 | 1.2986 | -3.4161 | -0.039889 | 3.6587 | 208.7136 | 3.4161 | 36.8353
0.398 | 3.631 | -14.4953 | -209.262 | 209.8216 | 1.303 | -3.4238 | -0.044321 | 3.6676 | 209.262 | 3.4238 | 37.0026
0.3998 | 3.6842 | -14.4111 | -209.7751 | 210.3302 | 1.3086 | -3.4338 | -0.047645 | 3.6787 | 209.7751 | 3.4338 | 37.17
0.4016 | 3.7352 | -14.3169 | -210.3147 | 210.8632 | 1.313 | -3.4449 | -0.05097 | 3.6909 | 210.3147 | 3.4449 | 37.3373
0.4034 | 3.785 | -14.2238 | -211.0371 | 211.5778 | 1.3175 | -3.456 | -0.055402 | 3.703 | 211.0371 | 3.456 | 37.5046
0.4051 | 3.8427 | -14.1163 | -211.8781 | 212.4122 | 1.3219 | -3.4681 | -0.059834 | 3.7163 | 211.8781 | 3.4681 | 37.6627
0.4069 | 3.8958 | -14.0122 | -212.759 | 213.2853 | 1.3263 | -3.4781 | -0.065374 | 3.7274 | 212.759 | 3.4781 | 37.83
0.4087 | 3.9535 | -13.9014 | -213.7163 | 214.236 | 1.3307 | -3.4903 | -0.070914 | 3.7407 | 213.7163 | 3.4903 | 37.9974
0.4105 | 4.0222 | -13.7695 | -214.7324 | 215.2443 | 1.3363 | -3.5036 | -0.077562 | 3.7551 | 214.7324 | 3.5036 | 38.1647
0.4123 | 4.0942 | -13.6587 | -215.662 | 216.1651 | 1.3407 | -3.5158 | -0.083102 | 3.7684 | 215.662 | 3.5158 | 38.3321
0.4141 | 4.1751 | -13.5778 | -216.4997 | 216.9972 | 1.3452 | -3.528 | -0.087535 | 3.7806 | 216.4997 | 3.528 | 38.4994
0.4159 | 4.2693 | -13.5069 | -217.3452 | 217.8393 | 1.3474 | -3.5402 | -0.093075 | 3.7939 | 217.3452 | 3.5402 | 38.6668
0.4177 | 4.3634 | -13.4316 | -218.2526 | 218.7424 | 1.3507 | -3.5524 | -0.097507 | 3.8072 | 218.2526 | 3.5524 | 38.8341
0.4195 | 4.4443 | -13.3485 | -219.1213 | 219.6055 | 1.3551 | -3.5645 | -0.10305 | 3.8194 | 219.1213 | 3.5645 | 39.0015
0.4213 | 4.523 | -13.2731 | -219.8837 | 220.3634 | 1.3607 | -3.5745 | -0.10748 | 3.8305 | 219.8837 | 3.5745 | 39.1688
0.4231 | 4.6006 | -13.1867 | -220.533 | 221.0072 | 1.3662 | -3.5834 | -0.11191 | 3.8404 | 220.533 | 3.5834 | 39.3362
0.4249 | 4.6626 | -13.0825 | -221.0859 | 221.5535 | 1.3717 | -3.5889 | -0.11634 | 3.8482 | 221.0859 | 3.5889 | 39.5035
0.4267 | 4.7247 | -12.995 | -221.4039 | 221.867 | 1.3784 | -3.5956 | -0.11967 | 3.8571 | 221.4039 | 3.5956 | 39.6709
0.4285 | 4.7834 | -12.8964 | -221.6576 | 222.1152 | 1.3861 | -3.6033 | -0.12188 | 3.8659 | 221.6576 | 3.6033 | 39.8382
0.4302 | 4.8343 | -12.759 | -221.9535 | 222.4055 | 1.3939 | -3.6122 | -0.1241 | 3.877 | 221.9535 | 3.6122 | 39.9963
0.432 | 4.8986 | -12.6183 | -222.338 | 222.7812 | 1.3994 | -3.6222 | -0.12632 | 3.8892 | 222.338 | 3.6222 | 40.1636
0.4338 | 4.974 | -12.4787 | -222.8066 | 223.2432 | 1.405 | -3.6332 | -0.12853 | 3.9014 | 222.8066 | 3.6332 | 40.331
0.4356 | 5.0515 | -12.3391 | -223.4781 | 223.908 | 1.4105 | -3.6465 | -0.13186 | 3.9158 | 223.4781 | 3.6465 | 40.4983
0.4374 | 5.1258 | -12.1961 | -224.2925 | 224.7147 | 1.415 | -3.6598 | -0.13629 | 3.9302 | 224.2925 | 3.6598 | 40.6657
0.4392 | 5.2011 | -12.0443 | -225.2321 | 225.6465 | 1.4194 | -3.6742 | -0.14072 | 3.9457 | 225.2321 | 3.6742 | 40.833
0.441 | 5.2898 | -11.8947 | -226.215 | 226.6227 | 1.4238 | -3.6886 | -0.14515 | 3.9612 | 226.215 | 3.6886 | 41.0004
0.4428 | 5.3673 | -11.7307 | -227.2665 | 227.6654 | 1.4283 | -3.7042 | -0.14958 | 3.9767 | 227.2665 | 3.7042 | 41.1677
0.4446 | 5.4382 | -11.5645 | -228.3579 | 228.7501 | 1.4316 | -3.7197 | -0.15512 | 3.9922 | 228.3579 | 3.7197 | 41.3351
0.4464 | 5.5091 | -11.4006 | -229.4604 | 229.8449 | 1.436 | -3.7341 | -0.15956 | 4.0078 | 229.4604 | 3.7341 | 41.5024
0.4482 | 5.5801 | -11.2244 | -230.5474 | 230.9241 | 1.4393 | -3.7485 | -0.1651 | 4.0233 | 230.5474 | 3.7485 | 41.6698
0.45 | 5.6499 | -11.046 | -231.61 | 231.9789 | 1.4438 | -3.764 | -0.16953 | 4.0388 | 231.61 | 3.764 | 41.8371
0.4518 | 5.7252 | -10.8654 | -232.7147 | 233.077 | 1.4471 | -3.7784 | -0.17396 | 4.0532 | 232.7147 | 3.7784 | 42.0045
0.4536 | 5.8194 | -10.6958 | -233.8681 | 234.2238 | 1.4504 | -3.7917 | -0.17839 | 4.0676 | 233.8681 | 3.7917 | 42.1718
0.4553 | 5.9302 | -10.5418 | -235.0183 | 235.3684 | 1.4537 | -3.8061 | -0.18283 | 4.082 | 235.0183 | 3.8061 | 42.3299
0.4571 | 6.041 | -10.3922 | -236.1152 | 236.4609 | 1.456 | -3.8205 | -0.18726 | 4.0964 | 236.1152 | 3.8205 | 42.4972
0.4589 | 6.1562 | -10.2427 | -237.1402 | 237.4814 | 1.4582 | -3.8338 | -0.19058 | 4.1108 | 237.1402 | 3.8338 | 42.6646
0.4607 | 6.2704 | -10.0964 | -238.0421 | 238.3778 | 1.4604 | -3.8482 | -0.1928 | 4.1252 | 238.0421 | 3.8482 | 42.8319
0.4625 | 6.3812 | -9.9512 | -238.8222 | 239.1535 | 1.4626 | -3.8626 | -0.19501 | 4.1385 | 238.8222 | 3.8626 | 42.9993
0.4643 | 6.4964 | -9.805 | -239.5934 | 239.9224 | 1.4648 | -3.877 | -0.19723 | 4.1529 | 239.5934 | 3.877 | 43.1666
0.4661 | 6.6105 | -9.6565 | -240.4255 | 240.7512 | 1.467 | -3.8892 | -0.19945 | 4.1651 | 240.4255 | 3.8892 | 43.334
0.4679 | 6.708 | -9.5025 | -241.3906 | 241.7119 | 1.4693 | -3.9014 | -0.20277 | 4.1773 | 241.3906 | 3.9014 | 43.5013
0.4697 | 6.8044 | -9.3485 | -242.4022 | 242.7202 | 1.4715 | -3.9125 | -0.2072 | 4.1895 | 242.4022 | 3.9125 | 43.6687
0.4715 | 6.8964 | -9.1967 | -243.4316 | 243.744 | 1.4737 | -3.9235 | -0.21053 | 4.2006 | 243.4316 | 3.9235 | 43.836
0.4733 | 6.9895 | -9.0504 | -244.4543 | 244.7634 | 1.4748 | -3.9346 | -0.21496 | 4.2116 | 244.4543 | 3.9346 | 44.0033
0.4751 | 7.0748 | -8.8997 | -245.4681 | 245.7729 | 1.477 | -3.9446 | -0.21939 | 4.2216 | 245.4681 | 3.9446 | 44.1707
0.4769 | 7.1535 | -8.7457 | -246.492 | 246.7934 | 1.4781 | -3.9546 | -0.22382 | 4.2316 | 246.492 | 3.9546 | 44.338
0.4786 | 7.231 | -8.5895 | -247.4992 | 247.7972 | 1.4803 | -3.9645 | -0.22936 | 4.2416 | 247.4992 | 3.9645 | 44.4961
0.4804 | 7.3075 | -8.4277 | -248.4377 | 248.7302 | 1.4814 | -3.9734 | -0.2338 | 4.2515 | 248.4377 | 3.9734 | 44.6634
0.4822 | 7.385 | -8.2726 | -249.3152 | 249.6044 | 1.4825 | -3.9823 | -0.23712 | 4.2604 | 249.3152 | 3.9823 | 44.8308
0.484 | 7.4637 | -8.1241 | -250.144 | 250.4288 | 1.4837 | -3.9911 | -0.24155 | 4.2681 | 250.144 | 3.9911 | 44.9981
0.4858 | 7.5413 | -7.9823 | -250.9263 | 251.2089 | 1.4848 | -3.9989 | -0.24488 | 4.277 | 250.9263 | 3.9989 | 45.1655
0.4876 | 7.6166 | -7.8471 | -251.662 | 251.9413 | 1.4859 | -4.0066 | -0.2482 | 4.2848 | 251.662 | 4.0066 | 45.3328
0.4894 | 7.6853 | -7.7152 | -252.3235 | 252.6006 | 1.4859 | -4.0144 | -0.25042 | 4.2914 | 252.3235 | 4.0144 | 45.5002
0.4912 | 7.7496 | -7.5778 | -252.8875 | 253.1601 | 1.4848 | -4.0222 | -0.25374 | 4.2981 | 252.8875 | 4.0222 | 45.6675
0.493 | 7.8061 | -7.4349 | -253.4316 | 253.7019 | 1.4848 | -4.0288 | -0.25596 | 4.3047 | 253.4316 | 4.0288 | 45.8349
0.4948 | 7.8504 | -7.2853 | -253.9823 | 254.2493 | 1.4837 | -4.0355 | -0.25928 | 4.3114 | 253.9823 | 4.0355 | 46.0022
0.4966 | 7.8881 | -7.1346 | -254.5152 | 254.7767 | 1.4837 | -4.041 | -0.2626 | 4.3169 | 254.5152 | 4.041 | 46.1696
0.4984 | 7.9424 | -6.9784 | -254.9374 | 255.1956 | 1.4848 | -4.0476 | -0.26593 | 4.3224 | 254.9374 | 4.0476 | 46.3369
0.5002 | 8.0443 | -6.8188 | -255.2975 | 255.5546 | 1.4881 | -4.0543 | -0.26704 | 4.3302 | 255.2975 | 4.0543 | 46.5043
0.502 | 8.1075 | -6.6183 | -255.6521 | 255.9036 | 1.4936 | -4.0587 | -0.27036 | 4.3368 | 255.6521 | 4.0587 | 46.6716
0.5037 | 8.1274 | -6.4211 | -256.0576 | 256.3025 | 1.497 | -4.062 | -0.27368 | 4.3413 | 256.0576 | 4.062 | 46.8297
0.5055 | 8.1274 | -6.2349 | -256.451 | 256.6914 | 1.4981 | -4.0654 | -0.27812 | 4.3446 | 256.451 | 4.0654 | 46.997
0.5073 | 8.1019 | -6.0476 | -256.7657 | 256.9994 | 1.4992 | -4.0676 | -0.28255 | 4.3479 | 256.7657 | 4.0676 | 47.1644
0.5091 | 8.0643 | -5.8427 | -257.0271 | 257.2554 | 1.4992 | -4.0698 | -0.28698 | 4.3501 | 257.0271 | 4.0698 | 47.3317
0.5109 | 8.0222 | -5.6399 | -257.241 | 257.4637 | 1.5003 | -4.072 | -0.29252 | 4.3524 | 257.241 | 4.072 | 47.4991
0.5127 | 7.9712 | -5.4294 | -257.3972 | 257.6133 | 1.5003 | -4.0731 | -0.29806 | 4.3546 | 257.3972 | 4.0731 | 47.6664
0.5145 | 7.8992 | -5.2022 | -257.4936 | 257.7019 | 1.5014 | -4.0731 | -0.3036 | 4.3557 | 257.4936 | 4.0731 | 47.8338
0.5163 | 7.8127 | -4.9607 | -257.5313 | 257.7307 | 1.5014 | -4.0731 | -0.31025 | 4.3557 | 257.5313 | 4.0731 | 48.0011
0.5181 | 7.7285 | -4.7435 | -257.4992 | 257.6931 | 1.5003 | -4.0731 | -0.31579 | 4.3546 | 257.4992 | 4.0731 | 48.1685
0.5199 | 7.6443 | -4.523 | -257.4083 | 257.5967 | 1.5003 | -4.0731 | -0.32022 | 4.3557 | 257.4083 | 4.0731 | 48.3358
0.5217 | 7.5424 | -4.267 | -257.2476 | 257.4283 | 1.4992 | -4.0731 | -0.32465 | 4.3546 | 257.2476 | 4.0731 | 48.5032
0.5235 | 7.4183 | -3.9934 | -257.0305 | 257.2066 | 1.4981 | -4.0709 | -0.3313 | 4.3535 | 257.0305 | 4.0709 | 48.6705
0.5253 | 7.2931 | -3.7285 | -256.7579 | 256.9296 | 1.4958 | -4.0676 | -0.33684 | 4.3501 | 256.7579 | 4.0676 | 48.8379
0.5271 | 7.1679 | -3.4704 | -256.431 | 256.5972 | 1.4925 | -4.0654 | -0.34238 | 4.3468 | 256.431 | 4.0654 | 49.0052
0.5288 | 7.0482 | -3.2055 | -256.0632 | 256.2249 | 1.4903 | -4.062 | -0.34792 | 4.3435 | 256.0632 | 4.062 | 49.1633
0.5306 | 6.8986 | -2.8753 | -255.6565 | 255.8139 | 1.4881 | -4.0554 | -0.35568 | 4.3368 | 255.6565 | 4.0554 | 49.3306
0.5324 | 6.7357 | -2.5163 | -255.2055 | 255.3573 | 1.487 | -4.0488 | -0.36233 | 4.3313 | 255.2055 | 4.0488 | 49.498
0.5342 | 6.5706 | -2.1607 | -254.6737 | 254.8211 | 1.4848 | -4.0432 | -0.36898 | 4.3247 | 254.6737 | 4.0432 | 49.6653
0.536 | 6.3934 | -1.8006 | -254.0687 | 254.2083 | 1.4837 | -4.0377 | -0.37452 | 4.3202 | 254.0687 | 4.0377 | 49.8327
0.5378 | 6.2083 | -1.4316 | -253.4327 | 253.5634 | 1.4814 | -4.0321 | -0.38006 | 4.3158 | 253.4327 | 4.0321 | 50
0.5396 | 6 | -1.0294 | -252.7557 | 252.8787 | 1.4792 | -4.0255 | -0.3856 | 4.3091 | 252.7557 | 4.0255 | 50.1673
0.5414 | 5.7729 | -0.60499 | -252.0366 | 252.1518 | 1.477 | -4.0177 | -0.39224 | 4.3003 | 252.0366 | 4.0177 | 50.3347
0.5432 | 5.5224 | -0.1651 | -251.2886 | 251.3961 | 1.4748 | -4.0089 | -0.39778 | 4.2925 | 251.2886 | 4.0089 | 50.502
0.545 | 5.2554 | 0.30914 | -250.4654 | 250.5662 | 1.4726 | -3.9978 | -0.40443 | 4.2825 | 250.4654 | 3.9978 | 50.6694
0.5468 | 4.9873 | 0.8 | -249.5834 | 249.6798 | 1.4704 | -3.9856 | -0.41219 | 4.2704 | 249.5834 | 3.9856 | 50.8367
0.5486 | 4.718 | 1.2997 | -248.6726 | 248.7645 | 1.467 | -3.9734 | -0.41884 | 4.2582 | 248.6726 | 3.9734 | 51.0041
0.5504 | 4.4499 | 1.8061 | -247.7374 | 247.826 | 1.4637 | -3.959 | -0.42659 | 4.2449 | 247.7374 | 3.959 | 51.1714
0.5521 | 4.1895 | 2.3235 | -246.7645 | 246.851 | 1.4582 | -3.9435 | -0.43324 | 4.2294 | 246.7645 | 3.9435 | 51.3295
0.5539 | 3.9424 | 2.8432 | -245.7629 | 245.8504 | 1.4526 | -3.9291 | -0.441 | 4.2139 | 245.7629 | 3.9291 | 51.4968
0.5557 | 3.7086 | 3.3651 | -244.7457 | 244.8321 | 1.4449 | -3.9136 | -0.44765 | 4.1983 | 244.7457 | 3.9136 | 51.6642
0.5575 | 3.4781 | 3.908 | -243.6942 | 243.7828 | 1.4371 | -3.897 | -0.4554 | 4.1806 | 243.6942 | 3.897 | 51.8315
0.5593 | 3.2299 | 4.4698 | -242.5884 | 242.6803 | 1.4271 | -3.8781 | -0.46205 | 4.1607 | 242.5884 | 3.8781 | 51.9989
0.5611 | 2.9562 | 5.0737 | -241.2997 | 241.3972 | 1.4172 | -3.856 | -0.46981 | 4.1374 | 241.2997 | 3.856 | 52.1662
0.5629 | 2.6914 | 5.6875 | -240.0399 | 240.1452 | 1.4061 | -3.8327 | -0.47645 | 4.113 | 240.0399 | 3.8327 | 52.3336
0.5647 | 2.4576 | 6.2992 | -238.9274 | 239.0449 | 1.3939 | -3.8072 | -0.48532 | 4.0853 | 238.9274 | 3.8072 | 52.5009
0.5665 | 2.2947 | 6.8532 | -238.1285 | 238.2593 | 1.3806 | -3.7795 | -0.4964 | 4.0565 | 238.1285 | 3.7795 | 52.6683
0.5683 | 2.1695 | 7.3839 | -237.5125 | 237.6576 | 1.364 | -3.7507 | -0.50748 | 4.0255 | 237.5125 | 3.7507 | 52.8356
0.5701 | 2.0975 | 7.8925 | -237.1934 | 237.354 | 1.344 | -3.7175 | -0.52188 | 3.99 | 237.1934 | 3.7175 | 53.003
0.5719 | 2.1385 | 8.3834 | -237.5102 | 237.6886 | 1.3208 | -3.6798 | -0.54072 | 3.9501 | 237.5102 | 3.6798 | 53.1703
0.5737 | 2.2338 | 8.8366 | -238.0753 | 238.2715 | 1.2953 | -3.6388 | -0.56177 | 3.9069 | 238.0753 | 3.6388 | 53.3377
0.5755 | 2.3435 | 9.272 | -238.6449 | 238.8587 | 1.2687 | -3.5945 | -0.58283 | 3.8604 | 238.6449 | 3.5945 | 53.505
0.5772 | 2.4199 | 9.7075 | -238.8731 | 239.1036 | 1.2432 | -3.5457 | -0.60277 | 3.8105 | 238.8731 | 3.5457 | 53.6631
0.579 | 2.5075 | 10.1241 | -238.6471 | 238.8953 | 1.2166 | -3.4936 | -0.6205 | 3.7562 | 238.6471 | 3.4936 | 53.8304
0.5808 | 2.5972 | 10.533 | -238.1828 | 238.4499 | 1.19 | -3.4349 | -0.63712 | 3.6964 | 238.1828 | 3.4349 | 53.9978
0.5826 | 2.6981 | 10.8831 | -237.5701 | 237.856 | 1.1623 | -3.3706 | -0.65374 | 3.6321 | 237.5701 | 3.3706 | 54.1651
0.5844 | 2.8177 | 11.2155 | -236.7956 | 237.1003 | 1.1335 | -3.3042 | -0.67036 | 3.5645 | 236.7956 | 3.3042 | 54.3325
0.5862 | 2.9474 | 11.487 | -235.8726 | 236.1928 | 1.1025 | -3.2355 | -0.68476 | 3.4936 | 235.8726 | 3.2355 | 54.4998
0.588 | 3.0715 | 11.6986 | -234.7911 | 235.1247 | 1.0715 | -3.1612 | -0.69917 | 3.4205 | 234.7911 | 3.1612 | 54.6672
0.5898 | 3.1867 | 11.8604 | -233.5291 | 233.8748 | 1.0393 | -3.0859 | -0.71136 | 3.344 | 233.5291 | 3.0859 | 54.8345
0.5916 | 3.303 | 11.9889 | -232.1064 | 232.4609 | 1.0072 | -3.0094 | -0.72244 | 3.2665 | 232.1064 | 3.0094 | 55.0019
0.5934 | 3.4083 | 12.1241 | -230.5141 | 230.8798 | 0.97507 | -2.9319 | -0.73241 | 3.1878 | 230.5141 | 2.9319 | 55.1692
0.5952 | 3.4981 | 12.2715 | -228.6936 | 229.0704 | 0.94294 | -2.8532 | -0.74238 | 3.108 | 228.6936 | 2.8532 | 55.3366
0.597 | 3.59 | 12.3956 | -226.6659 | 227.0548 | 0.9097 | -2.7734 | -0.75014 | 3.0271 | 226.6659 | 2.7734 | 55.5039
0.5988 | 3.7019 | 12.4643 | -224.5861 | 224.9861 | 0.87645 | -2.6936 | -0.75789 | 2.9463 | 224.5861 | 2.6936 | 55.6713
0.6006 | 3.8205 | 12.4222 | -222.4598 | 222.8643 | 0.841 | -2.6094 | -0.76343 | 2.8609 | 222.4598 | 2.6094 | 55.8386
0.6023 | 3.9269 | 12.369 | -220.2471 | 220.656 | 0.80554 | -2.523 | -0.76898 | 2.7734 | 220.2471 | 2.523 | 55.9967
0.6041 | 4.0355 | 12.3069 | -217.9845 | 218.3967 | 0.76898 | -2.4366 | -0.77452 | 2.6859 | 217.9845 | 2.4366 | 56.164
0.6059 | 4.1474 | 12.2404 | -215.631 | 216.0465 | 0.73241 | -2.3524 | -0.77895 | 2.5994 | 215.631 | 2.3524 | 56.3313
0.6077 | 4.2748 | 12.0864 | -213.241 | 213.6554 | 0.69252 | -2.2648 | -0.78227 | 2.5097 | 213.241 | 2.2648 | 56.4987
0.6095 | 4.4199 | 11.8438 | -210.9108 | 211.3186 | 0.65263 | -2.174 | -0.7856 | 2.4177 | 210.9108 | 2.174 | 56.666
0.6113 | 4.544 | 11.5535 | -208.4177 | 208.8166 | 0.61274 | -2.0798 | -0.7867 | 2.3224 | 208.4177 | 2.0798 | 56.8334
0.6131 | 4.6504 | 11.2355 | -205.7352 | 206.1241 | 0.57285 | -1.9834 | -0.78781 | 2.226 | 205.7352 | 1.9834 | 57.0007
0.6149 | 4.7247 | 10.8798 | -202.687 | 203.0637 | 0.53518 | -1.8848 | -0.78449 | 2.1263 | 202.687 | 1.8848 | 57.1681
0.6167 | 4.7778 | 10.513 | -199.4537 | 199.8205 | 0.49861 | -1.7861 | -0.78116 | 2.0277 | 199.4537 | 1.7861 | 57.3354
0.6185 | 4.8421 | 10.0709 | -196.1551 | 196.5075 | 0.46205 | -1.6875 | -0.77452 | 1.928 | 196.1551 | 1.6875 | 57.5028
0.6203 | 4.8997 | 9.5878 | -192.7446 | 193.0803 | 0.42548 | -1.5867 | -0.76676 | 1.8283 | 192.7446 | 1.5867 | 57.6701
0.6221 | 4.944 | 9.0161 | -189.1391 | 189.4526 | 0.39003 | -1.4848 | -0.75789 | 1.7263 | 189.1391 | 1.4848 | 57.8375
0.6239 | 4.9452 | 8.4266 | -185.3097 | 185.6 | 0.35568 | -1.3806 | -0.74571 | 1.6233 | 185.3097 | 1.3806 | 58.0048
0.6256 | 4.9108 | 7.815 | -181.3385 | 181.6033 | 0.32355 | -1.2776 | -0.73241 | 1.5202 | 181.3385 | 1.2776 | 58.1629
0.6274 | 4.8443 | 7.1402 | -177.1125 | 177.3496 | 0.29474 | -1.1734 | -0.71579 | 1.4194 | 177.1125 | 1.1734 | 58.3302
0.6292 | 4.7158 | 6.4632 | -172.554 | 172.7634 | 0.27147 | -1.0704 | -0.69584 | 1.3197 | 172.554 | 1.0704 | 58.4976
0.631 | 4.5352 | 5.7828 | -167.6776 | 167.8604 | 0.25042 | -0.9662 | -0.67368 | 1.2211 | 167.6776 | 0.9662 | 58.6649
0.6328 | 4.2936 | 5.1136 | -162.4488 | 162.6061 | 0.2338 | -0.86316 | -0.6482 | 1.1235 | 162.4488 | 0.86316 | 58.8323
0.6346 | 3.9778 | 4.4853 | -156.9219 | 157.056 | 0.2205 | -0.76122 | -0.62161 | 1.0316 | 156.9219 | 0.76122 | 58.9996
0.6364 | 3.5934 | 3.8859 | -151.0626 | 151.1756 | 0.21163 | -0.66482 | -0.59169 | 0.94626 | 151.0626 | 0.66482 | 59.167
0.6382 | 3.1801 | 3.3008 | -145.0416 | 145.1335 | 0.20499 | -0.56953 | -0.56066 | 0.86427 | 145.0416 | 0.56953 | 59.3343
0.64 | 2.7368 | 2.7302 | -138.9307 | 139.005 | 0.20055 | -0.47645 | -0.52853 | 0.78338 | 138.9307 | 0.47645 | 59.5017
0.6418 | 2.2604 | 2.1828 | -132.7856 | 132.8421 | 0.19612 | -0.3856 | -0.4964 | 0.70693 | 132.7856 | 0.3856 | 59.669
0.6436 | 1.7474 | 1.6986 | -126.831 | 126.8731 | 0.19612 | -0.29917 | -0.46648 | 0.64044 | 126.831 | 0.29917 | 59.8364
0.6454 | 1.2022 | 1.2753 | -120.9452 | 120.9773 | 0.19723 | -0.21828 | -0.43657 | 0.58615 | 120.9452 | 0.21828 | 60.0037
0.6472 | 0.65042 | 0.88532 | -115.1546 | 115.1812 | 0.20055 | -0.14404 | -0.40776 | 0.53961 | 115.1546 | 0.14404 | 60.1711
0.649 | 0.091967 | 0.52742 | -109.4305 | 109.4537 | 0.20499 | -0.072022 | -0.38006 | 0.49972 | 109.4305 | 0.072022 | 60.3384
0.6507 | -0.45651 | 0.20609 | -103.8615 | 103.8837 | 0.20942 | -0.0044321 | -0.35235 | 0.46759 | 103.8615 | 0.0044321 | 60.4965
0.6525 | -0.99723 | -0.059834 | -98.6116 | 98.6371 | 0.21496 | 0.05651 | -0.32798 | 0.44654 | 98.6116 | -0.05651 | 60.6638
0.6543 | -1.4992 | -0.28809 | -93.8604 | 93.8925 | 0.22161 | 0.11302 | -0.30582 | 0.43435 | 93.8604 | -0.11302 | 60.8312
0.6561 | -1.9202 | -0.49861 | -89.744 | 89.7839 | 0.22715 | 0.16177 | -0.28698 | 0.42992 | 89.744 | -0.16177 | 60.9985
0.6579 | -2.2626 | -0.71357 | -86.1629 | 86.2139 | 0.23047 | 0.20388 | -0.27147 | 0.43324 | 86.1629 | -0.20388 | 61.1659
0.6597 | -2.5263 | -0.92853 | -83.0139 | 83.0759 | 0.23269 | 0.24044 | -0.25706 | 0.441 | 83.0139 | -0.24044 | 61.3332
0.6615 | -2.7335 | -1.128 | -80.1551 | 80.226 | 0.2349 | 0.27258 | -0.24488 | 0.45097 | 80.1551 | -0.27258 | 61.5006
0.6633 | -2.8742 | -1.3042 | -77.6576 | 77.7363 | 0.23601 | 0.30139 | -0.2338 | 0.46094 | 77.6576 | -0.30139 | 61.6679
0.6651 | -2.9706 | -1.4416 | -75.4504 | 75.5368 | 0.23934 | 0.32576 | -0.22382 | 0.47313 | 75.4504 | -0.32576 | 61.8353
0.6669 | -3.0471 | -1.5446 | -73.3994 | 73.4925 | 0.24266 | 0.34792 | -0.21385 | 0.48421 | 73.3994 | -0.34792 | 62.0026
0.6687 | -3.1136 | -1.6255 | -71.4116 | 71.5114 | 0.24598 | 0.36898 | -0.20388 | 0.4964 | 71.4116 | -0.36898 | 62.17
0.6705 | -3.169 | -1.682 | -69.5136 | 69.6188 | 0.25042 | 0.38781 | -0.19501 | 0.50637 | 69.5136 | -0.38781 | 62.3373
0.6723 | -3.1978 | -1.692 | -67.8416 | 67.9501 | 0.25485 | 0.40443 | -0.18615 | 0.51745 | 67.8416 | -0.40443 | 62.5046
0.6741 | -3.2133 | -1.6709 | -66.3712 | 66.482 | 0.25928 | 0.41994 | -0.17839 | 0.52853 | 66.3712 | -0.41994 | 62.672
0.6758 | -3.2321 | -1.6222 | -65.0338 | 65.1468 | 0.26371 | 0.43324 | -0.17175 | 0.5385 | 65.0338 | -0.43324 | 62.83
0.6776 | -3.2576 | -1.5524 | -63.7629 | 63.877 | 0.26925 | 0.44765 | -0.16399 | 0.54958 | 63.7629 | -0.44765 | 62.9974
0.6794 | -3.2853 | -1.4637 | -62.6216 | 62.7368 | 0.27368 | 0.45873 | -0.15845 | 0.55956 | 62.6216 | -0.45873 | 63.1647
0.6812 | -3.3197 | -1.3507 | -61.5579 | 61.6753 | 0.27812 | 0.46981 | -0.15291 | 0.56842 | 61.5579 | -0.46981 | 63.3321
0.683 | -3.3562 | -1.2277 | -60.5551 | 60.6737 | 0.28144 | 0.47867 | -0.14737 | 0.57618 | 60.5551 | -0.47867 | 63.4994
0.6848 | -3.3861 | -1.0981 | -59.6399 | 59.7607 | 0.28476 | 0.48753 | -0.14183 | 0.58393 | 59.6399 | -0.48753 | 63.6668
0.6866 | -3.4194 | -0.96731 | -58.7823 | 58.9053 | 0.28698 | 0.49529 | -0.13629 | 0.58947 | 58.7823 | -0.49529 | 63.8341
0.6884 | -3.4482 | -0.85208 | -57.9468 | 58.0709 | 0.28809 | 0.50305 | -0.13075 | 0.59612 | 57.9468 | -0.50305 | 64.0015
0.6902 | -3.4582 | -0.75014 | -57.0792 | 57.2055 | 0.2892 | 0.50859 | -0.12521 | 0.60055 | 57.0792 | -0.50859 | 64.1688
0.692 | -3.4571 | -0.65596 | -56.2127 | 56.3391 | 0.2892 | 0.51302 | -0.11967 | 0.60388 | 56.2127 | -0.51302 | 64.3362
0.6938 | -3.4659 | -0.56842 | -55.3884 | 55.5169 | 0.2892 | 0.51856 | -0.11413 | 0.60609 | 55.3884 | -0.51856 | 64.5035
0.6956 | -3.4715 | -0.48199 | -54.6006 | 54.7302 | 0.28809 | 0.52299 | -0.10859 | 0.60831 | 54.6006 | -0.52299 | 64.6709
0.6974 | -3.4837 | -0.39889 | -53.8593 | 53.9911 | 0.28698 | 0.52632 | -0.10416 | 0.61053 | 53.8593 | -0.52632 | 64.8382
0.6992 | -3.5014 | -0.31911 | -53.1734 | 53.3064 | 0.28476 | 0.53075 | -0.099723 | 0.61274 | 53.1734 | -0.53075 | 65.0056
0.7009 | -3.5058 | -0.24488 | -52.523 | 52.6571 | 0.28255 | 0.53407 | -0.095291 | 0.61274 | 52.523 | -0.53407 | 65.1636
0.7027 | -3.5003 | -0.17507 | -51.8881 | 52.0222 | 0.28033 | 0.53629 | -0.090859 | 0.61274 | 51.8881 | -0.53629 | 65.331
0.7045 | -3.4881 | -0.097507 | -51.2366 | 51.3695 | 0.27922 | 0.5374 | -0.086427 | 0.61385 | 51.2366 | -0.5374 | 65.4983
0.7063 | -3.4615 | -0.022161 | -50.5596 | 50.6925 | 0.27701 | 0.5385 | -0.081994 | 0.61163 | 50.5596 | -0.5385 | 65.6657
0.7081 | -3.4271 | 0.054294 | -49.9114 | 50.0432 | 0.27368 | 0.5385 | -0.077562 | 0.61053 | 49.9114 | -0.5385 | 65.833
0.7099 | -3.4028 | 0.13075 | -49.3596 | 49.4892 | 0.27147 | 0.5374 | -0.07313 | 0.6072 | 49.3596 | -0.5374 | 66.0004
0.7117 | -3.3784 | 0.19612 | -48.8964 | 49.026 | 0.26704 | 0.53629 | -0.070914 | 0.60499 | 48.8964 | -0.53629 | 66.1677
0.7135 | -3.3529 | 0.2626 | -48.4554 | 48.5828 | 0.26371 | 0.53518 | -0.06759 | 0.60166 | 48.4554 | -0.53518 | 66.3351
0.7153 | -3.3341 | 0.31801 | -48.0044 | 48.1319 | 0.26039 | 0.53296 | -0.065374 | 0.59834 | 48.0044 | -0.53296 | 66.5024
0.7171 | -3.3241 | 0.35679 | -47.5125 | 47.6399 | 0.25706 | 0.53296 | -0.06205 | 0.59501 | 47.5125 | -0.53296 | 66.6698
0.7189 | -3.3152 | 0.38116 | -46.9163 | 47.0449 | 0.25263 | 0.53186 | -0.058726 | 0.5928 | 46.9163 | -0.53186 | 66.8371
0.7207 | -3.3086 | 0.4 | -46.1994 | 46.3302 | 0.2482 | 0.53075 | -0.054294 | 0.58947 | 46.1994 | -0.53075 | 67.0045
0.7225 | -3.3086 | 0.42105 | -45.4305 | 45.5634 | 0.24488 | 0.52964 | -0.048753 | 0.58615 | 45.4305 | -0.52964 | 67.1718
0.7242 | -3.2997 | 0.44986 | -44.636 | 44.7701 | 0.24044 | 0.52742 | -0.044321 | 0.58283 | 44.636 | -0.52742 | 67.3299
0.726 | -3.2831 | 0.47535 | -43.8416 | 43.9767 | 0.23712 | 0.52521 | -0.038781 | 0.57839 | 43.8416 | -0.52521 | 67.4972
0.7278 | -3.2576 | 0.49972 | -43.0947 | 43.231 | 0.2338 | 0.52188 | -0.034349 | 0.57285 | 43.0947 | -0.52188 | 67.6646
0.7296 | -3.2288 | 0.53407 | -42.4321 | 42.5684 | 0.23047 | 0.51745 | -0.031025 | 0.56731 | 42.4321 | -0.51745 | 67.8319
0.7314 | -3.1978 | 0.57507 | -41.8704 | 42.0066 | 0.22715 | 0.51302 | -0.027701 | 0.56177 | 41.8704 | -0.51302 | 67.9993
0.7332 | -3.1612 | 0.62604 | -41.426 | 41.5612 | 0.22271 | 0.50748 | -0.025485 | 0.55623 | 41.426 | -0.50748 | 68.1666
0.735 | -3.1324 | 0.67701 | -41.0648 | 41.2 | 0.21939 | 0.50305 | -0.024377 | 0.55069 | 41.0648 | -0.50305 | 68.334
0.7368 | -3.1058 | 0.70914 | -40.769 | 40.9042 | 0.21607 | 0.49972 | -0.023269 | 0.54515 | 40.769 | -0.49972 | 68.5013
0.7386 | -3.0726 | 0.7313 | -40.4343 | 40.5684 | 0.21163 | 0.49529 | -0.022161 | 0.53961 | 40.4343 | -0.49529 | 68.6687
0.7404 | -3.0227 | 0.74238 | -39.9368 | 40.0676 | 0.2072 | 0.48864 | -0.018837 | 0.53296 | 39.9368 | -0.48864 | 68.836
0.7422 | -2.9717 | 0.73684 | -39.3584 | 39.4881 | 0.20277 | 0.4831 | -0.015512 | 0.52521 | 39.3584 | -0.4831 | 69.0033
0.744 | -2.9241 | 0.73352 | -38.7357 | 38.8632 | 0.19834 | 0.47756 | -0.012188 | 0.51745 | 38.7357 | -0.47756 | 69.1707
0.7458 | -2.8898 | 0.73795 | -38.1163 | 38.2438 | 0.19391 | 0.47091 | -0.0088643 | 0.5108 | 38.1163 | -0.47091 | 69.338
0.7476 | -2.8587 | 0.74238 | -37.5457 | 37.672 | 0.19058 | 0.46427 | -0.0055402 | 0.50305 | 37.5457 | -0.46427 | 69.5054
0.7493 | -2.8144 | 0.74903 | -37.0371 | 37.1634 | 0.18615 | 0.45762 | -0.0033241 | 0.49418 | 37.0371 | -0.45762 | 69.6634
0.7511 | -2.7623 | 0.76122 | -36.6139 | 36.7368 | 0.18172 | 0.45097 | -0.001108 | 0.48643 | 36.6139 | -0.45097 | 69.8308
0.7529 | -2.7158 | 0.77452 | -36.2748 | 36.3967 | 0.17839 | 0.44432 | 0 | 0.47867 | 36.2748 | -0.44432 | 69.9981
0.7547 | -2.6693 | 0.79557 | -35.9922 | 36.1119 | 0.17396 | 0.43878 | 0.001108 | 0.47202 | 35.9922 | -0.43878 | 70.1655
0.7565 | -2.6216 | 0.81773 | -35.7307 | 35.8482 | 0.17064 | 0.43324 | 0.0022161 | 0.46648 | 35.7307 | -0.43324 | 70.3328
0.7583 | -2.5762 | 0.83989 | -35.5014 | 35.6177 | 0.16731 | 0.4277 | 0.0033241 | 0.45983 | 35.5014 | -0.4277 | 70.5002
0.7601 | -2.5363 | 0.86648 | -35.3285 | 35.4416 | 0.16399 | 0.42216 | 0.0044321 | 0.45319 | 35.3285 | -0.42216 | 70.6675
0.7619 | -2.5008 | 0.89751 | -35.2144 | 35.3263 | 0.16066 | 0.41662 | 0.0033241 | 0.44765 | 35.2144 | -0.41662 | 70.8349
0.7637 | -2.4654 | 0.93075 | -35.0925 | 35.2033 | 0.15734 | 0.41108 | 0.0033241 | 0.441 | 35.0925 | -0.41108 | 71.0022
0.7655 | -2.4255 | 0.96066 | -34.8532 | 34.9629 | 0.15402 | 0.40554 | 0.0044321 | 0.43435 | 34.8532 | -0.40554 | 71.1696
0.7673 | -2.3812 | 0.98393 | -34.513 | 34.6205 | 0.15069 | 0.39889 | 0.0055402 | 0.4277 | 34.513 | -0.39889 | 71.3369
0.7691 | -2.338 | 0.99501 | -34.092 | 34.1972 | 0.14737 | 0.39335 | 0.0088643 | 0.41994 | 34.092 | -0.39335 | 71.5043
0.7709 | -2.2992 | 0.99945 | -33.6033 | 33.7086 | 0.14404 | 0.3867 | 0.01108 | 0.4133 | 33.6033 | -0.3867 | 71.6716
0.7727 | -2.2648 | 1.0061 | -33.1435 | 33.2476 | 0.14072 | 0.38006 | 0.014404 | 0.40554 | 33.1435 | -0.38006 | 71.839
0.7744 | -2.2416 | 1.0238 | -32.8299 | 32.9341 | 0.1374 | 0.37452 | 0.015512 | 0.39889 | 32.8299 | -0.37452 | 71.997
0.7762 | -2.2271 | 1.0526 | -32.718 | 32.8222 | 0.13407 | 0.36898 | 0.015512 | 0.39335 | 32.718 | -0.36898 | 72.1644
0.778 | -2.2205 | 1.0892 | -32.6981 | 32.8033 | 0.13186 | 0.36565 | 0.015512 | 0.38892 | 32.6981 | -0.36565 | 72.3317
0.7798 | -2.2283 | 1.1402 | -32.769 | 32.8753 | 0.12964 | 0.36122 | 0.014404 | 0.38449 | 32.769 | -0.36122 | 72.4991
0.7816 | -2.2438 | 1.2033 | -32.9119 | 33.0216 | 0.12742 | 0.359 | 0.012188 | 0.38116 | 32.9119 | -0.359 | 72.6664
0.7834 | -2.2582 | 1.2665 | -33.0726 | 33.1845 | 0.12521 | 0.35568 | 0.0099723 | 0.37784 | 33.0726 | -0.35568 | 72.8338
0.7852 | -2.2537 | 1.3197 | -33.1213 | 33.2343 | 0.12299 | 0.35235 | 0.0099723 | 0.37341 | 33.1213 | -0.35235 | 73.0011
0.787 | -2.226 | 1.3429 | -33.0438 | 33.1557 | 0.11967 | 0.34792 | 0.0099723 | 0.36898 | 33.0438 | -0.34792 | 73.1685
0.7888 | -2.1839 | 1.3485 | -32.8731 | 32.9828 | 0.11634 | 0.34349 | 0.012188 | 0.36343 | 32.8731 | -0.34349 | 73.3358
0.7906 | -2.1341 | 1.3452 | -32.6693 | 32.7756 | 0.11302 | 0.33795 | 0.014404 | 0.35789 | 32.6693 | -0.33795 | 73.5032
0.7924 | -2.0831 | 1.3429 | -32.4521 | 32.5562 | 0.1097 | 0.33352 | 0.01662 | 0.35235 | 32.4521 | -0.33352 | 73.6705
0.7942 | -2.0399 | 1.3385 | -32.2825 | 32.3834 | 0.10637 | 0.32909 | 0.018837 | 0.34681 | 32.2825 | -0.32909 | 73.8379
0.796 | -2.0133 | 1.3474 | -32.226 | 32.3258 | 0.10305 | 0.32465 | 0.019945 | 0.34238 | 32.226 | -0.32465 | 74.0052
0.7977 | -2.0089 | 1.3751 | -32.2914 | 32.3911 | 0.10083 | 0.32244 | 0.019945 | 0.33906 | 32.2914 | -0.32244 | 74.1633
0.7995 | -2.0155 | 1.4116 | -32.4122 | 32.5141 | 0.098615 | 0.32022 | 0.018837 | 0.33573 | 32.4122 | -0.32022 | 74.3306
0.8013 | -2.0343 | 1.4615 | -32.605 | 32.7091 | 0.096399 | 0.31911 | 0.017729 | 0.33463 | 32.605 | -0.31911 | 74.498
0.8031 | -2.051 | 1.5125 | -32.8377 | 32.944 | 0.095291 | 0.31911 | 0.01662 | 0.33352 | 32.8377 | -0.31911 | 74.6653
0.8049 | -2.0787 | 1.5723 | -33.0958 | 33.2055 | 0.093075 | 0.31801 | 0.014404 | 0.33241 | 33.0958 | -0.31801 | 74.8327
0.8067 | -2.1108 | 1.6355 | -33.3551 | 33.4693 | 0.091967 | 0.31911 | 0.013296 | 0.33241 | 33.3551 | -0.31911 | 75
0.8085 | -2.1274 | 1.6953 | -33.5634 | 33.6798 | 0.090859 | 0.31801 | 0.012188 | 0.3313 | 33.5634 | -0.31801 | 75.1673
0.8103 | -2.123 | 1.7352 | -33.6399 | 33.7573 | 0.088643 | 0.3169 | 0.012188 | 0.33019 | 33.6399 | -0.3169 | 75.3347
0.8121 | -2.0975 | 1.7551 | -33.5823 | 33.6997 | 0.086427 | 0.31579 | 0.013296 | 0.32798 | 33.5823 | -0.31579 | 75.502
0.8139 | -2.0665 | 1.7651 | -33.4726 | 33.59 | 0.083102 | 0.31357 | 0.015512 | 0.32576 | 33.4726 | -0.31357 | 75.6694
0.8157 | -2.0321 | 1.7706 | -33.3562 | 33.4715 | 0.080886 | 0.31247 | 0.017729 | 0.32355 | 33.3562 | -0.31247 | 75.8367
0.8175 | -2.0022 | 1.7762 | -33.292 | 33.4061 | 0.07867 | 0.31025 | 0.019945 | 0.32133 | 33.292 | -0.31025 | 76.0041
0.8193 | -1.9701 | 1.7961 | -33.2986 | 33.4116 | 0.076454 | 0.30803 | 0.021053 | 0.31911 | 33.2986 | -0.30803 | 76.1714
0.8211 | -1.9512 | 1.8305 | -33.4249 | 33.5391 | 0.074238 | 0.30693 | 0.021053 | 0.3169 | 33.4249 | -0.30693 | 76.3388
0.8228 | -1.9512 | 1.8825 | -33.6753 | 33.7906 | 0.07313 | 0.30693 | 0.019945 | 0.3169 | 33.6753 | -0.30693 | 76.4968
0.8246 | -1.9668 | 1.9446 | -33.9867 | 34.1064 | 0.070914 | 0.30803 | 0.018837 | 0.3169 | 33.9867 | -0.30803 | 76.6642
0.8264 | -1.9945 | 2.0144 | -34.3457 | 34.4698 | 0.069806 | 0.30914 | 0.01662 | 0.31801 | 34.3457 | -0.30914 | 76.8315
0.8282 | -2.0222 | 2.0831 | -34.6848 | 34.8144 | 0.068698 | 0.31136 | 0.015512 | 0.32022 | 34.6848 | -0.31136 | 76.9989
0.83 | -2.0399 | 2.1396 | -34.9319 | 35.0648 | 0.06759 | 0.31247 | 0.014404 | 0.32133 | 34.9319 | -0.31247 | 77.1662
0.8318 | -2.0488 | 2.1762 | -35.087 | 35.2211 | 0.065374 | 0.31357 | 0.015512 | 0.32133 | 35.087 | -0.31357 | 77.3336
0.8336 | -2.0465 | 2.2116 | -35.1712 | 35.3086 | 0.064266 | 0.31468 | 0.015512 | 0.32244 | 35.1712 | -0.31468 | 77.5009
0.8354 | -2.0521 | 2.2427 | -35.2199 | 35.3596 | 0.06205 | 0.31579 | 0.01662 | 0.32355 | 35.2199 | -0.31579 | 77.6683
0.8372 | -2.0643 | 2.2693 | -35.2155 | 35.3573 | 0.060942 | 0.3169 | 0.017729 | 0.32355 | 35.2155 | -0.3169 | 77.8356
0.839 | -2.0787 | 2.287 | -35.1202 | 35.2643 | 0.058726 | 0.31911 | 0.019945 | 0.32576 | 35.1202 | -0.31911 | 78.003
0.8408 | -2.0787 | 2.2992 | -34.9762 | 35.1224 | 0.05651 | 0.32022 | 0.023269 | 0.32576 | 34.9762 | -0.32022 | 78.1703
0.8426 | -2.0787 | 2.308 | -34.8222 | 34.9684 | 0.054294 | 0.32133 | 0.025485 | 0.32687 | 34.8222 | -0.32133 | 78.3377
0.8444 | -2.0731 | 2.3091 | -34.6238 | 34.7712 | 0.052078 | 0.32133 | 0.028809 | 0.32798 | 34.6238 | -0.32133 | 78.505
0.8462 | -2.0676 | 2.3147 | -34.4465 | 34.5939 | 0.05097 | 0.32244 | 0.031025 | 0.32798 | 34.4465 | -0.32244 | 78.6724
0.8479 | -2.0731 | 2.3391 | -34.3834 | 34.533 | 0.048753 | 0.32355 | 0.033241 | 0.32909 | 34.3834 | -0.32355 | 78.8304
0.8497 | -2.0931 | 2.3801 | -34.4654 | 34.6183 | 0.048753 | 0.32465 | 0.033241 | 0.3313 | 34.4654 | -0.32465 | 78.9978
0.8515 | -2.123 | 2.4366 | -34.6681 | 34.8255 | 0.047645 | 0.32687 | 0.033241 | 0.33352 | 34.6681 | -0.32687 | 79.1651
0.8533 | -2.1562 | 2.4997 | -34.9363 | 35.0992 | 0.047645 | 0.33019 | 0.032133 | 0.33573 | 34.9363 | -0.33019 | 79.3325
0.8551 | -2.1928 | 2.5695 | -35.2166 | 35.385 | 0.046537 | 0.33241 | 0.031025 | 0.33795 | 35.2166 | -0.33241 | 79.4998
0.8569 | -2.2316 | 2.6327 | -35.4748 | 35.6488 | 0.046537 | 0.33573 | 0.029917 | 0.34127 | 35.4748 | -0.33573 | 79.6672
0.8587 | -2.2626 | 2.6704 | -35.6454 | 35.8227 | 0.045429 | 0.33795 | 0.029917 | 0.34349 | 35.6454 | -0.33795 | 79.8345
0.8605 | -2.2825 | 2.687 | -35.713 | 35.8936 | 0.043213 | 0.34127 | 0.029917 | 0.34571 | 35.713 | -0.34127 | 80.0019
0.8623 | -2.3047 | 2.6936 | -35.682 | 35.8648 | 0.042105 | 0.34349 | 0.032133 | 0.34792 | 35.682 | -0.34349 | 80.1692
0.8641 | -2.3501 | 2.6848 | -35.5479 | 35.7341 | 0.039889 | 0.34571 | 0.033241 | 0.35125 | 35.5479 | -0.34571 | 80.3366
0.8659 | -2.3911 | 2.6648 | -35.2898 | 35.4781 | 0.037673 | 0.34792 | 0.035457 | 0.35346 | 35.2898 | -0.34792 | 80.5039
0.8677 | -2.4222 | 2.6349 | -34.9551 | 35.1457 | 0.036565 | 0.35014 | 0.038781 | 0.35457 | 34.9551 | -0.35014 | 80.6713
0.8695 | -2.4499 | 2.6006 | -34.5884 | 34.7801 | 0.034349 | 0.35125 | 0.042105 | 0.35568 | 34.5884 | -0.35125 | 80.8386
0.8712 | -2.4665 | 2.5729 | -34.2194 | 34.4122 | 0.033241 | 0.35125 | 0.045429 | 0.35679 | 34.2194 | -0.35125 | 80.9967
0.873 | -2.4765 | 2.544 | -33.9202 | 34.1141 | 0.032133 | 0.35125 | 0.047645 | 0.35679 | 33.9202 | -0.35125 | 81.164
0.8748 | -2.4776 | 2.5219 | -33.7374 | 33.9302 | 0.031025 | 0.35125 | 0.048753 | 0.35568 | 33.7374 | -0.35125 | 81.3313
0.8766 | -2.4776 | 2.523 | -33.7773 | 33.9701 | 0.029917 | 0.35125 | 0.048753 | 0.35568 | 33.7773 | -0.35125 | 81.4987
0.8784 | -2.4765 | 2.5396 | -33.9501 | 34.1429 | 0.028809 | 0.35125 | 0.047645 | 0.35679 | 33.9501 | -0.35125 | 81.666
0.8802 | -2.4687 | 2.5452 | -34.154 | 34.3457 | 0.027701 | 0.35235 | 0.046537 | 0.35679 | 34.154 | -0.35235 | 81.8334
0.882 | -2.4554 | 2.5385 | -34.308 | 34.4975 | 0.026593 | 0.35235 | 0.046537 | 0.35679 | 34.308 | -0.35235 | 82.0007
0.8838 | -2.4521 | 2.5352 | -34.4787 | 34.667 | 0.025485 | 0.35346 | 0.045429 | 0.35789 | 34.4787 | -0.35346 | 82.1681
0.8856 | -2.4598 | 2.5363 | -34.667 | 34.8554 | 0.024377 | 0.35457 | 0.044321 | 0.359 | 34.667 | -0.35457 | 82.3354
0.8874 | -2.4565 | 2.5219 | -34.7778 | 34.964 | 0.023269 | 0.35457 | 0.044321 | 0.359 | 34.7778 | -0.35457 | 82.5028
0.8892 | -2.4543 | 2.4898 | -34.7391 | 34.923 | 0.022161 | 0.35457 | 0.044321 | 0.359 | 34.7391 | -0.35457 | 82.6701
0.891 | -2.4521 | 2.4488 | -34.5961 | 34.7778 | 0.019945 | 0.35568 | 0.045429 | 0.359 | 34.5961 | -0.35568 | 82.8375
0.8928 | -2.4432 | 2.3956 | -34.3801 | 34.5596 | 0.018837 | 0.35457 | 0.047645 | 0.359 | 34.3801 | -0.35457 | 83.0048
0.8946 | -2.4288 | 2.3435 | -34.1208 | 34.297 | 0.017729 | 0.35346 | 0.048753 | 0.35789 | 34.1208 | -0.35346 | 83.1722
0.8963 | -2.41 | 2.3003 | -33.8582 | 34.031 | 0.015512 | 0.35235 | 0.05097 | 0.35789 | 33.8582 | -0.35235 | 83.3302
0.8981 | -2.3934 | 2.267 | -33.6687 | 33.8393 | 0.015512 | 0.35125 | 0.052078 | 0.35568 | 33.6687 | -0.35125 | 83.4976
0.8999 | -2.3712 | 2.2382 | -33.5546 | 33.723 | 0.014404 | 0.34903 | 0.053186 | 0.35457 | 33.5546 | -0.34903 | 83.6649
0.9017 | -2.3446 | 2.2094 | -33.497 | 33.6609 | 0.014404 | 0.34792 | 0.053186 | 0.35346 | 33.497 | -0.34792 | 83.8323
0.9035 | -2.3202 | 2.1939 | -33.5291 | 33.692 | 0.014404 | 0.34681 | 0.053186 | 0.35125 | 33.5291 | -0.34681 | 83.9996
0.9053 | -2.3147 | 2.1983 | -33.682 | 33.8438 | 0.013296 | 0.34571 | 0.052078 | 0.35125 | 33.682 | -0.34571 | 84.167
0.9071 | -2.3269 | 2.2072 | -33.8958 | 34.0587 | 0.014404 | 0.34571 | 0.049861 | 0.35014 | 33.8958 | -0.34571 | 84.3343
0.9089 | -2.3402 | 2.2139 | -34.1452 | 34.3069 | 0.014404 | 0.34571 | 0.046537 | 0.35014 | 34.1452 | -0.34571 | 84.5017
0.9107 | -2.3512 | 2.2139 | -34.3302 | 34.492 | 0.014404 | 0.34571 | 0.044321 | 0.35014 | 34.3302 | -0.34571 | 84.669
0.9125 | -2.3601 | 2.1983 | -34.3457 | 34.5064 | 0.014404 | 0.3446 | 0.043213 | 0.34903 | 34.3457 | -0.3446 | 84.8364
0.9143 | -2.3745 | 2.164 | -34.236 | 34.3967 | 0.014404 | 0.34349 | 0.043213 | 0.34792 | 34.236 | -0.34349 | 85.0037
0.9161 | -2.3834 | 2.1197 | -34.0033 | 34.1629 | 0.013296 | 0.34238 | 0.044321 | 0.34571 | 34.0033 | -0.34238 | 85.1711
0.9179 | -2.3834 | 2.0676 | -33.6222 | 33.7806 | 0.013296 | 0.34017 | 0.046537 | 0.3446 | 33.6222 | -0.34017 | 85.3384
0.9197 | -2.3734 | 2.0166 | -33.1723 | 33.3285 | 0.013296 | 0.33684 | 0.048753 | 0.34127 | 33.1723 | -0.33684 | 85.5058
0.9214 | -2.359 | 1.9645 | -32.718 | 32.8731 | 0.012188 | 0.33352 | 0.05097 | 0.33795 | 32.718 | -0.33352 | 85.6638
0.9232 | -2.3512 | 1.918 | -32.3911 | 32.544 | 0.012188 | 0.33019 | 0.053186 | 0.33573 | 32.3911 | -0.33019 | 85.8312
0.925 | -2.3424 | 1.8825 | -32.1618 | 32.3136 | 0.012188 | 0.32687 | 0.053186 | 0.33241 | 32.1618 | -0.32687 | 85.9985
0.9268 | -2.328 | 1.8416 | -31.9845 | 32.1341 | 0.01108 | 0.32355 | 0.053186 | 0.32909 | 31.9845 | -0.32355 | 86.1659
0.9286 | -2.3291 | 1.8183 | -31.9047 | 32.0532 | 0.01108 | 0.32022 | 0.053186 | 0.32687 | 31.9047 | -0.32022 | 86.3332
0.9304 | -2.3391 | 1.8105 | -31.9058 | 32.0543 | 0.012188 | 0.31801 | 0.05097 | 0.32355 | 31.9058 | -0.31801 | 86.5006
0.9322 | -2.3424 | 1.8094 | -31.9612 | 32.1097 | 0.012188 | 0.31579 | 0.048753 | 0.32133 | 31.9612 | -0.31579 | 86.6679
0.934 | -2.3468 | 1.826 | -32.0942 | 32.2427 | 0.013296 | 0.31247 | 0.046537 | 0.31801 | 32.0942 | -0.31247 | 86.8353
0.9358 | -2.3501 | 1.8515 | -32.2482 | 32.3978 | 0.015512 | 0.31025 | 0.044321 | 0.31468 | 32.2482 | -0.31025 | 87.0026
0.9376 | -2.3546 | 1.8681 | -32.3535 | 32.503 | 0.01662 | 0.30693 | 0.042105 | 0.31247 | 32.3535 | -0.30693 | 87.17
0.9394 | -2.3601 | 1.8626 | -32.369 | 32.5186 | 0.01662 | 0.30471 | 0.040997 | 0.30914 | 32.369 | -0.30471 | 87.3373
0.9412 | -2.3457 | 1.846 | -32.2781 | 32.4266 | 0.01662 | 0.30139 | 0.039889 | 0.30582 | 32.2781 | -0.30139 | 87.5046
0.943 | -2.3224 | 1.8205 | -32.144 | 32.2903 | 0.01662 | 0.29806 | 0.039889 | 0.30249 | 32.144 | -0.29806 | 87.672
0.9447 | -2.3058 | 1.8017 | -31.9634 | 32.1086 | 0.017729 | 0.29252 | 0.039889 | 0.29695 | 31.9634 | -0.29252 | 87.83
0.9465 | -2.2992 | 1.7906 | -31.7019 | 31.8471 | 0.017729 | 0.28698 | 0.040997 | 0.29141 | 31.7019 | -0.28698 | 87.9974
0.9483 | -2.2903 | 1.7828 | -31.3429 | 31.4892 | 0.018837 | 0.28033 | 0.042105 | 0.28587 | 31.3429 | -0.28033 | 88.1647
0.9501 | -2.277 | 1.7795 | -30.9706 | 31.1191 | 0.018837 | 0.27368 | 0.043213 | 0.27922 | 30.9706 | -0.27368 | 88.3321
0.9519 | -2.2537 | 1.785 | -30.7058 | 30.8554 | 0.019945 | 0.26704 | 0.043213 | 0.27258 | 30.7058 | -0.26704 | 88.4994
0.9537 | -2.2072 | 1.8039 | -30.5983 | 30.7457 | 0.021053 | 0.26039 | 0.043213 | 0.26593 | 30.5983 | -0.26039 | 88.6668
0.9555 | -2.1629 | 1.8449 | -30.5817 | 30.7291 | 0.022161 | 0.25263 | 0.042105 | 0.25928 | 30.5817 | -0.25263 | 88.8341
0.9573 | -2.133 | 1.9114 | -30.6482 | 30.7978 | 0.023269 | 0.24598 | 0.039889 | 0.25152 | 30.6482 | -0.24598 | 89.0015
0.9591 | -2.1108 | 1.9845 | -30.7623 | 30.9152 | 0.025485 | 0.23934 | 0.037673 | 0.24488 | 30.7623 | -0.23934 | 89.1688
0.9609 | -2.0964 | 2.0476 | -30.8676 | 31.0249 | 0.026593 | 0.23269 | 0.035457 | 0.23934 | 30.8676 | -0.23269 | 89.3362
0.9627 | -2.0731 | 2.0964 | -30.882 | 31.0404 | 0.027701 | 0.22604 | 0.034349 | 0.23269 | 30.882 | -0.22604 | 89.5035
0.9645 | -2.0421 | 2.1341 | -30.8133 | 30.974 | 0.028809 | 0.21939 | 0.034349 | 0.22604 | 30.8133 | -0.21939 | 89.6709
0.9663 | -2.0078 | 2.1873 | -30.6814 | 30.8432 | 0.028809 | 0.21163 | 0.034349 | 0.21939 | 30.6814 | -0.21163 | 89.8382
0.9681 | -1.9789 | 2.2593 | -30.5828 | 30.749 | 0.029917 | 0.20388 | 0.034349 | 0.21053 | 30.5828 | -0.20388 | 90.0056
0.9698 | -1.9657 | 2.3645 | -30.6072 | 30.7823 | 0.031025 | 0.19501 | 0.033241 | 0.20277 | 30.6072 | -0.19501 | 90.1636
0.9716 | -1.9623 | 2.492 | -30.6471 | 30.8332 | 0.033241 | 0.18504 | 0.032133 | 0.19391 | 30.6471 | -0.18504 | 90.331
0.9734 | -1.9679 | 2.6172 | -30.482 | 30.6814 | 0.034349 | 0.17507 | 0.032133 | 0.18393 | 30.482 | -0.17507 | 90.4983
0.9752 | -1.9657 | 2.7391 | -30.1219 | 30.3357 | 0.036565 | 0.16177 | 0.034349 | 0.17396 | 30.1219 | -0.16177 | 90.6657
0.977 | -1.9745 | 2.851 | -29.672 | 29.9003 | 0.038781 | 0.14848 | 0.037673 | 0.16399 | 29.672 | -0.14848 | 90.833
0.9788 | -2.0033 | 2.9662 | -29.159 | 29.4061 | 0.039889 | 0.13407 | 0.040997 | 0.15291 | 29.159 | -0.13407 | 91.0004
0.9806 | -2.0332 | 3.0914 | -28.5019 | 28.7734 | 0.040997 | 0.11634 | 0.046537 | 0.14294 | 28.5019 | -0.11634 | 91.1677
0.9824 | -2.0654 | 3.2211 | -27.8371 | 28.1341 | 0.043213 | 0.098615 | 0.053186 | 0.13407 | 27.8371 | -0.098615 | 91.3351
0.9842 | -2.1042 | 3.3784 | -27.2078 | 27.5368 | 0.044321 | 0.07867 | 0.058726 | 0.12521 | 27.2078 | -0.07867 | 91.5024
0.986 | -2.1307 | 3.5612 | -26.749 | 27.1102 | 0.047645 | 0.059834 | 0.06205 | 0.11967 | 26.749 | -0.059834 | 91.6698
0.9878 | -2.154 | 3.7551 | -26.4133 | 26.8078 | 0.049861 | 0.038781 | 0.065374 | 0.11634 | 26.4133 | -0.038781 | 91.8371
0.9896 | -2.164 | 3.9911 | -26.1939 | 26.6283 | 0.053186 | 0.015512 | 0.068698 | 0.11967 | 26.1939 | -0.015512 | 92.0045
0.9914 | -2.1762 | 4.2515 | -26.0488 | 26.5274 | 0.055402 | -0.0088643 | 0.072022 | 0.1241 | 26.0488 | 0.0088643 | 92.1718
0.9932 | -2.1839 | 4.5307 | -26.0288 | 26.5562 | 0.058726 | -0.034349 | 0.074238 | 0.13186 | 26.0288 | 0.034349 | 92.3392
0.9949 | -2.1751 | 4.8078 | -26.1463 | 26.718 | 0.059834 | -0.059834 | 0.076454 | 0.14072 | 26.1463 | 0.059834 | 92.4972
0.9967 | -2.1485 | 5.0604 | -26.4332 | 27.0404 | 0.059834 | -0.086427 | 0.077562 | 0.15291 | 26.4332 | 0.086427 | 92.6646
0.9985 | -2.0643 | 5.3175 | -27.0947 | 27.7307 | 0.057618 | -0.11745 | 0.077562 | 0.16953 | 27.0947 | 0.11745 | 92.8319
1.0003 | -1.9025 | 5.5745 | -28.2515 | 28.9053 | 0.053186 | -0.14958 | 0.075346 | 0.1928 | 28.2515 | 0.14958 | 92.9993
1.0021 | -1.6909 | 5.846 | -29.8105 | 30.4787 | 0.046537 | -0.18393 | 0.072022 | 0.22161 | 29.8105 | 0.18393 | 93.1666
1.0039 | -1.4316 | 6.0942 | -31.6078 | 32.2814 | 0.037673 | -0.21828 | 0.06759 | 0.25042 | 31.6078 | 0.21828 | 93.334
1.0057 | -1.1191 | 6.3091 | -33.6111 | 34.2792 | 0.026593 | -0.25485 | 0.064266 | 0.28033 | 33.6111 | 0.25485 | 93.5013
1.0075 | -0.76343 | 6.5075 | -35.7463 | 36.4033 | 0.014404 | -0.2903 | 0.058726 | 0.31136 | 35.7463 | 0.2903 | 93.6687
1.0093 | -0.39003 | 6.7136 | -37.9479 | 38.595 | 0.0033241 | -0.32687 | 0.054294 | 0.34238 | 37.9479 | 0.32687 | 93.836
1.0111 | -0.017729 | 6.9274 | -40.154 | 40.7934 | -0.0088643 | -0.36233 | 0.048753 | 0.37341 | 40.154 | 0.36233 | 94.0033
1.0129 | 0.3313 | 7.128 | -42.2648 | 42.9019 | -0.021053 | -0.39557 | 0.044321 | 0.40443 | 42.2648 | 0.39557 | 94.1707
1.0147 | 0.65596 | 7.3097 | -44.2737 | 44.9141 | -0.033241 | -0.42659 | 0.040997 | 0.43435 | 44.2737 | 0.42659 | 94.338
1.0165 | 0.97175 | 7.4648 | -46.2493 | 46.8931 | -0.046537 | -0.4554 | 0.037673 | 0.46427 | 46.2493 | 0.4554 | 94.5054
1.0182 | 1.272 | 7.6122 | -48.1773 | 48.8255 | -0.058726 | -0.48421 | 0.034349 | 0.49307 | 48.1773 | 0.48421 | 94.6634
1.02 | 1.5579 | 7.7407 | -50.1374 | 50.7889 | -0.070914 | -0.5108 | 0.031025 | 0.52078 | 50.1374 | 0.5108 | 94.8308
1.0218 | 1.8393 | 7.8526 | -52.123 | 52.7756 | -0.084211 | -0.5374 | 0.027701 | 0.54848 | 52.123 | 0.5374 | 94.9981
1.0236 | 2.1252 | 7.9313 | -54.1485 | 54.7978 | -0.097507 | -0.56288 | 0.024377 | 0.57618 | 54.1485 | 0.56288 | 95.1655
1.0254 | 2.4078 | 7.949 | -56.215 | 56.8543 | -0.11191 | -0.58837 | 0.021053 | 0.60277 | 56.215 | 0.58837 | 95.3328
1.0272 | 2.6892 | 7.9136 | -58.3091 | 58.9352 | -0.12632 | -0.61163 | 0.01662 | 0.62825 | 58.3091 | 0.61163 | 95.5002
1.029 | 2.9762 | 7.836 | -60.4255 | 61.0327 | -0.14072 | -0.6338 | 0.013296 | 0.65374 | 60.4255 | 0.6338 | 95.6675
1.0308 | 3.2676 | 7.7341 | -62.5773 | 63.1645 | -0.15623 | -0.65485 | 0.0088643 | 0.67812 | 62.5773 | 0.65485 | 95.8349
1.0326 | 3.5579 | 7.5989 | -64.6637 | 65.2288 | -0.17064 | -0.6759 | 0.0055402 | 0.70139 | 64.6637 | 0.6759 | 96.0022
1.0344 | 3.8338 | 7.4116 | -66.615 | 67.159 | -0.18393 | -0.69474 | 0.0022161 | 0.72355 | 66.615 | 0.69474 | 96.1696
1.0362 | 4.0997 | 7.1845 | -68.4886 | 69.0105 | -0.19612 | -0.71247 | -0.0022161 | 0.7446 | 68.4886 | 0.71247 | 96.3369
1.038 | 4.359 | 6.933 | -70.2227 | 70.7235 | -0.20831 | -0.72909 | -0.0044321 | 0.76343 | 70.2227 | 0.72909 | 96.5043
1.0398 | 4.6127 | 6.6848 | -71.7562 | 72.2393 | -0.21607 | -0.74571 | -0.0077562 | 0.78116 | 71.7562 | 0.74571 | 96.6716
1.0416 | 4.831 | 6.4765 | -73.0748 | 73.5446 | -0.21828 | -0.76011 | -0.0088643 | 0.79668 | 73.0748 | 0.76011 | 96.839
1.0433 | 5.0105 | 6.2992 | -74.2349 | 74.6947 | -0.21828 | -0.77341 | -0.012188 | 0.80886 | 74.2349 | 0.77341 | 96.997
1.0451 | 5.1679 | 6.1374 | -75.2499 | 75.7008 | -0.21607 | -0.78449 | -0.014404 | 0.81884 | 75.2499 | 0.78449 | 97.1644
1.0469 | 5.3252 | 5.9812 | -76.1673 | 76.6105 | -0.21385 | -0.79557 | -0.01662 | 0.82881 | 76.1673 | 0.79557 | 97.3317
1.0487 | 5.4537 | 5.8737 | -77.0028 | 77.4404 | -0.20831 | -0.80554 | -0.018837 | 0.83657 | 77.0028 | 0.80554 | 97.4991
1.0505 | 5.5679 | 5.7817 | -77.7551 | 78.1884 | -0.20055 | -0.8133 | -0.022161 | 0.84321 | 77.7551 | 0.8133 | 97.6664
1.0523 | 5.672 | 5.6753 | -78.4033 | 78.831 | -0.19391 | -0.82105 | -0.023269 | 0.84875 | 78.4033 | 0.82105 | 97.8338
1.0541 | 5.7529 | 5.5668 | -78.9817 | 79.4039 | -0.18504 | -0.82659 | -0.025485 | 0.85208 | 78.9817 | 0.82659 | 98.0011
1.0559 | 5.8105 | 5.4526 | -79.4726 | 79.8881 | -0.17729 | -0.83102 | -0.027701 | 0.85429 | 79.4726 | 0.83102 | 98.1685
1.0577 | 5.8294 | 5.3452 | -79.8194 | 80.2271 | -0.16731 | -0.83324 | -0.028809 | 0.85429 | 79.8194 | 0.83324 | 98.3358
1.0595 | 5.7928 | 5.2343 | -79.9601 | 80.3568 | -0.15512 | -0.83435 | -0.028809 | 0.85319 | 79.9601 | 0.83435 | 98.5032
1.0613 | 5.7385 | 5.1091 | -79.9501 | 80.3346 | -0.14404 | -0.83546 | -0.027701 | 0.85208 | 79.9501 | 0.83546 | 98.6705
1.0631 | 5.6765 | 4.9673 | -79.7994 | 80.1717 | -0.13296 | -0.83546 | -0.026593 | 0.84986 | 79.7994 | 0.83546 | 98.8379
1.0649 | 5.5767 | 4.8078 | -79.4825 | 79.8393 | -0.12078 | -0.83324 | -0.023269 | 0.84543 | 79.4825 | 0.83324 | 99.0052
1.0667 | 5.4582 | 4.6227 | -79.1479 | 79.4881 | -0.10859 | -0.82992 | -0.019945 | 0.841 | 79.1479 | 0.82992 | 99.1726
1.0684 | 5.3341 | 4.4078 | -78.8875 | 79.21 | -0.096399 | -0.8277 | -0.015512 | 0.83767 | 78.8875 | 0.8277 | 99.3306
1.0702 | 5.2022 | 4.1828 | -78.7402 | 79.0438 | -0.085319 | -0.82548 | -0.01108 | 0.83435 | 78.7402 | 0.82548 | 99.498
1.072 | 5.0781 | 3.9645 | -78.7125 | 78.9983 | -0.07313 | -0.82327 | -0.0077562 | 0.83102 | 78.7125 | 0.82327 | 99.6653
1.0738 | 4.9762 | 3.7518 | -78.7812 | 79.0493 | -0.06205 | -0.82105 | -0.0044321 | 0.82881 | 78.7812 | 0.82105 | 99.8327
1.0756 | 4.8964 | 3.5745 | -78.8465 | 79.1003 | -0.053186 | -0.82105 | -0.001108 | 0.82659 | 78.8465 | 0.82105 | 100
| Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My
Minimum | -3.5058 | -31.7241 | -257.5313 | 26.5274 | -0.21828 | -4.0731 | -0.78781 | 0.11634 | 26.0288 | -0.5385
Maximum | 12.0244 | 12.4643 | -26.0288 | 257.7307 | 1.9102 | 0.5385 | 0.69806 | 4.3557 | 257.5313 | 4.0731

## Diagr_BW
Loads in %BW / %BWm
Contact forces and moments, measured telemetrically in knee implants of 8 subjects.
Averages of one individual subject.
Fx  lateral, Fy  anterior, -Fz  distal, Fres = Resultant Force
Mx around x, My around y, Mz around z, Mres = Resultant Moment
z-axis = axis of implant stem!
From Bergmann et al, PlosOne (2014) "Standardized Loads Acting in Knee Implants"
Subject K2L
Subject K2L

## Dat_AVER75
Data File: | k2l_2242_0_21_i1.akf
Curve Alignment Program: | 1.5.7
------------------------MEASURING INFOS------------------------
Diagram Title #1 | Forces and Moments at Knee Joint
Diagram Title #2 | Measured Relative to Implant
Comment #1 | Knee Joint; Gaitanalysis; Level Walking;
Comment #2 | K2L, 45,5 Months PO *
BodyWeight [N]: | 902.5
Patient(en): | k2l | [7 empty] | AVER75 loads in N (forces) and Nm (moments)
Implant Type: | Knee Joint
Varus Angle Leg [°] | 5 | [7 empty] | Subject | k2l
Axes Directions X, Y, Z: | Lateral, Anterior, Superior
Axes Origin: | Tibial Plateau at Stem Axis | [7 empty] | The individual peak value of Fres was adapted to the average peak of Fres from all subjects!
Transformation X Y Z (°): | 11 | -0.5 | 0.5 | [5 empty] | The AVER75 forces Fres of all subjects therefore have the same peak values.
Displacement X, Y, Z (mm): | 0 | 0 | 0 | [5 empty] | All 6 load components were then adapted by the same Multiplication Factor.
Activity Code: | KneeActivities | 1P_1_15_2
[9 empty] | Multiplication Factor = |  | Max(Fres, AVER75 of 8 Subjects) / |  |  | Max(Fres, Individual Subject)
[9 empty] | Max(Fres, AVER75 of 8 Subjects) = |  |  | 1960.21 | N
[9 empty] | Max(Fres, Individual Subject) = |  |  | 2326.02 | N
---------------------------DATA--------------------------------
Number of Data Sets = | 601
Number of analog channels = | 0
Time-Offset(sec) = | 0
Time | Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My | Time
[s] | [N] | [N] | [N] | [N] | [Nm] | [Nm] | [Nm] | [Nm] | [N] | (Nm] | [%Cycle]
0 | 26.9927 | 8.933 | -609.1936 | 610.3145 | 0.26967 | -5.7811 | 0.2781 | 5.8654 | 609.1936 | 5.7811 | 0
0.0018 | 25.7117 | 6.5227 | -611.2668 | 612.3455 | 0.36237 | -5.7811 | 0.33709 | 5.8823 | 611.2668 | 5.7811 | 0.16735
0.0036 | 24.1358 | 3.8007 | -613.7697 | 614.781 | 0.47193 | -5.7811 | 0.39608 | 5.8823 | 613.7697 | 5.7811 | 0.3347
0.0054 | 22.3745 | 1.0366 | -616.6771 | 617.6294 | 0.58991 | -5.7643 | 0.4635 | 5.8907 | 616.6771 | 5.7643 | 0.50205
0.0072 | 20.6132 | -2.149 | -620.7306 | 621.6661 | 0.71632 | -5.7643 | 0.53935 | 5.9075 | 620.7306 | 5.7643 | 0.66939
0.009 | 19.0204 | -5.444 | -625.3067 | 626.2168 | 0.8343 | -5.7643 | 0.60677 | 5.9244 | 625.3067 | 5.7643 | 0.83674
0.0108 | 17.3181 | -8.8655 | -630.6749 | 631.585 | 0.96071 | -5.7811 | 0.69104 | 5.975 | 630.6749 | 5.7811 | 1.0041
0.0125 | 15.4725 | -12.6578 | -636.9111 | 637.8381 | 1.104 | -5.7896 | 0.78374 | 6.0171 | 636.9111 | 5.7896 | 1.1621
0.0143 | 13.332 | -16.4248 | -643.4085 | 644.3777 | 1.2557 | -5.8148 | 0.87644 | 6.0845 | 643.4085 | 5.8148 | 1.3295
0.0161 | 11.1072 | -19.9475 | -649.6953 | 650.7066 | 1.4326 | -5.857 | 0.986 | 6.1856 | 649.6953 | 5.857 | 1.4968
0.0179 | 8.9835 | -23.7903 | -657.3726 | 658.4681 | 1.6096 | -5.916 | 1.0956 | 6.3036 | 657.3726 | 5.916 | 1.6642
0.0197 | 7.1042 | -28.063 | -666.7775 | 667.9742 | 1.7697 | -5.9665 | 1.2051 | 6.4132 | 666.7775 | 5.9665 | 1.8315
0.0215 | 5.3851 | -32.5463 | -677.6993 | 679.0392 | 1.9383 | -6.034 | 1.3062 | 6.5396 | 677.6993 | 6.034 | 1.9989
0.0233 | 4.2726 | -36.9285 | -691.124 | 692.6072 | 2.1237 | -6.1351 | 1.3989 | 6.6997 | 691.124 | 6.1351 | 2.1662
0.0251 | 3.7754 | -41.1169 | -707.1443 | 708.8129 | 2.3175 | -6.2699 | 1.4748 | 6.8935 | 707.1443 | 6.2699 | 2.3336
0.0269 | 3.9271 | -45.4485 | -726.4934 | 728.3727 | 2.4861 | -6.4216 | 1.5253 | 7.1042 | 726.4934 | 6.4216 | 2.5009
0.0287 | 5.2755 | -50.2184 | -752.7024 | 754.8176 | 2.6125 | -6.5986 | 1.5338 | 7.3065 | 752.7024 | 6.5986 | 2.6683
0.0305 | 7.2896 | -55.0135 | -782.9733 | 785.3414 | 2.7304 | -6.7924 | 1.5169 | 7.5256 | 782.9733 | 6.7924 | 2.8356
0.0323 | 9.2279 | -59.4631 | -814.3145 | 816.9101 | 2.8484 | -6.9862 | 1.4916 | 7.7278 | 814.3145 | 6.9862 | 3.003
0.0341 | 11.5707 | -64.7723 | -846.6248 | 849.5322 | 2.9411 | -7.1632 | 1.4579 | 7.9217 | 846.6248 | 7.1632 | 3.1703
0.0359 | 14.2169 | -70.7979 | -879.8621 | 883.1909 | 3.0423 | -7.3318 | 1.4411 | 8.1071 | 879.8621 | 7.3318 | 3.3377
0.0376 | 17.4024 | -77.6071 | -912.3747 | 916.2007 | 3.1434 | -7.5172 | 1.4326 | 8.3009 | 912.3747 | 7.5172 | 3.4957
0.0394 | 20.7312 | -84.6439 | -943.1681 | 947.5587 | 3.2614 | -7.6941 | 1.4411 | 8.5032 | 943.1681 | 7.6941 | 3.6631
0.0412 | 23.8324 | -91.4532 | -970.3715 | 975.3183 | 3.3962 | -7.8543 | 1.4664 | 8.7138 | 970.3715 | 7.8543 | 3.8304
0.043 | 26.4786 | -98.0181 | -993.4623 | 998.9569 | 3.5732 | -8.0059 | 1.5169 | 8.933 | 993.4623 | 8.0059 | 3.9978
0.0448 | 28.2821 | -104.0099 | -1011.4799 | 1017.5307 | 3.7839 | -8.1324 | 1.5843 | 9.1521 | 1011.4799 | 8.1324 | 4.1651
0.0466 | 29.1922 | -109.2685 | -1024.6265 | 1031.2082 | 4.0451 | -8.2335 | 1.6855 | 9.3796 | 1024.6265 | 8.2335 | 4.3325
0.0484 | 29.3439 | -113.8951 | -1034.5539 | 1041.5907 | 4.3232 | -8.3262 | 1.7782 | 9.6156 | 1034.5539 | 8.3262 | 4.4998
0.0502 | 29.2344 | -118.2352 | -1041.8266 | 1049.3017 | 4.6182 | -8.4189 | 1.8961 | 9.86 | 1041.8266 | 8.4189 | 4.6672
0.052 | 28.931 | -122.4489 | -1047.869 | 1055.7654 | 4.9047 | -8.52 | 2.0141 | 10.1043 | 1047.869 | 8.52 | 4.8345
0.0538 | 28.2821 | -126.5361 | -1052.8664 | 1061.1926 | 5.2081 | -8.6211 | 2.1321 | 10.374 | 1052.8664 | 8.6211 | 5.0019
0.0556 | 27.3045 | -130.1093 | -1056.9031 | 1065.6001 | 5.5115 | -8.756 | 2.2585 | 10.6774 | 1056.9031 | 8.756 | 5.1692
0.0574 | 26.0657 | -133.4128 | -1059.524 | 1068.5834 | 5.8233 | -8.9077 | 2.3934 | 10.9976 | 1059.524 | 8.9077 | 5.3366
0.0592 | 24.8269 | -136.809 | -1060.628 | 1070.0919 | 6.1435 | -9.0762 | 2.545 | 11.3432 | 1060.628 | 9.0762 | 5.5039
0.061 | 23.6808 | -140.2811 | -1061.2347 | 1071.1284 | 6.4553 | -9.2448 | 2.7052 | 11.6887 | 1061.2347 | 9.2448 | 5.6713
0.0627 | 22.3408 | -143.652 | -1061.3106 | 1071.6256 | 6.7756 | -9.4133 | 2.8653 | 12.0426 | 1061.3106 | 9.4133 | 5.8293
0.0645 | 20.6048 | -146.6942 | -1060.1392 | 1070.8419 | 7.0874 | -9.5987 | 3.0254 | 12.3966 | 1060.1392 | 9.5987 | 5.9967
0.0663 | 18.835 | -149.7871 | -1059.3807 | 1070.4795 | 7.3823 | -9.801 | 3.1855 | 12.7674 | 1059.3807 | 9.801 | 6.164
0.0681 | 17.1664 | -153.1749 | -1060.1392 | 1071.6762 | 7.6773 | -10.0454 | 3.3372 | 13.155 | 1060.1392 | 10.0454 | 6.3313
0.0699 | 15.5821 | -156.8407 | -1062.4314 | 1074.4403 | 7.9722 | -10.2982 | 3.4889 | 13.568 | 1062.4314 | 10.2982 | 6.4987
0.0717 | 14.0736 | -160.633 | -1065.7602 | 1078.2832 | 8.2672 | -10.5679 | 3.6406 | 13.9893 | 1065.7602 | 10.5679 | 6.666
0.0735 | 12.7421 | -164.4337 | -1069.9065 | 1082.9182 | 8.5453 | -10.8797 | 3.7754 | 14.4191 | 1069.9065 | 10.8797 | 6.8334
0.0753 | 11.6803 | -168.3777 | -1074.6173 | 1088.1516 | 8.7981 | -11.2252 | 3.9187 | 14.8742 | 1074.6173 | 11.2252 | 7.0007
0.0771 | 10.9049 | -172.4903 | -1080.0529 | 1094.135 | 9.0509 | -11.6044 | 4.0535 | 15.3461 | 1080.0529 | 11.6044 | 7.1681
0.0789 | 10.3656 | -176.6786 | -1086.5335 | 1101.1886 | 9.3122 | -11.9752 | 4.1884 | 15.8181 | 1086.5335 | 11.9752 | 7.3354
0.0807 | 10.0285 | -180.8754 | -1093.8063 | 1109.026 | 9.5734 | -12.3629 | 4.3232 | 16.3069 | 1093.8063 | 12.3629 | 7.5028
0.0825 | 9.8431 | -184.9964 | -1102.419 | 1118.1865 | 9.8262 | -12.7758 | 4.4412 | 16.8041 | 1102.419 | 12.7758 | 7.6701
0.0843 | 9.9358 | -189.0499 | -1112.3464 | 1128.628 | 10.0706 | -13.214 | 4.5592 | 17.3181 | 1112.3464 | 13.214 | 7.8375
0.086 | 10.3403 | -193.095 | -1123.1165 | 1139.9206 | 10.315 | -13.6522 | 4.6687 | 17.8238 | 1123.1165 | 13.6522 | 7.9955
0.0878 | 10.9724 | -197.1149 | -1134.4007 | 1151.7272 | 10.551 | -14.0989 | 4.7699 | 18.3378 | 1134.4007 | 14.0989 | 8.1629
0.0896 | 11.9246 | -201.1263 | -1146.0304 | 1163.8794 | 10.787 | -14.5624 | 4.871 | 18.8603 | 1146.0304 | 14.5624 | 8.3302
0.0914 | 13.2646 | -205.1292 | -1158.3174 | 1176.6805 | 11.0061 | -15.0596 | 4.9721 | 19.3912 | 1158.3174 | 15.0596 | 8.4976
0.0932 | 14.9922 | -209.0985 | -1172.0708 | 1190.9311 | 11.2168 | -15.5568 | 5.0564 | 19.9222 | 1172.0708 | 15.5568 | 8.6649
0.095 | 16.9726 | -212.9161 | -1187.1304 | 1206.4458 | 11.419 | -16.054 | 5.1238 | 20.4447 | 1187.1304 | 16.054 | 8.8323
0.0968 | 19.4502 | -216.5988 | -1203.5468 | 1223.2836 | 11.6044 | -16.5765 | 5.1828 | 20.984 | 1203.5468 | 16.5765 | 8.9996
0.0986 | 22.2734 | -219.9697 | -1220.7469 | 1240.8461 | 11.7898 | -17.0906 | 5.2334 | 21.5149 | 1220.7469 | 17.0906 | 9.167
0.1004 | 25.1808 | -223.0878 | -1238.8235 | 1259.2429 | 11.9836 | -17.6047 | 5.2755 | 22.0459 | 1238.8235 | 17.6047 | 9.3343
0.1022 | 28.2989 | -225.8352 | -1257.8103 | 1278.4825 | 12.169 | -18.1019 | 5.3008 | 22.5599 | 1257.8103 | 18.1019 | 9.5017
0.104 | 31.653 | -228.2538 | -1277.7746 | 1298.649 | 12.346 | -18.5991 | 5.3092 | 23.0571 | 1277.7746 | 18.5991 | 9.669
0.1058 | 35.1756 | -230.3353 | -1299.1378 | 1320.1471 | 12.5146 | -19.0794 | 5.3008 | 23.5375 | 1299.1378 | 19.0794 | 9.8364
0.1076 | 38.8499 | -231.9281 | -1321.3438 | 1342.4121 | 12.6747 | -19.5429 | 5.2755 | 24.001 | 1321.3438 | 19.5429 | 10.0037
0.1094 | 42.4147 | -233.4029 | -1342.6902 | 1363.8174 | 12.8432 | -20.0064 | 5.2586 | 24.4645 | 1342.6902 | 20.0064 | 10.1711
0.1111 | 45.7772 | -234.9198 | -1362.4185 | 1383.6216 | 13.0033 | -20.4447 | 5.2418 | 24.9111 | 1362.4185 | 20.4447 | 10.3291
0.1129 | 49.3503 | -236.2176 | -1381.5991 | 1402.8612 | 13.1635 | -20.8829 | 5.2334 | 25.3494 | 1381.5991 | 20.8829 | 10.4965
0.1147 | 52.9657 | -237.2794 | -1400.1897 | 1421.4771 | 13.2983 | -21.2958 | 5.2081 | 25.7623 | 1400.1897 | 21.2958 | 10.6638
0.1165 | 56.5136 | -238.2486 | -1418.3 | 1439.6211 | 13.4331 | -21.7003 | 5.1912 | 26.1584 | 1418.3 | 21.7003 | 10.8312
0.1183 | 59.9266 | -239.0745 | -1435.3822 | 1456.7286 | 13.5595 | -22.0711 | 5.1659 | 26.5208 | 1435.3822 | 22.0711 | 10.9985
0.1201 | 63.011 | -239.8498 | -1451.8744 | 1473.2293 | 13.686 | -22.4167 | 5.1407 | 26.8747 | 1451.8744 | 22.4167 | 11.1659
0.1219 | 65.9437 | -240.3807 | -1467.8021 | 1489.1485 | 13.8039 | -22.7453 | 5.1154 | 27.2034 | 1467.8021 | 22.7453 | 11.3332
0.1237 | 68.9017 | -240.7768 | -1483.0808 | 1504.4103 | 13.9051 | -23.0656 | 5.0817 | 27.5068 | 1483.0808 | 23.0656 | 11.5006
0.1255 | 71.7839 | -241.0465 | -1497.5589 | 1518.8547 | 13.9893 | -23.3605 | 5.048 | 27.7933 | 1497.5589 | 23.3605 | 11.6679
0.1273 | 74.6323 | -241.2656 | -1511.9865 | 1533.2317 | 14.082 | -23.6555 | 5.0143 | 28.0714 | 1511.9865 | 23.6555 | 11.8353
0.1291 | 77.3712 | -241.2824 | -1527.2905 | 1548.443 | 14.1747 | -23.942 | 4.9637 | 28.3495 | 1527.2905 | 23.942 | 12.0026
0.1309 | 80.0511 | -241.1054 | -1543.1675 | 1564.1853 | 14.2843 | -24.2201 | 4.9216 | 28.6276 | 1543.1675 | 24.2201 | 12.17
0.1327 | 82.5118 | -240.5745 | -1557.6878 | 1578.537 | 14.3686 | -24.4561 | 4.871 | 28.8635 | 1557.6878 | 24.4561 | 12.3373
0.1345 | 84.5428 | -240.17 | -1569.5366 | 1590.2594 | 14.4276 | -24.6499 | 4.8289 | 29.049 | 1569.5366 | 24.6499 | 12.5046
0.1362 | 86.0934 | -240.0015 | -1580.0202 | 1600.6587 | 14.4697 | -24.7847 | 4.7867 | 29.1754 | 1580.0202 | 24.7847 | 12.6627
0.138 | 87.4755 | -239.9088 | -1589.4166 | 1609.9877 | 14.495 | -24.9027 | 4.7446 | 29.2849 | 1589.4166 | 24.9027 | 12.83
0.1398 | 88.6975 | -239.8582 | -1597.6585 | 1618.179 | 14.5118 | -25.0038 | 4.7109 | 29.3692 | 1597.6585 | 25.0038 | 12.9974
0.1416 | 89.734 | -239.7992 | -1604.6532 | 1625.1316 | 14.5203 | -25.0965 | 4.6856 | 29.4535 | 1604.6532 | 25.0965 | 13.1647
0.1434 | 90.5093 | -239.6391 | -1610.6366 | 1631.056 | 14.5287 | -25.1724 | 4.6603 | 29.5125 | 1610.6366 | 25.1724 | 13.3321
0.1452 | 91.0908 | -239.3357 | -1616.1565 | 1636.4916 | 14.5287 | -25.2398 | 4.635 | 29.5714 | 1616.1565 | 25.2398 | 13.4994
0.147 | 91.3858 | -238.9733 | -1621.2382 | 1641.489 | 14.5287 | -25.2904 | 4.6013 | 29.6052 | 1621.2382 | 25.2904 | 13.6668
0.1488 | 91.4532 | -238.7037 | -1626.0165 | 1646.1746 | 14.5287 | -25.3241 | 4.576 | 29.6304 | 1626.0165 | 25.3241 | 13.8341
0.1506 | 91.3352 | -238.5941 | -1630.5251 | 1650.5989 | 14.5203 | -25.3409 | 4.5423 | 29.6389 | 1630.5251 | 25.3409 | 14.0015
0.1524 | 91.1077 | -238.2739 | -1634.1488 | 1654.13 | 14.5118 | -25.3578 | 4.5086 | 29.6389 | 1634.1488 | 25.3578 | 14.1688
0.1542 | 90.6358 | -237.7682 | -1637.1658 | 1657.0121 | 14.5034 | -25.3662 | 4.4749 | 29.6389 | 1637.1658 | 25.3662 | 14.3362
0.156 | 90.0374 | -236.9929 | -1639.7277 | 1659.4055 | 14.5034 | -25.3662 | 4.4412 | 29.6389 | 1639.7277 | 25.3662 | 14.5035
0.1578 | 89.5149 | -236.0153 | -1641.2699 | 1660.7623 | 14.5034 | -25.3831 | 4.4159 | 29.6473 | 1641.2699 | 25.3831 | 14.6709
0.1595 | 88.6806 | -234.8102 | -1641.9441 | 1661.2005 | 14.495 | -25.3915 | 4.3906 | 29.6473 | 1641.9441 | 25.3915 | 14.8289
0.1613 | 87.5682 | -233.5377 | -1641.9272 | 1660.9561 | 14.4697 | -25.3999 | 4.3653 | 29.6389 | 1641.9272 | 25.3999 | 14.9963
0.1631 | 86.321 | -232.1725 | -1641.4553 | 1660.2313 | 14.4276 | -25.3999 | 4.3316 | 29.6136 | 1641.4553 | 25.3999 | 15.1636
0.1649 | 84.8883 | -230.5629 | -1640.0901 | 1658.5964 | 14.3854 | -25.3915 | 4.2979 | 29.5883 | 1640.0901 | 25.3915 | 15.331
0.1667 | 83.3293 | -228.987 | -1638.2192 | 1656.4559 | 14.3349 | -25.3831 | 4.2642 | 29.5462 | 1638.2192 | 25.3831 | 15.4983
0.1685 | 81.6775 | -227.3942 | -1635.6573 | 1653.6243 | 14.2843 | -25.3662 | 4.2305 | 29.504 | 1635.6573 | 25.3662 | 15.6657
0.1703 | 79.9921 | -225.7087 | -1632.5223 | 1650.2197 | 14.2169 | -25.3494 | 4.1968 | 29.445 | 1632.5223 | 25.3494 | 15.833
0.1721 | 78.3993 | -223.9221 | -1629.4127 | 1646.8319 | 14.141 | -25.3409 | 4.1462 | 29.3945 | 1629.4127 | 25.3409 | 16.0004
0.1739 | 76.7981 | -222.0429 | -1626.2271 | 1643.3599 | 14.0568 | -25.3325 | 4.1041 | 29.3439 | 1626.2271 | 25.3325 | 16.1677
0.1757 | 75.1379 | -220.0793 | -1622.6624 | 1639.5086 | 13.9725 | -25.3241 | 4.0535 | 29.2849 | 1622.6624 | 25.3241 | 16.3351
0.1775 | 73.6294 | -218.1747 | -1617.6397 | 1634.2415 | 13.8798 | -25.3072 | 4.0198 | 29.2259 | 1617.6397 | 25.3072 | 16.5024
0.1793 | 71.9777 | -216.2027 | -1611.0833 | 1627.4322 | 13.7955 | -25.2819 | 3.9861 | 29.1585 | 1611.0833 | 25.2819 | 16.6698
0.1811 | 70.4439 | -214.2476 | -1603.3048 | 1619.4179 | 13.7028 | -25.2567 | 3.9608 | 29.0827 | 1603.3048 | 25.2567 | 16.8371
0.1829 | 68.9017 | -212.1492 | -1593.7061 | 1609.5916 | 13.5933 | -25.2314 | 3.9356 | 29.0068 | 1593.7061 | 25.2314 | 17.0045
0.1846 | 67.3342 | -209.8654 | -1582.8686 | 1598.5181 | 13.4837 | -25.2061 | 3.9187 | 28.9225 | 1582.8686 | 25.2061 | 17.1625
0.1864 | 65.8173 | -207.3541 | -1572.9497 | 1588.3211 | 13.3741 | -25.1892 | 3.885 | 28.8467 | 1572.9497 | 25.1892 | 17.3299
0.1882 | 64.3173 | -204.8174 | -1564.4128 | 1579.4977 | 13.273 | -25.1724 | 3.8597 | 28.7708 | 1564.4128 | 25.1724 | 17.4972
0.19 | 62.8088 | -202.3735 | -1557.4518 | 1572.2333 | 13.1719 | -25.1387 | 3.8176 | 28.695 | 1557.4518 | 25.1387 | 17.6646
0.1918 | 61.2834 | -200.0897 | -1551.6033 | 1566.0983 | 13.0876 | -25.105 | 3.7754 | 28.6192 | 1551.6033 | 25.105 | 17.8319
0.1936 | 59.7159 | -198.0587 | -1547.1368 | 1561.3705 | 13.0202 | -25.0797 | 3.7333 | 28.5602 | 1547.1368 | 25.0797 | 17.9993
0.1954 | 58.1569 | -196.2553 | -1544.4232 | 1558.4041 | 12.9612 | -25.0544 | 3.6912 | 28.5096 | 1544.4232 | 25.0544 | 18.1666
0.1972 | 56.6147 | -194.5108 | -1543.631 | 1557.3507 | 12.9022 | -25.0375 | 3.6406 | 28.459 | 1543.631 | 25.0375 | 18.334
0.199 | 55.1146 | -192.7917 | -1543.4288 | 1556.8956 | 12.8348 | -25.0207 | 3.5816 | 28.4085 | 1543.4288 | 25.0207 | 18.5013
0.2008 | 53.5977 | -191.2326 | -1544.4569 | 1557.6625 | 12.759 | -25.0038 | 3.5142 | 28.3495 | 1544.4569 | 25.0038 | 18.6687
0.2026 | 52.1566 | -189.8421 | -1545.9823 | 1558.9603 | 12.6747 | -24.9786 | 3.4468 | 28.2905 | 1545.9823 | 24.9786 | 18.836
0.2044 | 50.901 | -188.2072 | -1546.2688 | 1559.0109 | 12.5904 | -24.9533 | 3.3794 | 28.2146 | 1546.2688 | 24.9533 | 19.0033
0.2062 | 49.8307 | -186.6229 | -1545.3249 | 1557.8563 | 12.5146 | -24.9196 | 3.3288 | 28.1557 | 1545.3249 | 24.9196 | 19.1707
0.208 | 48.8531 | -185.1649 | -1543.7322 | 1556.0866 | 12.4471 | -24.8943 | 3.2782 | 28.0882 | 1543.7322 | 24.8943 | 19.338
0.2097 | 47.6817 | -183.6396 | -1541.0438 | 1553.2045 | 12.3797 | -24.8606 | 3.2445 | 28.0292 | 1541.0438 | 24.8606 | 19.4961
0.2115 | 46.4851 | -181.9204 | -1537.3274 | 1549.2858 | 12.3039 | -24.8353 | 3.2024 | 27.9703 | 1537.3274 | 24.8353 | 19.6634
0.2133 | 45.162 | -180.117 | -1533.0716 | 1544.8109 | 12.228 | -24.7932 | 3.1602 | 27.8944 | 1533.0716 | 24.7932 | 19.8308
0.2151 | 43.7799 | -178.263 | -1528.7315 | 1540.2601 | 12.1606 | -24.7594 | 3.1265 | 27.827 | 1528.7315 | 24.7594 | 19.9981
0.2169 | 42.3725 | -176.4005 | -1524.2566 | 1535.5661 | 12.0848 | -24.7173 | 3.0928 | 27.7596 | 1524.2566 | 24.7173 | 20.1655
0.2187 | 40.8893 | -174.6814 | -1520.1947 | 1531.3019 | 12.0089 | -24.692 | 3.0507 | 27.6922 | 1520.1947 | 24.692 | 20.3328
0.2205 | 39.5831 | -173.2066 | -1517.144 | 1528.0826 | 11.9331 | -24.6667 | 3.0086 | 27.6332 | 1517.144 | 24.6667 | 20.5002
0.2223 | 38.3696 | -171.7571 | -1515.0034 | 1525.7651 | 11.8657 | -24.6415 | 2.9664 | 27.5826 | 1515.0034 | 24.6415 | 20.6675
0.2241 | 37.3414 | -170.1727 | -1513.2759 | 1523.8437 | 11.7898 | -24.6246 | 2.9159 | 27.5236 | 1513.2759 | 24.6246 | 20.8349
0.2259 | 36.4481 | -168.7401 | -1510.8909 | 1521.2818 | 11.7224 | -24.5993 | 2.8737 | 27.473 | 1510.8909 | 24.5993 | 21.0022
0.2277 | 35.4874 | -167.5013 | -1507.2925 | 1517.5316 | 11.6381 | -24.574 | 2.84 | 27.4056 | 1507.2925 | 24.574 | 21.1696
0.2295 | 34.4003 | -166.0939 | -1502.4383 | 1512.5174 | 11.5454 | -24.5319 | 2.8063 | 27.3298 | 1502.4383 | 24.5319 | 21.3369
0.2313 | 33.2626 | -164.2568 | -1496.5139 | 1506.3823 | 11.4527 | -24.4898 | 2.7726 | 27.2455 | 1496.5139 | 24.4898 | 21.5043
0.2331 | 32.1418 | -162.4196 | -1490.0502 | 1499.7163 | 11.3516 | -24.4476 | 2.7389 | 27.1612 | 1490.0502 | 24.4476 | 21.6716
0.2348 | 30.9282 | -160.515 | -1482.2212 | 1491.6851 | 11.242 | -24.4055 | 2.7136 | 27.0685 | 1482.2212 | 24.4055 | 21.8297
0.2366 | 29.7147 | -158.5852 | -1474.5861 | 1483.8392 | 11.1409 | -24.3549 | 2.6883 | 26.9843 | 1474.5861 | 24.3549 | 21.997
0.2384 | 28.6192 | -156.9756 | -1468.1307 | 1477.2154 | 11.0398 | -24.3128 | 2.663 | 26.9 | 1468.1307 | 24.3128 | 22.1644
0.2402 | 27.4056 | -155.2817 | -1463.0575 | 1471.9399 | 10.9471 | -24.2791 | 2.6293 | 26.8157 | 1463.0575 | 24.2791 | 22.3317
0.242 | 26.15 | -153.453 | -1458.3213 | 1466.993 | 10.8544 | -24.2285 | 2.5872 | 26.7314 | 1458.3213 | 24.2285 | 22.4991
0.2438 | 24.928 | -151.7085 | -1452.6666 | 1461.1529 | 10.7617 | -24.178 | 2.5535 | 26.6387 | 1452.6666 | 24.178 | 22.6664
0.2456 | 23.7313 | -150.0483 | -1446.3377 | 1454.6555 | 10.669 | -24.119 | 2.5282 | 26.546 | 1446.3377 | 24.119 | 22.8338
0.2474 | 22.5009 | -148.405 | -1439.697 | 1447.8546 | 10.5763 | -24.06 | 2.4945 | 26.4449 | 1439.697 | 24.06 | 23.0011
0.2492 | 21.3211 | -146.8122 | -1433.5787 | 1441.5763 | 10.4836 | -23.9926 | 2.4692 | 26.3522 | 1433.5787 | 23.9926 | 23.1685
0.251 | 20.2508 | -145.4049 | -1428.3538 | 1436.2081 | 10.3993 | -23.9504 | 2.4439 | 26.2679 | 1428.3538 | 23.9504 | 23.3358
0.2528 | 19.2733 | -144.1492 | -1423.7946 | 1431.5309 | 10.315 | -23.9167 | 2.4186 | 26.2005 | 1423.7946 | 23.9167 | 23.5032
0.2546 | 18.3126 | -142.9694 | -1419.8422 | 1427.4521 | 10.2476 | -23.8914 | 2.3934 | 26.15 | 1419.8422 | 23.8914 | 23.6705
0.2564 | 17.2928 | -141.739 | -1416.0162 | 1423.4997 | 10.1802 | -23.8662 | 2.3681 | 26.091 | 1416.0162 | 23.8662 | 23.8379
0.2581 | 16.3153 | -140.4833 | -1411.7099 | 1419.0753 | 10.1043 | -23.8409 | 2.3428 | 26.032 | 1411.7099 | 23.8409 | 23.9959
0.2599 | 15.4473 | -139.2782 | -1407.749 | 1414.9965 | 10.0285 | -23.8156 | 2.3175 | 25.973 | 1407.749 | 23.8156 | 24.1633
0.2617 | 14.7309 | -138.1742 | -1404.5382 | 1411.6762 | 9.9611 | -23.7819 | 2.2838 | 25.914 | 1404.5382 | 23.7819 | 24.3306
0.2635 | 14.0146 | -137.1292 | -1401.951 | 1408.9794 | 9.8937 | -23.7566 | 2.2501 | 25.855 | 1401.951 | 23.7566 | 24.498
0.2653 | 13.3067 | -136.059 | -1399.8779 | 1406.7883 | 9.8262 | -23.7229 | 2.2164 | 25.8044 | 1399.8779 | 23.7229 | 24.6653
0.2671 | 12.641 | -135.0224 | -1397.9649 | 1404.7658 | 9.7673 | -23.6976 | 2.1827 | 25.7539 | 1397.9649 | 23.6976 | 24.8327
0.2689 | 12.0258 | -134.1881 | -1396.0351 | 1402.7432 | 9.7083 | -23.6723 | 2.149 | 25.7033 | 1396.0351 | 23.6723 | 25
0.2707 | 11.3516 | -133.5139 | -1394.2063 | 1400.8471 | 9.6408 | -23.647 | 2.1153 | 25.6443 | 1394.2063 | 23.647 | 25.1673
0.2725 | 10.7533 | -132.9324 | -1392.6304 | 1399.2206 | 9.5819 | -23.6133 | 2.09 | 25.5938 | 1392.6304 | 23.6133 | 25.3347
0.2743 | 10.315 | -132.4605 | -1391.518 | 1398.0576 | 9.5313 | -23.5965 | 2.0563 | 25.56 | 1391.518 | 23.5965 | 25.502
0.2761 | 9.9274 | -132.0054 | -1390.7933 | 1397.2907 | 9.4723 | -23.5881 | 2.0226 | 25.5179 | 1390.7933 | 23.5881 | 25.6694
0.2779 | 9.5481 | -131.4745 | -1390.591 | 1397.0295 | 9.4217 | -23.5796 | 1.9888 | 25.4926 | 1390.591 | 23.5796 | 25.8367
0.2797 | 9.1858 | -130.8509 | -1390.9028 | 1397.2739 | 9.3712 | -23.5796 | 1.9467 | 25.4673 | 1390.9028 | 23.5796 | 26.0041
0.2815 | 8.8403 | -130.1009 | -1392.0489 | 1398.3441 | 9.329 | -23.5796 | 1.9046 | 25.4421 | 1392.0489 | 23.5796 | 26.1714
0.2832 | 8.5622 | -129.4183 | -1393.9451 | 1400.1644 | 9.2953 | -23.5796 | 1.854 | 25.4252 | 1393.9451 | 23.5796 | 26.3295
0.285 | 8.3262 | -128.8199 | -1396.3384 | 1402.4904 | 9.2616 | -23.5796 | 1.8034 | 25.4168 | 1396.3384 | 23.5796 | 26.4968
0.2868 | 8.0397 | -128.0699 | -1398.9172 | 1404.9849 | 9.2195 | -23.5796 | 1.7529 | 25.3915 | 1398.9172 | 23.5796 | 26.6642
0.2886 | 7.7194 | -127.2272 | -1401.4117 | 1407.3782 | 9.1773 | -23.5628 | 1.6855 | 25.3578 | 1401.4117 | 23.5628 | 26.8315
0.2904 | 7.3908 | -126.3591 | -1403.4427 | 1409.3081 | 9.1352 | -23.5459 | 1.6265 | 25.3241 | 1403.4427 | 23.5459 | 26.9989
0.2922 | 7.1885 | -125.6681 | -1404.1758 | 1409.9654 | 9.1099 | -23.5375 | 1.5843 | 25.2988 | 1404.1758 | 23.5375 | 27.1662
0.294 | 7.1379 | -125.213 | -1403.6281 | 1409.3755 | 9.0678 | -23.5375 | 1.5506 | 25.2819 | 1403.6281 | 23.5375 | 27.3336
0.2958 | 7.1548 | -124.8844 | -1402.187 | 1407.9007 | 9.0341 | -23.5375 | 1.5253 | 25.2651 | 1402.187 | 23.5375 | 27.5009
0.2976 | 7.1716 | -124.7242 | -1400.4173 | 1406.1226 | 8.9919 | -23.5375 | 1.5085 | 25.2567 | 1400.4173 | 23.5375 | 27.6683
0.2994 | 7.1801 | -124.6231 | -1399.288 | 1404.9764 | 8.9498 | -23.5459 | 1.4916 | 25.2482 | 1399.288 | 23.5459 | 27.8356
0.3012 | 7.1464 | -124.4798 | -1399.4313 | 1405.1029 | 8.9161 | -23.5543 | 1.4579 | 25.2398 | 1399.4313 | 23.5543 | 28.003
0.303 | 7.0115 | -124.1933 | -1400.8976 | 1406.5355 | 8.8908 | -23.5628 | 1.4242 | 25.2398 | 1400.8976 | 23.5628 | 28.1703
0.3048 | 6.9525 | -123.8899 | -1402.9876 | 1408.5918 | 8.8655 | -23.5628 | 1.3821 | 25.2314 | 1402.9876 | 23.5628 | 28.3377
0.3066 | 6.9694 | -123.5613 | -1405.8782 | 1411.4318 | 8.8487 | -23.5881 | 1.3399 | 25.2398 | 1405.8782 | 23.5881 | 28.505
0.3083 | 7.0368 | -123.2916 | -1409.662 | 1415.1819 | 8.8318 | -23.6133 | 1.2978 | 25.2482 | 1409.662 | 23.6133 | 28.6631
0.3101 | 7.1379 | -123.1146 | -1413.783 | 1419.2692 | 8.815 | -23.6386 | 1.2472 | 25.2735 | 1413.783 | 23.6386 | 28.8304
0.3119 | 7.3992 | -123.1652 | -1416.5724 | 1422.0502 | 8.7981 | -23.6808 | 1.2135 | 25.2988 | 1416.5724 | 23.6808 | 28.9978
0.3137 | 7.8374 | -123.4601 | -1418.0809 | 1423.5755 | 8.7897 | -23.7313 | 1.1967 | 25.3494 | 1418.0809 | 23.7313 | 29.1651
0.3155 | 8.3093 | -123.7888 | -1419.6737 | 1425.1936 | 8.7813 | -23.7903 | 1.1798 | 25.3999 | 1419.6737 | 23.7903 | 29.3325
0.3173 | 8.7054 | -123.9574 | -1422.3873 | 1427.9072 | 8.7728 | -23.8409 | 1.1545 | 25.4421 | 1422.3873 | 23.8409 | 29.4998
0.3191 | 9.0088 | -123.8562 | -1426.4071 | 1431.9017 | 8.7644 | -23.8914 | 1.1124 | 25.4842 | 1426.4071 | 23.8914 | 29.6672
0.3209 | 9.3712 | -123.6455 | -1431.6657 | 1437.1266 | 8.7813 | -23.942 | 1.0618 | 25.5348 | 1431.6657 | 23.942 | 29.8345
0.3227 | 9.8937 | -123.5781 | -1437.9188 | 1443.3544 | 8.7897 | -23.9926 | 1.0113 | 25.5853 | 1437.9188 | 23.9926 | 30.0019
0.3245 | 10.4836 | -123.654 | -1444.0623 | 1449.4895 | 8.8065 | -24.0431 | 0.96071 | 25.6359 | 1444.0623 | 24.0431 | 30.1692
0.3263 | 11.0988 | -123.7382 | -1449.6328 | 1455.0431 | 8.8234 | -24.0937 | 0.91015 | 25.6865 | 1449.6328 | 24.0937 | 30.3366
0.3281 | 11.7561 | -123.6287 | -1453.7706 | 1459.1725 | 8.8487 | -24.1443 | 0.87644 | 25.7454 | 1453.7706 | 24.1443 | 30.5039
0.3299 | 12.3544 | -123.5613 | -1457.2426 | 1462.6277 | 8.8571 | -24.1948 | 0.84273 | 25.796 | 1457.2426 | 24.1948 | 30.6713
0.3316 | 12.8179 | -123.3422 | -1460.7231 | 1466.0913 | 8.874 | -24.237 | 0.80059 | 25.8381 | 1460.7231 | 24.237 | 30.8293
0.3334 | 13.1129 | -122.9208 | -1465.3413 | 1470.6589 | 8.9077 | -24.2791 | 0.75846 | 25.8887 | 1465.3413 | 24.2791 | 30.9967
0.3352 | 13.2646 | -122.3814 | -1471.0634 | 1476.3137 | 8.9498 | -24.3212 | 0.70789 | 25.9393 | 1471.0634 | 24.3212 | 31.164
0.337 | 13.4584 | -121.859 | -1477.0047 | 1482.2043 | 8.9835 | -24.3634 | 0.65733 | 25.9983 | 1477.0047 | 24.3634 | 31.3313
0.3388 | 13.7281 | -121.4966 | -1482.2381 | 1487.3956 | 9.0172 | -24.4139 | 0.61519 | 26.0488 | 1482.2381 | 24.4139 | 31.4987
0.3406 | 14.1073 | -121.3533 | -1486.3422 | 1491.4744 | 9.0341 | -24.4561 | 0.58148 | 26.091 | 1486.3422 | 24.4561 | 31.666
0.3424 | 14.4528 | -121.1932 | -1489.3844 | 1494.5082 | 9.0509 | -24.4898 | 0.54778 | 26.1331 | 1489.3844 | 24.4898 | 31.8334
0.3442 | 14.7562 | -121.0836 | -1491.4575 | 1496.5645 | 9.0678 | -24.5151 | 0.52249 | 26.1668 | 1491.4575 | 24.5151 | 32.0007
0.346 | 15.0175 | -120.9994 | -1493.1767 | 1498.2837 | 9.0846 | -24.5488 | 0.49721 | 26.2089 | 1493.1767 | 24.5488 | 32.1681
0.3478 | 15.2197 | -120.8055 | -1495.4352 | 1500.5169 | 9.1015 | -24.5909 | 0.47193 | 26.2511 | 1495.4352 | 24.5909 | 32.3354
0.3496 | 15.4978 | -120.5864 | -1498.9157 | 1503.9805 | 9.1268 | -24.6415 | 0.43822 | 26.3016 | 1498.9157 | 24.6415 | 32.5028
0.3514 | 15.9445 | -120.4347 | -1503.7951 | 1508.8347 | 9.1436 | -24.692 | 0.40451 | 26.3606 | 1503.7951 | 24.692 | 32.6701
0.3532 | 16.3827 | -120.2072 | -1509.728 | 1514.7422 | 9.1773 | -24.751 | 0.35395 | 26.4281 | 1509.728 | 24.751 | 32.8375
0.355 | 16.526 | -119.5667 | -1516.3602 | 1521.3155 | 9.2279 | -24.81 | 0.31181 | 26.4955 | 1516.3602 | 24.81 | 33.0048
0.3567 | 16.7114 | -119.0105 | -1522.9841 | 1527.8888 | 9.27 | -24.8606 | 0.26125 | 26.5629 | 1522.9841 | 24.8606 | 33.1629
0.3585 | 16.981 | -118.5807 | -1528.8327 | 1533.6952 | 9.3038 | -24.9111 | 0.21911 | 26.6219 | 1528.8327 | 24.9111 | 33.3302
0.3603 | 17.2844 | -118.2184 | -1533.5688 | 1538.4061 | 9.3375 | -24.9617 | 0.17697 | 26.6809 | 1533.5688 | 24.9617 | 33.4976
0.3621 | 17.7563 | -118.033 | -1537.892 | 1542.7125 | 9.3543 | -25.0123 | 0.14326 | 26.7314 | 1537.892 | 25.0123 | 33.6649
0.3639 | 18.3884 | -117.7717 | -1542.881 | 1547.6677 | 9.3796 | -25.0713 | 0.10113 | 26.7904 | 1542.881 | 25.0713 | 33.8323
0.3657 | 19.0879 | -117.2914 | -1548.6621 | 1553.4067 | 9.4217 | -25.1471 | 0.058991 | 26.8747 | 1548.6621 | 25.1471 | 33.9996
0.3675 | 19.7452 | -116.8868 | -1553.5753 | 1558.2861 | 9.4554 | -25.2145 | 0.033709 | 26.9505 | 1553.5753 | 25.2145 | 34.167
0.3693 | 20.3857 | -116.4992 | -1557.8226 | 1562.4998 | 9.4807 | -25.2735 | 0 | 27.018 | 1557.8226 | 25.2735 | 34.3343
0.3711 | 21.0261 | -116.0778 | -1561.2778 | 1565.9213 | 9.5144 | -25.3325 | -0.033709 | 27.0854 | 1561.2778 | 25.3325 | 34.5017
0.3729 | 21.5486 | -115.7239 | -1563.7555 | 1568.3652 | 9.5313 | -25.3662 | -0.058991 | 27.1275 | 1563.7555 | 25.3662 | 34.669
0.3747 | 22.0627 | -115.4626 | -1565.6179 | 1570.2108 | 9.5566 | -25.4084 | -0.084273 | 27.1697 | 1565.6179 | 25.4084 | 34.8364
0.3765 | 22.5515 | -115.3952 | -1566.7893 | 1571.3822 | 9.5734 | -25.4336 | -0.10113 | 27.2034 | 1566.7893 | 25.4336 | 35.0037
0.3783 | 23.0403 | -115.3446 | -1567.3961 | 1571.9889 | 9.5903 | -25.4673 | -0.10956 | 27.2371 | 1567.3961 | 25.4673 | 35.1711
0.3801 | 23.5628 | -115.0665 | -1567.5056 | 1572.0901 | 9.624 | -25.5095 | -0.10956 | 27.2961 | 1567.5056 | 25.5095 | 35.3384
0.3818 | 24.0684 | -114.6789 | -1567.5983 | 1572.1659 | 9.6493 | -25.56 | -0.11798 | 27.3466 | 1567.5983 | 25.56 | 35.4965
0.3836 | 24.4392 | -114.1732 | -1568.4916 | 1573.0255 | 9.6746 | -25.6022 | -0.13484 | 27.3972 | 1568.4916 | 25.6022 | 35.6638
0.3854 | 24.7679 | -113.7603 | -1570.3035 | 1574.8205 | 9.6998 | -25.6443 | -0.15169 | 27.4478 | 1570.3035 | 25.6443 | 35.8312
0.3872 | 25.0038 | -113.2631 | -1572.621 | 1577.0959 | 9.7251 | -25.6949 | -0.17697 | 27.4983 | 1572.621 | 25.6949 | 35.9985
0.389 | 25.2314 | -112.7659 | -1575.0986 | 1579.5398 | 9.7673 | -25.7454 | -0.20226 | 27.5573 | 1575.0986 | 25.7454 | 36.1659
0.3908 | 25.6106 | -112.2265 | -1577.6268 | 1582.0259 | 9.801 | -25.796 | -0.23596 | 27.6163 | 1577.6268 | 25.796 | 36.3332
0.3926 | 26.1331 | -111.7378 | -1580.3573 | 1584.7226 | 9.8262 | -25.8466 | -0.26125 | 27.6837 | 1580.3573 | 25.8466 | 36.5006
0.3944 | 26.7146 | -111.3248 | -1583.6439 | 1587.984 | 9.8515 | -25.914 | -0.2781 | 27.7511 | 1583.6439 | 25.914 | 36.6679
0.3962 | 27.2287 | -110.8529 | -1587.4025 | 1591.7089 | 9.8768 | -25.9814 | -0.30338 | 27.827 | 1587.4025 | 25.9814 | 36.8353
0.398 | 27.6163 | -110.2461 | -1591.574 | 1595.8298 | 9.9105 | -26.0404 | -0.33709 | 27.8944 | 1591.574 | 26.0404 | 37.0026
0.3998 | 28.0208 | -109.6056 | -1595.4759 | 1599.698 | 9.9527 | -26.1162 | -0.36237 | 27.9787 | 1595.4759 | 26.1162 | 37.17
0.4016 | 28.4085 | -108.8893 | -1599.58 | 1603.7515 | 9.9864 | -26.2005 | -0.38766 | 28.0714 | 1599.58 | 26.2005 | 37.3373
0.4034 | 28.7877 | -108.1814 | -1605.0746 | 1609.1871 | 10.0201 | -26.2848 | -0.42137 | 28.1641 | 1605.0746 | 26.2848 | 37.5046
0.4051 | 29.2259 | -107.364 | -1611.4709 | 1615.5329 | 10.0538 | -26.3775 | -0.45507 | 28.2652 | 1611.4709 | 26.3775 | 37.6627
0.4069 | 29.6304 | -106.5718 | -1618.1706 | 1622.1736 | 10.0875 | -26.4533 | -0.49721 | 28.3495 | 1618.1706 | 26.4533 | 37.83
0.4087 | 30.0687 | -105.7291 | -1625.4518 | 1629.4042 | 10.1212 | -26.546 | -0.53935 | 28.4506 | 1625.4518 | 26.546 | 37.9974
0.4105 | 30.5911 | -104.7262 | -1633.1797 | 1637.0731 | 10.1633 | -26.6472 | -0.58991 | 28.5602 | 1633.1797 | 26.6472 | 38.1647
0.4123 | 31.1389 | -103.8835 | -1640.2502 | 1644.0762 | 10.197 | -26.7399 | -0.63205 | 28.6613 | 1640.2502 | 26.7399 | 38.3321
0.4141 | 31.7541 | -103.2683 | -1646.6212 | 1650.4051 | 10.2308 | -26.8326 | -0.66576 | 28.754 | 1646.6212 | 26.8326 | 38.4994
0.4159 | 32.4704 | -102.729 | -1653.0513 | 1656.8099 | 10.2476 | -26.9253 | -0.70789 | 28.8551 | 1653.0513 | 26.9253 | 38.6668
0.4177 | 33.1868 | -102.1559 | -1659.9532 | 1663.6781 | 10.2729 | -27.018 | -0.7416 | 28.9562 | 1659.9532 | 27.018 | 38.8341
0.4195 | 33.802 | -101.5238 | -1666.5603 | 1670.243 | 10.3066 | -27.1107 | -0.78374 | 29.049 | 1666.5603 | 27.1107 | 39.0015
0.4213 | 34.4003 | -100.9508 | -1672.3582 | 1676.0073 | 10.3487 | -27.1865 | -0.81745 | 29.1332 | 1672.3582 | 27.1865 | 39.1688
0.4231 | 34.9902 | -100.2935 | -1677.2967 | 1680.9035 | 10.3909 | -27.2539 | -0.85116 | 29.2091 | 1677.2967 | 27.2539 | 39.3362
0.4249 | 35.4621 | -99.5013 | -1681.5019 | 1685.0582 | 10.433 | -27.2961 | -0.88487 | 29.2681 | 1681.5019 | 27.2961 | 39.5035
0.4267 | 35.9341 | -98.8355 | -1683.9205 | 1687.4431 | 10.4836 | -27.3466 | -0.91015 | 29.3355 | 1683.9205 | 27.3466 | 39.6709
0.4285 | 36.3807 | -98.0855 | -1685.8504 | 1689.3309 | 10.5426 | -27.4056 | -0.927 | 29.4029 | 1685.8504 | 27.4056 | 39.8382
0.4302 | 36.7684 | -97.0405 | -1688.1005 | 1691.5388 | 10.6016 | -27.473 | -0.94386 | 29.4872 | 1688.1005 | 27.473 | 39.9963
0.432 | 37.2572 | -95.9702 | -1691.0247 | 1694.3957 | 10.6437 | -27.5489 | -0.96071 | 29.5799 | 1691.0247 | 27.5489 | 40.1636
0.4338 | 37.8302 | -94.9084 | -1694.5895 | 1697.9099 | 10.6858 | -27.6332 | -0.97757 | 29.6726 | 1694.5895 | 27.6332 | 40.331
0.4356 | 38.4201 | -93.8466 | -1699.6965 | 1702.9663 | 10.728 | -27.7343 | -1.0029 | 29.7821 | 1699.6965 | 27.7343 | 40.4983
0.4374 | 38.9848 | -92.7594 | -1705.8905 | 1709.1013 | 10.7617 | -27.8354 | -1.0366 | 29.8917 | 1705.8905 | 27.8354 | 40.6657
0.4392 | 39.5578 | -91.6049 | -1713.0369 | 1716.1887 | 10.7954 | -27.945 | -1.0703 | 30.0097 | 1713.0369 | 27.945 | 40.833
0.441 | 40.232 | -90.4672 | -1720.5119 | 1723.6132 | 10.8291 | -28.0545 | -1.104 | 30.1276 | 1720.5119 | 28.0545 | 41.0004
0.4428 | 40.8219 | -89.22 | -1728.5094 | 1731.5433 | 10.8628 | -28.1725 | -1.1377 | 30.2456 | 1728.5094 | 28.1725 | 41.1677
0.4446 | 41.3613 | -87.9559 | -1736.8103 | 1739.7936 | 10.8881 | -28.2905 | -1.1798 | 30.3636 | 1736.8103 | 28.2905 | 41.3351
0.4464 | 41.9006 | -86.7086 | -1745.1955 | 1748.1198 | 10.9218 | -28.4 | -1.2135 | 30.4816 | 1745.1955 | 28.4 | 41.5024
0.4482 | 42.44 | -85.3687 | -1753.4627 | 1756.328 | 10.9471 | -28.5096 | -1.2557 | 30.5996 | 1753.4627 | 28.5096 | 41.6698
0.45 | 42.9709 | -84.0119 | -1761.5445 | 1764.3508 | 10.9808 | -28.6276 | -1.2894 | 30.7176 | 1761.5445 | 28.6276 | 41.8371
0.4518 | 43.5439 | -82.6382 | -1769.9465 | 1772.7023 | 11.0061 | -28.7371 | -1.3231 | 30.8271 | 1769.9465 | 28.7371 | 42.0045
0.4536 | 44.2603 | -81.3489 | -1778.7194 | 1781.4245 | 11.0314 | -28.8383 | -1.3568 | 30.9367 | 1778.7194 | 28.8383 | 42.1718
0.4553 | 45.103 | -80.1775 | -1787.4669 | 1790.13 | 11.0566 | -28.9478 | -1.3905 | 31.0462 | 1787.4669 | 28.9478 | 42.3299
0.4571 | 45.9457 | -79.0398 | -1795.81 | 1798.4393 | 11.0735 | -29.0574 | -1.4242 | 31.1558 | 1795.81 | 29.0574 | 42.4972
0.4589 | 46.8222 | -77.9021 | -1803.6052 | 1806.2008 | 11.0903 | -29.1585 | -1.4495 | 31.2653 | 1803.6052 | 29.1585 | 42.6646
0.4607 | 47.6902 | -76.7897 | -1810.4651 | 1813.0185 | 11.1072 | -29.2681 | -1.4664 | 31.3749 | 1810.4651 | 29.2681 | 42.8319
0.4625 | 48.5329 | -75.6857 | -1816.3979 | 1818.9177 | 11.1241 | -29.3776 | -1.4832 | 31.476 | 1816.3979 | 29.3776 | 42.9993
0.4643 | 49.4093 | -74.5733 | -1822.2633 | 1824.7662 | 11.1409 | -29.4872 | -1.5001 | 31.5856 | 1822.2633 | 29.4872 | 43.1666
0.4661 | 50.2774 | -73.444 | -1828.5922 | 1831.0698 | 11.1578 | -29.5799 | -1.5169 | 31.6783 | 1828.5922 | 29.5799 | 43.334
0.4679 | 51.019 | -72.2726 | -1835.9324 | 1838.3763 | 11.1746 | -29.6726 | -1.5422 | 31.771 | 1835.9324 | 29.6726 | 43.5013
0.4697 | 51.7521 | -71.1012 | -1843.6265 | 1846.0452 | 11.1915 | -29.7568 | -1.5759 | 31.8637 | 1843.6265 | 29.7568 | 43.6687
0.4715 | 52.4516 | -69.9467 | -1851.4555 | 1853.832 | 11.2083 | -29.8411 | -1.6012 | 31.9479 | 1851.4555 | 29.8411 | 43.836
0.4733 | 53.1595 | -68.8343 | -1859.2339 | 1861.5851 | 11.2168 | -29.9254 | -1.6349 | 32.0322 | 1859.2339 | 29.9254 | 44.0033
0.4751 | 53.8084 | -67.6882 | -1866.9449 | 1869.2624 | 11.2336 | -30.0012 | -1.6686 | 32.1081 | 1866.9449 | 30.0012 | 44.1707
0.4769 | 54.4067 | -66.5168 | -1874.7318 | 1877.024 | 11.242 | -30.0771 | -1.7023 | 32.1839 | 1874.7318 | 30.0771 | 44.338
0.4786 | 54.9966 | -65.3285 | -1882.3922 | 1884.6591 | 11.2589 | -30.1529 | -1.7445 | 32.2598 | 1882.3922 | 30.1529 | 44.4961
0.4804 | 55.5781 | -64.0981 | -1889.5301 | 1891.7549 | 11.2673 | -30.2203 | -1.7782 | 32.3356 | 1889.5301 | 30.2203 | 44.6634
0.4822 | 56.168 | -62.9183 | -1896.2046 | 1898.4041 | 11.2757 | -30.2878 | -1.8034 | 32.403 | 1896.2046 | 30.2878 | 44.8308
0.484 | 56.7664 | -61.7891 | -1902.5082 | 1904.674 | 11.2842 | -30.3552 | -1.8372 | 32.462 | 1902.5082 | 30.3552 | 44.9981
0.4858 | 57.3563 | -60.7104 | -1908.4579 | 1910.6068 | 11.2926 | -30.4142 | -1.8624 | 32.5294 | 1908.4579 | 30.4142 | 45.1655
0.4876 | 57.9294 | -59.6822 | -1914.0536 | 1916.1773 | 11.301 | -30.4732 | -1.8877 | 32.5884 | 1914.0536 | 30.4732 | 45.3328
0.4894 | 58.4518 | -58.6794 | -1919.0847 | 1921.1915 | 11.301 | -30.5322 | -1.9046 | 32.639 | 1919.0847 | 30.5322 | 45.5002
0.4912 | 58.9406 | -57.6344 | -1923.3742 | 1925.4473 | 11.2926 | -30.5911 | -1.9299 | 32.6895 | 1923.3742 | 30.5911 | 45.6675
0.493 | 59.3704 | -56.5473 | -1927.512 | 1929.5683 | 11.2926 | -30.6417 | -1.9467 | 32.7401 | 1927.512 | 30.6417 | 45.8349
0.4948 | 59.7075 | -55.4096 | -1931.7004 | 1933.7314 | 11.2842 | -30.6923 | -1.972 | 32.7907 | 1931.7004 | 30.6923 | 46.0022
0.4966 | 59.994 | -54.2635 | -1935.7539 | 1937.7428 | 11.2842 | -30.7344 | -1.9973 | 32.8328 | 1935.7539 | 30.7344 | 46.1696
0.4984 | 60.407 | -53.0752 | -1938.9647 | 1940.9283 | 11.2926 | -30.785 | -2.0226 | 32.875 | 1938.9647 | 30.785 | 46.3369
0.5002 | 61.1823 | -51.8617 | -1941.7036 | 1943.6588 | 11.3179 | -30.8355 | -2.031 | 32.9339 | 1941.7036 | 30.8355 | 46.5043
0.502 | 61.6627 | -50.3363 | -1944.4004 | 1946.3134 | 11.36 | -30.8692 | -2.0563 | 32.9845 | 1944.4004 | 30.8692 | 46.6716
0.5037 | 61.8143 | -48.8363 | -1947.4848 | 1949.3472 | 11.3853 | -30.8945 | -2.0815 | 33.0182 | 1947.4848 | 30.8945 | 46.8297
0.5055 | 61.8143 | -47.4205 | -1950.4765 | 1952.3052 | 11.3937 | -30.9198 | -2.1153 | 33.0435 | 1950.4765 | 30.9198 | 46.997
0.5073 | 61.6205 | -45.9963 | -1952.8698 | 1954.648 | 11.4022 | -30.9367 | -2.149 | 33.0688 | 1952.8698 | 30.9367 | 47.1644
0.5091 | 61.334 | -44.4372 | -1954.8587 | 1956.5947 | 11.4022 | -30.9535 | -2.1827 | 33.0856 | 1954.8587 | 30.9535 | 47.3317
0.5109 | 61.0138 | -42.895 | -1956.4851 | 1958.179 | 11.4106 | -30.9704 | -2.2248 | 33.1025 | 1956.4851 | 30.9704 | 47.4991
0.5127 | 60.6261 | -41.2938 | -1957.6734 | 1959.3167 | 11.4106 | -30.9788 | -2.2669 | 33.1193 | 1957.6734 | 30.9788 | 47.6664
0.5145 | 60.0783 | -39.5662 | -1958.4066 | 1959.9909 | 11.419 | -30.9788 | -2.3091 | 33.1278 | 1958.4066 | 30.9788 | 47.8338
0.5163 | 59.421 | -37.7291 | -1958.6931 | 1960.21 | 11.419 | -30.9788 | -2.3596 | 33.1278 | 1958.6931 | 30.9788 | 48.0011
0.5181 | 58.7805 | -36.0773 | -1958.4487 | 1959.9235 | 11.4106 | -30.9788 | -2.4018 | 33.1193 | 1958.4487 | 30.9788 | 48.1685
0.5199 | 58.14 | -34.4003 | -1957.7577 | 1959.1903 | 11.4106 | -30.9788 | -2.4355 | 33.1278 | 1957.7577 | 30.9788 | 48.3358
0.5217 | 57.3647 | -32.4536 | -1956.5357 | 1957.9093 | 11.4022 | -30.9788 | -2.4692 | 33.1193 | 1956.5357 | 30.9788 | 48.5032
0.5235 | 56.4209 | -30.372 | -1954.8839 | 1956.2239 | 11.3937 | -30.962 | -2.5198 | 33.1109 | 1954.8839 | 30.962 | 48.6705
0.5253 | 55.4686 | -28.3579 | -1952.8108 | 1954.1171 | 11.3769 | -30.9367 | -2.5619 | 33.0856 | 1952.8108 | 30.9367 | 48.8379
0.5271 | 54.5163 | -26.3943 | -1950.3248 | 1951.5889 | 11.3516 | -30.9198 | -2.604 | 33.0604 | 1950.3248 | 30.9198 | 49.0052
0.5288 | 53.6061 | -24.3802 | -1947.5269 | 1948.7573 | 11.3347 | -30.8945 | -2.6462 | 33.0351 | 1947.5269 | 30.8945 | 49.1633
0.5306 | 52.4685 | -21.8689 | -1944.4341 | 1945.6307 | 11.3179 | -30.844 | -2.7052 | 32.9845 | 1944.4341 | 30.844 | 49.3306
0.5324 | 51.2296 | -19.1384 | -1941.0042 | 1942.1587 | 11.3095 | -30.7934 | -2.7557 | 32.9424 | 1941.0042 | 30.7934 | 49.498
0.5342 | 49.974 | -16.4333 | -1936.959 | 1938.0799 | 11.2926 | -30.7513 | -2.8063 | 32.8918 | 1936.959 | 30.7513 | 49.6653
0.536 | 48.6256 | -13.6944 | -1932.3577 | 1933.4196 | 11.2842 | -30.7091 | -2.8484 | 32.8581 | 1932.3577 | 30.7091 | 49.8327
0.5378 | 47.2182 | -10.8881 | -1927.5205 | 1928.5149 | 11.2673 | -30.667 | -2.8906 | 32.8244 | 1927.5205 | 30.667 | 50
0.5396 | 45.6339 | -7.829 | -1922.3714 | 1923.3068 | 11.2505 | -30.6164 | -2.9327 | 32.7738 | 1922.3714 | 30.6164 | 50.1673
0.5414 | 43.9063 | -4.6013 | -1916.902 | 1917.7785 | 11.2336 | -30.5574 | -2.9833 | 32.7064 | 1916.902 | 30.5574 | 50.3347
0.5432 | 42.0017 | -1.2557 | -1911.2136 | 1912.031 | 11.2168 | -30.49 | -3.0254 | 32.6474 | 1911.2136 | 30.49 | 50.502
0.545 | 39.9707 | 2.3512 | -1904.9521 | 1905.719 | 11.1999 | -30.4057 | -3.076 | 32.5716 | 1904.9521 | 30.4057 | 50.6694
0.5468 | 37.9313 | 6.0845 | -1898.244 | 1898.9771 | 11.183 | -30.313 | -3.135 | 32.4789 | 1898.244 | 30.313 | 50.8367
0.5486 | 35.8835 | 9.8852 | -1891.3167 | 1892.0162 | 11.1578 | -30.2203 | -3.1855 | 32.3862 | 1891.3167 | 30.2203 | 51.0041
0.5504 | 33.8441 | 13.7365 | -1884.2041 | 1884.8782 | 11.1325 | -30.1108 | -3.2445 | 32.285 | 1884.2041 | 30.1108 | 51.1714
0.5521 | 31.8637 | 17.6721 | -1876.8049 | 1877.4622 | 11.0903 | -29.9928 | -3.2951 | 32.1671 | 1876.8049 | 29.9928 | 51.3295
0.5539 | 29.9844 | 21.6245 | -1869.1866 | 1869.8523 | 11.0482 | -29.8833 | -3.3541 | 32.0491 | 1869.1866 | 29.8833 | 51.4968
0.5557 | 28.2062 | 25.5938 | -1861.4503 | 1862.1076 | 10.9892 | -29.7653 | -3.4046 | 31.9311 | 1861.4503 | 29.7653 | 51.6642
0.5575 | 26.4533 | 29.7231 | -1853.4528 | 1854.127 | 10.9302 | -29.6389 | -3.4636 | 31.7963 | 1853.4528 | 29.6389 | 51.8315
0.5593 | 24.5656 | 33.9958 | -1845.0423 | 1845.7418 | 10.8544 | -29.4956 | -3.5142 | 31.6446 | 1845.0423 | 29.4956 | 51.9989
0.5611 | 22.4841 | 38.5887 | -1835.2414 | 1835.983 | 10.7785 | -29.3271 | -3.5732 | 31.4676 | 1835.2414 | 29.3271 | 52.1662
0.5629 | 20.4699 | 43.2574 | -1825.6595 | 1826.4601 | 10.6943 | -29.1501 | -3.6237 | 31.2822 | 1825.6595 | 29.1501 | 52.3336
0.5647 | 18.6918 | 47.9093 | -1817.1985 | 1818.0918 | 10.6016 | -28.9562 | -3.6912 | 31.0715 | 1817.1985 | 28.9562 | 52.5009
0.5665 | 17.453 | 52.1229 | -1811.1224 | 1812.1168 | 10.5004 | -28.7456 | -3.7754 | 30.8524 | 1811.1224 | 28.7456 | 52.6683
0.5683 | 16.5007 | 56.1596 | -1806.4368 | 1807.5408 | 10.374 | -28.5265 | -3.8597 | 30.6164 | 1806.4368 | 28.5265 | 52.8356
0.5701 | 15.9529 | 60.0278 | -1804.0097 | 1805.2317 | 10.2223 | -28.2736 | -3.9693 | 30.3468 | 1804.0097 | 28.2736 | 53.003
0.5719 | 16.2647 | 63.7611 | -1806.42 | 1807.7768 | 10.0454 | -27.9871 | -4.1125 | 30.0434 | 1806.42 | 27.9871 | 53.1703
0.5737 | 16.9895 | 67.2078 | -1810.7179 | 1812.2095 | 9.8515 | -27.6753 | -4.2726 | 29.7147 | 1810.7179 | 27.6753 | 53.3377
0.5755 | 17.8238 | 70.5198 | -1815.0495 | 1816.676 | 9.6493 | -27.3382 | -4.4328 | 29.3608 | 1815.0495 | 27.3382 | 53.505
0.5772 | 18.4053 | 73.8317 | -1816.7855 | 1818.5384 | 9.4554 | -26.9674 | -4.5845 | 28.9815 | 1816.7855 | 26.9674 | 53.6631
0.579 | 19.071 | 77.0004 | -1815.0664 | 1816.9541 | 9.2532 | -26.5713 | -4.7193 | 28.5686 | 1815.0664 | 26.5713 | 53.8304
0.5808 | 19.7536 | 80.11 | -1811.5353 | 1813.5663 | 9.0509 | -26.1247 | -4.8457 | 28.1135 | 1811.5353 | 26.1247 | 53.9978
0.5826 | 20.5205 | 82.7731 | -1806.875 | 1809.0493 | 8.8403 | -25.6359 | -4.9721 | 27.6247 | 1806.875 | 25.6359 | 54.1651
0.5844 | 21.4307 | 85.3013 | -1800.9843 | 1803.3018 | 8.6211 | -25.1302 | -5.0985 | 27.1107 | 1800.9843 | 25.1302 | 54.3325
0.5862 | 22.4167 | 87.366 | -1793.9644 | 1796.3999 | 8.3852 | -24.6078 | -5.2081 | 26.5713 | 1793.9644 | 24.6078 | 54.4998
0.588 | 23.3605 | 88.9756 | -1785.7393 | 1788.2759 | 8.1492 | -24.0431 | -5.3176 | 26.0151 | 1785.7393 | 24.0431 | 54.6672
0.5898 | 24.237 | 90.206 | -1776.1406 | 1778.7699 | 7.9048 | -23.4701 | -5.4103 | 25.4336 | 1776.1406 | 23.4701 | 54.8345
0.5916 | 25.1218 | 91.1835 | -1765.3199 | 1768.0167 | 7.6604 | -22.8886 | -5.4946 | 24.8437 | 1765.3199 | 22.8886 | 55.0019
0.5934 | 25.9224 | 92.2117 | -1753.2099 | 1755.9909 | 7.416 | -22.2987 | -5.5705 | 24.2454 | 1753.2099 | 22.2987 | 55.1692
0.5952 | 26.605 | 93.3325 | -1739.3638 | 1742.2291 | 7.1716 | -21.7003 | -5.6463 | 23.6386 | 1739.3638 | 21.7003 | 55.3366
0.597 | 27.3045 | 94.2764 | -1723.9418 | 1726.8998 | 6.9188 | -21.0936 | -5.7053 | 23.0234 | 1723.9418 | 21.0936 | 55.5039
0.5988 | 28.1557 | 94.7989 | -1708.1238 | 1711.166 | 6.666 | -20.4868 | -5.7643 | 22.4082 | 1708.1238 | 20.4868 | 55.6713
0.6006 | 29.0574 | 94.4786 | -1691.9518 | 1695.0277 | 6.3963 | -19.8463 | -5.8064 | 21.7593 | 1691.9518 | 19.8463 | 55.8386
0.6023 | 29.8664 | 94.0741 | -1675.1224 | 1678.2321 | 6.1267 | -19.189 | -5.8486 | 21.0936 | 1675.1224 | 19.189 | 55.9967
0.6041 | 30.6923 | 93.6022 | -1657.9138 | 1661.0488 | 5.8486 | -18.5317 | -5.8907 | 20.4278 | 1657.9138 | 18.5317 | 56.164
0.6059 | 31.5434 | 93.0965 | -1640.0142 | 1643.1745 | 5.5705 | -17.8912 | -5.9244 | 19.7705 | 1640.0142 | 17.8912 | 56.3313
0.6077 | 32.5126 | 91.9251 | -1621.8365 | 1624.9883 | 5.2671 | -17.2254 | -5.9497 | 19.0879 | 1621.8365 | 17.2254 | 56.4987
0.6095 | 33.6166 | 90.0796 | -1604.1139 | 1607.2151 | 4.9637 | -16.5344 | -5.975 | 18.3884 | 1604.1139 | 16.5344 | 56.666
0.6113 | 34.5604 | 87.8716 | -1585.1524 | 1588.1862 | 4.6603 | -15.8181 | -5.9834 | 17.6636 | 1585.1524 | 15.8181 | 56.8334
0.6131 | 35.3694 | 85.453 | -1564.7499 | 1567.7079 | 4.3569 | -15.0849 | -5.9918 | 16.9305 | 1564.7499 | 15.0849 | 57.0007
0.6149 | 35.9341 | 82.7478 | -1541.5663 | 1544.4316 | 4.0704 | -14.3349 | -5.9665 | 16.172 | 1541.5663 | 14.3349 | 57.1681
0.6167 | 36.3386 | 79.9584 | -1516.9754 | 1519.7649 | 3.7923 | -13.5848 | -5.9413 | 15.422 | 1516.9754 | 13.5848 | 57.3354
0.6185 | 36.8274 | 76.5959 | -1491.8873 | 1494.5672 | 3.5142 | -12.8348 | -5.8907 | 14.6635 | 1491.8873 | 12.8348 | 57.5028
0.6203 | 37.2656 | 72.9215 | -1465.9481 | 1468.5015 | 3.2361 | -12.0679 | -5.8317 | 13.9051 | 1465.9481 | 12.0679 | 57.6701
0.6221 | 37.6027 | 68.5731 | -1438.5256 | 1440.9105 | 2.9664 | -11.2926 | -5.7643 | 13.1298 | 1438.5256 | 11.2926 | 57.8375
0.6239 | 37.6111 | 64.0897 | -1409.4008 | 1411.6087 | 2.7052 | -10.5004 | -5.6716 | 12.346 | 1409.4008 | 10.5004 | 58.0048
0.6256 | 37.3499 | 59.4378 | -1379.1973 | 1381.2114 | 2.4608 | -9.7167 | -5.5705 | 11.5623 | 1379.1973 | 9.7167 | 58.1629
0.6274 | 36.8442 | 54.3056 | -1347.0555 | 1348.859 | 2.2417 | -8.9245 | -5.444 | 10.7954 | 1347.0555 | 8.9245 | 58.3302
0.6292 | 35.8666 | 49.1565 | -1312.3855 | 1313.9783 | 2.0647 | -8.1408 | -5.2924 | 10.0369 | 1312.3855 | 8.1408 | 58.4976
0.631 | 34.493 | 43.9821 | -1275.2969 | 1276.6874 | 1.9046 | -7.3486 | -5.1238 | 9.2869 | 1275.2969 | 7.3486 | 58.6649
0.6328 | 32.6558 | 38.8921 | -1235.5284 | 1236.7251 | 1.7782 | -6.5649 | -4.93 | 8.5453 | 1235.5284 | 6.5649 | 58.8323
0.6346 | 30.2541 | 34.1138 | -1193.493 | 1194.5127 | 1.677 | -5.7896 | -4.7277 | 7.8458 | 1193.493 | 5.7896 | 58.9996
0.6364 | 27.3298 | 29.5546 | -1148.9294 | 1149.789 | 1.6096 | -5.0564 | -4.5002 | 7.1969 | 1148.9294 | 5.0564 | 59.167
0.6382 | 24.1864 | 25.105 | -1103.1354 | 1103.8348 | 1.5591 | -4.3316 | -4.2642 | 6.5733 | 1103.1354 | 4.3316 | 59.3343
0.64 | 20.8155 | 20.7649 | -1056.6587 | 1057.2233 | 1.5253 | -3.6237 | -4.0198 | 5.9581 | 1056.6587 | 3.6237 | 59.5017
0.6418 | 17.1917 | 16.6018 | -1009.9208 | 1010.3506 | 1.4916 | -2.9327 | -3.7754 | 5.3766 | 1009.9208 | 2.9327 | 59.669
0.6436 | 13.2899 | 12.9191 | -964.6325 | 964.9527 | 1.4916 | -2.2754 | -3.5479 | 4.871 | 964.6325 | 2.2754 | 59.8364
0.6454 | 9.1436 | 9.6998 | -919.8666 | 920.111 | 1.5001 | -1.6602 | -3.3204 | 4.458 | 919.8666 | 1.6602 | 60.0037
0.6472 | 4.9468 | 6.7334 | -875.8254 | 876.0277 | 1.5253 | -1.0956 | -3.1013 | 4.1041 | 875.8254 | 1.0956 | 60.1711
0.649 | 0.69947 | 4.0114 | -832.2899 | 832.4669 | 1.5591 | -0.54778 | -2.8906 | 3.8007 | 832.2899 | 0.54778 | 60.3384
0.6507 | -3.4721 | 1.5675 | -789.9342 | 790.1028 | 1.5928 | -0.033709 | -2.6799 | 3.5563 | 789.9342 | 0.033709 | 60.4965
0.6525 | -7.5846 | -0.45507 | -750.0056 | 750.1995 | 1.6349 | 0.42979 | -2.4945 | 3.3962 | 750.0056 | -0.42979 | 60.6638
0.6543 | -11.4022 | -2.1911 | -713.8693 | 714.1137 | 1.6855 | 0.85959 | -2.3259 | 3.3035 | 713.8693 | -0.85959 | 60.8312
0.6561 | -14.6045 | -3.7923 | -682.5618 | 682.8652 | 1.7276 | 1.2304 | -2.1827 | 3.2698 | 682.5618 | -1.2304 | 60.9985
0.6579 | -17.2086 | -5.4272 | -655.3248 | 655.7124 | 1.7529 | 1.5506 | -2.0647 | 3.2951 | 655.3248 | -1.5506 | 61.1659
0.6597 | -19.2143 | -7.0621 | -631.3743 | 631.8463 | 1.7697 | 1.8287 | -1.9551 | 3.3541 | 631.3743 | -1.8287 | 61.3332
0.6615 | -20.7902 | -8.579 | -609.6319 | 610.1712 | 1.7866 | 2.0731 | -1.8624 | 3.4299 | 609.6319 | -2.0731 | 61.5006
0.6633 | -21.8605 | -9.9189 | -590.6367 | 591.235 | 1.795 | 2.2922 | -1.7782 | 3.5058 | 590.6367 | -2.2922 | 61.6679
0.6651 | -22.5936 | -10.9639 | -573.8495 | 574.5068 | 1.8203 | 2.4776 | -1.7023 | 3.5985 | 573.8495 | -2.4776 | 61.8353
0.6669 | -23.1751 | -11.7477 | -558.2505 | 558.9584 | 1.8456 | 2.6462 | -1.6265 | 3.6827 | 558.2505 | -2.6462 | 62.0026
0.6687 | -23.6808 | -12.3629 | -543.1319 | 543.8904 | 1.8709 | 2.8063 | -1.5506 | 3.7754 | 543.1319 | -2.8063 | 62.17
0.6705 | -24.1021 | -12.7927 | -528.6959 | 529.4965 | 1.9046 | 2.9496 | -1.4832 | 3.8513 | 528.6959 | -2.9496 | 62.3373
0.6723 | -24.3212 | -12.8685 | -515.9791 | 516.805 | 1.9383 | 3.076 | -1.4158 | 3.9356 | 515.9791 | -3.076 | 62.5046
0.6741 | -24.4392 | -12.7084 | -504.7961 | 505.6388 | 1.972 | 3.194 | -1.3568 | 4.0198 | 504.7961 | -3.194 | 62.672
0.6758 | -24.5825 | -12.3376 | -494.6243 | 495.4839 | 2.0057 | 3.2951 | -1.3062 | 4.0957 | 494.6243 | -3.2951 | 62.83
0.6776 | -24.7763 | -11.8067 | -484.9582 | 485.8262 | 2.0478 | 3.4046 | -1.2472 | 4.1799 | 484.9582 | -3.4046 | 62.9974
0.6794 | -24.987 | -11.1325 | -476.2781 | 477.1545 | 2.0815 | 3.4889 | -1.2051 | 4.2558 | 476.2781 | -3.4889 | 63.1647
0.6812 | -25.2482 | -10.2729 | -468.1878 | 469.0811 | 2.1153 | 3.5732 | -1.163 | 4.3232 | 468.1878 | -3.5732 | 63.3321
0.683 | -25.5263 | -9.3375 | -460.5611 | 461.4628 | 2.1405 | 3.6406 | -1.1208 | 4.3822 | 460.5611 | -3.6406 | 63.4994
0.6848 | -25.7539 | -8.3515 | -453.6002 | 454.5187 | 2.1658 | 3.708 | -1.0787 | 4.4412 | 453.6002 | -3.708 | 63.6668
0.6866 | -26.0067 | -7.357 | -447.0774 | 448.0128 | 2.1827 | 3.767 | -1.0366 | 4.4833 | 447.0774 | -3.767 | 63.8341
0.6884 | -26.2258 | -6.4806 | -440.7232 | 441.6671 | 2.1911 | 3.826 | -0.99442 | 4.5339 | 440.7232 | -3.826 | 64.0015
0.6902 | -26.3016 | -5.7053 | -434.1246 | 435.0853 | 2.1995 | 3.8681 | -0.95229 | 4.5676 | 434.1246 | -3.8681 | 64.1688
0.692 | -26.2932 | -4.989 | -427.5345 | 428.4952 | 2.1995 | 3.9018 | -0.91015 | 4.5929 | 427.5345 | -3.9018 | 64.3362
0.6938 | -26.3606 | -4.3232 | -421.2646 | 422.2421 | 2.1995 | 3.944 | -0.86801 | 4.6097 | 421.2646 | -3.944 | 64.5035
0.6956 | -26.4028 | -3.6659 | -415.2727 | 416.2587 | 2.1911 | 3.9777 | -0.82588 | 4.6266 | 415.2727 | -3.9777 | 64.6709
0.6974 | -26.4955 | -3.0338 | -409.6349 | 410.6377 | 2.1827 | 4.003 | -0.79217 | 4.6434 | 409.6349 | -4.003 | 64.8382
0.6992 | -26.6303 | -2.4271 | -404.4184 | 405.4296 | 2.1658 | 4.0367 | -0.75846 | 4.6603 | 404.4184 | -4.0367 | 65.0056
0.7009 | -26.664 | -1.8624 | -399.4715 | 400.4912 | 2.149 | 4.062 | -0.72475 | 4.6603 | 399.4715 | -4.062 | 65.1636
0.7027 | -26.6219 | -1.3315 | -394.6427 | 395.6624 | 2.1321 | 4.0788 | -0.69104 | 4.6603 | 394.6427 | -4.0788 | 65.331
0.7045 | -26.5292 | -0.7416 | -389.6874 | 390.6987 | 2.1237 | 4.0872 | -0.65733 | 4.6687 | 389.6874 | -4.0872 | 65.4983
0.7063 | -26.3269 | -0.16855 | -384.5383 | 385.5496 | 2.1068 | 4.0957 | -0.62362 | 4.6519 | 384.5383 | -4.0957 | 65.6657
0.7081 | -26.0657 | 0.41294 | -379.6083 | 380.6112 | 2.0815 | 4.0957 | -0.58991 | 4.6434 | 379.6083 | -4.0957 | 65.833
0.7099 | -25.8803 | 0.99442 | -375.4115 | 376.3975 | 2.0647 | 4.0872 | -0.5562 | 4.6182 | 375.4115 | -4.0872 | 66.0004
0.7117 | -25.6949 | 1.4916 | -371.8889 | 372.8749 | 2.031 | 4.0788 | -0.53935 | 4.6013 | 371.8889 | -4.0788 | 66.1677
0.7135 | -25.5011 | 1.9973 | -368.5349 | 369.504 | 2.0057 | 4.0704 | -0.51407 | 4.576 | 368.5349 | -4.0704 | 66.3351
0.7153 | -25.3578 | 2.4186 | -365.1049 | 366.0741 | 1.9804 | 4.0535 | -0.49721 | 4.5507 | 365.1049 | -4.0535 | 66.5024
0.7171 | -25.2819 | 2.7136 | -361.3632 | 362.3323 | 1.9551 | 4.0535 | -0.47193 | 4.5255 | 361.3632 | -4.0535 | 66.6698
0.7189 | -25.2145 | 2.899 | -356.8293 | 357.8069 | 1.9214 | 4.0451 | -0.44665 | 4.5086 | 356.8293 | -4.0451 | 66.8371
0.7207 | -25.164 | 3.0423 | -351.3768 | 352.3713 | 1.8877 | 4.0367 | -0.41294 | 4.4833 | 351.3768 | -4.0367 | 67.0045
0.7225 | -25.164 | 3.2024 | -345.5283 | 346.5396 | 1.8624 | 4.0283 | -0.3708 | 4.458 | 345.5283 | -4.0283 | 67.1718
0.7242 | -25.0965 | 3.4215 | -339.4859 | 340.5056 | 1.8287 | 4.0114 | -0.33709 | 4.4328 | 339.4859 | -4.0114 | 67.3299
0.726 | -24.9701 | 3.6153 | -333.4435 | 334.4716 | 1.8034 | 3.9945 | -0.29496 | 4.3991 | 333.4435 | -3.9945 | 67.4972
0.7278 | -24.7763 | 3.8007 | -327.7635 | 328.8001 | 1.7782 | 3.9693 | -0.26125 | 4.3569 | 327.7635 | -3.9693 | 67.6646
0.7296 | -24.5572 | 4.062 | -322.724 | 323.7605 | 1.7529 | 3.9356 | -0.23596 | 4.3148 | 322.724 | -3.9356 | 67.8319
0.7314 | -24.3212 | 4.3738 | -318.4513 | 319.4879 | 1.7276 | 3.9018 | -0.21068 | 4.2726 | 318.4513 | -3.9018 | 67.9993
0.7332 | -24.0431 | 4.7614 | -315.072 | 316.1001 | 1.6939 | 3.8597 | -0.19383 | 4.2305 | 315.072 | -3.8597 | 68.1666
0.735 | -23.824 | 5.1491 | -312.3247 | 313.3528 | 1.6686 | 3.826 | -0.1854 | 4.1884 | 312.3247 | -3.826 | 68.334
0.7368 | -23.6218 | 5.3935 | -310.0746 | 311.1027 | 1.6433 | 3.8007 | -0.17697 | 4.1462 | 310.0746 | -3.8007 | 68.5013
0.7386 | -23.3689 | 5.562 | -307.5295 | 308.5492 | 1.6096 | 3.767 | -0.16855 | 4.1041 | 307.5295 | -3.767 | 68.6687
0.7404 | -22.9897 | 5.6463 | -303.7457 | 304.7401 | 1.5759 | 3.7164 | -0.14326 | 4.0535 | 303.7457 | -3.7164 | 68.836
0.7422 | -22.6021 | 5.6042 | -299.3466 | 300.3326 | 1.5422 | 3.6743 | -0.11798 | 3.9945 | 299.3466 | -3.6743 | 69.0033
0.744 | -22.2397 | 5.5789 | -294.6105 | 295.5796 | 1.5085 | 3.6322 | -0.0927 | 3.9356 | 294.6105 | -3.6322 | 69.1707
0.7458 | -21.9784 | 5.6126 | -289.8996 | 290.8687 | 1.4748 | 3.5816 | -0.067419 | 3.885 | 289.8996 | -3.5816 | 69.338
0.7476 | -21.7425 | 5.6463 | -285.5595 | 286.5202 | 1.4495 | 3.531 | -0.042137 | 3.826 | 285.5595 | -3.531 | 69.5054
0.7493 | -21.4054 | 5.6969 | -281.6914 | 282.6521 | 1.4158 | 3.4805 | -0.025282 | 3.7586 | 281.6914 | -3.4805 | 69.6634
0.7511 | -21.0093 | 5.7896 | -278.4722 | 279.4076 | 1.3821 | 3.4299 | -0.0084273 | 3.6996 | 278.4722 | -3.4299 | 69.8308
0.7529 | -20.6553 | 5.8907 | -275.8934 | 276.8204 | 1.3568 | 3.3794 | 0 | 3.6406 | 275.8934 | -3.3794 | 69.9981
0.7547 | -20.3014 | 6.0508 | -273.7444 | 274.6546 | 1.3231 | 3.3372 | 0.0084273 | 3.59 | 273.7444 | -3.3372 | 70.1655
0.7565 | -19.939 | 6.2194 | -271.7556 | 272.6489 | 1.2978 | 3.2951 | 0.016855 | 3.5479 | 271.7556 | -3.2951 | 70.3328
0.7583 | -19.5935 | 6.3879 | -270.0111 | 270.896 | 1.2725 | 3.2529 | 0.025282 | 3.4973 | 270.0111 | -3.2529 | 70.5002
0.7601 | -19.2901 | 6.5902 | -268.6965 | 269.5561 | 1.2472 | 3.2108 | 0.033709 | 3.4468 | 268.6965 | -3.2108 | 70.6675
0.7619 | -19.0204 | 6.8261 | -267.8285 | 268.6796 | 1.222 | 3.1687 | 0.025282 | 3.4046 | 267.8285 | -3.1687 | 70.8349
0.7637 | -18.7508 | 7.0789 | -266.9014 | 267.7442 | 1.1967 | 3.1265 | 0.025282 | 3.3541 | 266.9014 | -3.1265 | 71.0022
0.7655 | -18.4474 | 7.3065 | -265.0811 | 265.9155 | 1.1714 | 3.0844 | 0.033709 | 3.3035 | 265.0811 | -3.0844 | 71.1696
0.7673 | -18.1103 | 7.4835 | -262.494 | 263.3114 | 1.1461 | 3.0338 | 0.042137 | 3.2529 | 262.494 | -3.0338 | 71.3369
0.7691 | -17.7816 | 7.5677 | -259.2916 | 260.0922 | 1.1208 | 2.9917 | 0.067419 | 3.194 | 259.2916 | -2.9917 | 71.5043
0.7709 | -17.4867 | 7.6014 | -255.5751 | 256.3757 | 1.0956 | 2.9411 | 0.084273 | 3.1434 | 255.5751 | -2.9411 | 71.6716
0.7727 | -17.2254 | 7.652 | -252.0778 | 252.87 | 1.0703 | 2.8906 | 0.10956 | 3.0844 | 252.0778 | -2.8906 | 71.839
0.7744 | -17.0485 | 7.7868 | -249.6929 | 250.485 | 1.045 | 2.8484 | 0.11798 | 3.0338 | 249.6929 | -2.8484 | 71.997
0.7762 | -16.9389 | 8.0059 | -248.8417 | 249.6339 | 1.0197 | 2.8063 | 0.11798 | 2.9917 | 248.8417 | -2.8063 | 72.1644
0.778 | -16.8883 | 8.284 | -248.69 | 249.4906 | 1.0029 | 2.781 | 0.11798 | 2.958 | 248.69 | -2.781 | 72.3317
0.7798 | -16.9473 | 8.6717 | -249.2294 | 250.0384 | 0.986 | 2.7473 | 0.10956 | 2.9243 | 249.2294 | -2.7473 | 72.4991
0.7816 | -17.0653 | 9.1521 | -250.3165 | 251.1508 | 0.96914 | 2.7304 | 0.0927 | 2.899 | 250.3165 | -2.7304 | 72.6664
0.7834 | -17.1749 | 9.6324 | -251.5385 | 252.3896 | 0.95229 | 2.7052 | 0.075846 | 2.8737 | 251.5385 | -2.7052 | 72.8338
0.7852 | -17.1412 | 10.0369 | -251.9093 | 252.7688 | 0.93543 | 2.6799 | 0.075846 | 2.84 | 251.9093 | -2.6799 | 73.0011
0.787 | -16.9305 | 10.2139 | -251.3193 | 252.1705 | 0.91015 | 2.6462 | 0.075846 | 2.8063 | 251.3193 | -2.6462 | 73.1685
0.7888 | -16.6102 | 10.256 | -250.0215 | 250.8558 | 0.88487 | 2.6125 | 0.0927 | 2.7642 | 250.0215 | -2.6125 | 73.3358
0.7906 | -16.231 | 10.2308 | -248.4709 | 249.2799 | 0.85959 | 2.5703 | 0.10956 | 2.722 | 248.4709 | -2.5703 | 73.5032
0.7924 | -15.8433 | 10.2139 | -246.8192 | 247.6113 | 0.8343 | 2.5366 | 0.12641 | 2.6799 | 246.8192 | -2.5366 | 73.6705
0.7942 | -15.5147 | 10.1802 | -245.5298 | 246.2967 | 0.80902 | 2.5029 | 0.14326 | 2.6377 | 245.5298 | -2.5029 | 73.8379
0.796 | -15.3124 | 10.2476 | -245.1 | 245.8584 | 0.78374 | 2.4692 | 0.15169 | 2.604 | 245.1 | -2.4692 | 74.0052
0.7977 | -15.2787 | 10.4583 | -245.5972 | 246.3557 | 0.76689 | 2.4523 | 0.15169 | 2.5788 | 245.5972 | -2.4523 | 74.1633
0.7995 | -15.3293 | 10.7364 | -246.5158 | 247.2911 | 0.75003 | 2.4355 | 0.14326 | 2.5535 | 246.5158 | -2.4355 | 74.3306
0.8013 | -15.4725 | 11.1156 | -247.9821 | 248.7743 | 0.73318 | 2.4271 | 0.13484 | 2.545 | 247.9821 | -2.4271 | 74.498
0.8031 | -15.599 | 11.5033 | -249.7519 | 250.5609 | 0.72475 | 2.4271 | 0.12641 | 2.5366 | 249.7519 | -2.4271 | 74.6653
0.8049 | -15.8096 | 11.9584 | -251.7154 | 252.5497 | 0.70789 | 2.4186 | 0.10956 | 2.5282 | 251.7154 | -2.4186 | 74.8327
0.8067 | -16.054 | 12.4387 | -253.6874 | 254.5554 | 0.69947 | 2.4271 | 0.10113 | 2.5282 | 253.6874 | -2.4271 | 75
0.8085 | -16.1804 | 12.8938 | -255.2718 | 256.1566 | 0.69104 | 2.4186 | 0.0927 | 2.5198 | 255.2718 | -2.4186 | 75.1673
0.8103 | -16.1467 | 13.1972 | -255.8532 | 256.7465 | 0.67419 | 2.4102 | 0.0927 | 2.5113 | 255.8532 | -2.4102 | 75.3347
0.8121 | -15.9529 | 13.3489 | -255.415 | 256.3083 | 0.65733 | 2.4018 | 0.10113 | 2.4945 | 255.415 | -2.4018 | 75.502
0.8139 | -15.7169 | 13.4247 | -254.5807 | 255.474 | 0.63205 | 2.3849 | 0.11798 | 2.4776 | 254.5807 | -2.3849 | 75.6694
0.8157 | -15.4557 | 13.4668 | -253.6958 | 254.5723 | 0.61519 | 2.3765 | 0.13484 | 2.4608 | 253.6958 | -2.3765 | 75.8367
0.8175 | -15.2282 | 13.509 | -253.2071 | 254.0751 | 0.59834 | 2.3596 | 0.15169 | 2.4439 | 253.2071 | -2.3596 | 76.0041
0.8193 | -14.9838 | 13.6607 | -253.2576 | 254.1172 | 0.58148 | 2.3428 | 0.16012 | 2.4271 | 253.2576 | -2.3428 | 76.1714
0.8211 | -14.8405 | 13.9219 | -254.2183 | 255.0864 | 0.56463 | 2.3344 | 0.16012 | 2.4102 | 254.2183 | -2.3344 | 76.3388
0.8228 | -14.8405 | 14.318 | -256.1229 | 256.9994 | 0.5562 | 2.3344 | 0.15169 | 2.4102 | 256.1229 | -2.3344 | 76.4968
0.8246 | -14.9585 | 14.7899 | -258.491 | 259.4011 | 0.53935 | 2.3428 | 0.14326 | 2.4102 | 258.491 | -2.3428 | 76.6642
0.8264 | -15.1692 | 15.3209 | -261.2214 | 262.1653 | 0.53092 | 2.3512 | 0.12641 | 2.4186 | 261.2214 | -2.3512 | 76.8315
0.8282 | -15.3798 | 15.8433 | -263.8002 | 264.7862 | 0.52249 | 2.3681 | 0.11798 | 2.4355 | 263.8002 | -2.3681 | 76.9989
0.83 | -15.5147 | 16.2731 | -265.6795 | 266.6908 | 0.51407 | 2.3765 | 0.10956 | 2.4439 | 265.6795 | -2.3765 | 77.1662
0.8318 | -15.5821 | 16.5512 | -266.8593 | 267.879 | 0.49721 | 2.3849 | 0.11798 | 2.4439 | 266.8593 | -2.3849 | 77.3336
0.8336 | -15.5652 | 16.8209 | -267.4998 | 268.5448 | 0.48878 | 2.3934 | 0.11798 | 2.4523 | 267.4998 | -2.3934 | 77.5009
0.8354 | -15.6074 | 17.0569 | -267.8706 | 268.9324 | 0.47193 | 2.4018 | 0.12641 | 2.4608 | 267.8706 | -2.4018 | 77.6683
0.8372 | -15.7001 | 17.2591 | -267.8369 | 268.9156 | 0.4635 | 2.4102 | 0.13484 | 2.4608 | 267.8369 | -2.4102 | 77.8356
0.839 | -15.8096 | 17.394 | -267.1121 | 268.2077 | 0.44665 | 2.4271 | 0.15169 | 2.4776 | 267.1121 | -2.4271 | 78.003
0.8408 | -15.8096 | 17.4867 | -266.0166 | 267.129 | 0.42979 | 2.4355 | 0.17697 | 2.4776 | 266.0166 | -2.4355 | 78.1703
0.8426 | -15.8096 | 17.5541 | -264.8452 | 265.9576 | 0.41294 | 2.4439 | 0.19383 | 2.4861 | 264.8452 | -2.4439 | 78.3377
0.8444 | -15.7675 | 17.5625 | -263.3367 | 264.4575 | 0.39608 | 2.4439 | 0.21911 | 2.4945 | 263.3367 | -2.4439 | 78.505
0.8462 | -15.7254 | 17.6047 | -261.9883 | 263.1092 | 0.38766 | 2.4523 | 0.23596 | 2.4945 | 261.9883 | -2.4523 | 78.6724
0.8479 | -15.7675 | 17.7901 | -261.508 | 262.6457 | 0.3708 | 2.4608 | 0.25282 | 2.5029 | 261.508 | -2.4608 | 78.8304
0.8497 | -15.9192 | 18.1019 | -262.1316 | 263.2946 | 0.3708 | 2.4692 | 0.25282 | 2.5198 | 262.1316 | -2.4692 | 78.9978
0.8515 | -16.1467 | 18.5317 | -263.6738 | 264.8705 | 0.36237 | 2.4861 | 0.25282 | 2.5366 | 263.6738 | -2.4861 | 79.1651
0.8533 | -16.3996 | 19.012 | -265.7132 | 266.952 | 0.36237 | 2.5113 | 0.24439 | 2.5535 | 265.7132 | -2.5113 | 79.3325
0.8551 | -16.6777 | 19.5429 | -267.8453 | 269.1263 | 0.35395 | 2.5282 | 0.23596 | 2.5703 | 267.8453 | -2.5282 | 79.4998
0.8569 | -16.9726 | 20.0233 | -269.8089 | 271.132 | 0.35395 | 2.5535 | 0.22754 | 2.5956 | 269.8089 | -2.5535 | 79.6672
0.8587 | -17.2086 | 20.3098 | -271.1067 | 272.455 | 0.34552 | 2.5703 | 0.22754 | 2.6125 | 271.1067 | -2.5703 | 79.8345
0.8605 | -17.3603 | 20.4362 | -271.6207 | 272.9944 | 0.32867 | 2.5956 | 0.22754 | 2.6293 | 271.6207 | -2.5956 | 80.0019
0.8623 | -17.5288 | 20.4868 | -271.3848 | 272.7753 | 0.32024 | 2.6125 | 0.24439 | 2.6462 | 271.3848 | -2.6125 | 80.1692
0.8641 | -17.8743 | 20.4194 | -270.3651 | 271.7809 | 0.30338 | 2.6293 | 0.25282 | 2.6715 | 270.3651 | -2.6293 | 80.3366
0.8659 | -18.1861 | 20.2677 | -268.4015 | 269.8342 | 0.28653 | 2.6462 | 0.26967 | 2.6883 | 268.4015 | -2.6462 | 80.5039
0.8677 | -18.4221 | 20.0402 | -265.8565 | 267.306 | 0.2781 | 2.663 | 0.29496 | 2.6967 | 265.8565 | -2.663 | 80.6713
0.8695 | -18.6328 | 19.7789 | -263.067 | 264.5249 | 0.26125 | 2.6715 | 0.32024 | 2.7052 | 263.067 | -2.6715 | 80.8386
0.8712 | -18.7592 | 19.5682 | -260.2607 | 261.7271 | 0.25282 | 2.6715 | 0.34552 | 2.7136 | 260.2607 | -2.6715 | 80.9967
0.873 | -18.835 | 19.3491 | -257.9854 | 259.4601 | 0.24439 | 2.6715 | 0.36237 | 2.7136 | 257.9854 | -2.6715 | 81.164
0.8748 | -18.8435 | 19.1806 | -256.5948 | 258.0612 | 0.23596 | 2.6715 | 0.3708 | 2.7052 | 256.5948 | -2.6715 | 81.3313
0.8766 | -18.8435 | 19.189 | -256.8982 | 258.3646 | 0.22754 | 2.6715 | 0.3708 | 2.7052 | 256.8982 | -2.6715 | 81.4987
0.8784 | -18.835 | 19.3154 | -258.2129 | 259.6792 | 0.21911 | 2.6715 | 0.36237 | 2.7136 | 258.2129 | -2.6715 | 81.666
0.8802 | -18.7761 | 19.3575 | -259.7635 | 261.2214 | 0.21068 | 2.6799 | 0.35395 | 2.7136 | 259.7635 | -2.6799 | 81.8334
0.882 | -18.6749 | 19.307 | -260.9349 | 262.376 | 0.20226 | 2.6799 | 0.35395 | 2.7136 | 260.9349 | -2.6799 | 82.0007
0.8838 | -18.6496 | 19.2817 | -262.2327 | 263.6654 | 0.19383 | 2.6883 | 0.34552 | 2.722 | 262.2327 | -2.6883 | 82.1681
0.8856 | -18.7086 | 19.2901 | -263.6654 | 265.098 | 0.1854 | 2.6967 | 0.33709 | 2.7304 | 263.6654 | -2.6967 | 82.3354
0.8874 | -18.6834 | 19.1806 | -264.5081 | 265.9239 | 0.17697 | 2.6967 | 0.33709 | 2.7304 | 264.5081 | -2.6967 | 82.5028
0.8892 | -18.6665 | 18.9362 | -264.2131 | 265.6121 | 0.16855 | 2.6967 | 0.33709 | 2.7304 | 264.2131 | -2.6967 | 82.6701
0.891 | -18.6496 | 18.6244 | -263.126 | 264.5081 | 0.15169 | 2.7052 | 0.34552 | 2.7304 | 263.126 | -2.7052 | 82.8375
0.8928 | -18.5822 | 18.2199 | -261.4827 | 262.8479 | 0.14326 | 2.6967 | 0.36237 | 2.7304 | 261.4827 | -2.6967 | 83.0048
0.8946 | -18.4727 | 17.8238 | -259.5107 | 260.8506 | 0.13484 | 2.6883 | 0.3708 | 2.722 | 259.5107 | -2.6883 | 83.1722
0.8963 | -18.3294 | 17.4951 | -257.5134 | 258.8281 | 0.11798 | 2.6799 | 0.38766 | 2.722 | 257.5134 | -2.6799 | 83.3302
0.8981 | -18.203 | 17.2423 | -256.0724 | 257.3702 | 0.11798 | 2.6715 | 0.39608 | 2.7052 | 256.0724 | -2.6715 | 83.4976
0.8999 | -18.0345 | 17.0232 | -255.2043 | 256.4853 | 0.10956 | 2.6546 | 0.40451 | 2.6967 | 255.2043 | -2.6546 | 83.6649
0.9017 | -17.8322 | 16.8041 | -254.7661 | 256.0134 | 0.10956 | 2.6462 | 0.40451 | 2.6883 | 254.7661 | -2.6462 | 83.8323
0.9035 | -17.6468 | 16.6861 | -255.0105 | 256.2493 | 0.10956 | 2.6377 | 0.40451 | 2.6715 | 255.0105 | -2.6377 | 83.9996
0.9053 | -17.6047 | 16.7198 | -256.1735 | 257.4039 | 0.10113 | 2.6293 | 0.39608 | 2.6715 | 256.1735 | -2.6293 | 84.167
0.9071 | -17.6974 | 16.7872 | -257.8 | 259.0388 | 0.10956 | 2.6293 | 0.37923 | 2.663 | 257.8 | -2.6293 | 84.3343
0.9089 | -17.7985 | 16.8378 | -259.6961 | 260.9265 | 0.10956 | 2.6293 | 0.35395 | 2.663 | 259.6961 | -2.6293 | 84.5017
0.9107 | -17.8828 | 16.8378 | -261.1035 | 262.3338 | 0.10956 | 2.6293 | 0.33709 | 2.663 | 261.1035 | -2.6293 | 84.669
0.9125 | -17.9502 | 16.7198 | -261.2214 | 262.4434 | 0.10956 | 2.6209 | 0.32867 | 2.6546 | 261.2214 | -2.6209 | 84.8364
0.9143 | -18.0597 | 16.4585 | -260.3871 | 261.6091 | 0.10956 | 2.6125 | 0.32867 | 2.6462 | 260.3871 | -2.6125 | 85.0037
0.9161 | -18.1272 | 16.1215 | -258.6174 | 259.8309 | 0.10113 | 2.604 | 0.33709 | 2.6293 | 258.6174 | -2.604 | 85.1711
0.9179 | -18.1272 | 15.7254 | -255.7184 | 256.9235 | 0.10113 | 2.5872 | 0.35395 | 2.6209 | 255.7184 | -2.5872 | 85.3384
0.9197 | -18.0513 | 15.3377 | -252.2969 | 253.4852 | 0.10113 | 2.5619 | 0.3708 | 2.5956 | 252.2969 | -2.5619 | 85.5058
0.9214 | -17.9418 | 14.9416 | -248.8417 | 250.0215 | 0.0927 | 2.5366 | 0.38766 | 2.5703 | 248.8417 | -2.5366 | 85.6638
0.9232 | -17.8828 | 14.5877 | -246.3557 | 247.5186 | 0.0927 | 2.5113 | 0.40451 | 2.5535 | 246.3557 | -2.5113 | 85.8312
0.925 | -17.8153 | 14.318 | -244.6112 | 245.7657 | 0.0927 | 2.4861 | 0.40451 | 2.5282 | 244.6112 | -2.4861 | 85.9985
0.9268 | -17.7058 | 14.0062 | -243.2628 | 244.4005 | 0.084273 | 2.4608 | 0.40451 | 2.5029 | 243.2628 | -2.4608 | 86.1659
0.9286 | -17.7142 | 13.8292 | -242.6561 | 243.7853 | 0.084273 | 2.4355 | 0.40451 | 2.4861 | 242.6561 | -2.4355 | 86.3332
0.9304 | -17.7901 | 13.7702 | -242.6645 | 243.7938 | 0.0927 | 2.4186 | 0.38766 | 2.4608 | 242.6645 | -2.4186 | 86.5006
0.9322 | -17.8153 | 13.7618 | -243.0859 | 244.2151 | 0.0927 | 2.4018 | 0.3708 | 2.4439 | 243.0859 | -2.4018 | 86.6679
0.934 | -17.8491 | 13.8882 | -244.0971 | 245.2264 | 0.10113 | 2.3765 | 0.35395 | 2.4186 | 244.0971 | -2.3765 | 86.8353
0.9358 | -17.8743 | 14.082 | -245.2685 | 246.4062 | 0.11798 | 2.3596 | 0.33709 | 2.3934 | 245.2685 | -2.3596 | 87.0026
0.9376 | -17.908 | 14.2085 | -246.0691 | 247.2068 | 0.12641 | 2.3344 | 0.32024 | 2.3765 | 246.0691 | -2.3344 | 87.17
0.9394 | -17.9502 | 14.1663 | -246.1871 | 247.3248 | 0.12641 | 2.3175 | 0.31181 | 2.3512 | 246.1871 | -2.3175 | 87.3373
0.9412 | -17.8406 | 14.0399 | -245.4961 | 246.6253 | 0.12641 | 2.2922 | 0.30338 | 2.3259 | 245.4961 | -2.2922 | 87.5046
0.943 | -17.6636 | 13.8461 | -244.4764 | 245.5888 | 0.12641 | 2.2669 | 0.30338 | 2.3007 | 244.4764 | -2.2669 | 87.672
0.9447 | -17.5372 | 13.7028 | -243.1027 | 244.2067 | 0.13484 | 2.2248 | 0.30338 | 2.2585 | 243.1027 | -2.2248 | 87.83
0.9465 | -17.4867 | 13.6185 | -241.1139 | 242.2178 | 0.13484 | 2.1827 | 0.31181 | 2.2164 | 241.1139 | -2.1827 | 87.9974
0.9483 | -17.4193 | 13.5595 | -238.3834 | 239.4958 | 0.14326 | 2.1321 | 0.32024 | 2.1742 | 238.3834 | -2.1321 | 88.1647
0.9501 | -17.3181 | 13.5343 | -235.5518 | 236.6811 | 0.14326 | 2.0815 | 0.32867 | 2.1237 | 235.5518 | -2.0815 | 88.3321
0.9519 | -17.1412 | 13.5764 | -233.5377 | 234.6754 | 0.15169 | 2.031 | 0.32867 | 2.0731 | 233.5377 | -2.031 | 88.4994
0.9537 | -16.7872 | 13.7197 | -232.7203 | 233.8411 | 0.16012 | 1.9804 | 0.32867 | 2.0226 | 232.7203 | -1.9804 | 88.6668
0.9555 | -16.4501 | 14.0315 | -232.5939 | 233.7147 | 0.16855 | 1.9214 | 0.32024 | 1.972 | 232.5939 | -1.9214 | 88.8341
0.9573 | -16.2226 | 14.5371 | -233.0995 | 234.2372 | 0.17697 | 1.8709 | 0.30338 | 1.913 | 233.0995 | -1.8709 | 89.0015
0.9591 | -16.054 | 15.0933 | -233.9675 | 235.1305 | 0.19383 | 1.8203 | 0.28653 | 1.8624 | 233.9675 | -1.8203 | 89.1688
0.9609 | -15.9445 | 15.5737 | -234.7681 | 235.9648 | 0.20226 | 1.7697 | 0.26967 | 1.8203 | 234.7681 | -1.7697 | 89.3362
0.9627 | -15.7675 | 15.9445 | -234.8777 | 236.0828 | 0.21068 | 1.7192 | 0.26125 | 1.7697 | 234.8777 | -1.7192 | 89.5035
0.9645 | -15.5315 | 16.231 | -234.3552 | 235.5771 | 0.21911 | 1.6686 | 0.26125 | 1.7192 | 234.3552 | -1.6686 | 89.6709
0.9663 | -15.2703 | 16.6355 | -233.3523 | 234.5827 | 0.21911 | 1.6096 | 0.26125 | 1.6686 | 233.3523 | -1.6096 | 89.8382
0.9681 | -15.0512 | 17.1833 | -232.6023 | 233.8664 | 0.22754 | 1.5506 | 0.26125 | 1.6012 | 232.6023 | -1.5506 | 90.0056
0.9698 | -14.9501 | 17.9839 | -232.7877 | 234.1192 | 0.23596 | 1.4832 | 0.25282 | 1.5422 | 232.7877 | -1.4832 | 90.1636
0.9716 | -14.9248 | 18.953 | -233.0911 | 234.5069 | 0.25282 | 1.4074 | 0.24439 | 1.4748 | 233.0911 | -1.4074 | 90.331
0.9734 | -14.9669 | 19.9053 | -231.8354 | 233.3523 | 0.26125 | 1.3315 | 0.24439 | 1.3989 | 231.8354 | -1.3315 | 90.4983
0.9752 | -14.9501 | 20.8323 | -229.0965 | 230.723 | 0.2781 | 1.2304 | 0.26125 | 1.3231 | 229.0965 | -1.2304 | 90.6657
0.977 | -15.0175 | 21.6835 | -225.675 | 227.4111 | 0.29496 | 1.1293 | 0.28653 | 1.2472 | 225.675 | -1.1293 | 90.833
0.9788 | -15.2366 | 22.5599 | -221.7732 | 223.6525 | 0.30338 | 1.0197 | 0.31181 | 1.163 | 221.7732 | -1.0197 | 91.0004
0.9806 | -15.4641 | 23.5122 | -216.7758 | 218.8405 | 0.31181 | 0.88487 | 0.35395 | 1.0871 | 216.7758 | -0.88487 | 91.1677
0.9824 | -15.7085 | 24.4982 | -211.7194 | 213.9779 | 0.32867 | 0.75003 | 0.40451 | 1.0197 | 211.7194 | -0.75003 | 91.3351
0.9842 | -16.0035 | 25.6949 | -206.9327 | 209.4356 | 0.33709 | 0.59834 | 0.44665 | 0.95229 | 206.9327 | -0.59834 | 91.5024
0.986 | -16.2057 | 27.0854 | -203.4438 | 206.1911 | 0.36237 | 0.45507 | 0.47193 | 0.91015 | 203.4438 | -0.45507 | 91.6698
0.9878 | -16.3827 | 28.5602 | -200.8903 | 203.8904 | 0.37923 | 0.29496 | 0.49721 | 0.88487 | 200.8903 | -0.29496 | 91.8371
0.9896 | -16.4585 | 30.3552 | -199.2217 | 202.5252 | 0.40451 | 0.11798 | 0.52249 | 0.91015 | 199.2217 | -0.11798 | 92.0045
0.9914 | -16.5512 | 32.3356 | -198.1177 | 201.7583 | 0.42137 | -0.067419 | 0.54778 | 0.94386 | 198.1177 | 0.067419 | 92.1718
0.9932 | -16.6102 | 34.4593 | -197.966 | 201.9774 | 0.44665 | -0.26125 | 0.56463 | 1.0029 | 197.966 | 0.26125 | 92.3392
0.9949 | -16.5428 | 36.5661 | -198.8593 | 203.2078 | 0.45507 | -0.45507 | 0.58148 | 1.0703 | 198.8593 | 0.45507 | 92.4972
0.9967 | -16.3406 | 38.4875 | -201.042 | 205.6602 | 0.45507 | -0.65733 | 0.58991 | 1.163 | 201.042 | 0.65733 | 92.6646
0.9985 | -15.7001 | 40.4427 | -206.0731 | 210.9104 | 0.43822 | -0.8933 | 0.58991 | 1.2894 | 206.0731 | 0.8933 | 92.8319
1.0003 | -14.4697 | 42.3978 | -214.8712 | 219.8433 | 0.40451 | -1.1377 | 0.57306 | 1.4664 | 214.8712 | 1.1377 | 92.9993
1.0021 | -12.8601 | 44.4625 | -226.7284 | 231.8101 | 0.35395 | -1.3989 | 0.54778 | 1.6855 | 226.7284 | 1.3989 | 93.1666
1.0039 | -10.8881 | 46.3502 | -240.3975 | 245.5214 | 0.28653 | -1.6602 | 0.51407 | 1.9046 | 240.3975 | 1.6602 | 93.334
1.0057 | -8.5116 | 47.9851 | -255.6341 | 260.7158 | 0.20226 | -1.9383 | 0.48878 | 2.1321 | 255.6341 | 1.9383 | 93.5013
1.0075 | -5.8064 | 49.4936 | -271.8736 | 276.871 | 0.10956 | -2.208 | 0.44665 | 2.3681 | 271.8736 | 2.208 | 93.6687
1.0093 | -2.9664 | 51.0611 | -288.6186 | 293.5402 | 0.025282 | -2.4861 | 0.41294 | 2.604 | 288.6186 | 2.4861 | 93.836
1.0111 | -0.13484 | 52.6876 | -305.3974 | 310.26 | -0.067419 | -2.7557 | 0.3708 | 2.84 | 305.3974 | 2.7557 | 94.0033
1.0129 | 2.5198 | 54.2129 | -321.4515 | 326.2972 | -0.16012 | -3.0086 | 0.33709 | 3.076 | 321.4515 | 3.0086 | 94.1707
1.0147 | 4.989 | 55.595 | -336.7302 | 341.6012 | -0.25282 | -3.2445 | 0.31181 | 3.3035 | 336.7302 | 3.2445 | 94.338
1.0165 | 7.3908 | 56.7748 | -351.7561 | 356.6523 | -0.35395 | -3.4636 | 0.28653 | 3.531 | 351.7561 | 3.4636 | 94.5054
1.0182 | 9.6746 | 57.8956 | -366.4196 | 371.3496 | -0.44665 | -3.6827 | 0.26125 | 3.7502 | 366.4196 | 3.6827 | 94.6634
1.02 | 11.8488 | 58.8732 | -381.3275 | 386.2828 | -0.53935 | -3.885 | 0.23596 | 3.9608 | 381.3275 | 3.885 | 94.8308
1.0218 | 13.9893 | 59.7244 | -396.4293 | 401.3929 | -0.64048 | -4.0872 | 0.21068 | 4.1715 | 396.4293 | 4.0872 | 94.9981
1.0236 | 16.1636 | 60.3227 | -411.8344 | 416.7728 | -0.7416 | -4.2811 | 0.1854 | 4.3822 | 411.8344 | 4.2811 | 95.1655
1.0254 | 18.3126 | 60.4575 | -427.5513 | 432.4139 | -0.85116 | -4.4749 | 0.16012 | 4.5845 | 427.5513 | 4.4749 | 95.3328
1.0272 | 20.4531 | 60.1879 | -443.479 | 448.2404 | -0.96071 | -4.6519 | 0.12641 | 4.7783 | 443.479 | 4.6519 | 95.5002
1.029 | 22.6358 | 59.598 | -459.5751 | 464.1933 | -1.0703 | -4.8204 | 0.10113 | 4.9721 | 459.5751 | 4.8204 | 95.6675
1.0308 | 24.8521 | 58.8226 | -475.941 | 480.4074 | -1.1883 | -4.9805 | 0.067419 | 5.1575 | 475.941 | 4.9805 | 95.8349
1.0326 | 27.0601 | 57.7945 | -491.8096 | 496.1075 | -1.2978 | -5.1407 | 0.042137 | 5.3345 | 491.8096 | 5.1407 | 96.0022
1.0344 | 29.1585 | 56.3703 | -506.6501 | 510.7879 | -1.3989 | -5.2839 | 0.016855 | 5.503 | 506.6501 | 5.2839 | 96.1696
1.0362 | 31.1811 | 54.6427 | -520.9007 | 524.8699 | -1.4916 | -5.4188 | -0.016855 | 5.6632 | 520.9007 | 5.4188 | 96.3369
1.038 | 33.1531 | 52.7297 | -534.0894 | 537.8986 | -1.5843 | -5.5452 | -0.033709 | 5.8064 | 534.0894 | 5.5452 | 96.5043
1.0398 | 35.0829 | 50.842 | -545.7528 | 549.4271 | -1.6433 | -5.6716 | -0.058991 | 5.9413 | 545.7528 | 5.6716 | 96.6716
1.0416 | 36.7431 | 49.2576 | -555.7813 | 559.3545 | -1.6602 | -5.7811 | -0.067419 | 6.0592 | 555.7813 | 5.7811 | 96.839
1.0433 | 38.1083 | 47.9093 | -564.6047 | 568.1021 | -1.6602 | -5.8823 | -0.0927 | 6.1519 | 564.6047 | 5.8823 | 96.997
1.0451 | 39.305 | 46.6789 | -572.3241 | 575.7541 | -1.6433 | -5.9665 | -0.10956 | 6.2278 | 572.3241 | 5.9665 | 97.1644
1.0469 | 40.5017 | 45.4906 | -579.302 | 582.6729 | -1.6265 | -6.0508 | -0.12641 | 6.3036 | 579.302 | 6.0508 | 97.3317
1.0487 | 41.4792 | 44.6732 | -585.6562 | 588.9849 | -1.5843 | -6.1267 | -0.14326 | 6.3626 | 585.6562 | 6.1267 | 97.4991
1.0505 | 42.3473 | 43.9737 | -591.3783 | 594.6734 | -1.5253 | -6.1856 | -0.16855 | 6.4132 | 591.3783 | 6.1856 | 97.6664
1.0523 | 43.1394 | 43.1647 | -596.3083 | 599.5612 | -1.4748 | -6.2446 | -0.17697 | 6.4553 | 596.3083 | 6.2446 | 97.8338
1.0541 | 43.7546 | 42.3388 | -600.7073 | 603.9181 | -1.4074 | -6.2868 | -0.19383 | 6.4806 | 600.7073 | 6.2868 | 98.0011
1.0559 | 44.1928 | 41.4708 | -604.4406 | 607.6009 | -1.3484 | -6.3205 | -0.21068 | 6.4975 | 604.4406 | 6.3205 | 98.1685
1.0577 | 44.3361 | 40.6534 | -607.0784 | 610.1796 | -1.2725 | -6.3373 | -0.21911 | 6.4975 | 607.0784 | 6.3373 | 98.3358
1.0595 | 44.058 | 39.8106 | -608.1487 | 611.1656 | -1.1798 | -6.3458 | -0.21911 | 6.489 | 608.1487 | 6.3458 | 98.5032
1.0613 | 43.6451 | 38.8583 | -608.0728 | 610.9971 | -1.0956 | -6.3542 | -0.21068 | 6.4806 | 608.0728 | 6.3542 | 98.6705
1.0631 | 43.1731 | 37.7796 | -606.9267 | 609.7583 | -1.0113 | -6.3542 | -0.20226 | 6.4637 | 606.9267 | 6.3542 | 98.8379
1.0649 | 42.4147 | 36.5661 | -604.5165 | 607.2301 | -0.91858 | -6.3373 | -0.17697 | 6.43 | 604.5165 | 6.3373 | 99.0052
1.0667 | 41.5129 | 35.1588 | -601.9714 | 604.5586 | -0.82588 | -6.3121 | -0.15169 | 6.3963 | 601.9714 | 6.3121 | 99.1726
1.0684 | 40.5691 | 33.5239 | -599.991 | 602.4434 | -0.73318 | -6.2952 | -0.11798 | 6.371 | 599.991 | 6.2952 | 99.3306
1.0702 | 39.5662 | 31.8131 | -598.8702 | 601.1793 | -0.6489 | -6.2783 | -0.084273 | 6.3458 | 598.8702 | 6.2783 | 99.498
1.072 | 38.6224 | 30.1529 | -598.6595 | 600.8338 | -0.5562 | -6.2615 | -0.058991 | 6.3205 | 598.6595 | 6.2615 | 99.6653
1.0738 | 37.8471 | 28.5349 | -599.182 | 601.2214 | -0.47193 | -6.2446 | -0.033709 | 6.3036 | 599.182 | 6.2446 | 99.8327
1.0756 | 37.2403 | 27.1865 | -599.6792 | 601.6091 | -0.40451 | -6.2446 | -0.0084273 | 6.2868 | 599.6792 | 6.2446 | 100
| Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My
Minimum | -26.664 | -241.2824 | -1958.6931 | 201.7583 | -1.6602 | -30.9788 | -5.9918 | 0.88487 | 197.966 | -4.0957
Maximum | 91.4532 | 94.7989 | -197.966 | 1960.21 | 14.5287 | 4.0957 | 5.3092 | 33.1278 | 1958.6931 | 30.9788

## Diagr_AVER75
AVER75 = Loads at Average Level in Subject with 75 kg Body Weight.
Contact forces and moments, measured telemetrically in knee implants of 8 subjects.
Averages of one individual subject.
Fx  lateral, Fy  anterior, -Fz  distal, Fres = Resultant Force
Mx around x, My around y, Mz around z, Mres = Resultant Moment
z-axis = axis of implant stem!
From Bergmann et al, PlosOne (2014) "Standardized Loads Acting in Knee Implants"
Subject K2L
Subject K2L
G.Bergmann et al., PlosOne, 2014

## Dat_HIGH100
Data File: | k2l_2242_0_21_i1.akf
Curve Alignment Program: | 1.5.7
------------------------MEASURING INFOS------------------------
Diagram Title #1 | Forces and Moments at Knee Joint
Diagram Title #2 | Measured Relative to Implant
Comment #1 | Knee Joint; Gaitanalysis; Level Walking;
Comment #2 | K2L, 45,5 Months PO *
BodyWeight [N]: | 902.5
Patient(en): | k2l | [7 empty] | Subject | k2l
Implant Type: | Knee Joint
Varus Angle Leg [°] | 5 | [7 empty] | HIGH100 Loads in N (forces) and Nm (moments)
Axes Directions X, Y, Z: | Lateral, Anterior, Superior
Axes Origin: | Tibial Plateau at Stem Axis | [7 empty] | = AVER75 Loads * 1/Caver
Transformation X Y Z (°): | 11 | -0.5 | 0.5
Displacement X, Y, Z (mm): | 0 | 0 | 0 | [5 empty] | Caver = | 0.58
Activity Code: | KneeActivities | 1P_1_15_2
---------------------------DATA--------------------------------
Number of Data Sets = | 601
Number of analog channels = | 0
Time-Offset(sec) = | 0
Time | Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My | Time
[s] | [N] | [N] | [N] | [N] | [Nm] | [Nm] | [Nm] | [Nm] | [N] | (Nm] | [%Cycle]
0 | 46.5391 | 15.4016 | -1050.3339 | 1052.2663 | 0.46496 | -9.9675 | 0.47949 | 10.1128 | 1050.3339 | 9.9675 | 0
0.0018 | 44.3306 | 11.2461 | -1053.9082 | 1055.768 | 0.62478 | -9.9675 | 0.58119 | 10.1418 | 1053.9082 | 9.9675 | 0.16735
0.0036 | 41.6135 | 6.553 | -1058.2236 | 1059.9672 | 0.81367 | -9.9675 | 0.6829 | 10.1418 | 1058.2236 | 9.9675 | 0.3347
0.0054 | 38.5768 | 1.7872 | -1063.2364 | 1064.8783 | 1.0171 | -9.9384 | 0.79914 | 10.1564 | 1063.2364 | 9.9384 | 0.50205
0.0072 | 35.54 | -3.7051 | -1070.2252 | 1071.8381 | 1.235 | -9.9384 | 0.92991 | 10.1854 | 1070.2252 | 9.9384 | 0.66939
0.009 | 32.7939 | -9.3863 | -1078.1149 | 1079.6842 | 1.4385 | -9.9384 | 1.0461 | 10.2145 | 1078.1149 | 9.9384 | 0.83674
0.0108 | 29.8588 | -15.2854 | -1087.3705 | 1088.9397 | 1.6564 | -9.9675 | 1.1914 | 10.3017 | 1087.3705 | 9.9675 | 1.0041
0.0125 | 26.6768 | -21.8238 | -1098.1226 | 1099.7208 | 1.9034 | -9.982 | 1.3513 | 10.3743 | 1098.1226 | 9.982 | 1.1621
0.0143 | 22.9862 | -28.3187 | -1109.3251 | 1110.996 | 2.1649 | -10.0256 | 1.5111 | 10.4906 | 1109.3251 | 10.0256 | 1.3295
0.0161 | 19.1503 | -34.3922 | -1120.1643 | 1121.9079 | 2.4701 | -10.0982 | 1.7 | 10.6649 | 1120.1643 | 10.0982 | 1.4968
0.0179 | 15.4888 | -41.0178 | -1133.401 | 1135.2899 | 2.7752 | -10.2 | 1.8889 | 10.8683 | 1133.401 | 10.2 | 1.6642
0.0197 | 12.2487 | -48.3844 | -1149.6163 | 1151.6796 | 3.0513 | -10.2871 | 2.0778 | 11.0572 | 1149.6163 | 10.2871 | 1.8315
0.0215 | 9.2846 | -56.1143 | -1168.447 | 1170.7573 | 3.3419 | -10.4034 | 2.2521 | 11.2752 | 1168.447 | 10.4034 | 1.9989
0.0233 | 7.3666 | -63.6698 | -1191.5931 | 1194.1503 | 3.6615 | -10.5777 | 2.412 | 11.5512 | 1191.5931 | 10.5777 | 2.1662
0.0251 | 6.5094 | -70.8911 | -1219.2143 | 1222.0912 | 3.9957 | -10.8102 | 2.5427 | 11.8854 | 1219.2143 | 10.8102 | 2.3336
0.0269 | 6.7709 | -78.3595 | -1252.5749 | 1255.815 | 4.2863 | -11.0717 | 2.6299 | 12.2487 | 1252.5749 | 11.0717 | 2.5009
0.0287 | 9.0957 | -86.5834 | -1297.7627 | 1301.4097 | 4.5043 | -11.3769 | 2.6444 | 12.5974 | 1297.7627 | 11.3769 | 2.6683
0.0305 | 12.5683 | -94.8509 | -1349.9539 | 1354.0368 | 4.7077 | -11.7111 | 2.6154 | 12.9752 | 1349.9539 | 11.7111 | 2.8356
0.0323 | 15.9102 | -102.5226 | -1403.9904 | 1408.4656 | 4.9111 | -12.0452 | 2.5718 | 13.3239 | 1403.9904 | 12.0452 | 3.003
0.0341 | 19.9495 | -111.6764 | -1459.6979 | 1464.7107 | 5.0709 | -12.3504 | 2.5137 | 13.6581 | 1459.6979 | 12.3504 | 3.1703
0.0359 | 24.5119 | -122.0653 | -1517.0036 | 1522.7429 | 5.2453 | -12.641 | 2.4846 | 13.9777 | 1517.0036 | 12.641 | 3.3377
0.0376 | 30.0041 | -133.8054 | -1573.0598 | 1579.6563 | 5.4196 | -12.9606 | 2.4701 | 14.3119 | 1573.0598 | 12.9606 | 3.4957
0.0394 | 35.7434 | -145.9378 | -1626.1519 | 1633.7219 | 5.6231 | -13.2658 | 2.4846 | 14.6606 | 1626.1519 | 13.2658 | 3.6631
0.0412 | 41.0904 | -157.6779 | -1673.0542 | 1681.5833 | 5.8555 | -13.5418 | 2.5282 | 15.0239 | 1673.0542 | 13.5418 | 3.8304
0.043 | 45.6528 | -168.9967 | -1712.866 | 1722.3395 | 6.1607 | -13.8034 | 2.6154 | 15.4016 | 1712.866 | 13.8034 | 3.9978
0.0448 | 48.7622 | -179.3274 | -1743.9308 | 1754.3633 | 6.5239 | -14.0213 | 2.7316 | 15.7794 | 1743.9308 | 14.0213 | 4.1651
0.0466 | 50.3314 | -188.394 | -1766.5974 | 1777.9452 | 6.9743 | -14.1957 | 2.906 | 16.1717 | 1766.5974 | 14.1957 | 4.3325
0.0484 | 50.5929 | -196.3709 | -1783.7136 | 1795.846 | 7.4538 | -14.3555 | 3.0658 | 16.5786 | 1783.7136 | 14.3555 | 4.4998
0.0502 | 50.4041 | -203.8538 | -1796.2528 | 1809.1408 | 7.9624 | -14.5153 | 3.2692 | 16.9999 | 1796.2528 | 14.5153 | 4.6672
0.052 | 49.881 | -211.1187 | -1806.6707 | 1820.2852 | 8.4564 | -14.6897 | 3.4726 | 17.4213 | 1806.6707 | 14.6897 | 4.8345
0.0538 | 48.7622 | -218.1657 | -1815.2869 | 1829.6424 | 8.9794 | -14.864 | 3.6761 | 17.8862 | 1815.2869 | 14.864 | 5.0019
0.0556 | 47.0767 | -224.3264 | -1822.2467 | 1837.2415 | 9.5025 | -15.0965 | 3.894 | 18.4093 | 1822.2467 | 15.0965 | 5.1692
0.0574 | 44.9408 | -230.0221 | -1826.7655 | 1842.3851 | 10.0401 | -15.3581 | 4.1265 | 18.9615 | 1826.7655 | 15.3581 | 5.3366
0.0592 | 42.8049 | -235.8776 | -1828.6689 | 1844.986 | 10.5923 | -15.6486 | 4.388 | 19.5572 | 1828.6689 | 15.6486 | 5.5039
0.061 | 40.8289 | -241.8639 | -1829.7151 | 1846.7731 | 11.1299 | -15.9392 | 4.6641 | 20.1529 | 1829.7151 | 15.9392 | 5.6713
0.0627 | 38.5186 | -247.6758 | -1829.8459 | 1847.6304 | 11.682 | -16.2298 | 4.9401 | 20.7632 | 1829.8459 | 16.2298 | 5.8293
0.0645 | 35.5255 | -252.9211 | -1827.8262 | 1846.2791 | 12.2196 | -16.5495 | 5.2162 | 21.3734 | 1827.8262 | 16.5495 | 5.9967
0.0663 | 32.4742 | -258.2536 | -1826.5185 | 1845.6543 | 12.7281 | -16.8982 | 5.4923 | 22.0127 | 1826.5185 | 16.8982 | 6.164
0.0681 | 29.5973 | -264.0946 | -1827.8262 | 1847.7176 | 13.2367 | -17.3196 | 5.7538 | 22.6811 | 1827.8262 | 17.3196 | 6.3313
0.0699 | 26.8657 | -270.4151 | -1831.7783 | 1852.4834 | 13.7452 | -17.7555 | 6.0154 | 23.3931 | 1831.7783 | 17.7555 | 6.4987
0.0717 | 24.2649 | -276.9535 | -1837.5176 | 1859.109 | 14.2538 | -18.2204 | 6.2769 | 24.1196 | 1837.5176 | 18.2204 | 6.666
0.0735 | 21.9691 | -283.5065 | -1844.6663 | 1867.1004 | 14.7333 | -18.758 | 6.5094 | 24.8606 | 1844.6663 | 18.758 | 6.8334
0.0753 | 20.1384 | -290.3064 | -1852.7885 | 1876.1234 | 15.1692 | -19.3538 | 6.7564 | 25.6452 | 1852.7885 | 19.3538 | 7.0007
0.0771 | 18.8016 | -297.397 | -1862.1602 | 1886.4396 | 15.6051 | -20.0076 | 6.9889 | 26.4589 | 1862.1602 | 20.0076 | 7.1681
0.0789 | 17.8717 | -304.6183 | -1873.3337 | 1898.6011 | 16.0555 | -20.6469 | 7.2213 | 27.2725 | 1873.3337 | 20.6469 | 7.3354
0.0807 | 17.2905 | -311.8542 | -1885.873 | 1912.1139 | 16.5059 | -21.3153 | 7.4538 | 28.1153 | 1885.873 | 21.3153 | 7.5028
0.0825 | 16.9709 | -318.9593 | -1900.7225 | 1927.9078 | 16.9418 | -22.0273 | 7.6572 | 28.9725 | 1900.7225 | 22.0273 | 7.6701
0.0843 | 17.1307 | -325.9481 | -1917.8386 | 1945.9103 | 17.3632 | -22.7828 | 7.8606 | 29.8588 | 1917.8386 | 22.7828 | 7.8375
0.086 | 17.8281 | -332.9225 | -1936.4078 | 1965.3803 | 17.7845 | -23.5384 | 8.0495 | 30.7306 | 1936.4078 | 23.5384 | 7.9955
0.0878 | 18.9179 | -339.8532 | -1955.8632 | 1985.7366 | 18.1914 | -24.3084 | 8.2239 | 31.617 | 1955.8632 | 24.3084 | 8.1629
0.0896 | 20.5597 | -346.7694 | -1975.9144 | 2006.6887 | 18.5982 | -25.1076 | 8.3983 | 32.5178 | 1975.9144 | 25.1076 | 8.3302
0.0914 | 22.87 | -353.6711 | -1997.099 | 2028.7595 | 18.976 | -25.9648 | 8.5726 | 33.4332 | 1997.099 | 25.9648 | 8.4976
0.0932 | 25.8486 | -360.5147 | -2020.8117 | 2053.3295 | 19.3392 | -26.8221 | 8.7179 | 34.3486 | 2020.8117 | 26.8221 | 8.6649
0.095 | 29.2631 | -367.0967 | -2046.7765 | 2080.0789 | 19.6879 | -27.6794 | 8.8341 | 35.2494 | 2046.7765 | 27.6794 | 8.8323
0.0968 | 33.5349 | -373.4462 | -2075.0807 | 2109.1096 | 20.0076 | -28.5802 | 8.9359 | 36.1793 | 2075.0807 | 28.5802 | 8.9996
0.0986 | 38.4024 | -379.2582 | -2104.7361 | 2139.3898 | 20.3273 | -29.4665 | 9.023 | 37.0947 | 2104.7361 | 29.4665 | 9.167
0.1004 | 43.4152 | -384.6342 | -2135.9026 | 2171.1085 | 20.6614 | -30.3529 | 9.0957 | 38.0101 | 2135.9026 | 30.3529 | 9.3343
0.1022 | 48.7912 | -389.3709 | -2168.6384 | 2204.2801 | 20.9811 | -31.2101 | 9.1393 | 38.8964 | 2168.6384 | 31.2101 | 9.5017
0.104 | 54.5741 | -393.541 | -2203.0596 | 2239.05 | 21.2862 | -32.0674 | 9.1538 | 39.7537 | 2203.0596 | 32.0674 | 9.669
0.1058 | 60.6476 | -397.1299 | -2239.8928 | 2276.1157 | 21.5768 | -32.8956 | 9.1393 | 40.5819 | 2239.8928 | 32.8956 | 9.8364
0.1076 | 66.9826 | -399.876 | -2278.1789 | 2314.5036 | 21.8529 | -33.6947 | 9.0957 | 41.381 | 2278.1789 | 33.6947 | 10.0037
0.1094 | 73.1287 | -402.4188 | -2314.983 | 2351.4094 | 22.1435 | -34.4939 | 9.0666 | 42.1802 | 2314.983 | 34.4939 | 10.1711
0.1111 | 78.9262 | -405.0341 | -2348.9974 | 2385.5545 | 22.4196 | -35.2494 | 9.0376 | 42.9502 | 2348.9974 | 35.2494 | 10.3291
0.1129 | 85.0868 | -407.2717 | -2382.0674 | 2418.7262 | 22.6956 | -36.005 | 9.023 | 43.7058 | 2382.0674 | 36.005 | 10.4965
0.1147 | 91.3201 | -409.1025 | -2414.1202 | 2450.8226 | 22.9281 | -36.7169 | 8.9794 | 44.4178 | 2414.1202 | 36.7169 | 10.6638
0.1165 | 97.4372 | -410.7734 | -2445.3449 | 2482.1054 | 23.1606 | -37.4144 | 8.9504 | 45.1007 | 2445.3449 | 37.4144 | 10.8312
0.1183 | 103.3218 | -412.1973 | -2474.7969 | 2511.601 | 23.3785 | -38.0537 | 8.9068 | 45.7254 | 2474.7969 | 38.0537 | 10.9985
0.1201 | 108.6397 | -413.5341 | -2503.2318 | 2540.0504 | 23.5965 | -38.6494 | 8.8632 | 46.3357 | 2503.2318 | 38.6494 | 11.1659
0.1219 | 113.6961 | -414.4495 | -2530.6932 | 2567.4973 | 23.7999 | -39.2161 | 8.8196 | 46.9024 | 2530.6932 | 39.2161 | 11.3332
0.1237 | 118.7961 | -415.1324 | -2557.0358 | 2593.8109 | 23.9743 | -39.7682 | 8.7615 | 47.4254 | 2557.0358 | 39.7682 | 11.5006
0.1255 | 123.7653 | -415.5973 | -2581.9981 | 2618.7151 | 24.1196 | -40.2767 | 8.7034 | 47.9194 | 2581.9981 | 40.2767 | 11.6679
0.1273 | 128.6764 | -415.9751 | -2606.8732 | 2643.503 | 24.2794 | -40.7853 | 8.6453 | 48.3989 | 2606.8732 | 40.7853 | 11.8353
0.1291 | 133.3986 | -416.0042 | -2633.2594 | 2669.7294 | 24.4392 | -41.2793 | 8.5581 | 48.8784 | 2633.2594 | 41.2793 | 12.0026
0.1309 | 138.0191 | -415.699 | -2660.6337 | 2696.8711 | 24.6281 | -41.7588 | 8.4854 | 49.3579 | 2660.6337 | 41.7588 | 12.17
0.1327 | 142.2618 | -414.7837 | -2685.6686 | 2721.6155 | 24.7734 | -42.1656 | 8.3983 | 49.7647 | 2685.6686 | 42.1656 | 12.3373
0.1345 | 145.7635 | -414.0862 | -2706.0976 | 2741.8265 | 24.8751 | -42.4998 | 8.3256 | 50.0844 | 2706.0976 | 42.4998 | 12.5046
0.1362 | 148.437 | -413.7956 | -2724.1727 | 2759.7563 | 24.9478 | -42.7323 | 8.253 | 50.3023 | 2724.1727 | 42.7323 | 12.6627
0.138 | 150.8199 | -413.6358 | -2740.3735 | 2775.8409 | 24.9913 | -42.9357 | 8.1803 | 50.4912 | 2740.3735 | 42.9357 | 12.83
0.1398 | 152.9267 | -413.5486 | -2754.5837 | 2789.9639 | 25.0204 | -43.1101 | 8.1222 | 50.6365 | 2754.5837 | 43.1101 | 12.9974
0.1416 | 154.7139 | -413.4469 | -2766.6435 | 2801.951 | 25.0349 | -43.2699 | 8.0786 | 50.7818 | 2766.6435 | 43.2699 | 13.1647
0.1434 | 156.0506 | -413.1708 | -2776.9597 | 2812.1655 | 25.0495 | -43.4007 | 8.035 | 50.8835 | 2776.9597 | 43.4007 | 13.3321
0.1452 | 157.0532 | -412.6478 | -2786.4767 | 2821.5372 | 25.0495 | -43.5169 | 7.9914 | 50.9852 | 2786.4767 | 43.5169 | 13.4994
0.147 | 157.5617 | -412.023 | -2795.2382 | 2830.1534 | 25.0495 | -43.6041 | 7.9333 | 51.0434 | 2795.2382 | 43.6041 | 13.6668
0.1488 | 157.6779 | -411.558 | -2803.4766 | 2838.232 | 25.0495 | -43.6622 | 7.8897 | 51.087 | 2803.4766 | 43.6622 | 13.8341
0.1506 | 157.4745 | -411.3691 | -2811.2501 | 2845.8602 | 25.0349 | -43.6913 | 7.8316 | 51.1015 | 2811.2501 | 43.6913 | 14.0015
0.1524 | 157.0822 | -410.817 | -2817.4979 | 2851.9482 | 25.0204 | -43.7203 | 7.7735 | 51.1015 | 2817.4979 | 43.7203 | 14.1688
0.1542 | 156.2685 | -409.9452 | -2822.6996 | 2856.9174 | 25.0059 | -43.7349 | 7.7154 | 51.1015 | 2822.6996 | 43.7349 | 14.3362
0.156 | 155.2369 | -408.6085 | -2827.1167 | 2861.0439 | 25.0059 | -43.7349 | 7.6572 | 51.1015 | 2827.1167 | 43.7349 | 14.5035
0.1578 | 154.3361 | -406.923 | -2829.7757 | 2863.3832 | 25.0059 | -43.7639 | 7.6136 | 51.116 | 2829.7757 | 43.7639 | 14.6709
0.1595 | 152.8976 | -404.8452 | -2830.9381 | 2864.1388 | 24.9913 | -43.7784 | 7.5701 | 51.116 | 2830.9381 | 43.7784 | 14.8289
0.1613 | 150.9797 | -402.6512 | -2830.909 | 2863.7174 | 24.9478 | -43.793 | 7.5265 | 51.1015 | 2830.909 | 43.793 | 14.9963
0.1631 | 148.8293 | -400.2974 | -2830.0953 | 2862.4678 | 24.8751 | -43.793 | 7.4683 | 51.0579 | 2830.0953 | 43.793 | 15.1636
0.1649 | 146.3592 | -397.5222 | -2827.7415 | 2859.649 | 24.8025 | -43.7784 | 7.4102 | 51.0143 | 2827.7415 | 43.7784 | 15.331
0.1667 | 143.6712 | -394.8051 | -2824.5159 | 2855.9585 | 24.7153 | -43.7639 | 7.3521 | 50.9417 | 2824.5159 | 43.7639 | 15.4983
0.1685 | 140.8233 | -392.059 | -2820.0988 | 2851.0764 | 24.6281 | -43.7349 | 7.294 | 50.869 | 2820.0988 | 43.7349 | 15.6657
0.1703 | 137.9173 | -389.153 | -2814.6937 | 2845.2064 | 24.5119 | -43.7058 | 7.2359 | 50.7673 | 2814.6937 | 43.7058 | 15.833
0.1721 | 135.1712 | -386.0727 | -2809.3322 | 2839.3654 | 24.3811 | -43.6913 | 7.1487 | 50.6801 | 2809.3322 | 43.6913 | 16.0004
0.1739 | 132.4105 | -382.8325 | -2803.8399 | 2833.3791 | 24.2358 | -43.6767 | 7.076 | 50.5929 | 2803.8399 | 43.6767 | 16.1677
0.1757 | 129.5482 | -379.4471 | -2797.6938 | 2826.7389 | 24.0905 | -43.6622 | 6.9889 | 50.4912 | 2797.6938 | 43.6622 | 16.3351
0.1775 | 126.9473 | -376.1633 | -2789.034 | 2817.6578 | 23.9307 | -43.6331 | 6.9307 | 50.3895 | 2789.034 | 43.6331 | 16.5024
0.1793 | 124.0995 | -372.7633 | -2777.7297 | 2805.9177 | 23.7854 | -43.5896 | 6.8726 | 50.2733 | 2777.7297 | 43.5896 | 16.6698
0.1811 | 121.455 | -369.3924 | -2764.3187 | 2792.0998 | 23.6255 | -43.546 | 6.829 | 50.1425 | 2764.3187 | 43.546 | 16.8371
0.1829 | 118.7961 | -365.7745 | -2747.7692 | 2775.158 | 23.4366 | -43.5024 | 6.7854 | 50.0117 | 2747.7692 | 43.5024 | 17.0045
0.1846 | 116.0935 | -361.8369 | -2729.0838 | 2756.0657 | 23.2478 | -43.4588 | 6.7564 | 49.8664 | 2729.0838 | 43.4588 | 17.1625
0.1864 | 113.4781 | -357.507 | -2711.9822 | 2738.4846 | 23.0589 | -43.4297 | 6.6983 | 49.7357 | 2711.9822 | 43.4297 | 17.3299
0.1882 | 110.8918 | -353.1335 | -2697.2634 | 2723.2719 | 22.8845 | -43.4007 | 6.6547 | 49.6049 | 2697.2634 | 43.4007 | 17.4972
0.19 | 108.291 | -348.9198 | -2685.2618 | 2710.7471 | 22.7102 | -43.3425 | 6.582 | 49.4741 | 2685.2618 | 43.3425 | 17.6646
0.1918 | 105.6611 | -344.9823 | -2675.1781 | 2700.1694 | 22.5649 | -43.2844 | 6.5094 | 49.3434 | 2675.1781 | 43.2844 | 17.8319
0.1936 | 102.9585 | -341.4806 | -2667.4772 | 2692.0182 | 22.4486 | -43.2408 | 6.4367 | 49.2417 | 2667.4772 | 43.2408 | 17.9993
0.1954 | 100.2705 | -338.3712 | -2662.7986 | 2686.9036 | 22.3469 | -43.1972 | 6.3641 | 49.1545 | 2662.7986 | 43.1972 | 18.1666
0.1972 | 97.6115 | -335.3635 | -2661.4328 | 2685.0874 | 22.2452 | -43.1682 | 6.2769 | 49.0673 | 2661.4328 | 43.1682 | 18.334
0.199 | 95.0252 | -332.3994 | -2661.0841 | 2684.3028 | 22.129 | -43.1391 | 6.1752 | 48.9801 | 2661.0841 | 43.1391 | 18.5013
0.2008 | 92.4099 | -329.7114 | -2662.8567 | 2685.625 | 21.9982 | -43.1101 | 6.0589 | 48.8784 | 2662.8567 | 43.1101 | 18.6687
0.2026 | 89.9252 | -327.314 | -2665.4866 | 2687.8626 | 21.8529 | -43.0665 | 5.9427 | 48.7767 | 2665.4866 | 43.0665 | 18.836
0.2044 | 87.7603 | -324.4952 | -2665.9807 | 2687.9498 | 21.7076 | -43.0229 | 5.8265 | 48.6459 | 2665.9807 | 43.0229 | 19.0033
0.2062 | 85.915 | -321.7636 | -2664.3533 | 2685.9592 | 21.5768 | -42.9648 | 5.7393 | 48.5442 | 2664.3533 | 42.9648 | 19.1707
0.208 | 84.2295 | -319.2499 | -2661.6072 | 2682.9079 | 21.4606 | -42.9212 | 5.6521 | 48.428 | 2661.6072 | 42.9212 | 19.338
0.2097 | 82.2099 | -316.62 | -2656.9722 | 2677.9387 | 21.3444 | -42.8631 | 5.594 | 48.3263 | 2656.9722 | 42.8631 | 19.4961
0.2115 | 80.1467 | -313.6559 | -2650.5645 | 2671.1823 | 21.2136 | -42.8195 | 5.5213 | 48.2246 | 2650.5645 | 42.8195 | 19.6634
0.2133 | 77.8655 | -310.5465 | -2643.2269 | 2663.467 | 21.0828 | -42.7468 | 5.4487 | 48.0938 | 2643.2269 | 42.7468 | 19.8308
0.2151 | 75.4826 | -307.3499 | -2635.744 | 2655.6209 | 20.9666 | -42.6887 | 5.3906 | 47.9776 | 2635.744 | 42.6887 | 19.9981
0.2169 | 73.0561 | -304.1388 | -2628.0287 | 2647.5277 | 20.8358 | -42.6161 | 5.3325 | 47.8613 | 2628.0287 | 42.6161 | 20.1655
0.2187 | 70.4988 | -301.1748 | -2621.0253 | 2640.1756 | 20.705 | -42.5725 | 5.2598 | 47.7451 | 2621.0253 | 42.5725 | 20.3328
0.2205 | 68.2467 | -298.632 | -2615.7655 | 2634.6252 | 20.5743 | -42.5289 | 5.1872 | 47.6434 | 2615.7655 | 42.5289 | 20.5002
0.2223 | 66.1544 | -296.1329 | -2612.0749 | 2630.6295 | 20.458 | -42.4853 | 5.1145 | 47.5562 | 2612.0749 | 42.4853 | 20.6675
0.2241 | 64.3818 | -293.4013 | -2609.0963 | 2627.3167 | 20.3273 | -42.4562 | 5.0273 | 47.4545 | 2609.0963 | 42.4562 | 20.8349
0.2259 | 62.8416 | -290.9312 | -2604.9843 | 2622.8997 | 20.211 | -42.4126 | 4.9547 | 47.3673 | 2604.9843 | 42.4126 | 21.0022
0.2277 | 61.1852 | -288.7953 | -2598.7801 | 2616.4339 | 20.0657 | -42.369 | 4.8966 | 47.2511 | 2598.7801 | 42.369 | 21.1696
0.2295 | 59.3109 | -286.3688 | -2590.4109 | 2607.7886 | 19.9059 | -42.2964 | 4.8384 | 47.1203 | 2590.4109 | 42.2964 | 21.3369
0.2313 | 57.3493 | -283.2013 | -2580.1964 | 2597.2109 | 19.7461 | -42.2237 | 4.7803 | 46.975 | 2580.1964 | 42.2237 | 21.5043
0.2331 | 55.4169 | -280.0338 | -2569.052 | 2585.7178 | 19.5717 | -42.1511 | 4.7222 | 46.8297 | 2569.052 | 42.1511 | 21.6716
0.2348 | 53.3246 | -276.7501 | -2555.5538 | 2571.8708 | 19.3828 | -42.0784 | 4.6786 | 46.6699 | 2555.5538 | 42.0784 | 21.8297
0.2366 | 51.2323 | -273.4227 | -2542.3897 | 2558.3435 | 19.2085 | -41.9913 | 4.635 | 46.5246 | 2542.3897 | 41.9913 | 21.997
0.2384 | 49.3434 | -270.6475 | -2531.2599 | 2546.9231 | 19.0341 | -41.9186 | 4.5914 | 46.3793 | 2531.2599 | 41.9186 | 22.1644
0.2402 | 47.2511 | -267.727 | -2522.5129 | 2537.8274 | 18.8743 | -41.8605 | 4.5333 | 46.234 | 2522.5129 | 41.8605 | 22.3317
0.242 | 45.0861 | -264.5741 | -2514.3471 | 2529.2984 | 18.7144 | -41.7733 | 4.4607 | 46.0887 | 2514.3471 | 41.7733 | 22.4991
0.2438 | 42.9793 | -261.5664 | -2504.5976 | 2519.2292 | 18.5546 | -41.6861 | 4.4025 | 45.9289 | 2504.5976 | 41.6861 | 22.6664
0.2456 | 40.9161 | -258.704 | -2493.6857 | 2508.0266 | 18.3948 | -41.5844 | 4.359 | 45.769 | 2493.6857 | 41.5844 | 22.8338
0.2474 | 38.7947 | -255.8707 | -2482.2362 | 2496.3011 | 18.235 | -41.4827 | 4.3008 | 45.5947 | 2482.2362 | 41.4827 | 23.0011
0.2492 | 36.7605 | -253.1245 | -2471.6875 | 2485.4763 | 18.0751 | -41.3665 | 4.2572 | 45.4348 | 2471.6875 | 41.3665 | 23.1685
0.251 | 34.9152 | -250.6981 | -2462.679 | 2476.2208 | 17.9298 | -41.2938 | 4.2137 | 45.2895 | 2462.679 | 41.2938 | 23.3358
0.2528 | 33.2298 | -248.5331 | -2454.8183 | 2468.1567 | 17.7845 | -41.2357 | 4.1701 | 45.1733 | 2454.8183 | 41.2357 | 23.5032
0.2546 | 31.5734 | -246.4989 | -2448.0038 | 2461.1243 | 17.6683 | -41.1921 | 4.1265 | 45.0861 | 2448.0038 | 41.1921 | 23.6705
0.2564 | 29.8153 | -244.3776 | -2441.4073 | 2454.3098 | 17.5521 | -41.1485 | 4.0829 | 44.9844 | 2441.4073 | 41.1485 | 23.8379
0.2581 | 28.1298 | -242.2126 | -2433.9825 | 2446.6816 | 17.4213 | -41.1049 | 4.0393 | 44.8827 | 2433.9825 | 41.1049 | 23.9959
0.2599 | 26.6332 | -240.1348 | -2427.1535 | 2439.6492 | 17.2905 | -41.0614 | 3.9957 | 44.781 | 2427.1535 | 41.0614 | 24.1633
0.2617 | 25.3982 | -238.2314 | -2421.6176 | 2433.9244 | 17.1743 | -41.0032 | 3.9376 | 44.6793 | 2421.6176 | 41.0032 | 24.3306
0.2635 | 24.1631 | -236.4297 | -2417.157 | 2429.2749 | 17.058 | -40.9597 | 3.8795 | 44.5776 | 2417.157 | 40.9597 | 24.498
0.2653 | 22.9426 | -234.5844 | -2413.5826 | 2425.4971 | 16.9418 | -40.9015 | 3.8214 | 44.4904 | 2413.5826 | 40.9015 | 24.6653
0.2671 | 21.7948 | -232.7973 | -2410.2843 | 2422.0099 | 16.8401 | -40.8579 | 3.7632 | 44.4032 | 2410.2843 | 40.8579 | 24.8327
0.2689 | 20.7341 | -231.3588 | -2406.957 | 2418.5228 | 16.7384 | -40.8144 | 3.7051 | 44.316 | 2406.957 | 40.8144 | 25
0.2707 | 19.5717 | -230.1964 | -2403.804 | 2415.2536 | 16.6221 | -40.7708 | 3.647 | 44.2143 | 2403.804 | 40.7708 | 25.1673
0.2725 | 18.5401 | -229.1939 | -2401.0869 | 2412.4493 | 16.5204 | -40.7126 | 3.6034 | 44.1272 | 2401.0869 | 40.7126 | 25.3347
0.2743 | 17.7845 | -228.3802 | -2399.169 | 2410.4442 | 16.4333 | -40.6836 | 3.5453 | 44.069 | 2399.169 | 40.6836 | 25.502
0.2761 | 17.1162 | -227.5956 | -2397.9194 | 2409.122 | 16.3316 | -40.6691 | 3.4872 | 43.9964 | 2397.9194 | 40.6691 | 25.6694
0.2779 | 16.4623 | -226.6802 | -2397.5707 | 2408.6715 | 16.2444 | -40.6545 | 3.429 | 43.9528 | 2397.5707 | 40.6545 | 25.8367
0.2797 | 15.8375 | -225.605 | -2398.1083 | 2409.0929 | 16.1572 | -40.6545 | 3.3564 | 43.9092 | 2398.1083 | 40.6545 | 26.0041
0.2815 | 15.2418 | -224.3118 | -2400.0844 | 2410.9382 | 16.0845 | -40.6545 | 3.2837 | 43.8656 | 2400.0844 | 40.6545 | 26.1714
0.2832 | 14.7623 | -223.1349 | -2403.3536 | 2414.0766 | 16.0264 | -40.6545 | 3.1966 | 43.8366 | 2403.3536 | 40.6545 | 26.3295
0.285 | 14.3555 | -222.1033 | -2407.4801 | 2418.0869 | 15.9683 | -40.6545 | 3.1094 | 43.822 | 2407.4801 | 40.6545 | 26.4968
0.2868 | 13.8615 | -220.8101 | -2411.9262 | 2422.3877 | 15.8957 | -40.6545 | 3.0222 | 43.7784 | 2411.9262 | 40.6545 | 26.6642
0.2886 | 13.3093 | -219.3572 | -2416.2271 | 2426.5142 | 15.823 | -40.6255 | 2.906 | 43.7203 | 2416.2271 | 40.6255 | 26.8315
0.2904 | 12.7427 | -217.8606 | -2419.7287 | 2429.8415 | 15.7504 | -40.5964 | 2.8043 | 43.6622 | 2419.7287 | 40.5964 | 26.9989
0.2922 | 12.394 | -216.6691 | -2420.9928 | 2430.9748 | 15.7068 | -40.5819 | 2.7316 | 43.6186 | 2420.9928 | 40.5819 | 27.1662
0.294 | 12.3068 | -215.8845 | -2420.0484 | 2429.9578 | 15.6341 | -40.5819 | 2.6735 | 43.5896 | 2420.0484 | 40.5819 | 27.3336
0.2958 | 12.3358 | -215.3179 | -2417.5638 | 2427.415 | 15.576 | -40.5819 | 2.6299 | 43.5605 | 2417.5638 | 40.5819 | 27.5009
0.2976 | 12.3649 | -215.0418 | -2414.5125 | 2424.3492 | 15.5034 | -40.5819 | 2.6008 | 43.546 | 2414.5125 | 40.5819 | 27.6683
0.2994 | 12.3794 | -214.8674 | -2412.5655 | 2422.3732 | 15.4307 | -40.5964 | 2.5718 | 43.5314 | 2412.5655 | 40.5964 | 27.8356
0.3012 | 12.3213 | -214.6204 | -2412.8125 | 2422.5911 | 15.3726 | -40.6109 | 2.5137 | 43.5169 | 2412.8125 | 40.6109 | 28.003
0.303 | 12.0888 | -214.1264 | -2415.3407 | 2425.0612 | 15.329 | -40.6255 | 2.4555 | 43.5169 | 2415.3407 | 40.6255 | 28.1703
0.3048 | 11.9871 | -213.6033 | -2418.9441 | 2428.6065 | 15.2854 | -40.6255 | 2.3829 | 43.5024 | 2418.9441 | 40.6255 | 28.3377
0.3066 | 12.0162 | -213.0367 | -2423.9279 | 2433.503 | 15.2563 | -40.6691 | 2.3102 | 43.5169 | 2423.9279 | 40.6691 | 28.505
0.3083 | 12.1324 | -212.5717 | -2430.4518 | 2439.9688 | 15.2273 | -40.7126 | 2.2376 | 43.5314 | 2430.4518 | 40.7126 | 28.6631
0.3101 | 12.3068 | -212.2666 | -2437.5569 | 2447.0158 | 15.1982 | -40.7562 | 2.1504 | 43.575 | 2437.5569 | 40.7562 | 28.8304
0.3119 | 12.7572 | -212.3538 | -2442.3663 | 2451.8107 | 15.1692 | -40.8289 | 2.0923 | 43.6186 | 2442.3663 | 40.8289 | 28.9978
0.3137 | 13.5128 | -212.8623 | -2444.9671 | 2454.4406 | 15.1546 | -40.9161 | 2.0632 | 43.7058 | 2444.9671 | 40.9161 | 29.1651
0.3155 | 14.3264 | -213.429 | -2447.7132 | 2457.2303 | 15.1401 | -41.0178 | 2.0342 | 43.793 | 2447.7132 | 41.0178 | 29.3325
0.3173 | 15.0093 | -213.7196 | -2452.3918 | 2461.9089 | 15.1256 | -41.1049 | 1.9906 | 43.8656 | 2452.3918 | 41.1049 | 29.4998
0.3191 | 15.5324 | -213.5452 | -2459.3226 | 2468.7961 | 15.111 | -41.1921 | 1.9179 | 43.9383 | 2459.3226 | 41.1921 | 29.6672
0.3209 | 16.1572 | -213.182 | -2468.3892 | 2477.8046 | 15.1401 | -41.2793 | 1.8308 | 44.0254 | 2468.3892 | 41.2793 | 29.8345
0.3227 | 17.058 | -213.0657 | -2479.1704 | 2488.5421 | 15.1546 | -41.3665 | 1.7436 | 44.1126 | 2479.1704 | 41.3665 | 30.0019
0.3245 | 18.0751 | -213.1965 | -2489.7626 | 2499.1199 | 15.1837 | -41.4537 | 1.6564 | 44.1998 | 2489.7626 | 41.4537 | 30.1692
0.3263 | 19.1358 | -213.3418 | -2499.3669 | 2508.695 | 15.2128 | -41.5408 | 1.5692 | 44.287 | 2499.3669 | 41.5408 | 30.3366
0.3281 | 20.2691 | -213.1529 | -2506.501 | 2515.8146 | 15.2563 | -41.628 | 1.5111 | 44.3887 | 2506.501 | 41.628 | 30.5039
0.3299 | 21.3008 | -213.0367 | -2512.4873 | 2521.7719 | 15.2709 | -41.7152 | 1.453 | 44.4759 | 2512.4873 | 41.7152 | 30.6713
0.3316 | 22.0999 | -212.6589 | -2518.4881 | 2527.7437 | 15.2999 | -41.7879 | 1.3803 | 44.5485 | 2518.4881 | 41.7879 | 30.8293
0.3334 | 22.6084 | -211.9324 | -2526.4505 | 2535.6188 | 15.3581 | -41.8605 | 1.3077 | 44.6357 | 2526.4505 | 41.8605 | 30.9967
0.3352 | 22.87 | -211.0025 | -2536.3163 | 2545.3684 | 15.4307 | -41.9332 | 1.2205 | 44.7229 | 2536.3163 | 41.9332 | 31.164
0.337 | 23.2042 | -210.1016 | -2546.5598 | 2555.5247 | 15.4888 | -42.0058 | 1.1333 | 44.8246 | 2546.5598 | 42.0058 | 31.3313
0.3388 | 23.6691 | -209.4769 | -2555.5829 | 2564.4751 | 15.5469 | -42.093 | 1.0607 | 44.9118 | 2555.5829 | 42.093 | 31.4987
0.3406 | 24.323 | -209.2299 | -2562.6589 | 2571.5076 | 15.576 | -42.1656 | 1.0026 | 44.9844 | 2562.6589 | 42.1656 | 31.666
0.3424 | 24.9187 | -208.9538 | -2567.9042 | 2576.7383 | 15.6051 | -42.2237 | 0.94444 | 45.0571 | 2567.9042 | 42.2237 | 31.8334
0.3442 | 25.4418 | -208.7649 | -2571.4785 | 2580.2836 | 15.6341 | -42.2673 | 0.90085 | 45.1152 | 2571.4785 | 42.2673 | 32.0007
0.346 | 25.8922 | -208.6196 | -2574.4426 | 2583.2477 | 15.6632 | -42.3255 | 0.85726 | 45.1878 | 2574.4426 | 42.3255 | 32.1681
0.3478 | 26.2409 | -208.2854 | -2578.3366 | 2587.0981 | 15.6922 | -42.3981 | 0.81367 | 45.2605 | 2578.3366 | 42.3981 | 32.3354
0.3496 | 26.7204 | -207.9076 | -2584.3374 | 2593.0699 | 15.7358 | -42.4853 | 0.75555 | 45.3477 | 2584.3374 | 42.4853 | 32.5028
0.3514 | 27.4905 | -207.6461 | -2592.7502 | 2601.4391 | 15.7649 | -42.5725 | 0.69743 | 45.4494 | 2592.7502 | 42.5725 | 32.6701
0.3532 | 28.246 | -207.2538 | -2602.9792 | 2611.6245 | 15.823 | -42.6742 | 0.61025 | 45.5656 | 2602.9792 | 42.6742 | 32.8375
0.355 | 28.493 | -206.1495 | -2614.4142 | 2622.9578 | 15.9102 | -42.7759 | 0.5376 | 45.6819 | 2614.4142 | 42.7759 | 33.0048
0.3567 | 28.8127 | -205.1906 | -2625.8347 | 2634.2911 | 15.9828 | -42.8631 | 0.45043 | 45.7981 | 2625.8347 | 42.8631 | 33.1629
0.3585 | 29.2776 | -204.4495 | -2635.9184 | 2644.3021 | 16.041 | -42.9502 | 0.37778 | 45.8998 | 2635.9184 | 42.9502 | 33.3302
0.3603 | 29.8007 | -203.8248 | -2644.0842 | 2652.4243 | 16.0991 | -43.0374 | 0.30513 | 46.0015 | 2644.0842 | 43.0374 | 33.4976
0.3621 | 30.6144 | -203.5051 | -2651.538 | 2659.8491 | 16.1281 | -43.1246 | 0.24701 | 46.0887 | 2651.538 | 43.1246 | 33.6649
0.3639 | 31.7041 | -203.0547 | -2660.1397 | 2668.3926 | 16.1717 | -43.2263 | 0.17436 | 46.1904 | 2660.1397 | 43.2263 | 33.8323
0.3657 | 32.9101 | -202.2265 | -2670.1071 | 2678.2874 | 16.2444 | -43.3571 | 0.10171 | 46.3357 | 2670.1071 | 43.3571 | 33.9996
0.3675 | 34.0434 | -201.529 | -2678.578 | 2686.7002 | 16.3025 | -43.4733 | 0.058119 | 46.4665 | 2678.578 | 43.4733 | 34.167
0.3693 | 35.1477 | -200.8607 | -2685.9011 | 2693.9652 | 16.3461 | -43.575 | 0 | 46.5827 | 2685.9011 | 43.575 | 34.3343
0.3711 | 36.252 | -200.1342 | -2691.8583 | 2699.8643 | 16.4042 | -43.6767 | -0.058119 | 46.6989 | 2691.8583 | 43.6767 | 34.5017
0.3729 | 37.1528 | -199.5239 | -2696.1301 | 2704.0779 | 16.4333 | -43.7349 | -0.10171 | 46.7716 | 2696.1301 | 43.7349 | 34.669
0.3747 | 38.0392 | -199.0735 | -2699.3412 | 2707.26 | 16.4769 | -43.8075 | -0.1453 | 46.8442 | 2699.3412 | 43.8075 | 34.8364
0.3765 | 38.8819 | -198.9573 | -2701.3608 | 2709.2796 | 16.5059 | -43.8511 | -0.17436 | 46.9024 | 2701.3608 | 43.8511 | 35.0037
0.3783 | 39.7246 | -198.8701 | -2702.407 | 2710.3258 | 16.535 | -43.9092 | -0.18889 | 46.9605 | 2702.407 | 43.9092 | 35.1711
0.3801 | 40.6255 | -198.3906 | -2702.5959 | 2710.5001 | 16.5931 | -43.9819 | -0.18889 | 47.0622 | 2702.5959 | 43.9819 | 35.3384
0.3818 | 41.4973 | -197.7222 | -2702.7557 | 2710.6309 | 16.6367 | -44.069 | -0.20342 | 47.1494 | 2702.7557 | 44.069 | 35.4965
0.3836 | 42.1366 | -196.8504 | -2704.2959 | 2712.1129 | 16.6803 | -44.1417 | -0.23248 | 47.2365 | 2704.2959 | 44.1417 | 35.6638
0.3854 | 42.7032 | -196.1385 | -2707.4198 | 2715.2078 | 16.7239 | -44.2143 | -0.26154 | 47.3237 | 2707.4198 | 44.2143 | 35.8312
0.3872 | 43.1101 | -195.2812 | -2711.4155 | 2719.1309 | 16.7674 | -44.3015 | -0.30513 | 47.4109 | 2711.4155 | 44.3015 | 35.9985
0.389 | 43.5024 | -194.4239 | -2715.6873 | 2723.3445 | 16.8401 | -44.3887 | -0.34872 | 47.5126 | 2715.6873 | 44.3887 | 36.1659
0.3908 | 44.1562 | -193.494 | -2720.0462 | 2727.6308 | 16.8982 | -44.4759 | -0.40684 | 47.6143 | 2720.0462 | 44.4759 | 36.3332
0.3926 | 45.0571 | -192.6513 | -2724.7539 | 2732.2804 | 16.9418 | -44.5631 | -0.45043 | 47.7306 | 2724.7539 | 44.5631 | 36.5006
0.3944 | 46.0596 | -191.9393 | -2730.4205 | 2737.9034 | 16.9854 | -44.6793 | -0.47949 | 47.8468 | 2730.4205 | 44.6793 | 36.6679
0.3962 | 46.9459 | -191.1257 | -2736.9009 | 2744.3256 | 17.029 | -44.7955 | -0.52307 | 47.9776 | 2736.9009 | 44.7955 | 36.8353
0.398 | 47.6143 | -190.0795 | -2744.0931 | 2751.4307 | 17.0871 | -44.8972 | -0.58119 | 48.0938 | 2744.0931 | 44.8972 | 37.0026
0.3998 | 48.3118 | -188.9752 | -2750.8205 | 2758.0999 | 17.1598 | -45.028 | -0.62478 | 48.2391 | 2750.8205 | 45.028 | 37.17
0.4016 | 48.9801 | -187.7402 | -2757.8965 | 2765.0888 | 17.2179 | -45.1733 | -0.66837 | 48.3989 | 2757.8965 | 45.1733 | 37.3373
0.4034 | 49.634 | -186.5197 | -2767.37 | 2774.4605 | 17.276 | -45.3186 | -0.72649 | 48.5588 | 2767.37 | 45.3186 | 37.5046
0.4051 | 50.3895 | -185.1103 | -2778.3981 | 2785.4015 | 17.3341 | -45.4784 | -0.78461 | 48.7331 | 2778.3981 | 45.4784 | 37.6627
0.4069 | 51.087 | -183.7445 | -2789.9493 | 2796.851 | 17.3922 | -45.6092 | -0.85726 | 48.8784 | 2789.9493 | 45.6092 | 37.83
0.4087 | 51.8425 | -182.2915 | -2802.5031 | 2809.3176 | 17.4504 | -45.769 | -0.92991 | 49.0528 | 2802.5031 | 45.769 | 37.9974
0.4105 | 52.7434 | -180.5625 | -2815.827 | 2822.5398 | 17.523 | -45.9434 | -1.0171 | 49.2417 | 2815.827 | 45.9434 | 38.1647
0.4123 | 53.6878 | -179.1095 | -2828.0176 | 2834.6141 | 17.5811 | -46.1032 | -1.0897 | 49.416 | 2828.0176 | 46.1032 | 38.3321
0.4141 | 54.7485 | -178.0488 | -2839.0021 | 2845.526 | 17.6392 | -46.263 | -1.1479 | 49.5759 | 2839.0021 | 46.263 | 38.4994
0.4159 | 55.9835 | -177.1189 | -2850.0884 | 2856.5687 | 17.6683 | -46.4229 | -1.2205 | 49.7502 | 2850.0884 | 46.4229 | 38.6668
0.4177 | 57.2186 | -176.1309 | -2861.9884 | 2868.4105 | 17.7119 | -46.5827 | -1.2786 | 49.9246 | 2861.9884 | 46.5827 | 38.8341
0.4195 | 58.2792 | -175.0411 | -2873.3798 | 2879.7293 | 17.77 | -46.7425 | -1.3513 | 50.0844 | 2873.3798 | 46.7425 | 39.0015
0.4213 | 59.3109 | -174.0531 | -2883.3763 | 2889.6677 | 17.8427 | -46.8733 | -1.4094 | 50.2297 | 2883.3763 | 46.8733 | 39.1688
0.4231 | 60.3279 | -172.9198 | -2891.8908 | 2898.1096 | 17.9153 | -46.9895 | -1.4675 | 50.3605 | 2891.8908 | 46.9895 | 39.3362
0.4249 | 61.1416 | -171.554 | -2899.1412 | 2905.2728 | 17.988 | -47.0622 | -1.5256 | 50.4622 | 2899.1412 | 47.0622 | 39.5035
0.4267 | 61.9553 | -170.4061 | -2903.3112 | 2909.3847 | 18.0751 | -47.1494 | -1.5692 | 50.5784 | 2903.3112 | 47.1494 | 39.6709
0.4285 | 62.7254 | -169.1129 | -2906.6386 | 2912.6394 | 18.1768 | -47.2511 | -1.5983 | 50.6947 | 2906.6386 | 47.2511 | 39.8382
0.4302 | 63.3937 | -167.3112 | -2910.5181 | 2916.4462 | 18.2786 | -47.3673 | -1.6273 | 50.8399 | 2910.5181 | 47.3673 | 39.9963
0.432 | 64.2365 | -165.4659 | -2915.5599 | 2921.3719 | 18.3512 | -47.4981 | -1.6564 | 50.9998 | 2915.5599 | 47.4981 | 40.1636
0.4338 | 65.2245 | -163.6352 | -2921.706 | 2927.4308 | 18.4239 | -47.6434 | -1.6855 | 51.1596 | 2921.706 | 47.6434 | 40.331
0.4356 | 66.2416 | -161.8044 | -2930.5111 | 2936.1487 | 18.4965 | -47.8177 | -1.7291 | 51.3485 | 2930.5111 | 47.8177 | 40.4983
0.4374 | 67.2151 | -159.9301 | -2941.1906 | 2946.7264 | 18.5546 | -47.9921 | -1.7872 | 51.5374 | 2941.1906 | 47.9921 | 40.6657
0.4392 | 68.2031 | -157.9395 | -2953.5119 | 2958.946 | 18.6127 | -48.181 | -1.8453 | 51.7408 | 2953.5119 | 48.181 | 40.833
0.441 | 69.3655 | -155.978 | -2966.3999 | 2971.7468 | 18.6709 | -48.3699 | -1.9034 | 51.9442 | 2966.3999 | 48.3699 | 41.0004
0.4428 | 70.3826 | -153.8275 | -2980.1887 | 2985.4194 | 18.729 | -48.5733 | -1.9615 | 52.1476 | 2980.1887 | 48.5733 | 41.1677
0.4446 | 71.3125 | -151.6481 | -2994.5006 | 2999.6442 | 18.7726 | -48.7767 | -2.0342 | 52.3511 | 2994.5006 | 48.7767 | 41.3351
0.4464 | 72.2424 | -149.4976 | -3008.9578 | 3013.9997 | 18.8307 | -48.9656 | -2.0923 | 52.5545 | 3008.9578 | 48.9656 | 41.5024
0.4482 | 73.1723 | -147.1874 | -3023.2116 | 3028.1517 | 18.8743 | -49.1545 | -2.1649 | 52.7579 | 3023.2116 | 49.1545 | 41.6698
0.45 | 74.0877 | -144.8481 | -3037.1457 | 3041.9841 | 18.9324 | -49.3579 | -2.2231 | 52.9613 | 3037.1457 | 49.3579 | 41.8371
0.4518 | 75.0757 | -142.4797 | -3051.632 | 3056.3832 | 18.976 | -49.5468 | -2.2812 | 53.1502 | 3051.632 | 49.5468 | 42.0045
0.4536 | 76.3108 | -140.2567 | -3066.7575 | 3071.4216 | 19.0196 | -49.7212 | -2.3393 | 53.3391 | 3066.7575 | 49.7212 | 42.1718
0.4553 | 77.7638 | -138.237 | -3081.8395 | 3086.431 | 19.0632 | -49.91 | -2.3974 | 53.528 | 3081.8395 | 49.91 | 42.3299
0.4571 | 79.2167 | -136.2755 | -3096.2241 | 3100.7574 | 19.0922 | -50.0989 | -2.4555 | 53.7169 | 3096.2241 | 50.0989 | 42.4972
0.4589 | 80.7279 | -134.3139 | -3109.6642 | 3114.1394 | 19.1213 | -50.2733 | -2.4991 | 53.9057 | 3109.6642 | 50.2733 | 42.6646
0.4607 | 82.2244 | -132.396 | -3121.4915 | 3125.894 | 19.1503 | -50.4622 | -2.5282 | 54.0946 | 3121.4915 | 50.4622 | 42.8319
0.4625 | 83.6774 | -130.4926 | -3131.7205 | 3136.0649 | 19.1794 | -50.6511 | -2.5573 | 54.269 | 3131.7205 | 50.6511 | 42.9993
0.4643 | 85.1885 | -128.5747 | -3141.8333 | 3146.1486 | 19.2085 | -50.8399 | -2.5863 | 54.4579 | 3141.8333 | 50.8399 | 43.1666
0.4661 | 86.6851 | -126.6277 | -3152.7452 | 3157.017 | 19.2375 | -50.9998 | -2.6154 | 54.6177 | 3152.7452 | 50.9998 | 43.334
0.4679 | 87.9637 | -124.608 | -3165.4007 | 3169.6144 | 19.2666 | -51.1596 | -2.659 | 54.7775 | 3165.4007 | 51.1596 | 43.5013
0.4697 | 89.2278 | -122.5884 | -3178.6665 | 3182.8365 | 19.2956 | -51.3049 | -2.7171 | 54.9374 | 3178.6665 | 51.3049 | 43.6687
0.4715 | 90.4338 | -120.5978 | -3192.1647 | 3196.2621 | 19.3247 | -51.4502 | -2.7607 | 55.0827 | 3192.1647 | 51.4502 | 43.836
0.4733 | 91.6543 | -118.6798 | -3205.5757 | 3209.6296 | 19.3392 | -51.5955 | -2.8188 | 55.228 | 3205.5757 | 51.5955 | 44.0033
0.4751 | 92.7731 | -116.7038 | -3218.8706 | 3222.8663 | 19.3683 | -51.7263 | -2.8769 | 55.3587 | 3218.8706 | 51.7263 | 44.1707
0.4769 | 93.8047 | -114.6841 | -3232.2961 | 3236.2483 | 19.3828 | -51.857 | -2.935 | 55.4895 | 3232.2961 | 51.857 | 44.338
0.4786 | 94.8218 | -112.6354 | -3245.5038 | 3249.4123 | 19.4119 | -51.9878 | -3.0077 | 55.6203 | 3245.5038 | 51.9878 | 44.4961
0.4804 | 95.8244 | -110.514 | -3257.8106 | 3261.6464 | 19.4264 | -52.104 | -3.0658 | 55.751 | 3257.8106 | 52.104 | 44.6634
0.4822 | 96.8415 | -108.4799 | -3269.3182 | 3273.1105 | 19.4409 | -52.2203 | -3.1094 | 55.8673 | 3269.3182 | 52.2203 | 44.8308
0.484 | 97.8731 | -106.5329 | -3280.1865 | 3283.9207 | 19.4555 | -52.3365 | -3.1675 | 55.969 | 3280.1865 | 52.3365 | 44.9981
0.4858 | 98.8902 | -104.673 | -3290.4446 | 3294.1497 | 19.47 | -52.4382 | -3.2111 | 56.0852 | 3290.4446 | 52.4382 | 45.1655
0.4876 | 99.8782 | -102.9004 | -3300.0924 | 3303.7539 | 19.4845 | -52.5399 | -3.2547 | 56.1869 | 3300.0924 | 52.5399 | 45.3328
0.4894 | 100.779 | -101.1714 | -3308.7667 | 3312.3992 | 19.4845 | -52.6417 | -3.2837 | 56.2741 | 3308.7667 | 52.6417 | 45.5002
0.4912 | 101.6218 | -99.3697 | -3316.1624 | 3319.7368 | 19.47 | -52.7434 | -3.3273 | 56.3613 | 3316.1624 | 52.7434 | 45.6675
0.493 | 102.3628 | -97.4953 | -3323.2966 | 3326.8419 | 19.47 | -52.8305 | -3.3564 | 56.4485 | 3323.2966 | 52.8305 | 45.8349
0.4948 | 102.944 | -95.5338 | -3330.5179 | 3334.0196 | 19.4555 | -52.9177 | -3.4 | 56.5357 | 3330.5179 | 52.9177 | 46.0022
0.4966 | 103.438 | -93.5577 | -3337.5068 | 3340.9358 | 19.4555 | -52.9904 | -3.4436 | 56.6083 | 3337.5068 | 52.9904 | 46.1696
0.4984 | 104.15 | -91.509 | -3343.0427 | 3346.4281 | 19.47 | -53.0775 | -3.4872 | 56.6809 | 3343.0427 | 53.0775 | 46.3369
0.5002 | 105.4867 | -89.4167 | -3347.7649 | 3351.1358 | 19.5136 | -53.1647 | -3.5017 | 56.7827 | 3347.7649 | 53.1647 | 46.5043
0.502 | 106.3149 | -86.7868 | -3352.4144 | 3355.7127 | 19.5862 | -53.2228 | -3.5453 | 56.8698 | 3352.4144 | 53.2228 | 46.6716
0.5037 | 106.5765 | -84.2005 | -3357.7323 | 3360.9434 | 19.6298 | -53.2664 | -3.5889 | 56.928 | 3357.7323 | 53.2664 | 46.8297
0.5055 | 106.5765 | -81.7595 | -3362.8904 | 3366.0434 | 19.6444 | -53.31 | -3.647 | 56.9715 | 3362.8904 | 53.31 | 46.997
0.5073 | 106.2423 | -79.3039 | -3367.0169 | 3370.0827 | 19.6589 | -53.3391 | -3.7051 | 57.0151 | 3367.0169 | 53.3391 | 47.1644
0.5091 | 105.7483 | -76.6159 | -3370.446 | 3373.4391 | 19.6589 | -53.3681 | -3.7632 | 57.0442 | 3370.446 | 53.3681 | 47.3317
0.5109 | 105.1961 | -73.9569 | -3373.2502 | 3376.1707 | 19.6734 | -53.3972 | -3.8359 | 57.0733 | 3373.2502 | 53.3972 | 47.4991
0.5127 | 104.5277 | -71.1963 | -3375.2989 | 3378.1322 | 19.6734 | -53.4117 | -3.9085 | 57.1023 | 3375.2989 | 53.4117 | 47.6664
0.5145 | 103.5833 | -68.2177 | -3376.563 | 3379.2946 | 19.6879 | -53.4117 | -3.9812 | 57.1168 | 3376.563 | 53.4117 | 47.8338
0.5163 | 102.45 | -65.0501 | -3377.057 | 3379.6724 | 19.6879 | -53.4117 | -4.0684 | 57.1168 | 3377.057 | 53.4117 | 48.0011
0.5181 | 101.3457 | -62.2023 | -3376.6357 | 3379.1784 | 19.6734 | -53.4117 | -4.141 | 57.1023 | 3376.6357 | 53.4117 | 48.1685
0.5199 | 100.2414 | -59.3109 | -3375.4442 | 3377.9143 | 19.6734 | -53.4117 | -4.1991 | 57.1168 | 3375.4442 | 53.4117 | 48.3358
0.5217 | 98.9047 | -55.9545 | -3373.3374 | 3375.7058 | 19.6589 | -53.4117 | -4.2572 | 57.1023 | 3373.3374 | 53.4117 | 48.5032
0.5235 | 97.2774 | -52.3656 | -3370.4895 | 3372.7998 | 19.6444 | -53.3827 | -4.3444 | 57.0878 | 3370.4895 | 53.3827 | 48.6705
0.5253 | 95.6355 | -48.8929 | -3366.9152 | 3369.1673 | 19.6153 | -53.3391 | -4.4171 | 57.0442 | 3366.9152 | 53.3391 | 48.8379
0.5271 | 93.9936 | -45.5075 | -3362.6289 | 3364.8084 | 19.5717 | -53.31 | -4.4897 | 57.0006 | 3362.6289 | 53.31 | 49.0052
0.5288 | 92.4244 | -42.0349 | -3357.805 | 3359.9263 | 19.5426 | -53.2664 | -4.5624 | 56.957 | 3357.805 | 53.2664 | 49.1633
0.5306 | 90.4629 | -37.705 | -3352.4725 | 3354.5358 | 19.5136 | -53.1793 | -4.6641 | 56.8698 | 3352.4725 | 53.1793 | 49.3306
0.5324 | 88.327 | -32.9973 | -3346.5589 | 3348.5495 | 19.4991 | -53.0921 | -4.7513 | 56.7972 | 3346.5589 | 53.0921 | 49.498
0.5342 | 86.162 | -28.3332 | -3339.5846 | 3341.517 | 19.47 | -53.0194 | -4.8384 | 56.71 | 3339.5846 | 53.0194 | 49.6653
0.536 | 83.8372 | -23.611 | -3331.6513 | 3333.482 | 19.4555 | -52.9468 | -4.9111 | 56.6519 | 3331.6513 | 52.9468 | 49.8327
0.5378 | 81.4108 | -18.7726 | -3323.3111 | 3325.0256 | 19.4264 | -52.8741 | -4.9837 | 56.5938 | 3323.3111 | 52.8741 | 50
0.5396 | 78.6791 | -13.4982 | -3314.4334 | 3316.0462 | 19.3974 | -52.7869 | -5.0564 | 56.5066 | 3314.4334 | 52.7869 | 50.1673
0.5414 | 75.7005 | -7.9333 | -3305.0035 | 3306.5146 | 19.3683 | -52.6852 | -5.1436 | 56.3904 | 3305.0035 | 52.6852 | 50.3347
0.5432 | 72.4168 | -2.1649 | -3295.1959 | 3296.6053 | 19.3392 | -52.569 | -5.2162 | 56.2886 | 3295.1959 | 52.569 | 50.502
0.545 | 68.9151 | 4.0538 | -3284.4002 | 3285.7224 | 19.3102 | -52.4237 | -5.3034 | 56.1579 | 3284.4002 | 52.4237 | 50.6694
0.5468 | 65.3989 | 10.4906 | -3272.8344 | 3274.0985 | 19.2811 | -52.2639 | -5.4051 | 55.998 | 3272.8344 | 52.2639 | 50.8367
0.5486 | 61.8681 | 17.0435 | -3260.8909 | 3262.0969 | 19.2375 | -52.104 | -5.4923 | 55.8382 | 3260.8909 | 52.104 | 51.0041
0.5504 | 58.3519 | 23.6837 | -3248.6277 | 3249.7901 | 19.1939 | -51.9152 | -5.594 | 55.6639 | 3248.6277 | 51.9152 | 51.1714
0.5521 | 54.9374 | 30.4691 | -3235.8705 | 3237.0038 | 19.1213 | -51.7117 | -5.6812 | 55.4604 | 3235.8705 | 51.7117 | 51.3295
0.5539 | 51.6972 | 37.2836 | -3222.7355 | 3223.8834 | 19.0486 | -51.5229 | -5.7829 | 55.257 | 3222.7355 | 51.5229 | 51.4968
0.5557 | 48.6314 | 44.1272 | -3209.3971 | 3210.5304 | 18.9469 | -51.3194 | -5.8701 | 55.0536 | 3209.3971 | 51.3194 | 51.6642
0.5575 | 45.6092 | 51.2468 | -3195.6083 | 3196.7706 | 18.8452 | -51.1015 | -5.9718 | 54.8211 | 3195.6083 | 51.1015 | 51.8315
0.5593 | 42.3545 | 58.6134 | -3181.1075 | 3182.3134 | 18.7144 | -50.8545 | -6.0589 | 54.5596 | 3181.1075 | 50.8545 | 51.9989
0.5611 | 38.7656 | 66.5322 | -3164.2093 | 3165.4879 | 18.5837 | -50.5639 | -6.1607 | 54.2545 | 3164.2093 | 50.5639 | 52.1662
0.5629 | 35.293 | 74.5817 | -3147.6888 | 3149.0691 | 18.4384 | -50.2588 | -6.2478 | 53.9348 | 3147.6888 | 50.2588 | 52.3336
0.5647 | 32.2272 | 82.6022 | -3133.1008 | 3134.641 | 18.2786 | -49.9246 | -6.3641 | 53.5716 | 3133.1008 | 49.9246 | 52.5009
0.5665 | 30.0913 | 89.8671 | -3122.6248 | 3124.3393 | 18.1042 | -49.5613 | -6.5094 | 53.1938 | 3122.6248 | 49.5613 | 52.6683
0.5683 | 28.4494 | 96.8269 | -3114.5462 | 3116.4496 | 17.8862 | -49.1835 | -6.6547 | 52.7869 | 3114.5462 | 49.1835 | 52.8356
0.5701 | 27.505 | 103.4961 | -3110.3616 | 3112.4685 | 17.6247 | -48.7477 | -6.8436 | 52.322 | 3110.3616 | 48.7477 | 53.003
0.5719 | 28.0426 | 109.9329 | -3114.5172 | 3116.8565 | 17.3196 | -48.2536 | -7.0906 | 51.7989 | 3114.5172 | 48.2536 | 53.1703
0.5737 | 29.2922 | 115.8756 | -3121.9274 | 3124.4992 | 16.9854 | -47.716 | -7.3666 | 51.2323 | 3121.9274 | 47.716 | 53.3377
0.5755 | 30.7306 | 121.5858 | -3129.3957 | 3132.2 | 16.6367 | -47.1348 | -7.6427 | 50.622 | 3129.3957 | 47.1348 | 53.505
0.5772 | 31.7332 | 127.296 | -3132.3889 | 3135.4111 | 16.3025 | -46.4955 | -7.9042 | 49.9682 | 3132.3889 | 46.4955 | 53.6631
0.579 | 32.8811 | 132.7592 | -3129.4248 | 3132.6795 | 15.9538 | -45.8126 | -8.1367 | 49.2562 | 3129.4248 | 45.8126 | 53.8304
0.5808 | 34.058 | 138.1208 | -3123.3368 | 3126.8385 | 15.6051 | -45.0425 | -8.3547 | 48.4716 | 3123.3368 | 45.0425 | 53.9978
0.5826 | 35.3802 | 142.7122 | -3115.3018 | 3119.0505 | 15.2418 | -44.1998 | -8.5726 | 47.6289 | 3115.3018 | 44.1998 | 54.1651
0.5844 | 36.9494 | 147.0712 | -3105.1454 | 3109.1411 | 14.864 | -43.328 | -8.7906 | 46.7425 | 3105.1454 | 43.328 | 54.3325
0.5862 | 38.6494 | 150.631 | -3093.042 | 3097.2412 | 14.4572 | -42.4272 | -8.9794 | 45.8126 | 3093.042 | 42.4272 | 54.4998
0.588 | 40.2767 | 153.4062 | -3078.8609 | 3083.2344 | 14.0504 | -41.4537 | -9.1683 | 44.8537 | 3078.8609 | 41.4537 | 54.6672
0.5898 | 41.7879 | 155.5275 | -3062.3114 | 3066.8447 | 13.629 | -40.4656 | -9.3282 | 43.8511 | 3062.3114 | 40.4656 | 54.8345
0.5916 | 43.3135 | 157.213 | -3043.6551 | 3048.3046 | 13.2076 | -39.4631 | -9.4735 | 42.834 | 3043.6551 | 39.4631 | 55.0019
0.5934 | 44.6938 | 158.9856 | -3022.7757 | 3027.5705 | 12.7863 | -38.446 | -9.6042 | 41.8024 | 3022.7757 | 38.446 | 55.1692
0.5952 | 45.8707 | 160.9181 | -2998.9031 | 3003.8433 | 12.3649 | -37.4144 | -9.735 | 40.7562 | 2998.9031 | 37.4144 | 55.3366
0.597 | 47.0767 | 162.5454 | -2972.3135 | 2977.4135 | 11.929 | -36.3682 | -9.8367 | 39.6956 | 2972.3135 | 36.3682 | 55.5039
0.5988 | 48.5442 | 163.4463 | -2945.041 | 2950.2863 | 11.4931 | -35.3221 | -9.9384 | 38.6349 | 2945.041 | 35.3221 | 55.6713
0.6006 | 50.0989 | 162.8942 | -2917.1582 | 2922.4616 | 11.0282 | -34.2178 | -10.0111 | 37.5161 | 2917.1582 | 34.2178 | 55.8386
0.6023 | 51.4938 | 162.1967 | -2888.1421 | 2893.5036 | 10.5632 | -33.0845 | -10.0837 | 36.3682 | 2888.1421 | 33.0845 | 55.9967
0.6041 | 52.9177 | 161.3831 | -2858.4721 | 2863.8772 | 10.0837 | -31.9511 | -10.1564 | 35.2204 | 2858.4721 | 31.9511 | 56.164
0.6059 | 54.3852 | 160.5113 | -2827.6107 | 2833.0594 | 9.6042 | -30.8469 | -10.2145 | 34.087 | 2827.6107 | 30.8469 | 56.3313
0.6077 | 56.0562 | 158.4916 | -2796.2698 | 2801.704 | 9.0812 | -29.699 | -10.2581 | 32.9101 | 2796.2698 | 29.699 | 56.4987
0.6095 | 57.9596 | 155.3096 | -2765.7136 | 2771.0605 | 8.5581 | -28.5076 | -10.3017 | 31.7041 | 2765.7136 | 28.5076 | 56.666
0.6113 | 59.5869 | 151.5028 | -2733.0214 | 2738.2521 | 8.035 | -27.2725 | -10.3162 | 30.4546 | 2733.0214 | 27.2725 | 56.8334
0.6131 | 60.9818 | 147.3327 | -2697.8446 | 2702.9446 | 7.5119 | -26.0084 | -10.3307 | 29.1905 | 2697.8446 | 26.0084 | 57.0007
0.6149 | 61.9553 | 142.6686 | -2657.873 | 2662.8132 | 7.0179 | -24.7153 | -10.2871 | 27.8828 | 2657.873 | 24.7153 | 57.1681
0.6167 | 62.6527 | 137.8592 | -2615.4749 | 2620.2843 | 6.5384 | -23.4221 | -10.2435 | 26.5896 | 2615.4749 | 23.4221 | 57.3354
0.6185 | 63.4954 | 132.0618 | -2572.2195 | 2576.84 | 6.0589 | -22.129 | -10.1564 | 25.2819 | 2572.2195 | 22.129 | 57.5028
0.6203 | 64.251 | 125.7268 | -2527.4966 | 2531.8992 | 5.5795 | -20.8067 | -10.0547 | 23.9743 | 2527.4966 | 20.8067 | 57.6701
0.6221 | 64.8322 | 118.2294 | -2480.2165 | 2484.3285 | 5.1145 | -19.47 | -9.9384 | 22.6375 | 2480.2165 | 19.47 | 57.8375
0.6239 | 64.8467 | 110.4995 | -2430.0013 | 2433.8082 | 4.6641 | -18.1042 | -9.7786 | 21.2862 | 2430.0013 | 18.1042 | 58.0048
0.6256 | 64.3963 | 102.479 | -2377.9264 | 2381.399 | 4.2427 | -16.7529 | -9.6042 | 19.935 | 2377.9264 | 16.7529 | 58.1629
0.6274 | 63.5245 | 93.6304 | -2322.5095 | 2325.6189 | 3.8649 | -15.3871 | -9.3863 | 18.6127 | 2322.5095 | 15.3871 | 58.3302
0.6292 | 61.839 | 84.7526 | -2262.7337 | 2265.4798 | 3.5598 | -14.0358 | -9.1247 | 17.3051 | 2262.7337 | 14.0358 | 58.4976
0.631 | 59.4707 | 75.8313 | -2198.7878 | 2201.1853 | 3.2837 | -12.67 | -8.8341 | 16.0119 | 2198.7878 | 12.67 | 58.6649
0.6328 | 56.3032 | 67.0553 | -2130.2215 | 2132.2847 | 3.0658 | -11.3188 | -8.5 | 14.7333 | 2130.2215 | 11.3188 | 58.8323
0.6346 | 52.1622 | 58.8168 | -2057.7466 | 2059.5047 | 2.8914 | -9.982 | -8.1512 | 13.5273 | 2057.7466 | 9.982 | 58.9996
0.6364 | 47.1203 | 50.9562 | -1980.9127 | 1982.3948 | 2.7752 | -8.7179 | -7.7589 | 12.4085 | 1980.9127 | 8.7179 | 59.167
0.6382 | 41.7007 | 43.2844 | -1901.9575 | 1903.1635 | 2.688 | -7.4683 | -7.3521 | 11.3333 | 1901.9575 | 7.4683 | 59.3343
0.64 | 35.8887 | 35.8016 | -1821.8254 | 1822.7989 | 2.6299 | -6.2478 | -6.9307 | 10.2726 | 1821.8254 | 6.2478 | 59.5017
0.6418 | 29.6409 | 28.6238 | -1741.2428 | 1741.9838 | 2.5718 | -5.0564 | -6.5094 | 9.27 | 1741.2428 | 5.0564 | 59.669
0.6436 | 22.9136 | 22.2743 | -1663.1594 | 1663.7115 | 2.5718 | -3.9231 | -6.1171 | 8.3983 | 1663.1594 | 3.9231 | 59.8364
0.6454 | 15.7649 | 16.7239 | -1585.9768 | 1586.3982 | 2.5863 | -2.8624 | -5.7248 | 7.6863 | 1585.9768 | 2.8624 | 60.0037
0.6472 | 8.529 | 11.6094 | -1510.0438 | 1510.3925 | 2.6299 | -1.8889 | -5.347 | 7.076 | 1510.0438 | 1.8889 | 60.1711
0.649 | 1.206 | 6.9162 | -1434.9826 | 1435.2877 | 2.688 | -0.94444 | -4.9837 | 6.553 | 1434.9826 | 0.94444 | 60.3384
0.6507 | -5.9863 | 2.7026 | -1361.9556 | 1362.2462 | 2.7461 | -0.058119 | -4.6205 | 6.1316 | 1361.9556 | 0.058119 | 60.4965
0.6525 | -13.0769 | -0.78461 | -1293.1132 | 1293.4473 | 2.8188 | 0.74102 | -4.3008 | 5.8555 | 1293.1132 | -0.74102 | 60.6638
0.6543 | -19.6589 | -3.7778 | -1230.8092 | 1231.2305 | 2.906 | 1.482 | -4.0102 | 5.6957 | 1230.8092 | -1.482 | 60.8312
0.6561 | -25.1802 | -6.5384 | -1176.8308 | 1177.3538 | 2.9786 | 2.1214 | -3.7632 | 5.6376 | 1176.8308 | -2.1214 | 60.9985
0.6579 | -29.67 | -9.3572 | -1129.8703 | 1130.5387 | 3.0222 | 2.6735 | -3.5598 | 5.6812 | 1129.8703 | -2.6735 | 61.1659
0.6597 | -33.1281 | -12.176 | -1088.5764 | 1089.3901 | 3.0513 | 3.153 | -3.3709 | 5.7829 | 1088.5764 | -3.153 | 61.3332
0.6615 | -35.8451 | -14.7914 | -1051.0894 | 1052.0193 | 3.0803 | 3.5743 | -3.2111 | 5.9136 | 1051.0894 | -3.5743 | 61.5006
0.6633 | -37.6904 | -17.1016 | -1018.3391 | 1019.3708 | 3.0949 | 3.9521 | -3.0658 | 6.0444 | 1018.3391 | -3.9521 | 61.6679
0.6651 | -38.9545 | -18.9033 | -989.3957 | 990.529 | 3.1384 | 4.2718 | -2.935 | 6.2042 | 989.3957 | -4.2718 | 61.8353
0.6669 | -39.9571 | -20.2546 | -962.5009 | 963.7214 | 3.182 | 4.5624 | -2.8043 | 6.3495 | 962.5009 | -4.5624 | 62.0026
0.6687 | -40.8289 | -21.3153 | -936.4344 | 937.7421 | 3.2256 | 4.8384 | -2.6735 | 6.5094 | 936.4344 | -4.8384 | 62.17
0.6705 | -41.5554 | -22.0563 | -911.5447 | 912.9251 | 3.2837 | 5.0854 | -2.5573 | 6.6401 | 911.5447 | -5.0854 | 62.3373
0.6723 | -41.9332 | -22.1871 | -889.6192 | 891.0431 | 3.3419 | 5.3034 | -2.441 | 6.7854 | 889.6192 | -5.3034 | 62.5046
0.6741 | -42.1366 | -21.911 | -870.3381 | 871.7911 | 3.4 | 5.5068 | -2.3393 | 6.9307 | 870.3381 | -5.5068 | 62.672
0.6758 | -42.3836 | -21.2717 | -852.8005 | 854.2826 | 3.4581 | 5.6812 | -2.2521 | 7.0615 | 852.8005 | -5.6812 | 62.83
0.6776 | -42.7178 | -20.3563 | -836.1348 | 837.6314 | 3.5308 | 5.8701 | -2.1504 | 7.2068 | 836.1348 | -5.8701 | 62.9974
0.6794 | -43.081 | -19.1939 | -821.1691 | 822.6802 | 3.5889 | 6.0154 | -2.0778 | 7.3376 | 821.1691 | -6.0154 | 63.1647
0.6812 | -43.5314 | -17.7119 | -807.2204 | 808.7606 | 3.647 | 6.1607 | -2.0051 | 7.4538 | 807.2204 | -6.1607 | 63.3321
0.683 | -44.0109 | -16.0991 | -794.0709 | 795.6256 | 3.6906 | 6.2769 | -1.9325 | 7.5555 | 794.0709 | -6.2769 | 63.4994
0.6848 | -44.4032 | -14.3991 | -782.0692 | 783.653 | 3.7342 | 6.3931 | -1.8598 | 7.6572 | 782.0692 | -6.3931 | 63.6668
0.6866 | -44.8391 | -12.6846 | -770.8231 | 772.4359 | 3.7632 | 6.4948 | -1.7872 | 7.7299 | 770.8231 | -6.4948 | 63.8341
0.6884 | -45.2169 | -11.1735 | -759.8676 | 761.495 | 3.7778 | 6.5966 | -1.7145 | 7.8171 | 759.8676 | -6.5966 | 64.0015
0.6902 | -45.3477 | -9.8367 | -748.4907 | 750.1471 | 3.7923 | 6.6692 | -1.6419 | 7.8752 | 748.4907 | -6.6692 | 64.1688
0.692 | -45.3331 | -8.6017 | -737.1284 | 738.7848 | 3.7923 | 6.7273 | -1.5692 | 7.9188 | 737.1284 | -6.7273 | 64.3362
0.6938 | -45.4494 | -7.4538 | -726.3182 | 728.0037 | 3.7923 | 6.8 | -1.4966 | 7.9478 | 726.3182 | -6.8 | 64.5035
0.6956 | -45.522 | -6.3205 | -715.9875 | 717.6875 | 3.7778 | 6.8581 | -1.4239 | 7.9769 | 715.9875 | -6.8581 | 64.6709
0.6974 | -45.6819 | -5.2307 | -706.267 | 707.9961 | 3.7632 | 6.9017 | -1.3658 | 8.0059 | 706.267 | -6.9017 | 64.8382
0.6992 | -45.9143 | -4.1846 | -697.273 | 699.0166 | 3.7342 | 6.9598 | -1.3077 | 8.035 | 697.273 | -6.9598 | 65.0056
0.7009 | -45.9724 | -3.2111 | -688.744 | 690.5021 | 3.7051 | 7.0034 | -1.2496 | 8.035 | 688.744 | -7.0034 | 65.1636
0.7027 | -45.8998 | -2.2957 | -680.4184 | 682.1765 | 3.6761 | 7.0324 | -1.1914 | 8.035 | 680.4184 | -7.0324 | 65.331
0.7045 | -45.74 | -1.2786 | -671.8748 | 673.6184 | 3.6615 | 7.047 | -1.1333 | 8.0495 | 671.8748 | -7.047 | 65.4983
0.7063 | -45.3913 | -0.2906 | -662.9971 | 664.7407 | 3.6325 | 7.0615 | -1.0752 | 8.0205 | 662.9971 | -7.0615 | 65.6657
0.7081 | -44.9408 | 0.71196 | -654.4971 | 656.2262 | 3.5889 | 7.0615 | -1.0171 | 8.0059 | 654.4971 | -7.0615 | 65.833
0.7099 | -44.6212 | 1.7145 | -647.2613 | 648.9613 | 3.5598 | 7.047 | -0.95897 | 7.9624 | 647.2613 | -7.047 | 66.0004
0.7117 | -44.3015 | 2.5718 | -641.1878 | 642.8878 | 3.5017 | 7.0324 | -0.92991 | 7.9333 | 641.1878 | -7.0324 | 66.1677
0.7135 | -43.9673 | 3.4436 | -635.4049 | 637.0758 | 3.4581 | 7.0179 | -0.88632 | 7.8897 | 635.4049 | -7.0179 | 66.3351
0.7153 | -43.7203 | 4.1701 | -629.4913 | 631.1622 | 3.4145 | 6.9889 | -0.85726 | 7.8461 | 629.4913 | -6.9889 | 66.5024
0.7171 | -43.5896 | 4.6786 | -623.04 | 624.7109 | 3.3709 | 6.9889 | -0.81367 | 7.8025 | 623.04 | -6.9889 | 66.6698
0.7189 | -43.4733 | 4.9983 | -615.223 | 616.9084 | 3.3128 | 6.9743 | -0.77008 | 7.7735 | 615.223 | -6.9743 | 66.8371
0.7207 | -43.3861 | 5.2453 | -605.8221 | 607.5367 | 3.2547 | 6.9598 | -0.71196 | 7.7299 | 605.8221 | -6.9598 | 67.0045
0.7225 | -43.3861 | 5.5213 | -595.7384 | 597.482 | 3.2111 | 6.9453 | -0.63931 | 7.6863 | 595.7384 | -6.9453 | 67.1718
0.7242 | -43.2699 | 5.8991 | -585.3205 | 587.0786 | 3.153 | 6.9162 | -0.58119 | 7.6427 | 585.3205 | -6.9162 | 67.3299
0.726 | -43.0519 | 6.2333 | -574.9026 | 576.6753 | 3.1094 | 6.8871 | -0.50854 | 7.5846 | 574.9026 | -6.8871 | 67.4972
0.7278 | -42.7178 | 6.553 | -565.1095 | 566.8967 | 3.0658 | 6.8436 | -0.45043 | 7.5119 | 565.1095 | -6.8436 | 67.6646
0.7296 | -42.34 | 7.0034 | -556.4206 | 558.2078 | 3.0222 | 6.7854 | -0.40684 | 7.4393 | 556.4206 | -6.7854 | 67.8319
0.7314 | -41.9332 | 7.541 | -549.054 | 550.8412 | 2.9786 | 6.7273 | -0.36325 | 7.3666 | 549.054 | -6.7273 | 67.9993
0.7332 | -41.4537 | 8.2094 | -543.2275 | 545.0002 | 2.9205 | 6.6547 | -0.33419 | 7.294 | 543.2275 | -6.6547 | 68.1666
0.735 | -41.0759 | 8.8777 | -538.4908 | 540.2635 | 2.8769 | 6.5966 | -0.31966 | 7.2213 | 538.4908 | -6.5966 | 68.334
0.7368 | -40.7272 | 9.2991 | -534.6113 | 536.384 | 2.8333 | 6.553 | -0.30513 | 7.1487 | 534.6113 | -6.553 | 68.5013
0.7386 | -40.2913 | 9.5897 | -530.2233 | 531.9814 | 2.7752 | 6.4948 | -0.2906 | 7.076 | 530.2233 | -6.4948 | 68.6687
0.7404 | -39.6374 | 9.735 | -523.6994 | 525.4139 | 2.7171 | 6.4077 | -0.24701 | 6.9889 | 523.6994 | -6.4077 | 68.836
0.7422 | -38.9691 | 9.6624 | -516.1148 | 517.8148 | 2.659 | 6.335 | -0.20342 | 6.8871 | 516.1148 | -6.335 | 69.0033
0.744 | -38.3443 | 9.6188 | -507.9491 | 509.62 | 2.6008 | 6.2624 | -0.15983 | 6.7854 | 507.9491 | -6.2624 | 69.1707
0.7458 | -37.8939 | 9.6769 | -499.8269 | 501.4978 | 2.5427 | 6.1752 | -0.11624 | 6.6983 | 499.8269 | -6.1752 | 69.338
0.7476 | -37.487 | 9.735 | -492.344 | 494.0004 | 2.4991 | 6.088 | -0.072649 | 6.5966 | 492.344 | -6.088 | 69.5054
0.7493 | -36.9058 | 9.8222 | -485.6748 | 487.3312 | 2.441 | 6.0008 | -0.04359 | 6.4803 | 485.6748 | -6.0008 | 69.6634
0.7511 | -36.2229 | 9.982 | -480.1244 | 481.7372 | 2.3829 | 5.9136 | -0.01453 | 6.3786 | 480.1244 | -5.9136 | 69.8308
0.7529 | -35.6127 | 10.1564 | -475.6783 | 477.2765 | 2.3393 | 5.8265 | 0 | 6.2769 | 475.6783 | -5.8265 | 69.9981
0.7547 | -35.0024 | 10.4324 | -471.9732 | 473.5424 | 2.2812 | 5.7538 | 0.01453 | 6.1897 | 471.9732 | -5.7538 | 70.1655
0.7565 | -34.3776 | 10.723 | -468.5441 | 470.0843 | 2.2376 | 5.6812 | 0.02906 | 6.1171 | 468.5441 | -5.6812 | 70.3328
0.7583 | -33.7819 | 11.0136 | -465.5364 | 467.0621 | 2.194 | 5.6085 | 0.04359 | 6.0299 | 465.5364 | -5.6085 | 70.5002
0.7601 | -33.2588 | 11.3623 | -463.2698 | 464.7518 | 2.1504 | 5.5359 | 0.058119 | 5.9427 | 463.2698 | -5.5359 | 70.6675
0.7619 | -32.7939 | 11.7692 | -461.7732 | 463.2407 | 2.1068 | 5.4632 | 0.04359 | 5.8701 | 461.7732 | -5.4632 | 70.8349
0.7637 | -32.3289 | 12.2051 | -460.1749 | 461.6279 | 2.0632 | 5.3906 | 0.04359 | 5.7829 | 460.1749 | -5.3906 | 71.0022
0.7655 | -31.8058 | 12.5974 | -457.0365 | 458.4749 | 2.0196 | 5.3179 | 0.058119 | 5.6957 | 457.0365 | -5.3179 | 71.1696
0.7673 | -31.2246 | 12.9025 | -452.5758 | 453.9852 | 1.9761 | 5.2307 | 0.072649 | 5.6085 | 452.5758 | -5.2307 | 71.3369
0.7691 | -30.658 | 13.0478 | -447.0545 | 448.4348 | 1.9325 | 5.1581 | 0.11624 | 5.5068 | 447.0545 | -5.1581 | 71.5043
0.7709 | -30.1494 | 13.1059 | -440.6468 | 442.0271 | 1.8889 | 5.0709 | 0.1453 | 5.4196 | 440.6468 | -5.0709 | 71.6716
0.7727 | -29.699 | 13.1931 | -434.6169 | 435.9827 | 1.8453 | 4.9837 | 0.18889 | 5.3179 | 434.6169 | -4.9837 | 71.839
0.7744 | -29.3939 | 13.4256 | -430.505 | 431.8708 | 1.8017 | 4.9111 | 0.20342 | 5.2307 | 430.505 | -4.9111 | 71.997
0.7762 | -29.205 | 13.8034 | -429.0374 | 430.4032 | 1.7581 | 4.8384 | 0.20342 | 5.1581 | 429.0374 | -4.8384 | 72.1644
0.778 | -29.1178 | 14.2828 | -428.7759 | 430.1562 | 1.7291 | 4.7949 | 0.20342 | 5.1 | 428.7759 | -4.7949 | 72.3317
0.7798 | -29.2195 | 14.9512 | -429.7058 | 431.1007 | 1.7 | 4.7367 | 0.18889 | 5.0419 | 429.7058 | -4.7367 | 72.4991
0.7816 | -29.4229 | 15.7794 | -431.5802 | 433.0186 | 1.6709 | 4.7077 | 0.15983 | 4.9983 | 431.5802 | -4.7077 | 72.6664
0.7834 | -29.6118 | 16.6076 | -433.687 | 435.1545 | 1.6419 | 4.6641 | 0.13077 | 4.9547 | 433.687 | -4.6641 | 72.8338
0.7852 | -29.5537 | 17.3051 | -434.3263 | 435.8084 | 1.6128 | 4.6205 | 0.13077 | 4.8966 | 434.3263 | -4.6205 | 73.0011
0.787 | -29.1905 | 17.6102 | -433.3092 | 434.7767 | 1.5692 | 4.5624 | 0.13077 | 4.8384 | 433.3092 | -4.5624 | 73.1685
0.7888 | -28.6383 | 17.6828 | -431.0716 | 432.5101 | 1.5256 | 4.5043 | 0.15983 | 4.7658 | 431.0716 | -4.5043 | 73.3358
0.7906 | -27.9845 | 17.6392 | -428.3981 | 429.793 | 1.482 | 4.4316 | 0.18889 | 4.6931 | 428.3981 | -4.4316 | 73.5032
0.7924 | -27.3161 | 17.6102 | -425.5503 | 426.9161 | 1.4385 | 4.3735 | 0.21795 | 4.6205 | 425.5503 | -4.3735 | 73.6705
0.7942 | -26.7495 | 17.5521 | -423.3272 | 424.6494 | 1.3949 | 4.3154 | 0.24701 | 4.5478 | 423.3272 | -4.3154 | 73.8379
0.796 | -26.4007 | 17.6683 | -422.5862 | 423.8939 | 1.3513 | 4.2572 | 0.26154 | 4.4897 | 422.5862 | -4.2572 | 74.0052
0.7977 | -26.3426 | 18.0315 | -423.4434 | 424.7511 | 1.3222 | 4.2282 | 0.26154 | 4.4461 | 423.4434 | -4.2282 | 74.1633
0.7995 | -26.4298 | 18.511 | -425.0272 | 426.3639 | 1.2932 | 4.1991 | 0.24701 | 4.4025 | 425.0272 | -4.1991 | 74.3306
0.8013 | -26.6768 | 19.1649 | -427.5554 | 428.9212 | 1.2641 | 4.1846 | 0.23248 | 4.388 | 427.5554 | -4.1846 | 74.498
0.8031 | -26.8948 | 19.8332 | -430.6067 | 432.0015 | 1.2496 | 4.1846 | 0.21795 | 4.3735 | 430.6067 | -4.1846 | 74.6653
0.8049 | -27.258 | 20.6179 | -433.9921 | 435.4306 | 1.2205 | 4.1701 | 0.18889 | 4.359 | 433.9921 | -4.1701 | 74.8327
0.8067 | -27.6794 | 21.4461 | -437.3921 | 438.8887 | 1.206 | 4.1846 | 0.17436 | 4.359 | 437.3921 | -4.1846 | 75
0.8085 | -27.8973 | 22.2307 | -440.1237 | 441.6494 | 1.1914 | 4.1701 | 0.15983 | 4.3444 | 440.1237 | -4.1701 | 75.1673
0.8103 | -27.8392 | 22.7537 | -441.1263 | 442.6664 | 1.1624 | 4.1555 | 0.15983 | 4.3299 | 441.1263 | -4.1555 | 75.3347
0.8121 | -27.505 | 23.0153 | -440.3707 | 441.9109 | 1.1333 | 4.141 | 0.17436 | 4.3008 | 440.3707 | -4.141 | 75.502
0.8139 | -27.0982 | 23.1461 | -438.9323 | 440.4724 | 1.0897 | 4.1119 | 0.20342 | 4.2718 | 438.9323 | -4.1119 | 75.6694
0.8157 | -26.6477 | 23.2187 | -437.4066 | 438.9177 | 1.0607 | 4.0974 | 0.23248 | 4.2427 | 437.4066 | -4.0974 | 75.8367
0.8175 | -26.2554 | 23.2914 | -436.5639 | 438.0605 | 1.0316 | 4.0684 | 0.26154 | 4.2137 | 436.5639 | -4.0684 | 76.0041
0.8193 | -25.8341 | 23.5529 | -436.6511 | 438.1331 | 1.0026 | 4.0393 | 0.27607 | 4.1846 | 436.6511 | -4.0393 | 76.1714
0.8211 | -25.5871 | 24.0033 | -438.3075 | 439.8041 | 0.9735 | 4.0248 | 0.27607 | 4.1555 | 438.3075 | -4.0248 | 76.3388
0.8228 | -25.5871 | 24.6862 | -441.5912 | 443.1023 | 0.95897 | 4.0248 | 0.26154 | 4.1555 | 441.5912 | -4.0248 | 76.4968
0.8246 | -25.7905 | 25.4999 | -445.6741 | 447.2433 | 0.92991 | 4.0393 | 0.24701 | 4.1555 | 445.6741 | -4.0393 | 76.6642
0.8264 | -26.1537 | 26.4153 | -450.3818 | 452.0091 | 0.91538 | 4.0538 | 0.21795 | 4.1701 | 450.3818 | -4.0538 | 76.8315
0.8282 | -26.517 | 27.3161 | -454.8279 | 456.5279 | 0.90085 | 4.0829 | 0.20342 | 4.1991 | 454.8279 | -4.0829 | 76.9989
0.83 | -26.7495 | 28.0571 | -458.0681 | 459.8117 | 0.88632 | 4.0974 | 0.18889 | 4.2137 | 458.0681 | -4.0974 | 77.1662
0.8318 | -26.8657 | 28.5366 | -460.1023 | 461.8604 | 0.85726 | 4.1119 | 0.20342 | 4.2137 | 460.1023 | -4.1119 | 77.3336
0.8336 | -26.8366 | 29.0016 | -461.2065 | 463.0082 | 0.84273 | 4.1265 | 0.20342 | 4.2282 | 461.2065 | -4.1265 | 77.5009
0.8354 | -26.9093 | 29.4084 | -461.8458 | 463.6766 | 0.81367 | 4.141 | 0.21795 | 4.2427 | 461.8458 | -4.141 | 77.6683
0.8372 | -27.0691 | 29.7571 | -461.7877 | 463.6475 | 0.79914 | 4.1555 | 0.23248 | 4.2427 | 461.7877 | -4.1555 | 77.8356
0.839 | -27.258 | 29.9896 | -460.5382 | 462.427 | 0.77008 | 4.1846 | 0.26154 | 4.2718 | 460.5382 | -4.1846 | 78.003
0.8408 | -27.258 | 30.1494 | -458.6493 | 460.5672 | 0.74102 | 4.1991 | 0.30513 | 4.2718 | 458.6493 | -4.1991 | 78.1703
0.8426 | -27.258 | 30.2657 | -456.6296 | 458.5476 | 0.71196 | 4.2137 | 0.33419 | 4.2863 | 456.6296 | -4.2137 | 78.3377
0.8444 | -27.1854 | 30.2802 | -454.0288 | 455.9613 | 0.6829 | 4.2137 | 0.37778 | 4.3008 | 454.0288 | -4.2137 | 78.505
0.8462 | -27.1127 | 30.3529 | -451.704 | 453.6365 | 0.66837 | 4.2282 | 0.40684 | 4.3008 | 451.704 | -4.2282 | 78.6724
0.8479 | -27.1854 | 30.6725 | -450.8758 | 452.8373 | 0.63931 | 4.2427 | 0.4359 | 4.3154 | 450.8758 | -4.2427 | 78.8304
0.8497 | -27.4469 | 31.2101 | -451.951 | 453.9561 | 0.63931 | 4.2572 | 0.4359 | 4.3444 | 451.951 | -4.2572 | 78.9978
0.8515 | -27.8392 | 31.9511 | -454.61 | 456.6732 | 0.62478 | 4.2863 | 0.4359 | 4.3735 | 454.61 | -4.2863 | 79.1651
0.8533 | -28.2751 | 32.7793 | -458.1262 | 460.2621 | 0.62478 | 4.3299 | 0.42137 | 4.4025 | 458.1262 | -4.3299 | 79.3325
0.8551 | -28.7546 | 33.6947 | -461.8023 | 464.0108 | 0.61025 | 4.359 | 0.40684 | 4.4316 | 461.8023 | -4.359 | 79.4998
0.8569 | -29.2631 | 34.5229 | -465.1877 | 467.4689 | 0.61025 | 4.4025 | 0.39231 | 4.4752 | 465.1877 | -4.4025 | 79.6672
0.8587 | -29.67 | 35.0169 | -467.4253 | 469.7501 | 0.59572 | 4.4316 | 0.39231 | 4.5043 | 467.4253 | -4.4316 | 79.8345
0.8605 | -29.9315 | 35.2349 | -468.3116 | 470.68 | 0.56666 | 4.4752 | 0.39231 | 4.5333 | 468.3116 | -4.4752 | 80.0019
0.8623 | -30.2221 | 35.3221 | -467.9048 | 470.3022 | 0.55213 | 4.5043 | 0.42137 | 4.5624 | 467.9048 | -4.5043 | 80.1692
0.8641 | -30.8178 | 35.2058 | -466.1467 | 468.5877 | 0.52307 | 4.5333 | 0.4359 | 4.606 | 466.1467 | -4.5333 | 80.3366
0.8659 | -31.3554 | 34.9443 | -462.7612 | 465.2313 | 0.49401 | 4.5624 | 0.46496 | 4.635 | 462.7612 | -4.5624 | 80.5039
0.8677 | -31.7623 | 34.552 | -458.3732 | 460.8723 | 0.47949 | 4.5914 | 0.50854 | 4.6496 | 458.3732 | -4.5914 | 80.6713
0.8695 | -32.1255 | 34.1016 | -453.5638 | 456.0775 | 0.45043 | 4.606 | 0.55213 | 4.6641 | 453.5638 | -4.606 | 80.8386
0.8712 | -32.3434 | 33.7383 | -448.7254 | 451.2536 | 0.4359 | 4.606 | 0.59572 | 4.6786 | 448.7254 | -4.606 | 80.9967
0.873 | -32.4742 | 33.3605 | -444.8023 | 447.3451 | 0.42137 | 4.606 | 0.62478 | 4.6786 | 444.8023 | -4.606 | 81.164
0.8748 | -32.4887 | 33.0699 | -442.4049 | 444.9331 | 0.40684 | 4.606 | 0.63931 | 4.6641 | 442.4049 | -4.606 | 81.3313
0.8766 | -32.4887 | 33.0845 | -442.928 | 445.4562 | 0.39231 | 4.606 | 0.63931 | 4.6641 | 442.928 | -4.606 | 81.4987
0.8784 | -32.4742 | 33.3024 | -445.1946 | 447.7228 | 0.37778 | 4.606 | 0.62478 | 4.6786 | 445.1946 | -4.606 | 81.666
0.8802 | -32.3725 | 33.3751 | -447.8681 | 450.3818 | 0.36325 | 4.6205 | 0.61025 | 4.6786 | 447.8681 | -4.6205 | 81.8334
0.882 | -32.1981 | 33.2879 | -449.8878 | 452.3724 | 0.34872 | 4.6205 | 0.61025 | 4.6786 | 449.8878 | -4.6205 | 82.0007
0.8838 | -32.1546 | 33.2443 | -452.1254 | 454.5954 | 0.33419 | 4.635 | 0.59572 | 4.6931 | 452.1254 | -4.635 | 82.1681
0.8856 | -32.2563 | 33.2588 | -454.5954 | 457.0655 | 0.31966 | 4.6496 | 0.58119 | 4.7077 | 454.5954 | -4.6496 | 82.3354
0.8874 | -32.2127 | 33.0699 | -456.0484 | 458.4894 | 0.30513 | 4.6496 | 0.58119 | 4.7077 | 456.0484 | -4.6496 | 82.5028
0.8892 | -32.1836 | 32.6486 | -455.5399 | 457.9518 | 0.2906 | 4.6496 | 0.58119 | 4.7077 | 455.5399 | -4.6496 | 82.6701
0.891 | -32.1546 | 32.111 | -453.6655 | 456.0484 | 0.26154 | 4.6641 | 0.59572 | 4.7077 | 453.6655 | -4.6641 | 82.8375
0.8928 | -32.0383 | 31.4135 | -450.8322 | 453.1861 | 0.24701 | 4.6496 | 0.62478 | 4.7077 | 450.8322 | -4.6496 | 83.0048
0.8946 | -31.8494 | 30.7306 | -447.4322 | 449.7425 | 0.23248 | 4.635 | 0.63931 | 4.6931 | 447.4322 | -4.635 | 83.1722
0.8963 | -31.6024 | 30.164 | -443.9887 | 446.2553 | 0.20342 | 4.6205 | 0.66837 | 4.6931 | 443.9887 | -4.6205 | 83.3302
0.8981 | -31.3845 | 29.7281 | -441.5041 | 443.7417 | 0.20342 | 4.606 | 0.6829 | 4.6641 | 441.5041 | -4.606 | 83.4976
0.8999 | -31.0939 | 29.3503 | -440.0075 | 442.216 | 0.18889 | 4.5769 | 0.69743 | 4.6496 | 440.0075 | -4.5769 | 83.6649
0.9017 | -30.7452 | 28.9725 | -439.2519 | 441.4023 | 0.18889 | 4.5624 | 0.69743 | 4.635 | 439.2519 | -4.5624 | 83.8323
0.9035 | -30.4255 | 28.7691 | -439.6733 | 441.8092 | 0.18889 | 4.5478 | 0.69743 | 4.606 | 439.6733 | -4.5478 | 83.9996
0.9053 | -30.3529 | 28.8272 | -441.6784 | 443.7998 | 0.17436 | 4.5333 | 0.6829 | 4.606 | 441.6784 | -4.5333 | 84.167
0.9071 | -30.5127 | 28.9435 | -444.4827 | 446.6186 | 0.18889 | 4.5333 | 0.65384 | 4.5914 | 444.4827 | -4.5333 | 84.3343
0.9089 | -30.687 | 29.0306 | -447.7519 | 449.8732 | 0.18889 | 4.5333 | 0.61025 | 4.5914 | 447.7519 | -4.5333 | 84.5017
0.9107 | -30.8323 | 29.0306 | -450.1784 | 452.2997 | 0.18889 | 4.5333 | 0.58119 | 4.5914 | 450.1784 | -4.5333 | 84.669
0.9125 | -30.9486 | 28.8272 | -450.3818 | 452.4886 | 0.18889 | 4.5188 | 0.56666 | 4.5769 | 450.3818 | -4.5188 | 84.8364
0.9143 | -31.1375 | 28.3768 | -448.9433 | 451.0502 | 0.18889 | 4.5043 | 0.56666 | 4.5624 | 448.9433 | -4.5043 | 85.0037
0.9161 | -31.2537 | 27.7956 | -445.8921 | 447.9844 | 0.17436 | 4.4897 | 0.58119 | 4.5333 | 445.8921 | -4.4897 | 85.1711
0.9179 | -31.2537 | 27.1127 | -440.8938 | 442.9716 | 0.17436 | 4.4607 | 0.61025 | 4.5188 | 440.8938 | -4.4607 | 85.3384
0.9197 | -31.1229 | 26.4443 | -434.9947 | 437.0434 | 0.17436 | 4.4171 | 0.63931 | 4.4752 | 434.9947 | -4.4171 | 85.5058
0.9214 | -30.9341 | 25.7614 | -429.0374 | 431.0716 | 0.15983 | 4.3735 | 0.66837 | 4.4316 | 429.0374 | -4.3735 | 85.6638
0.9232 | -30.8323 | 25.1512 | -424.7511 | 426.7563 | 0.15983 | 4.3299 | 0.69743 | 4.4025 | 424.7511 | -4.3299 | 85.8312
0.925 | -30.7161 | 24.6862 | -421.7435 | 423.734 | 0.15983 | 4.2863 | 0.69743 | 4.359 | 421.7435 | -4.2863 | 85.9985
0.9268 | -30.5272 | 24.1486 | -419.4187 | 421.3802 | 0.1453 | 4.2427 | 0.69743 | 4.3154 | 419.4187 | -4.2427 | 86.1659
0.9286 | -30.5417 | 23.8435 | -418.3725 | 420.3195 | 0.1453 | 4.1991 | 0.69743 | 4.2863 | 418.3725 | -4.1991 | 86.3332
0.9304 | -30.6725 | 23.7418 | -418.3871 | 420.3341 | 0.15983 | 4.1701 | 0.66837 | 4.2427 | 418.3871 | -4.1701 | 86.5006
0.9322 | -30.7161 | 23.7272 | -419.1136 | 421.0606 | 0.15983 | 4.141 | 0.63931 | 4.2137 | 419.1136 | -4.141 | 86.6679
0.934 | -30.7742 | 23.9452 | -420.8571 | 422.8041 | 0.17436 | 4.0974 | 0.61025 | 4.1701 | 420.8571 | -4.0974 | 86.8353
0.9358 | -30.8178 | 24.2794 | -422.8768 | 424.8383 | 0.20342 | 4.0684 | 0.58119 | 4.1265 | 422.8768 | -4.0684 | 87.0026
0.9376 | -30.8759 | 24.4973 | -424.2571 | 426.2187 | 0.21795 | 4.0248 | 0.55213 | 4.0974 | 424.2571 | -4.0248 | 87.17
0.9394 | -30.9486 | 24.4247 | -424.4605 | 426.4221 | 0.21795 | 3.9957 | 0.5376 | 4.0538 | 424.4605 | -3.9957 | 87.3373
0.9412 | -30.7597 | 24.2067 | -423.2691 | 425.2161 | 0.21795 | 3.9521 | 0.52307 | 4.0102 | 423.2691 | -3.9521 | 87.5046
0.943 | -30.4546 | 23.8725 | -421.511 | 423.4289 | 0.21795 | 3.9085 | 0.52307 | 3.9666 | 421.511 | -3.9085 | 87.672
0.9447 | -30.2366 | 23.6255 | -419.1426 | 421.046 | 0.23248 | 3.8359 | 0.52307 | 3.894 | 419.1426 | -3.8359 | 87.83
0.9465 | -30.1494 | 23.4802 | -415.7136 | 417.617 | 0.23248 | 3.7632 | 0.5376 | 3.8214 | 415.7136 | -3.7632 | 87.9974
0.9483 | -30.0332 | 23.3785 | -411.0059 | 412.9238 | 0.24701 | 3.6761 | 0.55213 | 3.7487 | 411.0059 | -3.6761 | 88.1647
0.9501 | -29.8588 | 23.3349 | -406.1239 | 408.0709 | 0.24701 | 3.5889 | 0.56666 | 3.6615 | 406.1239 | -3.5889 | 88.3321
0.9519 | -29.5537 | 23.4076 | -402.6512 | 404.6128 | 0.26154 | 3.5017 | 0.56666 | 3.5743 | 402.6512 | -3.5017 | 88.4994
0.9537 | -28.9435 | 23.6546 | -401.2418 | 403.1743 | 0.27607 | 3.4145 | 0.56666 | 3.4872 | 401.2418 | -3.4145 | 88.6668
0.9555 | -28.3623 | 24.1922 | -401.0239 | 402.9564 | 0.2906 | 3.3128 | 0.55213 | 3.4 | 401.0239 | -3.3128 | 88.8341
0.9573 | -27.97 | 25.064 | -401.8957 | 403.8572 | 0.30513 | 3.2256 | 0.52307 | 3.2983 | 401.8957 | -3.2256 | 89.0015
0.9591 | -27.6794 | 26.023 | -403.3923 | 405.3974 | 0.33419 | 3.1384 | 0.49401 | 3.2111 | 403.3923 | -3.1384 | 89.1688
0.9609 | -27.4905 | 26.8512 | -404.7726 | 406.8358 | 0.34872 | 3.0513 | 0.46496 | 3.1384 | 404.7726 | -3.0513 | 89.3362
0.9627 | -27.1854 | 27.4905 | -404.9615 | 407.0392 | 0.36325 | 2.9641 | 0.45043 | 3.0513 | 404.9615 | -2.9641 | 89.5035
0.9645 | -26.7785 | 27.9845 | -404.0606 | 406.1675 | 0.37778 | 2.8769 | 0.45043 | 2.9641 | 404.0606 | -2.8769 | 89.6709
0.9663 | -26.3281 | 28.6819 | -402.3316 | 404.4529 | 0.37778 | 2.7752 | 0.45043 | 2.8769 | 402.3316 | -2.7752 | 89.8382
0.9681 | -25.9503 | 29.6264 | -401.0384 | 403.2179 | 0.39231 | 2.6735 | 0.45043 | 2.7607 | 401.0384 | -2.6735 | 90.0056
0.9698 | -25.776 | 31.0067 | -401.3581 | 403.6538 | 0.40684 | 2.5573 | 0.4359 | 2.659 | 401.3581 | -2.5573 | 90.1636
0.9716 | -25.7324 | 32.6776 | -401.8812 | 404.3222 | 0.4359 | 2.4265 | 0.42137 | 2.5427 | 401.8812 | -2.4265 | 90.331
0.9734 | -25.805 | 34.3195 | -399.7162 | 402.3316 | 0.45043 | 2.2957 | 0.42137 | 2.412 | 399.7162 | -2.2957 | 90.4983
0.9752 | -25.776 | 35.9178 | -394.994 | 397.7983 | 0.47949 | 2.1214 | 0.45043 | 2.2812 | 394.994 | -2.1214 | 90.6657
0.977 | -25.8922 | 37.3853 | -389.0949 | 392.088 | 0.50854 | 1.947 | 0.49401 | 2.1504 | 389.0949 | -1.947 | 90.833
0.9788 | -26.27 | 38.8964 | -382.3676 | 385.6077 | 0.52307 | 1.7581 | 0.5376 | 2.0051 | 382.3676 | -1.7581 | 91.0004
0.9806 | -26.6623 | 40.5383 | -373.7514 | 377.3112 | 0.5376 | 1.5256 | 0.61025 | 1.8744 | 373.7514 | -1.5256 | 91.1677
0.9824 | -27.0836 | 42.2383 | -365.0334 | 368.9274 | 0.56666 | 1.2932 | 0.69743 | 1.7581 | 365.0334 | -1.2932 | 91.3351
0.9842 | -27.5922 | 44.3015 | -356.7805 | 361.0959 | 0.58119 | 1.0316 | 0.77008 | 1.6419 | 356.7805 | -1.0316 | 91.5024
0.986 | -27.9409 | 46.6989 | -350.7651 | 355.5019 | 0.62478 | 0.78461 | 0.81367 | 1.5692 | 350.7651 | -0.78461 | 91.6698
0.9878 | -28.246 | 49.2417 | -346.3626 | 351.5352 | 0.65384 | 0.50854 | 0.85726 | 1.5256 | 346.3626 | -0.50854 | 91.8371
0.9896 | -28.3768 | 52.3365 | -343.4857 | 349.1814 | 0.69743 | 0.20342 | 0.90085 | 1.5692 | 343.4857 | -0.20342 | 92.0045
0.9914 | -28.5366 | 55.751 | -341.5823 | 347.8592 | 0.72649 | -0.11624 | 0.94444 | 1.6273 | 341.5823 | 0.11624 | 92.1718
0.9932 | -28.6383 | 59.4126 | -341.3207 | 348.2369 | 0.77008 | -0.45043 | 0.9735 | 1.7291 | 341.3207 | 0.45043 | 92.3392
0.9949 | -28.5221 | 63.045 | -342.8609 | 350.3583 | 0.78461 | -0.78461 | 1.0026 | 1.8453 | 342.8609 | 0.78461 | 92.4972
0.9967 | -28.1734 | 66.3578 | -346.6241 | 354.5865 | 0.78461 | -1.1333 | 1.0171 | 2.0051 | 346.6241 | 1.1333 | 92.6646
0.9985 | -27.0691 | 69.7288 | -355.2984 | 363.6386 | 0.75555 | -1.5402 | 1.0171 | 2.2231 | 355.2984 | 1.5402 | 92.8319
1.0003 | -24.9478 | 73.0997 | -370.4676 | 379.0402 | 0.69743 | -1.9615 | 0.98803 | 2.5282 | 370.4676 | 1.9615 | 92.9993
1.0021 | -22.1726 | 76.6595 | -390.9111 | 399.6726 | 0.61025 | -2.412 | 0.94444 | 2.906 | 390.9111 | 2.412 | 93.1666
1.0039 | -18.7726 | 79.9142 | -414.4785 | 423.3127 | 0.49401 | -2.8624 | 0.88632 | 3.2837 | 414.4785 | 2.8624 | 93.334
1.0057 | -14.6751 | 82.733 | -440.7485 | 449.51 | 0.34872 | -3.3419 | 0.84273 | 3.6761 | 440.7485 | 3.3419 | 93.5013
1.0075 | -10.0111 | 85.3338 | -468.7475 | 477.3637 | 0.18889 | -3.8068 | 0.77008 | 4.0829 | 468.7475 | 3.8068 | 93.6687
1.0093 | -5.1145 | 88.0364 | -497.6183 | 506.1038 | 0.04359 | -4.2863 | 0.71196 | 4.4897 | 497.6183 | 4.2863 | 93.836
1.0111 | -0.23248 | 90.8406 | -526.5473 | 534.931 | -0.11624 | -4.7513 | 0.63931 | 4.8966 | 526.5473 | 4.7513 | 94.0033
1.0129 | 4.3444 | 93.4705 | -554.2266 | 562.5813 | -0.27607 | -5.1872 | 0.58119 | 5.3034 | 554.2266 | 5.1872 | 94.1707
1.0147 | 8.6017 | 95.8534 | -580.5693 | 588.9675 | -0.4359 | -5.594 | 0.5376 | 5.6957 | 580.5693 | 5.594 | 94.338
1.0165 | 12.7427 | 97.8876 | -606.476 | 614.9178 | -0.61025 | -5.9718 | 0.49401 | 6.088 | 606.476 | 5.9718 | 94.5054
1.0182 | 16.6803 | 99.8201 | -631.7579 | 640.2579 | -0.77008 | -6.3495 | 0.45043 | 6.4658 | 631.7579 | 6.3495 | 94.6634
1.02 | 20.429 | 101.5055 | -657.4612 | 666.0048 | -0.92991 | -6.6983 | 0.40684 | 6.829 | 657.4612 | 6.6983 | 94.8308
1.0218 | 24.1196 | 102.9731 | -683.4987 | 692.0568 | -1.1043 | -7.047 | 0.36325 | 7.1923 | 683.4987 | 7.047 | 94.9981
1.0236 | 27.8683 | 104.0047 | -710.0593 | 718.5738 | -1.2786 | -7.3812 | 0.31966 | 7.5555 | 710.0593 | 7.3812 | 95.1655
1.0254 | 31.5734 | 104.2372 | -737.1575 | 745.5412 | -1.4675 | -7.7154 | 0.27607 | 7.9042 | 737.1575 | 7.7154 | 95.3328
1.0272 | 35.2639 | 103.7722 | -764.6189 | 772.8282 | -1.6564 | -8.0205 | 0.21795 | 8.2384 | 764.6189 | 8.0205 | 95.5002
1.029 | 39.0272 | 102.7551 | -792.3709 | 800.3333 | -1.8453 | -8.3111 | 0.17436 | 8.5726 | 792.3709 | 8.3111 | 95.6675
1.0308 | 42.8485 | 101.4184 | -820.5879 | 828.2887 | -2.0487 | -8.5871 | 0.11624 | 8.8923 | 820.5879 | 8.5871 | 95.8349
1.0326 | 46.6554 | 99.6457 | -847.9476 | 855.3578 | -2.2376 | -8.8632 | 0.072649 | 9.1974 | 847.9476 | 8.8632 | 96.0022
1.0344 | 50.2733 | 97.1902 | -873.5346 | 880.6688 | -2.412 | -9.1102 | 0.02906 | 9.488 | 873.5346 | 9.1102 | 96.1696
1.0362 | 53.7604 | 94.2116 | -898.1046 | 904.9482 | -2.5718 | -9.3427 | -0.02906 | 9.7641 | 898.1046 | 9.3427 | 96.3369
1.038 | 57.1604 | 90.9133 | -920.8438 | 927.4113 | -2.7316 | -9.5606 | -0.058119 | 10.0111 | 920.8438 | 9.5606 | 96.5043
1.0398 | 60.4878 | 87.6586 | -940.9532 | 947.2882 | -2.8333 | -9.7786 | -0.10171 | 10.2435 | 940.9532 | 9.7786 | 96.6716
1.0416 | 63.3502 | 84.927 | -958.2437 | 964.4043 | -2.8624 | -9.9675 | -0.11624 | 10.447 | 958.2437 | 9.9675 | 96.839
1.0433 | 65.704 | 82.6022 | -973.4564 | 979.4863 | -2.8624 | -10.1418 | -0.15983 | 10.6068 | 973.4564 | 10.1418 | 96.997
1.0451 | 67.7672 | 80.4808 | -986.7658 | 992.6794 | -2.8333 | -10.2871 | -0.18889 | 10.7376 | 986.7658 | 10.2871 | 97.1644
1.0469 | 69.8305 | 78.4321 | -998.7965 | 1004.6084 | -2.8043 | -10.4324 | -0.21795 | 10.8683 | 998.7965 | 10.4324 | 97.3317
1.0487 | 71.5159 | 77.0227 | -1009.752 | 1015.4913 | -2.7316 | -10.5632 | -0.24701 | 10.97 | 1009.752 | 10.5632 | 97.4991
1.0505 | 73.0125 | 75.8168 | -1019.6178 | 1025.2989 | -2.6299 | -10.6649 | -0.2906 | 11.0572 | 1019.6178 | 10.6649 | 97.6664
1.0523 | 74.3783 | 74.4219 | -1028.1177 | 1033.7263 | -2.5427 | -10.7666 | -0.30513 | 11.1299 | 1028.1177 | 10.7666 | 97.8338
1.0541 | 75.439 | 72.998 | -1035.7023 | 1041.2382 | -2.4265 | -10.8393 | -0.33419 | 11.1735 | 1035.7023 | 10.8393 | 98.0011
1.0559 | 76.1945 | 71.5014 | -1042.139 | 1047.5877 | -2.3248 | -10.8974 | -0.36325 | 11.2025 | 1042.139 | 10.8974 | 98.1685
1.0577 | 76.4415 | 70.092 | -1046.6869 | 1052.0339 | -2.194 | -10.9264 | -0.37778 | 11.2025 | 1046.6869 | 10.9264 | 98.3358
1.0595 | 75.9621 | 68.639 | -1048.5322 | 1053.7339 | -2.0342 | -10.941 | -0.37778 | 11.188 | 1048.5322 | 10.941 | 98.5032
1.0613 | 75.2501 | 66.9971 | -1048.4014 | 1053.4433 | -1.8889 | -10.9555 | -0.36325 | 11.1735 | 1048.4014 | 10.9555 | 98.6705
1.0631 | 74.4364 | 65.1373 | -1046.4253 | 1051.3074 | -1.7436 | -10.9555 | -0.34872 | 11.1444 | 1046.4253 | 10.9555 | 98.8379
1.0649 | 73.1287 | 63.045 | -1042.2698 | 1046.9484 | -1.5838 | -10.9264 | -0.30513 | 11.0863 | 1042.2698 | 10.9264 | 99.0052
1.0667 | 71.574 | 60.6185 | -1037.8818 | 1042.3425 | -1.4239 | -10.8829 | -0.26154 | 11.0282 | 1037.8818 | 10.8829 | 99.1726
1.0684 | 69.9467 | 57.7997 | -1034.4673 | 1038.6955 | -1.2641 | -10.8538 | -0.20342 | 10.9846 | 1034.4673 | 10.8538 | 99.3306
1.0702 | 68.2177 | 54.8502 | -1032.5348 | 1036.516 | -1.1188 | -10.8247 | -0.1453 | 10.941 | 1032.5348 | 10.8247 | 99.498
1.072 | 66.5903 | 51.9878 | -1032.1716 | 1035.9203 | -0.95897 | -10.7957 | -0.10171 | 10.8974 | 1032.1716 | 10.7957 | 99.6653
1.0738 | 65.2536 | 49.1981 | -1033.0724 | 1036.5886 | -0.81367 | -10.7666 | -0.058119 | 10.8683 | 1033.0724 | 10.7666 | 99.8327
1.0756 | 64.2074 | 46.8733 | -1033.9297 | 1037.257 | -0.69743 | -10.7666 | -0.01453 | 10.8393 | 1033.9297 | 10.7666 | 100
| Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My
Minimum | -45.9724 | -416.0042 | -3377.057 | 347.8592 | -2.8624 | -53.4117 | -10.3307 | 1.5256 | 341.3207 | -7.0615
Maximum | 157.6779 | 163.4463 | -341.3207 | 3379.6724 | 25.0495 | 7.0615 | 9.1538 | 57.1168 | 3377.057 | 53.4117

## Diagr_HIGH100
HIGH100 = Loads at High Level in Subject with 100kg Body Weight.
Contact forces and moments, measured telemetrically in knee implants of 8 subjects.
Averages of one individual subject.
Fx  lateral, Fy  anterior, -Fz  distal, Fres = Resultant Force
Mx around x, My around y, Mz around z, Mres = Resultant Moment
z-axis = axis of implant stem!
From Bergmann et al, PlosOne (2014) "Standardized Loads Acting in Knee Implants"
Subject K2L
Subject K2L
G.Bergmann et al., PlosOne, 2014
G.Bergmann et al., PlosOne, 2014