## Dat_Individ
Data File: | k5r_2242_0_13_i1.akf
Curve Alignment Program: | 1.5.7
------------------------MEASURING INFOS------------------------
Diagram Title #1 | Forces and Moments at Knee Joint
Diagram Title #2 | Measured Relative to Implant
Comment #1 | Knee Joint; Gaitanalysis; Level Walking;
Comment #2 | K5R, 30 Months PO *
BodyWeight [N]: | 942
Patient(en): | k5r | [7 empty] | Subject | k5r
Implant Type: | Knee Joint
Varus Angle Leg [°] | 1 | [7 empty] | Individual loads in N (forces) and Nm (moments)
Axes Directions X, Y, Z: | Lateral, Anterior, Superior
Axes Origin: | Tibial Plateau at Stem Axis | [7 empty] | BW [kg] = | 96.0245
Transformation X Y Z (°): | 7 | -2 | -1
Displacement X, Y, Z (mm): | 0 | 0 | 0
Activity Code: | KneeActivities | 1P_1_15_2
---------------------------DATA--------------------------------
Number of Data Sets = | 601
Number of analog channels = | 0
Time-Offset(sec) = | 0
Time | Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My | Time
[s] | [N] | [N] | [N] | [N] | [Nm] | [Nm] | [Nm] | [Nm] | [N] | (Nm] | [%Cycle]
0 | 11.62 | 29.16 | -603.26 | 604.22 | 1.88 | -3.84 | -1.05 | 4.55 | 603.26 | 3.84 | 0
0.0018 | 10.54 | 27.18 | -601.81 | 602.62 | 1.97 | -3.69 | -1 | 4.44 | 601.81 | 3.69 | 0.16735
0.0036 | 9.36 | 25.4 | -601.26 | 601.96 | 2.06 | -3.52 | -0.95 | 4.31 | 601.26 | 3.52 | 0.3347
0.0054 | 8.42 | 23.76 | -601.99 | 602.6 | 2.14 | -3.38 | -0.9 | 4.23 | 601.99 | 3.38 | 0.50205
0.0072 | 7.58 | 22.21 | -604.34 | 604.88 | 2.23 | -3.26 | -0.85 | 4.18 | 604.34 | 3.26 | 0.66939
0.009 | 6.78 | 20.75 | -609.34 | 609.83 | 2.32 | -3.15 | -0.8 | 4.17 | 609.34 | 3.15 | 0.83674
0.0108 | 6.29 | 19.18 | -618.14 | 618.57 | 2.44 | -3.08 | -0.75 | 4.2 | 618.14 | 3.08 | 1.0041
0.0125 | 5.82 | 17.5 | -629.4 | 629.77 | 2.56 | -3.02 | -0.7 | 4.22 | 629.4 | 3.02 | 1.1621
0.0143 | 5.4 | 15.98 | -642.58 | 642.91 | 2.69 | -2.98 | -0.65 | 4.27 | 642.58 | 2.98 | 1.3295
0.0161 | 5.11 | 14.32 | -660.09 | 660.38 | 2.84 | -2.98 | -0.58 | 4.37 | 660.09 | 2.98 | 1.4968
0.0179 | 4.89 | 12.6 | -680.25 | 680.51 | 2.99 | -2.98 | -0.53 | 4.48 | 680.25 | 2.98 | 1.6642
0.0197 | 4.76 | 10.82 | -702.4 | 702.63 | 3.15 | -2.99 | -0.48 | 4.61 | 702.4 | 2.99 | 1.8315
0.0215 | 4.8 | 9.03 | -726.52 | 726.72 | 3.32 | -3.03 | -0.42 | 4.76 | 726.52 | 3.03 | 1.9989
0.0233 | 4.89 | 7.47 | -751.19 | 751.38 | 3.52 | -3.08 | -0.37 | 4.93 | 751.19 | 3.08 | 2.1662
0.0251 | 4.94 | 6.08 | -775.1 | 775.27 | 3.72 | -3.14 | -0.31 | 5.12 | 775.1 | 3.14 | 2.3336
0.0269 | 5.03 | 4.68 | -796.8 | 796.96 | 3.91 | -3.2 | -0.26 | 5.31 | 796.8 | 3.2 | 2.5009
0.0287 | 5.26 | 2.92 | -816.58 | 816.74 | 4.09 | -3.26 | -0.22 | 5.49 | 816.58 | 3.26 | 2.6683
0.0305 | 5.55 | 0.57 | -833.23 | 833.4 | 4.26 | -3.32 | -0.19 | 5.64 | 833.23 | 3.32 | 2.8356
0.0323 | 5.74 | -2.36 | -847.78 | 847.97 | 4.42 | -3.35 | -0.16 | 5.78 | 847.78 | 3.35 | 3.003
0.0341 | 5.93 | -5.38 | -861.27 | 861.5 | 4.58 | -3.37 | -0.12 | 5.9 | 861.27 | 3.37 | 3.1703
0.0359 | 6.19 | -8.4 | -874.13 | 874.39 | 4.73 | -3.37 | -0.09 | 6.01 | 874.13 | 3.37 | 3.3377
0.0376 | 6.88 | -11.06 | -889.79 | 890.08 | 4.92 | -3.41 | -0.06 | 6.17 | 889.79 | 3.41 | 3.4957
0.0394 | 7.63 | -13.71 | -906.03 | 906.35 | 5.11 | -3.45 | -0.04 | 6.35 | 906.03 | 3.45 | 3.6631
0.0412 | 8.59 | -16.47 | -924.57 | 924.96 | 5.34 | -3.55 | -0.01 | 6.58 | 924.57 | 3.55 | 3.8304
0.043 | 9.69 | -19.27 | -944.15 | 944.61 | 5.58 | -3.65 | 0 | 6.83 | 944.15 | 3.65 | 3.9978
0.0448 | 10.73 | -21.97 | -964.35 | 964.88 | 5.83 | -3.78 | 0.02 | 7.11 | 964.35 | 3.78 | 4.1651
0.0466 | 11.56 | -24.71 | -984.48 | 985.08 | 6.1 | -3.92 | 0.04 | 7.4 | 984.48 | 3.92 | 4.3325
0.0484 | 12.28 | -27.54 | -1004.36 | 1005.04 | 6.37 | -4.1 | 0.07 | 7.72 | 1004.36 | 4.1 | 4.4998
0.0502 | 13.05 | -30.58 | -1025.17 | 1025.96 | 6.65 | -4.3 | 0.1 | 8.06 | 1025.17 | 4.3 | 4.6672
0.052 | 13.64 | -33.56 | -1045.14 | 1046.03 | 6.92 | -4.49 | 0.13 | 8.39 | 1045.14 | 4.49 | 4.8345
0.0538 | 14.16 | -36.54 | -1063.93 | 1064.92 | 7.18 | -4.72 | 0.17 | 8.72 | 1063.93 | 4.72 | 5.0019
0.0556 | 14.42 | -39.46 | -1080.91 | 1082 | 7.41 | -4.98 | 0.2 | 9.04 | 1080.91 | 4.98 | 5.1692
0.0574 | 14.55 | -42.41 | -1097.14 | 1098.32 | 7.62 | -5.25 | 0.22 | 9.36 | 1097.14 | 5.25 | 5.3366
0.0592 | 14.6 | -45.36 | -1112.79 | 1114.05 | 7.83 | -5.52 | 0.25 | 9.68 | 1112.79 | 5.52 | 5.5039
0.061 | 14.74 | -48.48 | -1129.02 | 1130.38 | 8.04 | -5.83 | 0.27 | 10.03 | 1129.02 | 5.83 | 5.6713
0.0627 | 15.03 | -51.81 | -1147.05 | 1148.54 | 8.27 | -6.18 | 0.3 | 10.42 | 1147.05 | 6.18 | 5.8293
0.0645 | 15.41 | -55.26 | -1167.04 | 1168.67 | 8.52 | -6.54 | 0.33 | 10.83 | 1167.04 | 6.54 | 5.9967
0.0663 | 15.91 | -58.92 | -1187.94 | 1189.73 | 8.77 | -6.93 | 0.36 | 11.27 | 1187.94 | 6.93 | 6.164
0.0681 | 16.51 | -62.79 | -1209.6 | 1211.57 | 9.02 | -7.33 | 0.4 | 11.73 | 1209.6 | 7.33 | 6.3313
0.0699 | 17.03 | -66.59 | -1232.19 | 1234.35 | 9.29 | -7.75 | 0.43 | 12.2 | 1232.19 | 7.75 | 6.4987
0.0717 | 17.58 | -70.48 | -1255.43 | 1257.79 | 9.55 | -8.2 | 0.46 | 12.69 | 1255.43 | 8.2 | 6.666
0.0735 | 18.1 | -74.32 | -1279.3 | 1281.85 | 9.81 | -8.66 | 0.47 | 13.18 | 1279.3 | 8.66 | 6.8334
0.0753 | 18.68 | -78.14 | -1303.58 | 1306.32 | 10.07 | -9.12 | 0.5 | 13.69 | 1303.58 | 9.12 | 7.0007
0.0771 | 19.38 | -82.14 | -1328.14 | 1331.1 | 10.34 | -9.58 | 0.52 | 14.19 | 1328.14 | 9.58 | 7.1681
0.0789 | 19.78 | -85.92 | -1353.49 | 1356.64 | 10.62 | -10.04 | 0.54 | 14.7 | 1353.49 | 10.04 | 7.3354
0.0807 | 20.17 | -89.62 | -1380 | 1383.35 | 10.91 | -10.53 | 0.55 | 15.25 | 1380 | 10.53 | 7.5028
0.0825 | 20.32 | -93.27 | -1407.53 | 1411.06 | 11.2 | -11.05 | 0.57 | 15.82 | 1407.53 | 11.05 | 7.6701
0.0843 | 20.52 | -96.99 | -1435.56 | 1439.27 | 11.49 | -11.6 | 0.57 | 16.42 | 1435.56 | 11.6 | 7.8375
0.086 | 20.77 | -100.71 | -1464.19 | 1468.07 | 11.8 | -12.15 | 0.58 | 17.03 | 1464.19 | 12.15 | 7.9955
0.0878 | 20.99 | -104.42 | -1491.45 | 1495.53 | 12.09 | -12.68 | 0.59 | 17.61 | 1491.45 | 12.68 | 8.1629
0.0896 | 21.3 | -107.86 | -1517.77 | 1522.02 | 12.37 | -13.23 | 0.59 | 18.2 | 1517.77 | 13.23 | 8.3302
0.0914 | 21.45 | -111.29 | -1542.97 | 1547.4 | 12.63 | -13.74 | 0.59 | 18.75 | 1542.97 | 13.74 | 8.4976
0.0932 | 21.61 | -114.72 | -1567.66 | 1572.27 | 12.88 | -14.27 | 0.59 | 19.31 | 1567.66 | 14.27 | 8.6649
0.095 | 21.72 | -118.19 | -1591.98 | 1596.78 | 13.13 | -14.79 | 0.6 | 19.86 | 1591.98 | 14.79 | 8.8323
0.0968 | 21.75 | -121.64 | -1615.56 | 1620.56 | 13.36 | -15.29 | 0.6 | 20.39 | 1615.56 | 15.29 | 8.9996
0.0986 | 21.71 | -125.06 | -1638.63 | 1643.82 | 13.59 | -15.79 | 0.6 | 20.92 | 1638.63 | 15.79 | 9.167
0.1004 | 21.66 | -128.39 | -1661.21 | 1666.59 | 13.81 | -16.27 | 0.6 | 21.43 | 1661.21 | 16.27 | 9.3343
0.1022 | 21.59 | -131.63 | -1683.52 | 1689.09 | 14.02 | -16.75 | 0.61 | 21.93 | 1683.52 | 16.75 | 9.5017
0.104 | 21.52 | -134.75 | -1705.53 | 1711.28 | 14.22 | -17.21 | 0.62 | 22.42 | 1705.53 | 17.21 | 9.669
0.1058 | 21.48 | -137.6 | -1726.87 | 1732.78 | 14.42 | -17.63 | 0.64 | 22.87 | 1726.87 | 17.63 | 9.8364
0.1076 | 21.42 | -140.1 | -1747.3 | 1753.35 | 14.6 | -18.02 | 0.65 | 23.29 | 1747.3 | 18.02 | 10.0037
0.1094 | 21.33 | -142.48 | -1767.39 | 1773.55 | 14.78 | -18.38 | 0.67 | 23.69 | 1767.39 | 18.38 | 10.1711
0.1111 | 21.23 | -144.75 | -1787.09 | 1793.38 | 14.97 | -18.73 | 0.69 | 24.07 | 1787.09 | 18.73 | 10.3291
0.1129 | 21.03 | -146.94 | -1806.19 | 1812.6 | 15.15 | -19.09 | 0.71 | 24.47 | 1806.19 | 19.09 | 10.4965
0.1147 | 20.77 | -148.91 | -1824.48 | 1830.98 | 15.32 | -19.44 | 0.73 | 24.85 | 1824.48 | 19.44 | 10.6638
0.1165 | 20.47 | -150.52 | -1841.5 | 1848.07 | 15.49 | -19.77 | 0.76 | 25.21 | 1841.5 | 19.77 | 10.8312
0.1183 | 20.14 | -151.84 | -1857.61 | 1864.23 | 15.64 | -20.1 | 0.79 | 25.57 | 1857.61 | 20.1 | 10.9985
0.1201 | 19.83 | -153.09 | -1873.17 | 1879.84 | 15.79 | -20.42 | 0.81 | 25.92 | 1873.17 | 20.42 | 11.1659
0.1219 | 19.53 | -154.33 | -1888.57 | 1895.29 | 15.93 | -20.74 | 0.84 | 26.26 | 1888.57 | 20.74 | 11.3332
0.1237 | 19.22 | -155.42 | -1903.12 | 1909.88 | 16.07 | -21.06 | 0.87 | 26.59 | 1903.12 | 21.06 | 11.5006
0.1255 | 18.94 | -156.35 | -1916.83 | 1923.61 | 16.19 | -21.37 | 0.9 | 26.91 | 1916.83 | 21.37 | 11.6679
0.1273 | 18.62 | -157.22 | -1929.53 | 1936.34 | 16.3 | -21.69 | 0.93 | 27.24 | 1929.53 | 21.69 | 11.8353
0.1291 | 18.35 | -157.92 | -1940.47 | 1947.3 | 16.38 | -21.99 | 0.96 | 27.53 | 1940.47 | 21.99 | 12.0026
0.1309 | 18.03 | -158.64 | -1950.01 | 1956.86 | 16.45 | -22.26 | 0.99 | 27.79 | 1950.01 | 22.26 | 12.17
0.1327 | 17.59 | -159.51 | -1958.4 | 1965.3 | 16.51 | -22.52 | 1.02 | 28.04 | 1958.4 | 22.52 | 12.3373
0.1345 | 17.11 | -160.35 | -1965.98 | 1972.91 | 16.57 | -22.76 | 1.04 | 28.27 | 1965.98 | 22.76 | 12.5046
0.1362 | 16.63 | -161.02 | -1972.8 | 1979.77 | 16.62 | -22.96 | 1.06 | 28.46 | 1972.8 | 22.96 | 12.6627
0.138 | 16.19 | -161.38 | -1979.1 | 1986.07 | 16.65 | -23.12 | 1.09 | 28.62 | 1979.1 | 23.12 | 12.83
0.1398 | 15.65 | -161.84 | -1983.24 | 1990.23 | 16.66 | -23.3 | 1.11 | 28.77 | 1983.24 | 23.3 | 12.9974
0.1416 | 15.03 | -162.14 | -1986.29 | 1993.29 | 16.65 | -23.48 | 1.14 | 28.91 | 1986.29 | 23.48 | 13.1647
0.1434 | 14.57 | -162.31 | -1989.37 | 1996.37 | 16.6 | -23.66 | 1.17 | 29.04 | 1989.37 | 23.66 | 13.3321
0.1452 | 14.06 | -162.55 | -1992.32 | 1999.32 | 16.57 | -23.83 | 1.2 | 29.17 | 1992.32 | 23.83 | 13.4994
0.147 | 13.53 | -162.76 | -1994.85 | 2001.85 | 16.54 | -23.98 | 1.23 | 29.28 | 1994.85 | 23.98 | 13.6668
0.1488 | 12.88 | -162.65 | -1996.51 | 2003.49 | 16.5 | -24.05 | 1.25 | 29.31 | 1996.51 | 24.05 | 13.8341
0.1506 | 12.12 | -162.38 | -1996.22 | 2003.19 | 16.42 | -24.15 | 1.28 | 29.36 | 1996.22 | 24.15 | 14.0015
0.1524 | 11.54 | -161.94 | -1995.67 | 2002.61 | 16.31 | -24.26 | 1.29 | 29.4 | 1995.67 | 24.26 | 14.1688
0.1542 | 10.81 | -161.33 | -1994.93 | 2001.83 | 16.21 | -24.34 | 1.31 | 29.42 | 1994.93 | 24.34 | 14.3362
0.156 | 9.96 | -160.62 | -1993.91 | 2000.78 | 16.12 | -24.41 | 1.32 | 29.43 | 1993.91 | 24.41 | 14.5035
0.1578 | 9.21 | -160.15 | -1992.19 | 1999.03 | 16.02 | -24.48 | 1.34 | 29.44 | 1992.19 | 24.48 | 14.6709
0.1595 | 8.68 | -159.65 | -1989.92 | 1996.73 | 15.89 | -24.54 | 1.35 | 29.42 | 1989.92 | 24.54 | 14.8289
0.1613 | 8.25 | -158.99 | -1987.17 | 1993.95 | 15.75 | -24.56 | 1.36 | 29.38 | 1987.17 | 24.56 | 14.9963
0.1631 | 7.62 | -158.07 | -1983.94 | 1990.67 | 15.61 | -24.56 | 1.38 | 29.3 | 1983.94 | 24.56 | 15.1636
0.1649 | 7 | -157.07 | -1980.58 | 1987.25 | 15.46 | -24.56 | 1.39 | 29.22 | 1980.58 | 24.56 | 15.331
0.1667 | 6.33 | -156.1 | -1976.92 | 1983.54 | 15.33 | -24.56 | 1.4 | 29.15 | 1976.92 | 24.56 | 15.4983
0.1685 | 5.56 | -154.96 | -1972.08 | 1978.63 | 15.17 | -24.54 | 1.42 | 29.06 | 1972.08 | 24.54 | 15.6657
0.1703 | 4.58 | -153.25 | -1966.24 | 1972.69 | 14.98 | -24.47 | 1.43 | 28.9 | 1966.24 | 24.47 | 15.833
0.1721 | 3.79 | -151.42 | -1960.32 | 1966.65 | 14.76 | -24.39 | 1.43 | 28.72 | 1960.32 | 24.39 | 16.0004
0.1739 | 3.45 | -149.8 | -1954.18 | 1960.41 | 14.5 | -24.33 | 1.42 | 28.54 | 1954.18 | 24.33 | 16.1677
0.1757 | 3.03 | -148.32 | -1946.89 | 1953.03 | 14.25 | -24.26 | 1.42 | 28.36 | 1946.89 | 24.26 | 16.3351
0.1775 | 2.82 | -146.52 | -1939.06 | 1945.09 | 13.99 | -24.2 | 1.4 | 28.18 | 1939.06 | 24.2 | 16.5024
0.1793 | 2.83 | -144.64 | -1931.14 | 1937.06 | 13.73 | -24.14 | 1.38 | 28.02 | 1931.14 | 24.14 | 16.6698
0.1811 | 2.92 | -142.87 | -1923.19 | 1929.01 | 13.48 | -24.09 | 1.37 | 27.86 | 1923.19 | 24.09 | 16.8371
0.1829 | 2.91 | -141.3 | -1914.51 | 1920.23 | 13.23 | -23.98 | 1.36 | 27.64 | 1914.51 | 23.98 | 17.0045
0.1846 | 2.68 | -139.85 | -1905.31 | 1910.96 | 12.96 | -23.91 | 1.36 | 27.44 | 1905.31 | 23.91 | 17.1625
0.1864 | 2.46 | -138.02 | -1894.59 | 1900.13 | 12.66 | -23.88 | 1.35 | 27.27 | 1894.59 | 23.88 | 17.3299
0.1882 | 2.5 | -136.19 | -1884.13 | 1889.56 | 12.38 | -23.82 | 1.34 | 27.07 | 1884.13 | 23.82 | 17.4972
0.19 | 2.53 | -134.6 | -1875.62 | 1880.96 | 12.13 | -23.79 | 1.33 | 26.93 | 1875.62 | 23.79 | 17.6646
0.1918 | 2.6 | -132.55 | -1867.14 | 1872.36 | 11.87 | -23.76 | 1.31 | 26.77 | 1867.14 | 23.76 | 17.8319
0.1936 | 2.81 | -130.56 | -1859.47 | 1864.57 | 11.6 | -23.75 | 1.29 | 26.64 | 1859.47 | 23.75 | 17.9993
0.1954 | 3.11 | -128.58 | -1852.21 | 1857.19 | 11.33 | -23.74 | 1.27 | 26.51 | 1852.21 | 23.74 | 18.1666
0.1972 | 3.52 | -126.73 | -1845.05 | 1849.93 | 11.09 | -23.75 | 1.24 | 26.41 | 1845.05 | 23.75 | 18.334
0.199 | 3.91 | -125.01 | -1837.47 | 1842.26 | 10.86 | -23.76 | 1.21 | 26.33 | 1837.47 | 23.76 | 18.5013
0.2008 | 4.29 | -123.23 | -1830.03 | 1834.73 | 10.64 | -23.77 | 1.19 | 26.25 | 1830.03 | 23.77 | 18.6687
0.2026 | 4.64 | -121.18 | -1823.34 | 1827.94 | 10.43 | -23.78 | 1.16 | 26.17 | 1823.34 | 23.78 | 18.836
0.2044 | 5.01 | -119.21 | -1817.31 | 1821.79 | 10.2 | -23.77 | 1.15 | 26.06 | 1817.31 | 23.77 | 19.0033
0.2062 | 5.59 | -117.66 | -1812.94 | 1817.34 | 9.99 | -23.74 | 1.12 | 25.95 | 1812.94 | 23.74 | 19.1707
0.208 | 6.34 | -116.39 | -1809.79 | 1814.13 | 9.82 | -23.69 | 1.1 | 25.84 | 1809.79 | 23.69 | 19.338
0.2097 | 7.15 | -114.85 | -1807.61 | 1811.87 | 9.64 | -23.66 | 1.08 | 25.74 | 1807.61 | 23.66 | 19.4961
0.2115 | 7.96 | -113.18 | -1805.38 | 1809.55 | 9.45 | -23.64 | 1.05 | 25.64 | 1805.38 | 23.64 | 19.6634
0.2133 | 8.86 | -111.5 | -1800.84 | 1804.92 | 9.27 | -23.69 | 1.03 | 25.63 | 1800.84 | 23.69 | 19.8308
0.2151 | 9.61 | -109.87 | -1795.41 | 1799.43 | 9.09 | -23.68 | 1 | 25.56 | 1795.41 | 23.68 | 19.9981
0.2169 | 10.33 | -107.96 | -1792.37 | 1796.31 | 8.96 | -23.73 | 0.97 | 25.57 | 1792.37 | 23.73 | 20.1655
0.2187 | 11.03 | -106.16 | -1789.56 | 1793.45 | 8.85 | -23.77 | 0.94 | 25.58 | 1789.56 | 23.77 | 20.3328
0.2205 | 11.75 | -104.51 | -1787.63 | 1791.45 | 8.76 | -23.81 | 0.91 | 25.58 | 1787.63 | 23.81 | 20.5002
0.2223 | 12.35 | -102.84 | -1785.63 | 1789.37 | 8.68 | -23.85 | 0.89 | 25.59 | 1785.63 | 23.85 | 20.6675
0.2241 | 12.63 | -101.19 | -1782.88 | 1786.56 | 8.6 | -23.9 | 0.86 | 25.61 | 1782.88 | 23.9 | 20.8349
0.2259 | 12.8 | -99.61 | -1779.03 | 1782.63 | 8.5 | -23.95 | 0.83 | 25.62 | 1779.03 | 23.95 | 21.0022
0.2277 | 13.06 | -98.25 | -1774.87 | 1778.4 | 8.39 | -24.01 | 0.79 | 25.64 | 1774.87 | 24.01 | 21.1696
0.2295 | 13.25 | -96.73 | -1770.51 | 1773.97 | 8.28 | -24.05 | 0.75 | 25.65 | 1770.51 | 24.05 | 21.3369
0.2313 | 13.47 | -95.01 | -1765.01 | 1768.37 | 8.14 | -24.08 | 0.71 | 25.62 | 1765.01 | 24.08 | 21.5043
0.2331 | 13.81 | -93.02 | -1757.83 | 1761.07 | 7.96 | -24.12 | 0.68 | 25.58 | 1757.83 | 24.12 | 21.6716
0.2348 | 14.06 | -90.88 | -1749.87 | 1752.99 | 7.78 | -24.16 | 0.64 | 25.55 | 1749.87 | 24.16 | 21.8297
0.2366 | 14.1 | -88.77 | -1741.42 | 1744.41 | 7.59 | -24.18 | 0.62 | 25.5 | 1741.42 | 24.18 | 21.997
0.2384 | 14.2 | -86.84 | -1734.42 | 1737.3 | 7.44 | -24.19 | 0.58 | 25.46 | 1734.42 | 24.19 | 22.1644
0.2402 | 14.22 | -84.75 | -1728.74 | 1731.53 | 7.32 | -24.18 | 0.54 | 25.41 | 1728.74 | 24.18 | 22.3317
0.242 | 14.22 | -82.75 | -1723.99 | 1726.71 | 7.22 | -24.15 | 0.51 | 25.36 | 1723.99 | 24.15 | 22.4991
0.2438 | 14.31 | -81.04 | -1719.77 | 1722.43 | 7.14 | -24.14 | 0.46 | 25.33 | 1719.77 | 24.14 | 22.6664
0.2456 | 14.65 | -79.84 | -1715.13 | 1717.74 | 7.04 | -24.14 | 0.41 | 25.3 | 1715.13 | 24.14 | 22.8338
0.2474 | 14.73 | -78.23 | -1709.96 | 1712.49 | 6.94 | -24.11 | 0.37 | 25.24 | 1709.96 | 24.11 | 23.0011
0.2492 | 14.83 | -76.18 | -1703.26 | 1705.68 | 6.81 | -24.1 | 0.33 | 25.18 | 1703.26 | 24.1 | 23.1685
0.251 | 14.83 | -74.03 | -1695.66 | 1697.97 | 6.67 | -24.09 | 0.3 | 25.13 | 1695.66 | 24.09 | 23.3358
0.2528 | 14.7 | -71.9 | -1687.92 | 1690.13 | 6.53 | -24.07 | 0.26 | 25.07 | 1687.92 | 24.07 | 23.5032
0.2546 | 14.63 | -69.8 | -1680.17 | 1682.29 | 6.39 | -24.06 | 0.22 | 25.01 | 1680.17 | 24.06 | 23.6705
0.2564 | 14.68 | -67.35 | -1672.08 | 1674.08 | 6.25 | -24.08 | 0.19 | 24.99 | 1672.08 | 24.08 | 23.8379
0.2581 | 14.83 | -64.94 | -1664.23 | 1666.11 | 6.11 | -24.09 | 0.15 | 24.96 | 1664.23 | 24.09 | 23.9959
0.2599 | 15.11 | -63.01 | -1657.15 | 1658.94 | 5.99 | -24.08 | 0.11 | 24.91 | 1657.15 | 24.08 | 24.1633
0.2617 | 15.36 | -61.13 | -1650.1 | 1651.8 | 5.85 | -24.08 | 0.07 | 24.87 | 1650.1 | 24.08 | 24.3306
0.2635 | 15.47 | -59.07 | -1643.71 | 1645.31 | 5.72 | -24.08 | 0.03 | 24.83 | 1643.71 | 24.08 | 24.498
0.2653 | 15.69 | -57.01 | -1638.27 | 1639.79 | 5.61 | -24.05 | -0.02 | 24.77 | 1638.27 | 24.05 | 24.6653
0.2671 | 15.87 | -54.87 | -1632.84 | 1634.26 | 5.51 | -24.01 | -0.06 | 24.71 | 1632.84 | 24.01 | 24.8327
0.2689 | 15.95 | -52.74 | -1627.34 | 1628.68 | 5.41 | -23.94 | -0.11 | 24.62 | 1627.34 | 23.94 | 25
0.2707 | 16.27 | -50.92 | -1622.53 | 1623.8 | 5.31 | -23.89 | -0.16 | 24.55 | 1622.53 | 23.89 | 25.1673
0.2725 | 16.62 | -49.24 | -1618.24 | 1619.46 | 5.23 | -23.82 | -0.2 | 24.46 | 1618.24 | 23.82 | 25.3347
0.2743 | 16.91 | -47.58 | -1614.13 | 1615.29 | 5.15 | -23.76 | -0.25 | 24.38 | 1614.13 | 23.76 | 25.502
0.2761 | 17.04 | -45.9 | -1610.65 | 1611.76 | 5.08 | -23.68 | -0.29 | 24.3 | 1610.65 | 23.68 | 25.6694
0.2779 | 17.14 | -44.26 | -1607.12 | 1608.18 | 5 | -23.62 | -0.34 | 24.22 | 1607.12 | 23.62 | 25.8367
0.2797 | 17.26 | -42.63 | -1603.48 | 1604.49 | 4.91 | -23.56 | -0.4 | 24.14 | 1603.48 | 23.56 | 26.0041
0.2815 | 17.27 | -41.03 | -1599.29 | 1600.23 | 4.82 | -23.48 | -0.45 | 24.05 | 1599.29 | 23.48 | 26.1714
0.2832 | 17.35 | -39.6 | -1595.49 | 1596.4 | 4.74 | -23.4 | -0.5 | 23.95 | 1595.49 | 23.4 | 26.3295
0.285 | 17.49 | -38.08 | -1592.08 | 1592.95 | 4.66 | -23.34 | -0.56 | 23.87 | 1592.08 | 23.34 | 26.4968
0.2868 | 17.82 | -36.56 | -1589.25 | 1590.08 | 4.59 | -23.27 | -0.61 | 23.8 | 1589.25 | 23.27 | 26.6642
0.2886 | 18.18 | -35.1 | -1586.75 | 1587.55 | 4.52 | -23.21 | -0.66 | 23.72 | 1586.75 | 23.21 | 26.8315
0.2904 | 18.43 | -33.66 | -1584.21 | 1584.99 | 4.46 | -23.14 | -0.72 | 23.65 | 1584.21 | 23.14 | 26.9989
0.2922 | 18.66 | -32.19 | -1581.49 | 1582.25 | 4.39 | -23.07 | -0.77 | 23.57 | 1581.49 | 23.07 | 27.1662
0.294 | 18.77 | -30.84 | -1578.72 | 1579.45 | 4.34 | -22.99 | -0.83 | 23.48 | 1578.72 | 22.99 | 27.3336
0.2958 | 18.96 | -29.44 | -1576.37 | 1577.07 | 4.29 | -22.91 | -0.88 | 23.4 | 1576.37 | 22.91 | 27.5009
0.2976 | 19.17 | -28 | -1574.33 | 1574.99 | 4.23 | -22.83 | -0.93 | 23.31 | 1574.33 | 22.83 | 27.6683
0.2994 | 19.3 | -26.48 | -1572.8 | 1573.44 | 4.18 | -22.77 | -0.99 | 23.24 | 1572.8 | 22.77 | 27.8356
0.3012 | 19.38 | -25 | -1571.45 | 1572.05 | 4.13 | -22.7 | -1.04 | 23.16 | 1571.45 | 22.7 | 28.003
0.303 | 19.49 | -23.59 | -1570.24 | 1570.81 | 4.09 | -22.61 | -1.09 | 23.07 | 1570.24 | 22.61 | 28.1703
0.3048 | 19.57 | -22.21 | -1569.09 | 1569.64 | 4.05 | -22.51 | -1.15 | 22.96 | 1569.09 | 22.51 | 28.3377
0.3066 | 19.6 | -20.83 | -1568.01 | 1568.54 | 4.01 | -22.37 | -1.2 | 22.83 | 1568.01 | 22.37 | 28.505
0.3083 | 19.63 | -19.56 | -1566.97 | 1567.47 | 3.99 | -22.26 | -1.25 | 22.71 | 1566.97 | 22.26 | 28.6631
0.3101 | 19.59 | -18.36 | -1566.15 | 1566.64 | 3.95 | -22.15 | -1.31 | 22.61 | 1566.15 | 22.15 | 28.8304
0.3119 | 19.49 | -17.21 | -1565.37 | 1565.85 | 3.92 | -22.06 | -1.37 | 22.51 | 1565.37 | 22.06 | 28.9978
0.3137 | 19.36 | -16.01 | -1564.64 | 1565.1 | 3.9 | -21.98 | -1.43 | 22.43 | 1564.64 | 21.98 | 29.1651
0.3155 | 19.34 | -14.56 | -1564.13 | 1564.56 | 3.87 | -21.9 | -1.49 | 22.35 | 1564.13 | 21.9 | 29.3325
0.3173 | 19.33 | -13.06 | -1563.66 | 1564.06 | 3.84 | -21.81 | -1.55 | 22.27 | 1563.66 | 21.81 | 29.4998
0.3191 | 19.25 | -11.68 | -1563.6 | 1563.98 | 3.82 | -21.74 | -1.61 | 22.19 | 1563.6 | 21.74 | 29.6672
0.3209 | 19.13 | -10.34 | -1563.56 | 1563.92 | 3.8 | -21.66 | -1.67 | 22.12 | 1563.56 | 21.66 | 29.8345
0.3227 | 19.03 | -8.89 | -1563.76 | 1564.11 | 3.78 | -21.59 | -1.73 | 22.05 | 1563.76 | 21.59 | 30.0019
0.3245 | 19 | -7.4 | -1564.06 | 1564.37 | 3.77 | -21.52 | -1.8 | 21.98 | 1564.06 | 21.52 | 30.1692
0.3263 | 18.93 | -5.9 | -1564.35 | 1564.64 | 3.75 | -21.45 | -1.88 | 21.91 | 1564.35 | 21.45 | 30.3366
0.3281 | 18.8 | -4.51 | -1564.96 | 1565.24 | 3.73 | -21.39 | -1.95 | 21.86 | 1564.96 | 21.39 | 30.5039
0.3299 | 18.64 | -3.22 | -1565.82 | 1566.09 | 3.71 | -21.34 | -2.02 | 21.81 | 1565.82 | 21.34 | 30.6713
0.3316 | 18.56 | -2.02 | -1566.74 | 1566.99 | 3.69 | -21.28 | -2.09 | 21.76 | 1566.74 | 21.28 | 30.8293
0.3334 | 18.52 | -0.79 | -1567.97 | 1568.21 | 3.68 | -21.22 | -2.15 | 21.69 | 1567.97 | 21.22 | 30.9967
0.3352 | 18.46 | 0.4 | -1569.41 | 1569.64 | 3.67 | -21.15 | -2.23 | 21.63 | 1569.41 | 21.15 | 31.164
0.337 | 18.43 | 1.52 | -1571.03 | 1571.25 | 3.66 | -21.07 | -2.29 | 21.57 | 1571.03 | 21.07 | 31.3313
0.3388 | 18.4 | 2.63 | -1572.7 | 1572.92 | 3.66 | -21 | -2.36 | 21.51 | 1572.7 | 21 | 31.4987
0.3406 | 18.37 | 3.79 | -1574.67 | 1574.89 | 3.65 | -20.94 | -2.42 | 21.46 | 1574.67 | 20.94 | 31.666
0.3424 | 18.41 | 4.88 | -1576.92 | 1577.14 | 3.64 | -20.89 | -2.49 | 21.41 | 1576.92 | 20.89 | 31.8334
0.3442 | 18.48 | 5.93 | -1579.42 | 1579.63 | 3.63 | -20.84 | -2.57 | 21.37 | 1579.42 | 20.84 | 32.0007
0.346 | 18.51 | 7.03 | -1582.08 | 1582.29 | 3.63 | -20.79 | -2.64 | 21.33 | 1582.08 | 20.79 | 32.1681
0.3478 | 18.47 | 8.09 | -1584.92 | 1585.13 | 3.62 | -20.73 | -2.7 | 21.28 | 1584.92 | 20.73 | 32.3354
0.3496 | 18.41 | 9.12 | -1587.94 | 1588.16 | 3.61 | -20.67 | -2.77 | 21.23 | 1587.94 | 20.67 | 32.5028
0.3514 | 18.36 | 10.11 | -1591.42 | 1591.64 | 3.6 | -20.61 | -2.85 | 21.18 | 1591.42 | 20.61 | 32.6701
0.3532 | 18.33 | 11.14 | -1595.58 | 1595.8 | 3.61 | -20.54 | -2.93 | 21.12 | 1595.58 | 20.54 | 32.8375
0.355 | 18.33 | 12.12 | -1599.92 | 1600.15 | 3.62 | -20.48 | -3.02 | 21.07 | 1599.92 | 20.48 | 33.0048
0.3567 | 18.35 | 13.02 | -1604.42 | 1604.65 | 3.63 | -20.42 | -3.1 | 21.03 | 1604.42 | 20.42 | 33.1629
0.3585 | 18.35 | 13.86 | -1608.9 | 1609.14 | 3.64 | -20.36 | -3.18 | 20.99 | 1608.9 | 20.36 | 33.3302
0.3603 | 18.3 | 14.62 | -1614.11 | 1614.36 | 3.66 | -20.31 | -3.27 | 20.96 | 1614.11 | 20.31 | 33.4976
0.3621 | 18.21 | 15.39 | -1620.15 | 1620.4 | 3.71 | -20.26 | -3.37 | 20.93 | 1620.15 | 20.26 | 33.6649
0.3639 | 18.14 | 16.26 | -1625.2 | 1625.47 | 3.73 | -20.21 | -3.45 | 20.91 | 1625.2 | 20.21 | 33.8323
0.3657 | 18.03 | 17.13 | -1630.16 | 1630.42 | 3.75 | -20.17 | -3.53 | 20.88 | 1630.16 | 20.17 | 33.9996
0.3675 | 17.87 | 18.12 | -1635.21 | 1635.48 | 3.77 | -20.11 | -3.61 | 20.85 | 1635.21 | 20.11 | 34.167
0.3693 | 17.72 | 19.11 | -1640.28 | 1640.56 | 3.8 | -20.06 | -3.68 | 20.82 | 1640.28 | 20.06 | 34.3343
0.3711 | 17.64 | 19.94 | -1645.76 | 1646.05 | 3.82 | -20.02 | -3.76 | 20.8 | 1645.76 | 20.02 | 34.5017
0.3729 | 17.6 | 20.57 | -1651.31 | 1651.61 | 3.85 | -19.97 | -3.84 | 20.77 | 1651.31 | 19.97 | 34.669
0.3747 | 17.52 | 21.16 | -1657.35 | 1657.65 | 3.88 | -19.94 | -3.92 | 20.76 | 1657.35 | 19.94 | 34.8364
0.3765 | 17.46 | 21.75 | -1663.26 | 1663.57 | 3.9 | -19.9 | -4.01 | 20.75 | 1663.26 | 19.9 | 35.0037
0.3783 | 17.29 | 22.11 | -1668.76 | 1669.07 | 3.92 | -19.86 | -4.08 | 20.72 | 1668.76 | 19.86 | 35.1711
0.3801 | 17.25 | 22.66 | -1673.62 | 1673.94 | 3.94 | -19.82 | -4.16 | 20.7 | 1673.62 | 19.82 | 35.3384
0.3818 | 17.11 | 23.25 | -1677.91 | 1678.22 | 3.96 | -19.78 | -4.23 | 20.68 | 1677.91 | 19.78 | 35.4965
0.3836 | 16.92 | 23.76 | -1681.33 | 1681.65 | 3.98 | -19.74 | -4.3 | 20.66 | 1681.33 | 19.74 | 35.6638
0.3854 | 16.75 | 24.19 | -1684.43 | 1684.74 | 4 | -19.7 | -4.38 | 20.63 | 1684.43 | 19.7 | 35.8312
0.3872 | 16.57 | 24.47 | -1687.23 | 1687.54 | 4.01 | -19.67 | -4.45 | 20.62 | 1687.23 | 19.67 | 35.9985
0.389 | 16.39 | 24.69 | -1690.04 | 1690.35 | 4.01 | -19.65 | -4.53 | 20.61 | 1690.04 | 19.65 | 36.1659
0.3908 | 16.18 | 25.01 | -1694.27 | 1694.58 | 4.01 | -19.63 | -4.6 | 20.62 | 1694.27 | 19.63 | 36.3332
0.3926 | 16.13 | 25.51 | -1699.21 | 1699.53 | 4.02 | -19.61 | -4.68 | 20.61 | 1699.21 | 19.61 | 36.5006
0.3944 | 16.12 | 25.98 | -1704.45 | 1704.78 | 4.05 | -19.59 | -4.76 | 20.62 | 1704.45 | 19.59 | 36.6679
0.3962 | 16.12 | 26.49 | -1710.15 | 1710.48 | 4.08 | -19.57 | -4.85 | 20.62 | 1710.15 | 19.57 | 36.8353
0.398 | 16.06 | 27.06 | -1716.37 | 1716.71 | 4.11 | -19.53 | -4.94 | 20.62 | 1716.37 | 19.53 | 37.0026
0.3998 | 16.06 | 27.81 | -1722.91 | 1723.27 | 4.15 | -19.51 | -5.03 | 20.63 | 1722.91 | 19.51 | 37.17
0.4016 | 16.06 | 28.74 | -1729.52 | 1729.89 | 4.19 | -19.49 | -5.12 | 20.64 | 1729.52 | 19.49 | 37.3373
0.4034 | 16.02 | 29.61 | -1736.34 | 1736.73 | 4.23 | -19.46 | -5.2 | 20.65 | 1736.34 | 19.46 | 37.5046
0.4051 | 15.98 | 30.33 | -1742.5 | 1742.9 | 4.27 | -19.44 | -5.29 | 20.66 | 1742.5 | 19.44 | 37.6627
0.4069 | 15.97 | 31.04 | -1749.23 | 1749.64 | 4.31 | -19.42 | -5.37 | 20.67 | 1749.23 | 19.42 | 37.83
0.4087 | 16.01 | 31.77 | -1757.01 | 1757.43 | 4.37 | -19.4 | -5.46 | 20.69 | 1757.01 | 19.4 | 37.9974
0.4105 | 16.12 | 32.6 | -1765.48 | 1765.91 | 4.43 | -19.38 | -5.56 | 20.72 | 1765.48 | 19.38 | 38.1647
0.4123 | 16.28 | 33.64 | -1774.92 | 1775.38 | 4.5 | -19.37 | -5.65 | 20.75 | 1774.92 | 19.37 | 38.3321
0.4141 | 16.46 | 34.56 | -1784.03 | 1784.51 | 4.58 | -19.36 | -5.73 | 20.79 | 1784.03 | 19.36 | 38.4994
0.4159 | 16.57 | 35.43 | -1792.64 | 1793.13 | 4.65 | -19.36 | -5.82 | 20.83 | 1792.64 | 19.36 | 38.6668
0.4177 | 16.58 | 36.1 | -1800.6 | 1801.1 | 4.71 | -19.35 | -5.91 | 20.87 | 1800.6 | 19.35 | 38.8341
0.4195 | 16.57 | 36.7 | -1808.09 | 1808.6 | 4.78 | -19.35 | -5.99 | 20.91 | 1808.09 | 19.35 | 39.0015
0.4213 | 16.52 | 37.08 | -1814.22 | 1814.74 | 4.82 | -19.37 | -6.08 | 20.95 | 1814.22 | 19.37 | 39.1688
0.4231 | 16.25 | 37.41 | -1820.36 | 1820.88 | 4.87 | -19.36 | -6.16 | 20.98 | 1820.36 | 19.36 | 39.3362
0.4249 | 15.95 | 37.73 | -1826.55 | 1827.08 | 4.92 | -19.36 | -6.25 | 21.01 | 1826.55 | 19.36 | 39.5035
0.4267 | 15.81 | 38.06 | -1834.06 | 1834.6 | 4.98 | -19.35 | -6.34 | 21.05 | 1834.06 | 19.35 | 39.6709
0.4285 | 15.67 | 38.33 | -1841.22 | 1841.75 | 5.04 | -19.37 | -6.43 | 21.1 | 1841.22 | 19.37 | 39.8382
0.4302 | 15.56 | 38.62 | -1848.56 | 1849.1 | 5.1 | -19.38 | -6.52 | 21.16 | 1848.56 | 19.38 | 39.9963
0.432 | 15.5 | 38.94 | -1856.61 | 1857.16 | 5.17 | -19.38 | -6.62 | 21.21 | 1856.61 | 19.38 | 40.1636
0.4338 | 15.51 | 39.34 | -1865.4 | 1865.95 | 5.26 | -19.38 | -6.72 | 21.27 | 1865.4 | 19.38 | 40.331
0.4356 | 15.5 | 39.79 | -1874.61 | 1875.17 | 5.36 | -19.4 | -6.81 | 21.33 | 1874.61 | 19.4 | 40.4983
0.4374 | 15.45 | 40.2 | -1884.18 | 1884.74 | 5.46 | -19.4 | -6.9 | 21.39 | 1884.18 | 19.4 | 40.6657
0.4392 | 15.34 | 40.67 | -1894.29 | 1894.86 | 5.58 | -19.38 | -7 | 21.44 | 1894.29 | 19.38 | 40.833
0.441 | 15.34 | 41.22 | -1904.99 | 1905.57 | 5.71 | -19.38 | -7.1 | 21.51 | 1904.99 | 19.38 | 41.0004
0.4428 | 15.37 | 41.83 | -1916.47 | 1917.07 | 5.85 | -19.38 | -7.2 | 21.58 | 1916.47 | 19.38 | 41.1677
0.4446 | 15.37 | 42.35 | -1928.17 | 1928.78 | 6.01 | -19.37 | -7.29 | 21.66 | 1928.17 | 19.37 | 41.3351
0.4464 | 15.31 | 42.75 | -1938.76 | 1939.38 | 6.17 | -19.36 | -7.38 | 21.74 | 1938.76 | 19.36 | 41.5024
0.4482 | 15.18 | 43.16 | -1948.02 | 1948.64 | 6.3 | -19.35 | -7.48 | 21.8 | 1948.02 | 19.35 | 41.6698
0.45 | 14.98 | 43.4 | -1956.54 | 1957.17 | 6.42 | -19.33 | -7.57 | 21.86 | 1956.54 | 19.33 | 41.8371
0.4518 | 14.72 | 43.45 | -1963.85 | 1964.47 | 6.52 | -19.32 | -7.66 | 21.9 | 1963.85 | 19.32 | 42.0045
0.4536 | 14.51 | 43.53 | -1971.31 | 1971.93 | 6.62 | -19.3 | -7.75 | 21.95 | 1971.31 | 19.3 | 42.1718
0.4553 | 14.53 | 43.73 | -1980.04 | 1980.67 | 6.75 | -19.29 | -7.84 | 22.01 | 1980.04 | 19.29 | 42.3299
0.4571 | 14.62 | 43.96 | -1990.52 | 1991.15 | 6.91 | -19.29 | -7.93 | 22.1 | 1990.52 | 19.29 | 42.4972
0.4589 | 14.75 | 44.15 | -2001.06 | 2001.7 | 7.08 | -19.3 | -8.02 | 22.19 | 2001.06 | 19.3 | 42.6646
0.4607 | 14.65 | 44.35 | -2010.44 | 2011.08 | 7.26 | -19.29 | -8.11 | 22.27 | 2010.44 | 19.29 | 42.8319
0.4625 | 14.4 | 44.4 | -2018.65 | 2019.29 | 7.43 | -19.27 | -8.2 | 22.34 | 2018.65 | 19.27 | 42.9993
0.4643 | 14.12 | 44.43 | -2026.4 | 2027.04 | 7.61 | -19.25 | -8.29 | 22.41 | 2026.4 | 19.25 | 43.1666
0.4661 | 13.83 | 44.52 | -2034.96 | 2035.6 | 7.79 | -19.23 | -8.38 | 22.49 | 2034.96 | 19.23 | 43.334
0.4679 | 13.59 | 44.7 | -2045.09 | 2045.74 | 7.98 | -19.19 | -8.47 | 22.56 | 2045.09 | 19.19 | 43.5013
0.4697 | 13.38 | 44.93 | -2055.95 | 2056.6 | 8.18 | -19.16 | -8.55 | 22.64 | 2055.95 | 19.16 | 43.6687
0.4715 | 13.11 | 45.28 | -2067.5 | 2068.15 | 8.38 | -19.11 | -8.63 | 22.71 | 2067.5 | 19.11 | 43.836
0.4733 | 12.81 | 45.54 | -2079.27 | 2079.93 | 8.59 | -19.06 | -8.72 | 22.79 | 2079.27 | 19.06 | 44.0033
0.4751 | 12.55 | 45.75 | -2089.89 | 2090.55 | 8.81 | -19.02 | -8.8 | 22.87 | 2089.89 | 19.02 | 44.1707
0.4769 | 12.28 | 45.88 | -2100.1 | 2100.76 | 9.02 | -18.97 | -8.88 | 22.95 | 2100.1 | 18.97 | 44.338
0.4786 | 11.87 | 45.84 | -2110.04 | 2110.7 | 9.23 | -18.91 | -8.96 | 23.02 | 2110.04 | 18.91 | 44.4961
0.4804 | 11.28 | 45.83 | -2120.35 | 2121 | 9.44 | -18.82 | -9.03 | 23.07 | 2120.35 | 18.82 | 44.6634
0.4822 | 10.6 | 45.97 | -2130.79 | 2131.43 | 9.66 | -18.72 | -9.1 | 23.12 | 2130.79 | 18.72 | 44.8308
0.484 | 9.92 | 46.05 | -2140.14 | 2140.78 | 9.88 | -18.64 | -9.18 | 23.19 | 2140.14 | 18.64 | 44.9981
0.4858 | 9.32 | 46.02 | -2148.6 | 2149.24 | 10.07 | -18.58 | -9.25 | 23.25 | 2148.6 | 18.58 | 45.1655
0.4876 | 8.74 | 46.1 | -2156.69 | 2157.33 | 10.27 | -18.53 | -9.33 | 23.31 | 2156.69 | 18.53 | 45.3328
0.4894 | 8.14 | 46.26 | -2164.81 | 2165.45 | 10.47 | -18.47 | -9.41 | 23.37 | 2164.81 | 18.47 | 45.5002
0.4912 | 7.44 | 46.46 | -2173.45 | 2174.09 | 10.67 | -18.39 | -9.49 | 23.43 | 2173.45 | 18.39 | 45.6675
0.493 | 6.66 | 46.61 | -2181.71 | 2182.35 | 10.85 | -18.31 | -9.56 | 23.48 | 2181.71 | 18.31 | 45.8349
0.4948 | 5.82 | 46.78 | -2188.67 | 2189.31 | 11.04 | -18.21 | -9.63 | 23.52 | 2188.67 | 18.21 | 46.0022
0.4966 | 4.94 | 47.01 | -2194.74 | 2195.39 | 11.24 | -18.1 | -9.7 | 23.57 | 2194.74 | 18.1 | 46.1696
0.4984 | 3.94 | 47.53 | -2200.32 | 2200.98 | 11.43 | -17.99 | -9.76 | 23.6 | 2200.32 | 17.99 | 46.3369
0.5002 | 2.88 | 48.14 | -2205.81 | 2206.49 | 11.6 | -17.87 | -9.83 | 23.63 | 2205.81 | 17.87 | 46.5043
0.502 | 1.85 | 48.79 | -2211.19 | 2211.88 | 11.78 | -17.76 | -9.88 | 23.65 | 2211.19 | 17.76 | 46.6716
0.5037 | 0.82 | 49.46 | -2216.8 | 2217.5 | 11.95 | -17.65 | -9.94 | 23.68 | 2216.8 | 17.65 | 46.8297
0.5055 | -0.18 | 49.98 | -2223.57 | 2224.29 | 12.1 | -17.53 | -10 | 23.69 | 2223.57 | 17.53 | 46.997
0.5073 | -1.25 | 50.34 | -2231.15 | 2231.87 | 12.25 | -17.38 | -10.07 | 23.69 | 2231.15 | 17.38 | 47.1644
0.5091 | -2.31 | 50.63 | -2239.27 | 2240 | 12.41 | -17.23 | -10.13 | 23.69 | 2239.27 | 17.23 | 47.3317
0.5109 | -3.45 | 51.01 | -2247.72 | 2248.46 | 12.57 | -17.07 | -10.19 | 23.69 | 2247.72 | 17.07 | 47.4991
0.5127 | -4.57 | 51.26 | -2255.03 | 2255.77 | 12.7 | -16.91 | -10.25 | 23.67 | 2255.03 | 16.91 | 47.6664
0.5145 | -5.71 | 51.46 | -2261.63 | 2262.37 | 12.82 | -16.75 | -10.31 | 23.65 | 2261.63 | 16.75 | 47.8338
0.5163 | -6.82 | 51.56 | -2267.82 | 2268.57 | 12.94 | -16.59 | -10.36 | 23.63 | 2267.82 | 16.59 | 48.0011
0.5181 | -8.04 | 51.52 | -2273.24 | 2273.99 | 13.05 | -16.43 | -10.42 | 23.61 | 2273.24 | 16.43 | 48.1685
0.5199 | -9.22 | 51.51 | -2277.88 | 2278.63 | 13.17 | -16.27 | -10.47 | 23.58 | 2277.88 | 16.27 | 48.3358
0.5217 | -10.24 | 51.54 | -2281.93 | 2282.68 | 13.28 | -16.11 | -10.52 | 23.56 | 2281.93 | 16.11 | 48.5032
0.5235 | -11.13 | 51.62 | -2285.64 | 2286.4 | 13.39 | -15.97 | -10.57 | 23.54 | 2285.64 | 15.97 | 48.6705
0.5253 | -12.09 | 51.81 | -2289.18 | 2289.94 | 13.48 | -15.83 | -10.63 | 23.51 | 2289.18 | 15.83 | 48.8379
0.5271 | -13.38 | 51.67 | -2291.68 | 2292.46 | 13.57 | -15.66 | -10.67 | 23.47 | 2291.68 | 15.66 | 49.0052
0.5288 | -14.64 | 51.47 | -2293.51 | 2294.28 | 13.66 | -15.49 | -10.71 | 23.43 | 2293.51 | 15.49 | 49.1633
0.5306 | -15.9 | 51.62 | -2295.33 | 2296.11 | 13.78 | -15.31 | -10.74 | 23.4 | 2295.33 | 15.31 | 49.3306
0.5324 | -17.16 | 51.72 | -2296.91 | 2297.69 | 13.87 | -15.12 | -10.78 | 23.35 | 2296.91 | 15.12 | 49.498
0.5342 | -18.35 | 51.72 | -2298.04 | 2298.83 | 13.94 | -14.93 | -10.81 | 23.29 | 2298.04 | 14.93 | 49.6653
0.536 | -19.28 | 51.75 | -2298.7 | 2299.49 | 13.99 | -14.76 | -10.83 | 23.22 | 2298.7 | 14.76 | 49.8327
0.5378 | -20.69 | 51.78 | -2298.58 | 2299.38 | 14.03 | -14.57 | -10.87 | 23.14 | 2298.58 | 14.57 | 50
0.5396 | -22.16 | 51.93 | -2298.21 | 2299.03 | 14.09 | -14.38 | -10.9 | 23.07 | 2298.21 | 14.38 | 50.1673
0.5414 | -23.4 | 51.73 | -2297.53 | 2298.36 | 14.12 | -14.17 | -10.92 | 22.97 | 2297.53 | 14.17 | 50.3347
0.5432 | -24.72 | 51.63 | -2296.14 | 2296.99 | 14.14 | -13.95 | -10.94 | 22.86 | 2296.14 | 13.95 | 50.502
0.545 | -26.02 | 51.47 | -2294.16 | 2295.01 | 14.16 | -13.73 | -10.95 | 22.74 | 2294.16 | 13.73 | 50.6694
0.5468 | -27.34 | 51.28 | -2291.77 | 2292.63 | 14.17 | -13.49 | -10.96 | 22.62 | 2291.77 | 13.49 | 50.8367
0.5486 | -28.85 | 51.31 | -2288.9 | 2289.78 | 14.18 | -13.26 | -10.98 | 22.49 | 2288.9 | 13.26 | 51.0041
0.5504 | -30.21 | 51.28 | -2285.41 | 2286.31 | 14.17 | -13.03 | -10.99 | 22.36 | 2285.41 | 13.03 | 51.1714
0.5521 | -31.56 | 51.29 | -2281.1 | 2282.02 | 14.17 | -12.79 | -11 | 22.22 | 2281.1 | 12.79 | 51.3295
0.5539 | -32.91 | 51.26 | -2276.27 | 2277.21 | 14.16 | -12.54 | -11.01 | 22.07 | 2276.27 | 12.54 | 51.4968
0.5557 | -34.41 | 51.47 | -2271.03 | 2271.99 | 14.13 | -12.29 | -11.02 | 21.92 | 2271.03 | 12.29 | 51.6642
0.5575 | -35.87 | 51.62 | -2264.4 | 2265.4 | 14.1 | -12 | -11.02 | 21.73 | 2264.4 | 12 | 51.8315
0.5593 | -37.16 | 51.48 | -2256.72 | 2257.73 | 14.05 | -11.72 | -11 | 21.54 | 2256.72 | 11.72 | 51.9989
0.5611 | -38.48 | 51.26 | -2248.6 | 2249.64 | 14 | -11.44 | -10.98 | 21.34 | 2248.6 | 11.44 | 52.1662
0.5629 | -39.88 | 51.07 | -2240.17 | 2241.22 | 13.95 | -11.17 | -10.96 | 21.15 | 2240.17 | 11.17 | 52.3336
0.5647 | -41.34 | 50.92 | -2230.9 | 2231.98 | 13.89 | -10.88 | -10.93 | 20.94 | 2230.9 | 10.88 | 52.5009
0.5665 | -42.81 | 50.88 | -2220.16 | 2221.27 | 13.81 | -10.56 | -10.88 | 20.71 | 2220.16 | 10.56 | 52.6683
0.5683 | -44.26 | 50.71 | -2208.48 | 2209.62 | 13.73 | -10.24 | -10.82 | 20.45 | 2208.48 | 10.24 | 52.8356
0.5701 | -45.73 | 50.39 | -2195.78 | 2196.95 | 13.63 | -9.91 | -10.77 | 20.19 | 2195.78 | 9.91 | 53.003
0.5719 | -47.21 | 49.8 | -2181.01 | 2182.2 | 13.51 | -9.57 | -10.7 | 19.91 | 2181.01 | 9.57 | 53.1703
0.5737 | -48.85 | 49.24 | -2165.14 | 2166.37 | 13.38 | -9.21 | -10.63 | 19.62 | 2165.14 | 9.21 | 53.3377
0.5755 | -50.61 | 48.28 | -2147.57 | 2148.82 | 13.24 | -8.78 | -10.55 | 19.27 | 2147.57 | 8.78 | 53.505
0.5772 | -52.39 | 47.46 | -2129.25 | 2130.54 | 13.08 | -8.35 | -10.46 | 18.92 | 2129.25 | 8.35 | 53.6631
0.579 | -54.26 | 46.65 | -2109.47 | 2110.8 | 12.92 | -7.91 | -10.36 | 18.57 | 2109.47 | 7.91 | 53.8304
0.5808 | -56.19 | 45.66 | -2087.71 | 2089.08 | 12.73 | -7.42 | -10.24 | 18.17 | 2087.71 | 7.42 | 53.9978
0.5826 | -58.1 | 44.82 | -2064.73 | 2066.14 | 12.54 | -6.9 | -10.11 | 17.78 | 2064.73 | 6.9 | 54.1651
0.5844 | -59.97 | 44.08 | -2040.59 | 2042.05 | 12.34 | -6.35 | -9.98 | 17.39 | 2040.59 | 6.35 | 54.3325
0.5862 | -61.75 | 43.44 | -2015.87 | 2017.39 | 12.15 | -5.8 | -9.85 | 16.98 | 2015.87 | 5.8 | 54.4998
0.588 | -63.39 | 43.01 | -1990.63 | 1992.21 | 11.95 | -5.23 | -9.7 | 16.57 | 1990.63 | 5.23 | 54.6672
0.5898 | -64.95 | 42.68 | -1964.9 | 1966.54 | 11.74 | -4.66 | -9.56 | 16.16 | 1964.9 | 4.66 | 54.8345
0.5916 | -66.45 | 42.38 | -1938.73 | 1940.43 | 11.53 | -4.07 | -9.41 | 15.75 | 1938.73 | 4.07 | 55.0019
0.5934 | -67.9 | 42.13 | -1912.12 | 1913.89 | 11.31 | -3.46 | -9.25 | 15.34 | 1912.12 | 3.46 | 55.1692
0.5952 | -69.32 | 42.02 | -1884.58 | 1886.42 | 11.08 | -2.83 | -9.09 | 14.92 | 1884.58 | 2.83 | 55.3366
0.597 | -70.78 | 42.09 | -1855.77 | 1857.71 | 10.84 | -2.19 | -8.91 | 14.53 | 1855.77 | 2.19 | 55.5039
0.5988 | -72.25 | 42.25 | -1826.3 | 1828.33 | 10.61 | -1.54 | -8.72 | 14.13 | 1826.3 | 1.54 | 55.6713
0.6006 | -73.69 | 42.51 | -1795.78 | 1797.91 | 10.36 | -0.9 | -8.52 | 13.75 | 1795.78 | 0.9 | 55.8386
0.6023 | -75.07 | 42.89 | -1764.37 | 1766.6 | 10.11 | -0.27 | -8.31 | 13.4 | 1764.37 | 0.27 | 55.9967
0.6041 | -76.34 | 43.4 | -1732.11 | 1734.46 | 9.85 | 0.35 | -8.1 | 13.08 | 1732.11 | -0.35 | 56.164
0.6059 | -77.5 | 44.01 | -1699.44 | 1701.9 | 9.59 | 0.97 | -7.87 | 12.79 | 1699.44 | -0.97 | 56.3313
0.6077 | -78.59 | 44.66 | -1666.17 | 1668.75 | 9.33 | 1.59 | -7.64 | 12.52 | 1666.17 | -1.59 | 56.4987
0.6095 | -79.63 | 45.29 | -1632.09 | 1634.79 | 9.06 | 2.2 | -7.4 | 12.25 | 1632.09 | -2.2 | 56.666
0.6113 | -80.49 | 45.89 | -1597.67 | 1600.47 | 8.79 | 2.8 | -7.15 | 12 | 1597.67 | -2.8 | 56.8334
0.6131 | -81.19 | 46.49 | -1563.2 | 1566.11 | 8.51 | 3.39 | -6.89 | 11.78 | 1563.2 | -3.39 | 57.0007
0.6149 | -81.69 | 47.03 | -1528.08 | 1531.1 | 8.23 | 3.94 | -6.63 | 11.59 | 1528.08 | -3.94 | 57.1681
0.6167 | -81.98 | 47.44 | -1493.12 | 1496.24 | 7.95 | 4.45 | -6.36 | 11.41 | 1493.12 | -4.45 | 57.3354
0.6185 | -82.28 | 47.9 | -1457.89 | 1461.12 | 7.67 | 4.96 | -6.09 | 11.25 | 1457.89 | -4.96 | 57.5028
0.6203 | -82.52 | 48.44 | -1422.49 | 1425.83 | 7.4 | 5.46 | -5.82 | 11.13 | 1422.49 | -5.46 | 57.6701
0.6221 | -82.63 | 48.95 | -1386.93 | 1390.37 | 7.12 | 5.92 | -5.53 | 11.04 | 1386.93 | -5.92 | 57.8375
0.6239 | -82.74 | 49.36 | -1352.16 | 1355.7 | 6.86 | 6.36 | -5.24 | 10.95 | 1352.16 | -6.36 | 58.0048
0.6256 | -82.77 | 49.74 | -1317.06 | 1320.71 | 6.6 | 6.78 | -4.95 | 10.9 | 1317.06 | -6.78 | 58.1629
0.6274 | -82.73 | 49.98 | -1282.14 | 1285.89 | 6.34 | 7.18 | -4.65 | 10.85 | 1282.14 | -7.18 | 58.3302
0.6292 | -82.55 | 50.05 | -1247.43 | 1251.27 | 6.09 | 7.54 | -4.36 | 10.8 | 1247.43 | -7.54 | 58.4976
0.631 | -82.23 | 49.89 | -1213.24 | 1217.15 | 5.84 | 7.87 | -4.08 | 10.77 | 1213.24 | -7.87 | 58.6649
0.6328 | -81.75 | 49.32 | -1178.81 | 1182.77 | 5.6 | 8.18 | -3.79 | 10.77 | 1178.81 | -8.18 | 58.8323
0.6346 | -81.17 | 48.65 | -1144.09 | 1148.1 | 5.37 | 8.46 | -3.5 | 10.76 | 1144.09 | -8.46 | 58.9996
0.6364 | -80.52 | 47.94 | -1109.59 | 1113.64 | 5.14 | 8.72 | -3.22 | 10.76 | 1109.59 | -8.72 | 59.167
0.6382 | -79.82 | 47.19 | -1075.4 | 1079.49 | 4.92 | 8.96 | -2.93 | 10.75 | 1075.4 | -8.96 | 59.3343
0.64 | -79.04 | 46.36 | -1041.65 | 1045.77 | 4.71 | 9.16 | -2.65 | 10.74 | 1041.65 | -9.16 | 59.5017
0.6418 | -78.18 | 45.32 | -1008.84 | 1012.98 | 4.52 | 9.32 | -2.38 | 10.71 | 1008.84 | -9.32 | 59.669
0.6436 | -77.31 | 44.12 | -976.86 | 981.01 | 4.36 | 9.45 | -2.13 | 10.69 | 976.86 | -9.45 | 59.8364
0.6454 | -76.41 | 42.9 | -945.54 | 949.69 | 4.2 | 9.55 | -1.88 | 10.66 | 945.54 | -9.55 | 60.0037
0.6472 | -75.53 | 41.54 | -914.83 | 918.98 | 4.05 | 9.62 | -1.64 | 10.62 | 914.83 | -9.62 | 60.1711
0.649 | -74.7 | 39.97 | -884.43 | 888.57 | 3.92 | 9.65 | -1.41 | 10.55 | 884.43 | -9.65 | 60.3384
0.6507 | -73.88 | 38.34 | -854.34 | 858.48 | 3.81 | 9.64 | -1.21 | 10.47 | 854.34 | -9.64 | 60.4965
0.6525 | -73.16 | 36.83 | -825 | 829.15 | 3.72 | 9.6 | -1.02 | 10.38 | 825 | -9.6 | 60.6638
0.6543 | -72.64 | 35.41 | -796.42 | 800.61 | 3.64 | 9.53 | -0.85 | 10.26 | 796.42 | -9.53 | 60.8312
0.6561 | -72.18 | 34.02 | -768.23 | 772.47 | 3.57 | 9.44 | -0.69 | 10.14 | 768.23 | -9.44 | 60.9985
0.6579 | -71.81 | 32.67 | -740.85 | 745.14 | 3.51 | 9.34 | -0.54 | 10.01 | 740.85 | -9.34 | 61.1659
0.6597 | -71.48 | 31.36 | -714.8 | 719.14 | 3.45 | 9.21 | -0.42 | 9.86 | 714.8 | -9.21 | 61.3332
0.6615 | -70.91 | 29.78 | -689.98 | 694.34 | 3.39 | 9.03 | -0.33 | 9.66 | 689.98 | -9.03 | 61.5006
0.6633 | -70.05 | 27.77 | -666.64 | 670.97 | 3.33 | 8.79 | -0.28 | 9.42 | 666.64 | -8.79 | 61.6679
0.6651 | -69.19 | 25.71 | -643.93 | 648.23 | 3.27 | 8.57 | -0.22 | 9.18 | 643.93 | -8.57 | 61.8353
0.6669 | -68.18 | 23.32 | -622.01 | 626.26 | 3.21 | 8.34 | -0.19 | 8.94 | 622.01 | -8.34 | 62.0026
0.6687 | -67.03 | 20.68 | -600.51 | 604.67 | 3.14 | 8.1 | -0.17 | 8.69 | 600.51 | -8.1 | 62.17
0.6705 | -65.86 | 18.09 | -579.82 | 583.9 | 3.07 | 7.85 | -0.15 | 8.44 | 579.82 | -7.85 | 62.3373
0.6723 | -64.74 | 15.56 | -559.95 | 563.97 | 3 | 7.62 | -0.14 | 8.19 | 559.95 | -7.62 | 62.5046
0.6741 | -63.52 | 13.04 | -541.46 | 545.41 | 2.94 | 7.39 | -0.13 | 7.96 | 541.46 | -7.39 | 62.672
0.6758 | -62.27 | 10.6 | -524.26 | 528.15 | 2.88 | 7.17 | -0.14 | 7.73 | 524.26 | -7.17 | 62.83
0.6776 | -61.14 | 8.16 | -507.92 | 511.75 | 2.84 | 6.97 | -0.14 | 7.53 | 507.92 | -6.97 | 62.9974
0.6794 | -59.99 | 5.83 | -491.86 | 495.63 | 2.79 | 6.77 | -0.15 | 7.33 | 491.86 | -6.77 | 63.1647
0.6812 | -59.04 | 3.74 | -477.1 | 480.85 | 2.74 | 6.59 | -0.15 | 7.14 | 477.1 | -6.59 | 63.3321
0.683 | -58.23 | 1.82 | -463.46 | 467.19 | 2.7 | 6.42 | -0.16 | 6.97 | 463.46 | -6.42 | 63.4994
0.6848 | -57.55 | 0.16 | -451.28 | 455.01 | 2.66 | 6.28 | -0.17 | 6.82 | 451.28 | -6.28 | 63.6668
0.6866 | -56.77 | -1.41 | -439.62 | 443.35 | 2.63 | 6.13 | -0.19 | 6.68 | 439.62 | -6.13 | 63.8341
0.6884 | -55.98 | -2.88 | -428.26 | 431.98 | 2.6 | 5.99 | -0.2 | 6.53 | 428.26 | -5.99 | 64.0015
0.6902 | -55.34 | -4.16 | -417.85 | 421.6 | 2.56 | 5.87 | -0.2 | 6.41 | 417.85 | -5.87 | 64.1688
0.692 | -54.79 | -5.35 | -408.07 | 411.84 | 2.53 | 5.75 | -0.21 | 6.29 | 408.07 | -5.75 | 64.3362
0.6938 | -54.28 | -6.35 | -399.13 | 402.91 | 2.49 | 5.64 | -0.22 | 6.17 | 399.13 | -5.64 | 64.5035
0.6956 | -53.59 | -7.31 | -390.32 | 394.12 | 2.46 | 5.53 | -0.22 | 6.06 | 390.32 | -5.53 | 64.6709
0.6974 | -53.15 | -7.93 | -382.27 | 386.09 | 2.42 | 5.43 | -0.22 | 5.95 | 382.27 | -5.43 | 64.8382
0.6992 | -52.81 | -8.4 | -374.93 | 378.79 | 2.38 | 5.34 | -0.22 | 5.85 | 374.93 | -5.34 | 65.0056
0.7009 | -52.5 | -8.9 | -368.15 | 372.04 | 2.35 | 5.26 | -0.22 | 5.77 | 368.15 | -5.26 | 65.1636
0.7027 | -52.25 | -9.37 | -361.92 | 365.84 | 2.31 | 5.18 | -0.21 | 5.68 | 361.92 | -5.18 | 65.331
0.7045 | -51.95 | -9.65 | -356.17 | 360.12 | 2.28 | 5.11 | -0.21 | 5.6 | 356.17 | -5.11 | 65.4983
0.7063 | -51.53 | -9.86 | -350.94 | 354.88 | 2.24 | 5.05 | -0.2 | 5.53 | 350.94 | -5.05 | 65.6657
0.7081 | -51.08 | -10.01 | -346.24 | 350.17 | 2.2 | 4.98 | -0.2 | 5.45 | 346.24 | -4.98 | 65.833
0.7099 | -50.75 | -9.88 | -342.03 | 345.96 | 2.16 | 4.93 | -0.18 | 5.38 | 342.03 | -4.93 | 66.0004
0.7117 | -50.35 | -9.75 | -338.03 | 341.94 | 2.12 | 4.88 | -0.17 | 5.32 | 338.03 | -4.88 | 66.1677
0.7135 | -49.75 | -9.54 | -334.04 | 337.9 | 2.09 | 4.82 | -0.15 | 5.26 | 334.04 | -4.82 | 66.3351
0.7153 | -49.29 | -9.29 | -330.66 | 334.48 | 2.05 | 4.77 | -0.14 | 5.2 | 330.66 | -4.77 | 66.5024
0.7171 | -48.82 | -8.88 | -327.77 | 331.55 | 2.02 | 4.73 | -0.12 | 5.14 | 327.77 | -4.73 | 66.6698
0.7189 | -48.25 | -8.46 | -325.03 | 328.75 | 1.98 | 4.68 | -0.1 | 5.09 | 325.03 | -4.68 | 66.8371
0.7207 | -47.9 | -7.91 | -322.45 | 326.14 | 1.95 | 4.64 | -0.08 | 5.04 | 322.45 | -4.64 | 67.0045
0.7225 | -47.51 | -7.24 | -319.97 | 323.61 | 1.92 | 4.6 | -0.06 | 4.99 | 319.97 | -4.6 | 67.1718
0.7242 | -47.42 | -6.52 | -318.02 | 321.66 | 1.89 | 4.58 | -0.03 | 4.96 | 318.02 | -4.58 | 67.3299
0.726 | -47.33 | -5.77 | -316.45 | 320.08 | 1.87 | 4.57 | -0.01 | 4.94 | 316.45 | -4.57 | 67.4972
0.7278 | -47.07 | -5.22 | -314.44 | 318.05 | 1.84 | 4.54 | 0.01 | 4.9 | 314.44 | -4.54 | 67.6646
0.7296 | -46.6 | -4.89 | -311.88 | 315.45 | 1.8 | 4.5 | 0.03 | 4.85 | 311.88 | -4.5 | 67.8319
0.7314 | -46.1 | -4.55 | -309.12 | 312.64 | 1.77 | 4.45 | 0.05 | 4.79 | 309.12 | -4.45 | 67.9993
0.7332 | -45.51 | -4.32 | -306.29 | 309.75 | 1.73 | 4.4 | 0.07 | 4.73 | 306.29 | -4.4 | 68.1666
0.735 | -44.66 | -4.23 | -303.22 | 306.59 | 1.69 | 4.33 | 0.08 | 4.66 | 303.22 | -4.33 | 68.334
0.7368 | -43.62 | -4.23 | -299.88 | 303.13 | 1.65 | 4.26 | 0.1 | 4.57 | 299.88 | -4.26 | 68.5013
0.7386 | -42.39 | -4.32 | -296.34 | 299.46 | 1.6 | 4.17 | 0.11 | 4.48 | 296.34 | -4.17 | 68.6687
0.7404 | -41.08 | -4.47 | -292.53 | 295.5 | 1.55 | 4.08 | 0.12 | 4.37 | 292.53 | -4.08 | 68.836
0.7422 | -39.79 | -4.62 | -288.52 | 291.36 | 1.51 | 3.99 | 0.13 | 4.27 | 288.52 | -3.99 | 69.0033
0.744 | -38.48 | -4.8 | -284.36 | 287.06 | 1.46 | 3.89 | 0.14 | 4.16 | 284.36 | -3.89 | 69.1707
0.7458 | -37.13 | -5.01 | -280 | 282.57 | 1.41 | 3.78 | 0.15 | 4.05 | 280 | -3.78 | 69.338
0.7476 | -35.77 | -5.19 | -275.43 | 277.86 | 1.37 | 3.68 | 0.15 | 3.93 | 275.43 | -3.68 | 69.5054
0.7493 | -34.42 | -5.28 | -270.82 | 273.12 | 1.33 | 3.57 | 0.16 | 3.82 | 270.82 | -3.57 | 69.6634
0.7511 | -33.16 | -5.37 | -266.45 | 268.63 | 1.29 | 3.47 | 0.16 | 3.71 | 266.45 | -3.47 | 69.8308
0.7529 | -31.94 | -5.33 | -262.25 | 264.32 | 1.26 | 3.36 | 0.17 | 3.6 | 262.25 | -3.36 | 69.9981
0.7547 | -30.73 | -5.3 | -258.06 | 260 | 1.22 | 3.26 | 0.17 | 3.49 | 258.06 | -3.26 | 70.1655
0.7565 | -29.56 | -5.25 | -253.86 | 255.69 | 1.19 | 3.16 | 0.18 | 3.38 | 253.86 | -3.16 | 70.3328
0.7583 | -28.46 | -5.17 | -249.68 | 251.42 | 1.16 | 3.05 | 0.18 | 3.28 | 249.68 | -3.05 | 70.5002
0.7601 | -27.51 | -5.11 | -245.57 | 247.22 | 1.13 | 2.96 | 0.19 | 3.18 | 245.57 | -2.96 | 70.6675
0.7619 | -26.63 | -5.05 | -241.56 | 243.14 | 1.11 | 2.87 | 0.19 | 3.09 | 241.56 | -2.87 | 70.8349
0.7637 | -25.88 | -5 | -237.53 | 239.05 | 1.08 | 2.78 | 0.19 | 2.99 | 237.53 | -2.78 | 71.0022
0.7655 | -25.18 | -4.96 | -233.52 | 234.99 | 1.05 | 2.7 | 0.2 | 2.91 | 233.52 | -2.7 | 71.1696
0.7673 | -24.56 | -4.86 | -229.69 | 231.12 | 1.03 | 2.61 | 0.2 | 2.82 | 229.69 | -2.61 | 71.3369
0.7691 | -24.17 | -4.57 | -226.01 | 227.41 | 1.01 | 2.55 | 0.2 | 2.76 | 226.01 | -2.55 | 71.5043
0.7709 | -23.71 | -4.34 | -222.55 | 223.92 | 0.99 | 2.48 | 0.21 | 2.69 | 222.55 | -2.48 | 71.6716
0.7727 | -23.23 | -4.13 | -219.22 | 220.56 | 0.97 | 2.41 | 0.21 | 2.61 | 219.22 | -2.41 | 71.839
0.7744 | -22.78 | -3.9 | -216.12 | 217.42 | 0.96 | 2.34 | 0.22 | 2.55 | 216.12 | -2.34 | 71.997
0.7762 | -22.46 | -3.58 | -213.49 | 214.76 | 0.94 | 2.28 | 0.22 | 2.49 | 213.49 | -2.28 | 72.1644
0.778 | -22.29 | -3.14 | -211.43 | 212.69 | 0.93 | 2.23 | 0.23 | 2.44 | 211.43 | -2.23 | 72.3317
0.7798 | -22.17 | -2.9 | -209.24 | 210.51 | 0.91 | 2.19 | 0.24 | 2.39 | 209.24 | -2.19 | 72.4991
0.7816 | -22.03 | -2.72 | -206.67 | 207.94 | 0.9 | 2.14 | 0.24 | 2.35 | 206.67 | -2.14 | 72.6664
0.7834 | -21.82 | -2.59 | -203.67 | 204.93 | 0.88 | 2.1 | 0.24 | 2.3 | 203.67 | -2.1 | 72.8338
0.7852 | -21.61 | -2.37 | -200.57 | 201.83 | 0.86 | 2.05 | 0.24 | 2.25 | 200.57 | -2.05 | 73.0011
0.787 | -21.52 | -2.01 | -197.54 | 198.81 | 0.85 | 2.01 | 0.24 | 2.21 | 197.54 | -2.01 | 73.1685
0.7888 | -21.42 | -1.65 | -194.63 | 195.9 | 0.84 | 1.97 | 0.24 | 2.17 | 194.63 | -1.97 | 73.3358
0.7906 | -21.31 | -1.34 | -191.99 | 193.27 | 0.83 | 1.94 | 0.24 | 2.14 | 191.99 | -1.94 | 73.5032
0.7924 | -21.41 | -0.93 | -189.79 | 191.09 | 0.82 | 1.92 | 0.25 | 2.11 | 189.79 | -1.92 | 73.6705
0.7942 | -21.72 | -0.48 | -188.11 | 189.47 | 0.81 | 1.91 | 0.25 | 2.11 | 188.11 | -1.91 | 73.8379
0.796 | -22.09 | -0.19 | -186.47 | 187.89 | 0.8 | 1.91 | 0.26 | 2.1 | 186.47 | -1.91 | 74.0052
0.7977 | -22.11 | -0.05 | -184.93 | 186.38 | 0.78 | 1.91 | 0.26 | 2.09 | 184.93 | -1.91 | 74.1633
0.7995 | -21.76 | 0.02 | -182.92 | 184.34 | 0.76 | 1.89 | 0.25 | 2.07 | 182.92 | -1.89 | 74.3306
0.8013 | -21.72 | 0.16 | -180.4 | 181.83 | 0.74 | 1.87 | 0.26 | 2.04 | 180.4 | -1.87 | 74.498
0.8031 | -21.6 | 0.3 | -177.82 | 179.26 | 0.72 | 1.84 | 0.26 | 2.01 | 177.82 | -1.84 | 74.6653
0.8049 | -21.23 | 0.28 | -174.71 | 176.13 | 0.69 | 1.8 | 0.26 | 1.96 | 174.71 | -1.8 | 74.8327
0.8067 | -21.03 | 0.32 | -171.77 | 173.2 | 0.66 | 1.77 | 0.26 | 1.93 | 171.77 | -1.77 | 75
0.8085 | -20.99 | 0.49 | -169.02 | 170.46 | 0.64 | 1.76 | 0.26 | 1.91 | 169.02 | -1.76 | 75.1673
0.8103 | -20.86 | 0.67 | -166.84 | 168.28 | 0.62 | 1.75 | 0.26 | 1.89 | 166.84 | -1.75 | 75.3347
0.8121 | -20.93 | 0.98 | -164.98 | 166.45 | 0.6 | 1.75 | 0.26 | 1.88 | 164.98 | -1.75 | 75.502
0.8139 | -20.98 | 1.32 | -163.28 | 164.78 | 0.59 | 1.74 | 0.26 | 1.87 | 163.28 | -1.74 | 75.6694
0.8157 | -21.07 | 1.71 | -161.73 | 163.26 | 0.57 | 1.74 | 0.26 | 1.87 | 161.73 | -1.74 | 75.8367
0.8175 | -21.35 | 2.16 | -160.41 | 161.99 | 0.56 | 1.75 | 0.27 | 1.87 | 160.41 | -1.75 | 76.0041
0.8193 | -21.51 | 2.44 | -158.99 | 160.62 | 0.55 | 1.76 | 0.27 | 1.88 | 158.99 | -1.76 | 76.1714
0.8211 | -21.47 | 2.62 | -157.45 | 159.09 | 0.53 | 1.76 | 0.27 | 1.87 | 157.45 | -1.76 | 76.3388
0.8228 | -21.35 | 2.85 | -156.06 | 157.69 | 0.51 | 1.76 | 0.27 | 1.87 | 156.06 | -1.76 | 76.4968
0.8246 | -21.23 | 3.17 | -154.87 | 156.5 | 0.49 | 1.76 | 0.27 | 1.86 | 154.87 | -1.76 | 76.6642
0.8264 | -21.1 | 3.53 | -153.81 | 155.43 | 0.47 | 1.76 | 0.28 | 1.86 | 153.81 | -1.76 | 76.8315
0.8282 | -20.89 | 3.81 | -152.83 | 154.41 | 0.45 | 1.77 | 0.28 | 1.86 | 152.83 | -1.77 | 76.9989
0.83 | -20.5 | 3.92 | -151.82 | 153.36 | 0.42 | 1.76 | 0.29 | 1.86 | 151.82 | -1.76 | 77.1662
0.8318 | -20.36 | 4.29 | -150.98 | 152.51 | 0.4 | 1.77 | 0.29 | 1.86 | 150.98 | -1.77 | 77.3336
0.8336 | -20.32 | 4.73 | -150.22 | 151.77 | 0.38 | 1.78 | 0.3 | 1.86 | 150.22 | -1.78 | 77.5009
0.8354 | -20.54 | 5.35 | -149.74 | 151.35 | 0.37 | 1.81 | 0.31 | 1.89 | 149.74 | -1.81 | 77.6683
0.8372 | -20.68 | 5.88 | -149.21 | 150.88 | 0.36 | 1.84 | 0.31 | 1.91 | 149.21 | -1.84 | 77.8356
0.839 | -20.71 | 6.32 | -148.76 | 150.47 | 0.35 | 1.86 | 0.32 | 1.94 | 148.76 | -1.86 | 78.003
0.8408 | -20.68 | 6.73 | -148.39 | 150.12 | 0.33 | 1.89 | 0.33 | 1.96 | 148.39 | -1.89 | 78.1703
0.8426 | -20.66 | 7.23 | -148.48 | 150.22 | 0.31 | 1.91 | 0.34 | 1.98 | 148.48 | -1.91 | 78.3377
0.8444 | -20.84 | 7.83 | -148.96 | 150.74 | 0.3 | 1.94 | 0.34 | 2 | 148.96 | -1.94 | 78.505
0.8462 | -21.19 | 8.51 | -149.55 | 151.4 | 0.3 | 1.97 | 0.35 | 2.04 | 149.55 | -1.97 | 78.6724
0.8479 | -21.42 | 9.07 | -150.14 | 152.04 | 0.28 | 2 | 0.36 | 2.07 | 150.14 | -2 | 78.8304
0.8497 | -21.45 | 9.45 | -150.65 | 152.57 | 0.26 | 2.03 | 0.37 | 2.09 | 150.65 | -2.03 | 78.9978
0.8515 | -21.46 | 9.93 | -151.23 | 153.17 | 0.25 | 2.05 | 0.38 | 2.11 | 151.23 | -2.05 | 79.1651
0.8533 | -21.26 | 10.21 | -151.71 | 153.63 | 0.23 | 2.07 | 0.38 | 2.13 | 151.71 | -2.07 | 79.3325
0.8551 | -21.25 | 10.65 | -152.07 | 154.02 | 0.21 | 2.1 | 0.39 | 2.15 | 152.07 | -2.1 | 79.4998
0.8569 | -21.1 | 10.96 | -152.25 | 154.2 | 0.18 | 2.12 | 0.4 | 2.17 | 152.25 | -2.12 | 79.6672
0.8587 | -20.56 | 11.06 | -152.16 | 154.04 | 0.14 | 2.12 | 0.4 | 2.17 | 152.16 | -2.12 | 79.8345
0.8605 | -19.95 | 11.14 | -152.06 | 153.87 | 0.11 | 2.12 | 0.41 | 2.17 | 152.06 | -2.12 | 80.0019
0.8623 | -19.43 | 11.28 | -152.4 | 154.15 | 0.08 | 2.13 | 0.41 | 2.17 | 152.4 | -2.13 | 80.1692
0.8641 | -18.86 | 11.44 | -152.86 | 154.54 | 0.04 | 2.13 | 0.42 | 2.18 | 152.86 | -2.13 | 80.3366
0.8659 | -18.19 | 11.63 | -153.57 | 155.18 | 0.02 | 2.13 | 0.42 | 2.18 | 153.57 | -2.13 | 80.5039
0.8677 | -17.59 | 11.88 | -154.31 | 155.86 | -0.01 | 2.12 | 0.43 | 2.17 | 154.31 | -2.12 | 80.6713
0.8695 | -16.99 | 12.13 | -155.06 | 156.55 | -0.03 | 2.11 | 0.44 | 2.16 | 155.06 | -2.11 | 80.8386
0.8712 | -16.49 | 12.52 | -156.1 | 157.56 | -0.04 | 2.11 | 0.45 | 2.16 | 156.1 | -2.11 | 80.9967
0.873 | -16.03 | 12.91 | -157.52 | 158.94 | -0.05 | 2.1 | 0.46 | 2.16 | 157.52 | -2.1 | 81.164
0.8748 | -15.8 | 13.43 | -159.08 | 160.5 | -0.05 | 2.11 | 0.46 | 2.17 | 159.08 | -2.11 | 81.3313
0.8766 | -15.73 | 14.08 | -160.83 | 162.28 | -0.04 | 2.12 | 0.47 | 2.18 | 160.83 | -2.12 | 81.4987
0.8784 | -15.61 | 14.81 | -162.81 | 164.29 | -0.03 | 2.14 | 0.48 | 2.2 | 162.81 | -2.14 | 81.666
0.8802 | -15.69 | 15.57 | -165.04 | 166.59 | -0.03 | 2.15 | 0.49 | 2.22 | 165.04 | -2.15 | 81.8334
0.882 | -15.77 | 16.3 | -167.37 | 168.98 | -0.02 | 2.17 | 0.49 | 2.24 | 167.37 | -2.17 | 82.0007
0.8838 | -16.07 | 17.1 | -169.86 | 171.56 | -0.01 | 2.2 | 0.5 | 2.27 | 169.86 | -2.2 | 82.1681
0.8856 | -16.25 | 17.85 | -172.29 | 174.06 | -0.01 | 2.23 | 0.51 | 2.3 | 172.29 | -2.23 | 82.3354
0.8874 | -16.35 | 18.56 | -174.85 | 176.68 | 0 | 2.25 | 0.52 | 2.32 | 174.85 | -2.25 | 82.5028
0.8892 | -16.57 | 19.22 | -177.55 | 179.45 | 0.01 | 2.27 | 0.52 | 2.34 | 177.55 | -2.27 | 82.6701
0.891 | -16.74 | 19.78 | -180.09 | 182.04 | 0.03 | 2.28 | 0.53 | 2.35 | 180.09 | -2.28 | 82.8375
0.8928 | -16.82 | 20.21 | -182.04 | 184.02 | 0.03 | 2.28 | 0.53 | 2.35 | 182.04 | -2.28 | 83.0048
0.8946 | -16.96 | 20.59 | -183.83 | 185.85 | 0.05 | 2.28 | 0.53 | 2.35 | 183.83 | -2.28 | 83.1722
0.8963 | -17.09 | 20.98 | -185.7 | 187.76 | 0.06 | 2.28 | 0.52 | 2.35 | 185.7 | -2.28 | 83.3302
0.8981 | -17.39 | 21.35 | -187.35 | 189.46 | 0.07 | 2.29 | 0.52 | 2.36 | 187.35 | -2.29 | 83.4976
0.8999 | -17.57 | 21.62 | -188.69 | 190.82 | 0.08 | 2.28 | 0.51 | 2.35 | 188.69 | -2.28 | 83.6649
0.9017 | -17.67 | 21.72 | -189.74 | 191.88 | 0.1 | 2.26 | 0.49 | 2.32 | 189.74 | -2.26 | 83.8323
0.9035 | -17.71 | 21.63 | -190.56 | 192.69 | 0.1 | 2.23 | 0.48 | 2.3 | 190.56 | -2.23 | 83.9996
0.9053 | -17.53 | 21.4 | -190.97 | 193.06 | 0.11 | 2.19 | 0.46 | 2.26 | 190.97 | -2.19 | 84.167
0.9071 | -17.23 | 21.09 | -191.29 | 193.32 | 0.11 | 2.15 | 0.44 | 2.22 | 191.29 | -2.15 | 84.3343
0.9089 | -16.83 | 20.71 | -191.58 | 193.53 | 0.11 | 2.09 | 0.42 | 2.16 | 191.58 | -2.09 | 84.5017
0.9107 | -16.47 | 20.28 | -191.75 | 193.62 | 0.1 | 2.04 | 0.4 | 2.11 | 191.75 | -2.04 | 84.669
0.9125 | -15.93 | 19.7 | -191.67 | 193.45 | 0.1 | 1.97 | 0.38 | 2.04 | 191.67 | -1.97 | 84.8364
0.9143 | -15.49 | 19 | -191.35 | 193.04 | 0.1 | 1.9 | 0.36 | 1.96 | 191.35 | -1.9 | 85.0037
0.9161 | -15.18 | 18.37 | -191.03 | 192.65 | 0.09 | 1.83 | 0.34 | 1.88 | 191.03 | -1.83 | 85.1711
0.9179 | -15.05 | 17.81 | -190.88 | 192.45 | 0.1 | 1.76 | 0.31 | 1.82 | 190.88 | -1.76 | 85.3384
0.9197 | -14.78 | 17.07 | -190.55 | 192.05 | 0.1 | 1.7 | 0.29 | 1.76 | 190.55 | -1.7 | 85.5058
0.9214 | -14.22 | 16.2 | -190.07 | 191.48 | 0.09 | 1.63 | 0.27 | 1.7 | 190.07 | -1.63 | 85.6638
0.9232 | -13.66 | 15.29 | -189.24 | 190.55 | 0.09 | 1.57 | 0.25 | 1.64 | 189.24 | -1.57 | 85.8312
0.925 | -13.08 | 14.53 | -188.15 | 189.37 | 0.08 | 1.5 | 0.23 | 1.59 | 188.15 | -1.5 | 85.9985
0.9268 | -12.6 | 13.84 | -186.77 | 187.91 | 0.08 | 1.44 | 0.21 | 1.55 | 186.77 | -1.44 | 86.1659
0.9286 | -11.94 | 13.14 | -186.14 | 187.19 | 0.08 | 1.36 | 0.19 | 1.48 | 186.14 | -1.36 | 86.3332
0.9304 | -11.19 | 12.4 | -185.87 | 186.8 | 0.07 | 1.27 | 0.17 | 1.39 | 185.87 | -1.27 | 86.5006
0.9322 | -10.66 | 11.81 | -185.92 | 186.76 | 0.08 | 1.19 | 0.15 | 1.31 | 185.92 | -1.19 | 86.6679
0.934 | -10.35 | 11.42 | -186.18 | 186.96 | 0.1 | 1.12 | 0.13 | 1.26 | 186.18 | -1.12 | 86.8353
0.9358 | -10.22 | 11.17 | -186.65 | 187.39 | 0.11 | 1.06 | 0.12 | 1.21 | 186.65 | -1.06 | 87.0026
0.9376 | -10.2 | 10.97 | -187.48 | 188.21 | 0.13 | 1.01 | 0.11 | 1.17 | 187.48 | -1.01 | 87.17
0.9394 | -10.21 | 10.78 | -188.39 | 189.1 | 0.15 | 0.96 | 0.11 | 1.13 | 188.39 | -0.96 | 87.3373
0.9412 | -10.4 | 10.64 | -189.56 | 190.27 | 0.17 | 0.94 | 0.1 | 1.12 | 189.56 | -0.94 | 87.5046
0.943 | -10.68 | 10.48 | -190.91 | 191.62 | 0.2 | 0.92 | 0.09 | 1.11 | 190.91 | -0.92 | 87.672
0.9447 | -10.86 | 10.2 | -192.27 | 192.97 | 0.22 | 0.88 | 0.09 | 1.09 | 192.27 | -0.88 | 87.83
0.9465 | -10.83 | 9.79 | -193.31 | 193.98 | 0.23 | 0.83 | 0.08 | 1.06 | 193.31 | -0.83 | 87.9974
0.9483 | -10.6 | 9.27 | -193.83 | 194.45 | 0.24 | 0.77 | 0.07 | 1.05 | 193.83 | -0.77 | 88.1647
0.9501 | -10.15 | 8.55 | -193.9 | 194.48 | 0.24 | 0.7 | 0.06 | 1.05 | 193.9 | -0.7 | 88.3321
0.9519 | -9.67 | 7.85 | -193.55 | 194.08 | 0.24 | 0.64 | 0.06 | 1.08 | 193.55 | -0.64 | 88.4994
0.9537 | -9.39 | 7.49 | -193.53 | 194.05 | 0.24 | 0.58 | 0.06 | 1.1 | 193.53 | -0.58 | 88.6668
0.9555 | -9.42 | 7.18 | -193.13 | 193.64 | 0.25 | 0.56 | 0.06 | 1.12 | 193.13 | -0.56 | 88.8341
0.9573 | -9.37 | 6.94 | -193.28 | 193.79 | 0.26 | 0.53 | 0.06 | 1.11 | 193.28 | -0.53 | 89.0015
0.9591 | -9.31 | 6.85 | -193.59 | 194.09 | 0.27 | 0.5 | 0.05 | 1.09 | 193.59 | -0.5 | 89.1688
0.9609 | -9.34 | 7.02 | -194.3 | 194.81 | 0.29 | 0.47 | 0.05 | 1.09 | 194.3 | -0.47 | 89.3362
0.9627 | -9.45 | 7.37 | -195.43 | 195.94 | 0.31 | 0.43 | 0.05 | 1.09 | 195.43 | -0.43 | 89.5035
0.9645 | -9.68 | 7.85 | -197.23 | 197.77 | 0.33 | 0.4 | 0.05 | 1.09 | 197.23 | -0.4 | 89.6709
0.9663 | -9.96 | 8.54 | -199.47 | 200.04 | 0.36 | 0.37 | 0.05 | 1.1 | 199.47 | -0.37 | 89.8382
0.9681 | -10.22 | 9.15 | -201.91 | 202.5 | 0.38 | 0.35 | 0.05 | 1.1 | 201.91 | -0.35 | 90.0056
0.9698 | -10.4 | 9.69 | -204.4 | 205.02 | 0.4 | 0.34 | 0.04 | 1.1 | 204.4 | -0.34 | 90.1636
0.9716 | -10.64 | 10.24 | -207.18 | 207.82 | 0.42 | 0.31 | 0.04 | 1.09 | 207.18 | -0.31 | 90.331
0.9734 | -10.74 | 10.83 | -210.49 | 211.16 | 0.44 | 0.26 | 0.03 | 1.09 | 210.49 | -0.26 | 90.4983
0.9752 | -10.7 | 11.39 | -214.13 | 214.82 | 0.45 | 0.21 | 0.02 | 1.08 | 214.13 | -0.21 | 90.6657
0.977 | -10.54 | 11.89 | -217.83 | 218.54 | 0.46 | 0.17 | 0.01 | 1.08 | 217.83 | -0.17 | 90.833
0.9788 | -10.23 | 12.5 | -221.76 | 222.47 | 0.47 | 0.11 | 0 | 1.09 | 221.76 | -0.11 | 91.0004
0.9806 | -9.97 | 13.33 | -226.12 | 226.84 | 0.48 | 0.04 | -0.01 | 1.1 | 226.12 | -0.04 | 91.1677
0.9824 | -9.75 | 14.31 | -230.89 | 231.64 | 0.49 | -0.03 | -0.02 | 1.12 | 230.89 | 0.03 | 91.3351
0.9842 | -9.53 | 15.4 | -235.97 | 236.76 | 0.5 | -0.11 | -0.04 | 1.14 | 235.97 | 0.11 | 91.5024
0.986 | -9.41 | 16.3 | -241 | 241.83 | 0.51 | -0.2 | -0.05 | 1.13 | 241 | 0.2 | 91.6698
0.9878 | -9.07 | 17.13 | -246.08 | 246.93 | 0.51 | -0.29 | -0.07 | 1.14 | 246.08 | 0.29 | 91.8371
0.9896 | -8.73 | 18.13 | -251.52 | 252.4 | 0.52 | -0.39 | -0.09 | 1.16 | 251.52 | 0.39 | 92.0045
0.9914 | -8.32 | 19.31 | -257.59 | 258.52 | 0.52 | -0.5 | -0.11 | 1.19 | 257.59 | 0.5 | 92.1718
0.9932 | -7.78 | 20.51 | -264.31 | 265.29 | 0.53 | -0.62 | -0.12 | 1.21 | 264.31 | 0.62 | 92.3392
0.9949 | -7.26 | 21.87 | -271.51 | 272.55 | 0.53 | -0.74 | -0.14 | 1.26 | 271.51 | 0.74 | 92.4972
0.9967 | -6.97 | 23.43 | -278.96 | 280.1 | 0.55 | -0.85 | -0.16 | 1.32 | 278.96 | 0.85 | 92.6646
0.9985 | -6.95 | 25.18 | -286.75 | 288 | 0.56 | -0.96 | -0.18 | 1.39 | 286.75 | 0.96 | 92.8319
1.0003 | -7.01 | 27.08 | -294.88 | 296.27 | 0.58 | -1.06 | -0.2 | 1.47 | 294.88 | 1.06 | 92.9993
1.0021 | -6.91 | 29 | -303.5 | 305.03 | 0.58 | -1.18 | -0.23 | 1.56 | 303.5 | 1.18 | 93.1666
1.0039 | -6.57 | 30.83 | -312.59 | 314.25 | 0.59 | -1.33 | -0.26 | 1.67 | 312.59 | 1.33 | 93.334
1.0057 | -6.13 | 32.64 | -321.94 | 323.72 | 0.59 | -1.48 | -0.29 | 1.78 | 321.94 | 1.48 | 93.5013
1.0075 | -5.58 | 34.34 | -331.54 | 333.44 | 0.59 | -1.63 | -0.33 | 1.9 | 331.54 | 1.63 | 93.6687
1.0093 | -4.93 | 35.96 | -341.11 | 343.11 | 0.57 | -1.8 | -0.37 | 2.04 | 341.11 | 1.8 | 93.836
1.0111 | -4.35 | 37.77 | -350.63 | 352.76 | 0.56 | -1.97 | -0.41 | 2.19 | 350.63 | 1.97 | 94.0033
1.0129 | -3.69 | 39.6 | -361.06 | 363.31 | 0.54 | -2.16 | -0.45 | 2.36 | 361.06 | 2.16 | 94.1707
1.0147 | -3.06 | 41.51 | -371.94 | 374.33 | 0.53 | -2.34 | -0.5 | 2.53 | 371.94 | 2.34 | 94.338
1.0165 | -2.36 | 43.39 | -383.25 | 385.78 | 0.51 | -2.54 | -0.54 | 2.71 | 383.25 | 2.54 | 94.5054
1.0182 | -1.54 | 45.23 | -394.84 | 397.5 | 0.48 | -2.73 | -0.6 | 2.9 | 394.84 | 2.73 | 94.6634
1.02 | -0.8 | 47.11 | -406.28 | 409.09 | 0.45 | -2.92 | -0.65 | 3.08 | 406.28 | 2.92 | 94.8308
1.0218 | -0.14 | 48.92 | -417.58 | 420.52 | 0.42 | -3.11 | -0.7 | 3.27 | 417.58 | 3.11 | 94.9981
1.0236 | 0.65 | 50.6 | -428.86 | 431.93 | 0.38 | -3.31 | -0.76 | 3.46 | 428.86 | 3.31 | 95.1655
1.0254 | 1.56 | 52.21 | -440.2 | 443.38 | 0.34 | -3.51 | -0.81 | 3.67 | 440.2 | 3.51 | 95.3328
1.0272 | 2.56 | 53.68 | -451.68 | 454.96 | 0.29 | -3.72 | -0.87 | 3.87 | 451.68 | 3.72 | 95.5002
1.029 | 3.82 | 55.09 | -463.7 | 467.09 | 0.24 | -3.92 | -0.94 | 4.08 | 463.7 | 3.92 | 95.6675
1.0308 | 5.24 | 56.4 | -476.06 | 479.54 | 0.2 | -4.11 | -1 | 4.28 | 476.06 | 4.11 | 95.8349
1.0326 | 6.68 | 57.75 | -489.01 | 492.59 | 0.15 | -4.31 | -1.07 | 4.49 | 489.01 | 4.31 | 96.0022
1.0344 | 8.05 | 59.43 | -502.92 | 506.62 | 0.11 | -4.54 | -1.13 | 4.71 | 502.92 | 4.54 | 96.1696
1.0362 | 9.51 | 60.72 | -515.91 | 519.69 | 0.06 | -4.75 | -1.2 | 4.94 | 515.91 | 4.75 | 96.3369
1.038 | 10.92 | 61.75 | -528.09 | 531.93 | 0.02 | -4.94 | -1.27 | 5.14 | 528.09 | 4.94 | 96.5043
1.0398 | 12.38 | 62.48 | -539.32 | 543.21 | -0.01 | -5.13 | -1.32 | 5.33 | 539.32 | 5.13 | 96.6716
1.0416 | 14.09 | 62.91 | -550.03 | 553.94 | -0.03 | -5.32 | -1.37 | 5.52 | 550.03 | 5.32 | 96.839
1.0433 | 15.88 | 63.2 | -560.58 | 564.51 | -0.05 | -5.49 | -1.41 | 5.7 | 560.58 | 5.49 | 96.997
1.0451 | 17.51 | 63.71 | -571 | 574.97 | -0.08 | -5.65 | -1.46 | 5.87 | 571 | 5.65 | 97.1644
1.0469 | 19.29 | 63.79 | -580.86 | 584.82 | -0.1 | -5.79 | -1.52 | 6.02 | 580.86 | 5.79 | 97.3317
1.0487 | 20.87 | 63.72 | -589.69 | 593.63 | -0.12 | -5.93 | -1.57 | 6.16 | 589.69 | 5.93 | 97.4991
1.0505 | 22.42 | 63.4 | -597.92 | 601.82 | -0.13 | -6.07 | -1.61 | 6.31 | 597.92 | 6.07 | 97.6664
1.0523 | 23.92 | 62.74 | -605.3 | 609.12 | -0.11 | -6.18 | -1.65 | 6.43 | 605.3 | 6.18 | 97.8338
1.0541 | 25.54 | 61.89 | -612.38 | 616.13 | -0.08 | -6.28 | -1.68 | 6.53 | 612.38 | 6.28 | 98.0011
1.0559 | 27.15 | 60.97 | -619.12 | 622.8 | -0.02 | -6.35 | -1.71 | 6.61 | 619.12 | 6.35 | 98.1685
1.0577 | 28.56 | 60.07 | -625.06 | 628.67 | 0.04 | -6.4 | -1.74 | 6.66 | 625.06 | 6.4 | 98.3358
1.0595 | 29.16 | 58.76 | -627.1 | 630.61 | 0.15 | -6.39 | -1.74 | 6.66 | 627.1 | 6.39 | 98.5032
1.0613 | 28.88 | 57.54 | -626.81 | 630.2 | 0.26 | -6.35 | -1.72 | 6.62 | 626.81 | 6.35 | 98.6705
1.0631 | 27.99 | 56.19 | -624.12 | 627.37 | 0.37 | -6.24 | -1.7 | 6.52 | 624.12 | 6.24 | 98.8379
1.0649 | 27.12 | 54.87 | -621.34 | 624.45 | 0.48 | -6.14 | -1.68 | 6.43 | 621.34 | 6.14 | 99.0052
1.0667 | 26.31 | 53.44 | -618.4 | 621.37 | 0.58 | -6.03 | -1.66 | 6.33 | 618.4 | 6.03 | 99.1726
1.0684 | 25.1 | 51.76 | -613.44 | 616.23 | 0.67 | -5.87 | -1.62 | 6.18 | 613.44 | 5.87 | 99.3306
1.0702 | 23.89 | 50.08 | -608.51 | 611.13 | 0.76 | -5.71 | -1.59 | 6.02 | 608.51 | 5.71 | 99.498
1.072 | 22.86 | 48.41 | -604.53 | 607 | 0.85 | -5.55 | -1.55 | 5.87 | 604.53 | 5.55 | 99.6653
1.0738 | 22.05 | 46.79 | -602.21 | 604.53 | 0.95 | -5.42 | -1.51 | 5.75 | 602.21 | 5.42 | 99.8327
1.0756 | 21.27 | 45.11 | -600.81 | 602.99 | 1.04 | -5.3 | -1.47 | 5.64 | 600.81 | 5.3 | 100
| Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My
Minimum | -82.77 | -162.76 | -2298.7 | 150.12 | -0.13 | -24.56 | -11.02 | 1.05 | 148.39 | -9.65
Maximum | 29.16 | 63.79 | -148.39 | 2299.49 | 16.66 | 9.65 | 1.43 | 29.44 | 2298.7 | 24.56

## Diagr_Individ
Individual Loads
Contact forces and moments, measured telemetrically in knee implants of 8 subjects.
Averages of one individual subject.
Fx  lateral, Fy  anterior, -Fz  distal, Fres = Resultant Force
Mx around x, My around y, Mz around z, Mres = Resultant Moment
z-axis = axis of implant stem!
From Bergmann et al, PlosOne (2014) "Standardized Loads Acting in Knee Implants"
Subject K5R
Subject K5R

## Dat_BW
Data File: | k5r_2242_0_13_i1.akf
Curve Alignment Program: | 1.5.7
------------------------MEASURING INFOS------------------------
Diagram Title #1 | Forces and Moments at Knee Joint
Diagram Title #2 | Measured Relative to Implant
Comment #1 | Knee Joint; Gaitanalysis; Level Walking;
Comment #2 | K5R, 30 Months PO *
BodyWeight [N]: | 942
Patient(en): | k5r | [7 empty] | Subject | k5r
Implant Type: | Knee Joint
Varus Angle Leg [°] | 1 | [7 empty] | Normalized Loads in %BW (Forces) and %BWm (moments)
Axes Directions X, Y, Z: | Lateral, Anterior, Superior
Axes Origin: | Tibial Plateau at Stem Axis | [7 empty] | BW [kg] = | 96.0245
Transformation X Y Z (°): | 7 | -2 | -1
Displacement X, Y, Z (mm): | 0 | 0 | 0
Activity Code: | KneeActivities | 1P_1_15_2
---------------------------DATA--------------------------------
Number of Data Sets = | 601
Number of analog channels = | 0
Time-Offset(sec) = | 0
Time | Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My | Time
[s] | [%BW] | [%BW] | [%BW] | [%BW] | [%BWm] | [%BWm] | [%BWm] | [%BWm] | [%BW] | [%BWm] | [%Cycle]
0 | 1.2335 | 3.0955 | -64.0403 | 64.1423 | 0.19958 | -0.40764 | -0.11146 | 0.48301 | 64.0403 | 0.40764 | 0
0.0018 | 1.1189 | 2.8854 | -63.8864 | 63.9724 | 0.20913 | -0.39172 | -0.10616 | 0.47134 | 63.8864 | 0.39172 | 0.16735
0.0036 | 0.99363 | 2.6964 | -63.828 | 63.9023 | 0.21868 | -0.37367 | -0.10085 | 0.45754 | 63.828 | 0.37367 | 0.3347
0.0054 | 0.89384 | 2.5223 | -63.9055 | 63.9703 | 0.22718 | -0.35881 | -0.095541 | 0.44904 | 63.9055 | 0.35881 | 0.50205
0.0072 | 0.80467 | 2.3577 | -64.155 | 64.2123 | 0.23673 | -0.34607 | -0.090234 | 0.44374 | 64.155 | 0.34607 | 0.66939
0.009 | 0.71975 | 2.2028 | -64.6858 | 64.7378 | 0.24628 | -0.33439 | -0.084926 | 0.44268 | 64.6858 | 0.33439 | 0.83674
0.0108 | 0.66773 | 2.0361 | -65.62 | 65.6656 | 0.25902 | -0.32696 | -0.079618 | 0.44586 | 65.62 | 0.32696 | 1.0041
0.0125 | 0.61783 | 1.8577 | -66.8153 | 66.8546 | 0.27176 | -0.32059 | -0.07431 | 0.44798 | 66.8153 | 0.32059 | 1.1621
0.0143 | 0.57325 | 1.6964 | -68.2144 | 68.2495 | 0.28556 | -0.31635 | -0.069002 | 0.45329 | 68.2144 | 0.31635 | 1.3295
0.0161 | 0.54246 | 1.5202 | -70.0732 | 70.104 | 0.30149 | -0.31635 | -0.061571 | 0.46391 | 70.0732 | 0.31635 | 1.4968
0.0179 | 0.51911 | 1.3376 | -72.2134 | 72.241 | 0.31741 | -0.31635 | -0.056263 | 0.47558 | 72.2134 | 0.31635 | 1.6642
0.0197 | 0.50531 | 1.1486 | -74.5648 | 74.5892 | 0.33439 | -0.31741 | -0.050955 | 0.48938 | 74.5648 | 0.31741 | 1.8315
0.0215 | 0.50955 | 0.9586 | -77.1253 | 77.1465 | 0.35244 | -0.32166 | -0.044586 | 0.50531 | 77.1253 | 0.32166 | 1.9989
0.0233 | 0.51911 | 0.79299 | -79.7442 | 79.7643 | 0.37367 | -0.32696 | -0.039278 | 0.52335 | 79.7442 | 0.32696 | 2.1662
0.0251 | 0.52442 | 0.64544 | -82.2824 | 82.3004 | 0.3949 | -0.33333 | -0.032909 | 0.54352 | 82.2824 | 0.33333 | 2.3336
0.0269 | 0.53397 | 0.49682 | -84.586 | 84.603 | 0.41507 | -0.3397 | -0.027601 | 0.56369 | 84.586 | 0.3397 | 2.5009
0.0287 | 0.55839 | 0.30998 | -86.6858 | 86.7028 | 0.43418 | -0.34607 | -0.023355 | 0.5828 | 86.6858 | 0.34607 | 2.6683
0.0305 | 0.58917 | 0.06051 | -88.4533 | 88.4713 | 0.45223 | -0.35244 | -0.02017 | 0.59873 | 88.4533 | 0.35244 | 2.8356
0.0323 | 0.60934 | -0.25053 | -89.9979 | 90.018 | 0.46921 | -0.35563 | -0.016985 | 0.61359 | 89.9979 | 0.35563 | 3.003
0.0341 | 0.62951 | -0.57113 | -91.4299 | 91.4544 | 0.4862 | -0.35775 | -0.012739 | 0.62633 | 91.4299 | 0.35775 | 3.1703
0.0359 | 0.65711 | -0.89172 | -92.7951 | 92.8227 | 0.50212 | -0.35775 | -0.0095541 | 0.638 | 92.7951 | 0.35775 | 3.3377
0.0376 | 0.73036 | -1.1741 | -94.4575 | 94.4883 | 0.52229 | -0.362 | -0.0063694 | 0.65499 | 94.4575 | 0.362 | 3.4957
0.0394 | 0.80998 | -1.4554 | -96.1815 | 96.2155 | 0.54246 | -0.36624 | -0.0042463 | 0.6741 | 96.1815 | 0.36624 | 3.6631
0.0412 | 0.91189 | -1.7484 | -98.1497 | 98.1911 | 0.56688 | -0.37686 | -0.0010616 | 0.69851 | 98.1497 | 0.37686 | 3.8304
0.043 | 1.0287 | -2.0456 | -100.2282 | 100.2771 | 0.59236 | -0.38747 | 0 | 0.72505 | 100.2282 | 0.38747 | 3.9978
0.0448 | 1.1391 | -2.3323 | -102.3726 | 102.4289 | 0.6189 | -0.40127 | 0.0021231 | 0.75478 | 102.3726 | 0.40127 | 4.1651
0.0466 | 1.2272 | -2.6231 | -104.5096 | 104.5732 | 0.64756 | -0.41614 | 0.0042463 | 0.78556 | 104.5096 | 0.41614 | 4.3325
0.0484 | 1.3036 | -2.9236 | -106.62 | 106.6921 | 0.67622 | -0.43524 | 0.007431 | 0.81953 | 106.62 | 0.43524 | 4.4998
0.0502 | 1.3854 | -3.2463 | -108.8291 | 108.913 | 0.70594 | -0.45648 | 0.010616 | 0.85563 | 108.8291 | 0.45648 | 4.6672
0.052 | 1.448 | -3.5626 | -110.949 | 111.0435 | 0.73461 | -0.47665 | 0.0138 | 0.89066 | 110.949 | 0.47665 | 4.8345
0.0538 | 1.5032 | -3.879 | -112.9437 | 113.0488 | 0.76221 | -0.50106 | 0.018047 | 0.92569 | 112.9437 | 0.50106 | 5.0019
0.0556 | 1.5308 | -4.189 | -114.7463 | 114.862 | 0.78662 | -0.52866 | 0.021231 | 0.95966 | 114.7463 | 0.52866 | 5.1692
0.0574 | 1.5446 | -4.5021 | -116.4692 | 116.5945 | 0.80892 | -0.55732 | 0.023355 | 0.99363 | 116.4692 | 0.55732 | 5.3366
0.0592 | 1.5499 | -4.8153 | -118.1306 | 118.2643 | 0.83121 | -0.58599 | 0.026539 | 1.0276 | 118.1306 | 0.58599 | 5.5039
0.061 | 1.5648 | -5.1465 | -119.8535 | 119.9979 | 0.8535 | -0.6189 | 0.028662 | 1.0648 | 119.8535 | 0.6189 | 5.6713
0.0627 | 1.5955 | -5.5 | -121.7675 | 121.9257 | 0.87792 | -0.65605 | 0.031847 | 1.1062 | 121.7675 | 0.65605 | 5.8293
0.0645 | 1.6359 | -5.8662 | -123.8896 | 124.0626 | 0.90446 | -0.69427 | 0.035032 | 1.1497 | 123.8896 | 0.69427 | 5.9967
0.0663 | 1.689 | -6.2548 | -126.1083 | 126.2983 | 0.931 | -0.73567 | 0.038217 | 1.1964 | 126.1083 | 0.73567 | 6.164
0.0681 | 1.7527 | -6.6656 | -128.4076 | 128.6168 | 0.95754 | -0.77813 | 0.042463 | 1.2452 | 128.4076 | 0.77813 | 6.3313
0.0699 | 1.8079 | -7.069 | -130.8057 | 131.035 | 0.9862 | -0.82272 | 0.045648 | 1.2951 | 130.8057 | 0.82272 | 6.4987
0.0717 | 1.8662 | -7.482 | -133.2728 | 133.5234 | 1.0138 | -0.87049 | 0.048832 | 1.3471 | 133.2728 | 0.87049 | 6.666
0.0735 | 1.9214 | -7.8896 | -135.8068 | 136.0775 | 1.0414 | -0.91932 | 0.049894 | 1.3992 | 135.8068 | 0.91932 | 6.8334
0.0753 | 1.983 | -8.2951 | -138.3843 | 138.6752 | 1.069 | -0.96815 | 0.053079 | 1.4533 | 138.3843 | 0.96815 | 7.0007
0.0771 | 2.0573 | -8.7197 | -140.9915 | 141.3057 | 1.0977 | -1.017 | 0.055202 | 1.5064 | 140.9915 | 1.017 | 7.1681
0.0789 | 2.0998 | -9.121 | -143.6826 | 144.017 | 1.1274 | -1.0658 | 0.057325 | 1.5605 | 143.6826 | 1.0658 | 7.3354
0.0807 | 2.1412 | -9.5138 | -146.4968 | 146.8524 | 1.1582 | -1.1178 | 0.058386 | 1.6189 | 146.4968 | 1.1178 | 7.5028
0.0825 | 2.1571 | -9.9013 | -149.4193 | 149.7941 | 1.189 | -1.173 | 0.06051 | 1.6794 | 149.4193 | 1.173 | 7.6701
0.0843 | 2.1783 | -10.2962 | -152.3949 | 152.7887 | 1.2197 | -1.2314 | 0.06051 | 1.7431 | 152.3949 | 1.2314 | 7.8375
0.086 | 2.2049 | -10.6911 | -155.4342 | 155.8461 | 1.2527 | -1.2898 | 0.061571 | 1.8079 | 155.4342 | 1.2898 | 7.9955
0.0878 | 2.2282 | -11.0849 | -158.328 | 158.7611 | 1.2834 | -1.3461 | 0.062633 | 1.8694 | 158.328 | 1.3461 | 8.1629
0.0896 | 2.2611 | -11.4501 | -161.1221 | 161.5732 | 1.3132 | -1.4045 | 0.062633 | 1.9321 | 161.1221 | 1.4045 | 8.3302
0.0914 | 2.2771 | -11.8142 | -163.7972 | 164.2675 | 1.3408 | -1.4586 | 0.062633 | 1.9904 | 163.7972 | 1.4586 | 8.4976
0.0932 | 2.2941 | -12.1783 | -166.4183 | 166.9076 | 1.3673 | -1.5149 | 0.062633 | 2.0499 | 166.4183 | 1.5149 | 8.6649
0.095 | 2.3057 | -12.5467 | -169 | 169.5096 | 1.3938 | -1.5701 | 0.063694 | 2.1083 | 169 | 1.5701 | 8.8323
0.0968 | 2.3089 | -12.913 | -171.5032 | 172.034 | 1.4183 | -1.6231 | 0.063694 | 2.1645 | 171.5032 | 1.6231 | 8.9996
0.0986 | 2.3047 | -13.276 | -173.9522 | 174.5032 | 1.4427 | -1.6762 | 0.063694 | 2.2208 | 173.9522 | 1.6762 | 9.167
0.1004 | 2.2994 | -13.6295 | -176.3493 | 176.9204 | 1.466 | -1.7272 | 0.063694 | 2.2749 | 176.3493 | 1.7272 | 9.3343
0.1022 | 2.2919 | -13.9735 | -178.7176 | 179.3089 | 1.4883 | -1.7781 | 0.064756 | 2.328 | 178.7176 | 1.7781 | 9.5017
0.104 | 2.2845 | -14.3047 | -181.0541 | 181.6645 | 1.5096 | -1.827 | 0.065817 | 2.38 | 181.0541 | 1.827 | 9.669
0.1058 | 2.2803 | -14.6072 | -183.3195 | 183.9469 | 1.5308 | -1.8715 | 0.067941 | 2.4278 | 183.3195 | 1.8715 | 9.8364
0.1076 | 2.2739 | -14.8726 | -185.4883 | 186.1306 | 1.5499 | -1.913 | 0.069002 | 2.4724 | 185.4883 | 1.913 | 10.0037
0.1094 | 2.2643 | -15.1253 | -187.621 | 188.2749 | 1.569 | -1.9512 | 0.071125 | 2.5149 | 187.621 | 1.9512 | 10.1711
0.1111 | 2.2537 | -15.3662 | -189.7123 | 190.38 | 1.5892 | -1.9883 | 0.073248 | 2.5552 | 189.7123 | 1.9883 | 10.3291
0.1129 | 2.2325 | -15.5987 | -191.7399 | 192.4204 | 1.6083 | -2.0265 | 0.075372 | 2.5977 | 191.7399 | 2.0265 | 10.4965
0.1147 | 2.2049 | -15.8079 | -193.6815 | 194.3715 | 1.6263 | -2.0637 | 0.077495 | 2.638 | 193.6815 | 2.0637 | 10.6638
0.1165 | 2.173 | -15.9788 | -195.4883 | 196.1858 | 1.6444 | -2.0987 | 0.080679 | 2.6762 | 195.4883 | 2.0987 | 10.8312
0.1183 | 2.138 | -16.1189 | -197.1985 | 197.9013 | 1.6603 | -2.1338 | 0.083864 | 2.7144 | 197.1985 | 2.1338 | 10.9985
0.1201 | 2.1051 | -16.2516 | -198.8503 | 199.5584 | 1.6762 | -2.1677 | 0.085987 | 2.7516 | 198.8503 | 2.1677 | 11.1659
0.1219 | 2.0732 | -16.3832 | -200.4851 | 201.1985 | 1.6911 | -2.2017 | 0.089172 | 2.7877 | 200.4851 | 2.2017 | 11.3332
0.1237 | 2.0403 | -16.4989 | -202.0297 | 202.7473 | 1.7059 | -2.2357 | 0.092357 | 2.8227 | 202.0297 | 2.2357 | 11.5006
0.1255 | 2.0106 | -16.5977 | -203.4851 | 204.2049 | 1.7187 | -2.2686 | 0.095541 | 2.8567 | 203.4851 | 2.2686 | 11.6679
0.1273 | 1.9766 | -16.69 | -204.8333 | 205.5563 | 1.7304 | -2.3025 | 0.098726 | 2.8917 | 204.8333 | 2.3025 | 11.8353
0.1291 | 1.948 | -16.7643 | -205.9947 | 206.7197 | 1.7389 | -2.3344 | 0.10191 | 2.9225 | 205.9947 | 2.3344 | 12.0026
0.1309 | 1.914 | -16.8408 | -207.0074 | 207.7346 | 1.7463 | -2.3631 | 0.1051 | 2.9501 | 207.0074 | 2.3631 | 12.17
0.1327 | 1.8673 | -16.9331 | -207.8981 | 208.6306 | 1.7527 | -2.3907 | 0.10828 | 2.9766 | 207.8981 | 2.3907 | 12.3373
0.1345 | 1.8163 | -17.0223 | -208.7028 | 209.4384 | 1.759 | -2.4161 | 0.1104 | 3.0011 | 208.7028 | 2.4161 | 12.5046
0.1362 | 1.7654 | -17.0934 | -209.4268 | 210.1667 | 1.7643 | -2.4374 | 0.11253 | 3.0212 | 209.4268 | 2.4374 | 12.6627
0.138 | 1.7187 | -17.1316 | -210.0955 | 210.8355 | 1.7675 | -2.4544 | 0.11571 | 3.0382 | 210.0955 | 2.4544 | 12.83
0.1398 | 1.6614 | -17.1805 | -210.535 | 211.2771 | 1.7686 | -2.4735 | 0.11783 | 3.0541 | 210.535 | 2.4735 | 12.9974
0.1416 | 1.5955 | -17.2123 | -210.8588 | 211.6019 | 1.7675 | -2.4926 | 0.12102 | 3.069 | 210.8588 | 2.4926 | 13.1647
0.1434 | 1.5467 | -17.2304 | -211.1858 | 211.9289 | 1.7622 | -2.5117 | 0.1242 | 3.0828 | 211.1858 | 2.5117 | 13.3321
0.1452 | 1.4926 | -17.2558 | -211.4989 | 212.242 | 1.759 | -2.5297 | 0.12739 | 3.0966 | 211.4989 | 2.5297 | 13.4994
0.147 | 1.4363 | -17.2781 | -211.7675 | 212.5106 | 1.7558 | -2.5456 | 0.13057 | 3.1083 | 211.7675 | 2.5456 | 13.6668
0.1488 | 1.3673 | -17.2665 | -211.9437 | 212.6847 | 1.7516 | -2.5531 | 0.1327 | 3.1115 | 211.9437 | 2.5531 | 13.8341
0.1506 | 1.2866 | -17.2378 | -211.913 | 212.6529 | 1.7431 | -2.5637 | 0.13588 | 3.1168 | 211.913 | 2.5637 | 14.0015
0.1524 | 1.2251 | -17.1911 | -211.8546 | 212.5913 | 1.7314 | -2.5754 | 0.13694 | 3.121 | 211.8546 | 2.5754 | 14.1688
0.1542 | 1.1476 | -17.1263 | -211.776 | 212.5085 | 1.7208 | -2.5839 | 0.13907 | 3.1231 | 211.776 | 2.5839 | 14.3362
0.156 | 1.0573 | -17.051 | -211.6677 | 212.397 | 1.7113 | -2.5913 | 0.14013 | 3.1242 | 211.6677 | 2.5913 | 14.5035
0.1578 | 0.97771 | -17.0011 | -211.4851 | 212.2113 | 1.7006 | -2.5987 | 0.14225 | 3.1253 | 211.4851 | 2.5987 | 14.6709
0.1595 | 0.92144 | -16.948 | -211.2442 | 211.9671 | 1.6868 | -2.6051 | 0.14331 | 3.1231 | 211.2442 | 2.6051 | 14.8289
0.1613 | 0.8758 | -16.8779 | -210.9522 | 211.672 | 1.672 | -2.6072 | 0.14437 | 3.1189 | 210.9522 | 2.6072 | 14.9963
0.1631 | 0.80892 | -16.7803 | -210.6093 | 211.3238 | 1.6571 | -2.6072 | 0.1465 | 3.1104 | 210.6093 | 2.6072 | 15.1636
0.1649 | 0.7431 | -16.6741 | -210.2527 | 210.9607 | 1.6412 | -2.6072 | 0.14756 | 3.1019 | 210.2527 | 2.6072 | 15.331
0.1667 | 0.67197 | -16.5711 | -209.8641 | 210.5669 | 1.6274 | -2.6072 | 0.14862 | 3.0945 | 209.8641 | 2.6072 | 15.4983
0.1685 | 0.59023 | -16.4501 | -209.3503 | 210.0456 | 1.6104 | -2.6051 | 0.15074 | 3.0849 | 209.3503 | 2.6051 | 15.6657
0.1703 | 0.4862 | -16.2686 | -208.7304 | 209.4151 | 1.5902 | -2.5977 | 0.1518 | 3.0679 | 208.7304 | 2.5977 | 15.833
0.1721 | 0.40234 | -16.0743 | -208.1019 | 208.7739 | 1.5669 | -2.5892 | 0.1518 | 3.0488 | 208.1019 | 2.5892 | 16.0004
0.1739 | 0.36624 | -15.9023 | -207.4501 | 208.1115 | 1.5393 | -2.5828 | 0.15074 | 3.0297 | 207.4501 | 2.5828 | 16.1677
0.1757 | 0.32166 | -15.7452 | -206.6762 | 207.328 | 1.5127 | -2.5754 | 0.15074 | 3.0106 | 206.6762 | 2.5754 | 16.3351
0.1775 | 0.29936 | -15.5541 | -205.845 | 206.4851 | 1.4851 | -2.569 | 0.14862 | 2.9915 | 205.845 | 2.569 | 16.5024
0.1793 | 0.30042 | -15.3546 | -205.0042 | 205.6327 | 1.4575 | -2.5626 | 0.1465 | 2.9745 | 205.0042 | 2.5626 | 16.6698
0.1811 | 0.30998 | -15.1667 | -204.1603 | 204.7781 | 1.431 | -2.5573 | 0.14544 | 2.9575 | 204.1603 | 2.5573 | 16.8371
0.1829 | 0.30892 | -15 | -203.2389 | 203.8461 | 1.4045 | -2.5456 | 0.14437 | 2.9342 | 203.2389 | 2.5456 | 17.0045
0.1846 | 0.2845 | -14.8461 | -202.2622 | 202.862 | 1.3758 | -2.5382 | 0.14437 | 2.913 | 202.2622 | 2.5382 | 17.1625
0.1864 | 0.26115 | -14.6518 | -201.1242 | 201.7123 | 1.3439 | -2.535 | 0.14331 | 2.8949 | 201.1242 | 2.535 | 17.3299
0.1882 | 0.26539 | -14.4575 | -200.0138 | 200.5902 | 1.3142 | -2.5287 | 0.14225 | 2.8737 | 200.0138 | 2.5287 | 17.4972
0.19 | 0.26858 | -14.2887 | -199.1104 | 199.6773 | 1.2877 | -2.5255 | 0.14119 | 2.8588 | 199.1104 | 2.5255 | 17.6646
0.1918 | 0.27601 | -14.0711 | -198.2102 | 198.7643 | 1.2601 | -2.5223 | 0.13907 | 2.8418 | 198.2102 | 2.5223 | 17.8319
0.1936 | 0.2983 | -13.8599 | -197.396 | 197.9374 | 1.2314 | -2.5212 | 0.13694 | 2.828 | 197.396 | 2.5212 | 17.9993
0.1954 | 0.33015 | -13.6497 | -196.6253 | 197.1539 | 1.2028 | -2.5202 | 0.13482 | 2.8142 | 196.6253 | 2.5202 | 18.1666
0.1972 | 0.37367 | -13.4533 | -195.8652 | 196.3832 | 1.1773 | -2.5212 | 0.13163 | 2.8036 | 195.8652 | 2.5212 | 18.334
0.199 | 0.41507 | -13.2707 | -195.0605 | 195.569 | 1.1529 | -2.5223 | 0.12845 | 2.7951 | 195.0605 | 2.5223 | 18.5013
0.2008 | 0.45541 | -13.0817 | -194.2707 | 194.7696 | 1.1295 | -2.5234 | 0.12633 | 2.7866 | 194.2707 | 2.5234 | 18.6687
0.2026 | 0.49257 | -12.8641 | -193.5605 | 194.0488 | 1.1072 | -2.5244 | 0.12314 | 2.7781 | 193.5605 | 2.5244 | 18.836
0.2044 | 0.53185 | -12.655 | -192.9204 | 193.396 | 1.0828 | -2.5234 | 0.12208 | 2.7665 | 192.9204 | 2.5234 | 19.0033
0.2062 | 0.59342 | -12.4904 | -192.4565 | 192.9236 | 1.0605 | -2.5202 | 0.1189 | 2.7548 | 192.4565 | 2.5202 | 19.1707
0.208 | 0.67304 | -12.3556 | -192.1221 | 192.5828 | 1.0425 | -2.5149 | 0.11677 | 2.7431 | 192.1221 | 2.5149 | 19.338
0.2097 | 0.75902 | -12.1921 | -191.8907 | 192.3429 | 1.0234 | -2.5117 | 0.11465 | 2.7325 | 191.8907 | 2.5117 | 19.4961
0.2115 | 0.84501 | -12.0149 | -191.6539 | 192.0966 | 1.0032 | -2.5096 | 0.11146 | 2.7219 | 191.6539 | 2.5096 | 19.6634
0.2133 | 0.94055 | -11.8365 | -191.172 | 191.6051 | 0.98408 | -2.5149 | 0.10934 | 2.7208 | 191.172 | 2.5149 | 19.8308
0.2151 | 1.0202 | -11.6635 | -190.5955 | 191.0223 | 0.96497 | -2.5138 | 0.10616 | 2.7134 | 190.5955 | 2.5138 | 19.9981
0.2169 | 1.0966 | -11.4607 | -190.2728 | 190.6911 | 0.95117 | -2.5191 | 0.10297 | 2.7144 | 190.2728 | 2.5191 | 20.1655
0.2187 | 1.1709 | -11.2696 | -189.9745 | 190.3875 | 0.93949 | -2.5234 | 0.099788 | 2.7155 | 189.9745 | 2.5234 | 20.3328
0.2205 | 1.2473 | -11.0945 | -189.7696 | 190.1752 | 0.92994 | -2.5276 | 0.096603 | 2.7155 | 189.7696 | 2.5276 | 20.5002
0.2223 | 1.311 | -10.9172 | -189.5573 | 189.9544 | 0.92144 | -2.5318 | 0.09448 | 2.7166 | 189.5573 | 2.5318 | 20.6675
0.2241 | 1.3408 | -10.742 | -189.2654 | 189.6561 | 0.91295 | -2.5372 | 0.091295 | 2.7187 | 189.2654 | 2.5372 | 20.8349
0.2259 | 1.3588 | -10.5743 | -188.8567 | 189.2389 | 0.90234 | -2.5425 | 0.08811 | 2.7197 | 188.8567 | 2.5425 | 21.0022
0.2277 | 1.3864 | -10.4299 | -188.4151 | 188.7898 | 0.89066 | -2.5488 | 0.083864 | 2.7219 | 188.4151 | 2.5488 | 21.1696
0.2295 | 1.4066 | -10.2686 | -187.9522 | 188.3195 | 0.87898 | -2.5531 | 0.079618 | 2.7229 | 187.9522 | 2.5531 | 21.3369
0.2313 | 1.4299 | -10.086 | -187.3684 | 187.7251 | 0.86412 | -2.5563 | 0.075372 | 2.7197 | 187.3684 | 2.5563 | 21.5043
0.2331 | 1.466 | -9.8747 | -186.6062 | 186.9501 | 0.84501 | -2.5605 | 0.072187 | 2.7155 | 186.6062 | 2.5605 | 21.6716
0.2348 | 1.4926 | -9.6476 | -185.7611 | 186.0924 | 0.8259 | -2.5648 | 0.067941 | 2.7123 | 185.7611 | 2.5648 | 21.8297
0.2366 | 1.4968 | -9.4236 | -184.8641 | 185.1815 | 0.80573 | -2.5669 | 0.065817 | 2.707 | 184.8641 | 2.5669 | 21.997
0.2384 | 1.5074 | -9.2187 | -184.121 | 184.4268 | 0.78981 | -2.5679 | 0.061571 | 2.7028 | 184.121 | 2.5679 | 22.1644
0.2402 | 1.5096 | -8.9968 | -183.518 | 183.8142 | 0.77707 | -2.5669 | 0.057325 | 2.6975 | 183.518 | 2.5669 | 22.3317
0.242 | 1.5096 | -8.7845 | -183.0138 | 183.3025 | 0.76645 | -2.5637 | 0.05414 | 2.6921 | 183.0138 | 2.5637 | 22.4991
0.2438 | 1.5191 | -8.603 | -182.5658 | 182.8482 | 0.75796 | -2.5626 | 0.048832 | 2.689 | 182.5658 | 2.5626 | 22.6664
0.2456 | 1.5552 | -8.4756 | -182.0732 | 182.3503 | 0.74735 | -2.5626 | 0.043524 | 2.6858 | 182.0732 | 2.5626 | 22.8338
0.2474 | 1.5637 | -8.3047 | -181.5244 | 181.793 | 0.73673 | -2.5594 | 0.039278 | 2.6794 | 181.5244 | 2.5594 | 23.0011
0.2492 | 1.5743 | -8.087 | -180.8132 | 181.0701 | 0.72293 | -2.5584 | 0.035032 | 2.673 | 180.8132 | 2.5584 | 23.1685
0.251 | 1.5743 | -7.8588 | -180.0064 | 180.2516 | 0.70807 | -2.5573 | 0.031847 | 2.6677 | 180.0064 | 2.5573 | 23.3358
0.2528 | 1.5605 | -7.6327 | -179.1847 | 179.4193 | 0.69321 | -2.5552 | 0.027601 | 2.6614 | 179.1847 | 2.5552 | 23.5032
0.2546 | 1.5531 | -7.4098 | -178.362 | 178.587 | 0.67834 | -2.5541 | 0.023355 | 2.655 | 178.362 | 2.5541 | 23.6705
0.2564 | 1.5584 | -7.1497 | -177.5032 | 177.7155 | 0.66348 | -2.5563 | 0.02017 | 2.6529 | 177.5032 | 2.5563 | 23.8379
0.2581 | 1.5743 | -6.8938 | -176.6699 | 176.8694 | 0.64862 | -2.5573 | 0.015924 | 2.6497 | 176.6699 | 2.5573 | 23.9959
0.2599 | 1.604 | -6.689 | -175.9183 | 176.1083 | 0.63588 | -2.5563 | 0.011677 | 2.6444 | 175.9183 | 2.5563 | 24.1633
0.2617 | 1.6306 | -6.4894 | -175.1699 | 175.3503 | 0.62102 | -2.5563 | 0.007431 | 2.6401 | 175.1699 | 2.5563 | 24.3306
0.2635 | 1.6423 | -6.2707 | -174.4915 | 174.6614 | 0.60722 | -2.5563 | 0.0031847 | 2.6359 | 174.4915 | 2.5563 | 24.498
0.2653 | 1.6656 | -6.052 | -173.914 | 174.0754 | 0.59554 | -2.5531 | -0.0021231 | 2.6295 | 173.914 | 2.5531 | 24.6653
0.2671 | 1.6847 | -5.8248 | -173.3376 | 173.4883 | 0.58493 | -2.5488 | -0.0063694 | 2.6231 | 173.3376 | 2.5488 | 24.8327
0.2689 | 1.6932 | -5.5987 | -172.7537 | 172.896 | 0.57431 | -2.5414 | -0.011677 | 2.6136 | 172.7537 | 2.5414 | 25
0.2707 | 1.7272 | -5.4055 | -172.2431 | 172.3779 | 0.56369 | -2.5361 | -0.016985 | 2.6062 | 172.2431 | 2.5361 | 25.1673
0.2725 | 1.7643 | -5.2272 | -171.7877 | 171.9172 | 0.5552 | -2.5287 | -0.021231 | 2.5966 | 171.7877 | 2.5287 | 25.3347
0.2743 | 1.7951 | -5.051 | -171.3514 | 171.4745 | 0.54671 | -2.5223 | -0.026539 | 2.5881 | 171.3514 | 2.5223 | 25.502
0.2761 | 1.8089 | -4.8726 | -170.982 | 171.0998 | 0.53928 | -2.5138 | -0.030786 | 2.5796 | 170.982 | 2.5138 | 25.6694
0.2779 | 1.8195 | -4.6985 | -170.6072 | 170.7197 | 0.53079 | -2.5074 | -0.036093 | 2.5711 | 170.6072 | 2.5074 | 25.8367
0.2797 | 1.8323 | -4.5255 | -170.2208 | 170.328 | 0.52123 | -2.5011 | -0.042463 | 2.5626 | 170.2208 | 2.5011 | 26.0041
0.2815 | 1.8333 | -4.3556 | -169.776 | 169.8758 | 0.51168 | -2.4926 | -0.047771 | 2.5531 | 169.776 | 2.4926 | 26.1714
0.2832 | 1.8418 | -4.2038 | -169.3726 | 169.4692 | 0.50318 | -2.4841 | -0.053079 | 2.5425 | 169.3726 | 2.4841 | 26.3295
0.285 | 1.8567 | -4.0425 | -169.0106 | 169.103 | 0.49469 | -2.4777 | -0.059448 | 2.534 | 169.0106 | 2.4777 | 26.4968
0.2868 | 1.8917 | -3.8811 | -168.7102 | 168.7983 | 0.48726 | -2.4703 | -0.064756 | 2.5265 | 168.7102 | 2.4703 | 26.6642
0.2886 | 1.9299 | -3.7261 | -168.4448 | 168.5297 | 0.47983 | -2.4639 | -0.070064 | 2.518 | 168.4448 | 2.4639 | 26.8315
0.2904 | 1.9565 | -3.5732 | -168.1752 | 168.258 | 0.47346 | -2.4565 | -0.076433 | 2.5106 | 168.1752 | 2.4565 | 26.9989
0.2922 | 1.9809 | -3.4172 | -167.8864 | 167.9671 | 0.46603 | -2.449 | -0.081741 | 2.5021 | 167.8864 | 2.449 | 27.1662
0.294 | 1.9926 | -3.2739 | -167.5924 | 167.6699 | 0.46072 | -2.4406 | -0.08811 | 2.4926 | 167.5924 | 2.4406 | 27.3336
0.2958 | 2.0127 | -3.1253 | -167.3429 | 167.4172 | 0.45541 | -2.4321 | -0.093418 | 2.4841 | 167.3429 | 2.4321 | 27.5009
0.2976 | 2.035 | -2.9724 | -167.1263 | 167.1964 | 0.44904 | -2.4236 | -0.098726 | 2.4745 | 167.1263 | 2.4236 | 27.6683
0.2994 | 2.0488 | -2.811 | -166.9639 | 167.0318 | 0.44374 | -2.4172 | -0.1051 | 2.4671 | 166.9639 | 2.4172 | 27.8356
0.3012 | 2.0573 | -2.6539 | -166.8206 | 166.8843 | 0.43843 | -2.4098 | -0.1104 | 2.4586 | 166.8206 | 2.4098 | 28.003
0.303 | 2.069 | -2.5042 | -166.6921 | 166.7527 | 0.43418 | -2.4002 | -0.11571 | 2.449 | 166.6921 | 2.4002 | 28.1703
0.3048 | 2.0775 | -2.3577 | -166.5701 | 166.6285 | 0.42994 | -2.3896 | -0.12208 | 2.4374 | 166.5701 | 2.3896 | 28.3377
0.3066 | 2.0807 | -2.2113 | -166.4554 | 166.5117 | 0.42569 | -2.3747 | -0.12739 | 2.4236 | 166.4554 | 2.3747 | 28.505
0.3083 | 2.0839 | -2.0764 | -166.345 | 166.3981 | 0.42357 | -2.3631 | -0.1327 | 2.4108 | 166.345 | 2.3631 | 28.6631
0.3101 | 2.0796 | -1.949 | -166.258 | 166.31 | 0.41932 | -2.3514 | -0.13907 | 2.4002 | 166.258 | 2.3514 | 28.8304
0.3119 | 2.069 | -1.827 | -166.1752 | 166.2261 | 0.41614 | -2.3418 | -0.14544 | 2.3896 | 166.1752 | 2.3418 | 28.9978
0.3137 | 2.0552 | -1.6996 | -166.0977 | 166.1465 | 0.41401 | -2.3333 | -0.1518 | 2.3811 | 166.0977 | 2.3333 | 29.1651
0.3155 | 2.0531 | -1.5456 | -166.0435 | 166.0892 | 0.41083 | -2.3248 | -0.15817 | 2.3726 | 166.0435 | 2.3248 | 29.3325
0.3173 | 2.052 | -1.3864 | -165.9936 | 166.0361 | 0.40764 | -2.3153 | -0.16454 | 2.3641 | 165.9936 | 2.3153 | 29.4998
0.3191 | 2.0435 | -1.2399 | -165.9873 | 166.0276 | 0.40552 | -2.3079 | -0.17091 | 2.3556 | 165.9873 | 2.3079 | 29.6672
0.3209 | 2.0308 | -1.0977 | -165.983 | 166.0212 | 0.4034 | -2.2994 | -0.17728 | 2.3482 | 165.983 | 2.2994 | 29.8345
0.3227 | 2.0202 | -0.94374 | -166.0042 | 166.0414 | 0.40127 | -2.2919 | -0.18365 | 2.3408 | 166.0042 | 2.2919 | 30.0019
0.3245 | 2.017 | -0.78556 | -166.0361 | 166.069 | 0.40021 | -2.2845 | -0.19108 | 2.3333 | 166.0361 | 2.2845 | 30.1692
0.3263 | 2.0096 | -0.62633 | -166.0669 | 166.0977 | 0.39809 | -2.2771 | -0.19958 | 2.3259 | 166.0669 | 2.2771 | 30.3366
0.3281 | 1.9958 | -0.47877 | -166.1316 | 166.1614 | 0.39597 | -2.2707 | -0.20701 | 2.3206 | 166.1316 | 2.2707 | 30.5039
0.3299 | 1.9788 | -0.34183 | -166.2229 | 166.2516 | 0.39384 | -2.2654 | -0.21444 | 2.3153 | 166.2229 | 2.2654 | 30.6713
0.3316 | 1.9703 | -0.21444 | -166.3206 | 166.3471 | 0.39172 | -2.259 | -0.22187 | 2.31 | 166.3206 | 2.259 | 30.8293
0.3334 | 1.966 | -0.083864 | -166.4512 | 166.4766 | 0.39066 | -2.2527 | -0.22824 | 2.3025 | 166.4512 | 2.2527 | 30.9967
0.3352 | 1.9597 | 0.042463 | -166.604 | 166.6285 | 0.3896 | -2.2452 | -0.23673 | 2.2962 | 166.604 | 2.2452 | 31.164
0.337 | 1.9565 | 0.16136 | -166.776 | 166.7994 | 0.38854 | -2.2367 | -0.2431 | 2.2898 | 166.776 | 2.2367 | 31.3313
0.3388 | 1.9533 | 0.27919 | -166.9533 | 166.9766 | 0.38854 | -2.2293 | -0.25053 | 2.2834 | 166.9533 | 2.2293 | 31.4987
0.3406 | 1.9501 | 0.40234 | -167.1624 | 167.1858 | 0.38747 | -2.2229 | -0.2569 | 2.2781 | 167.1624 | 2.2229 | 31.666
0.3424 | 1.9544 | 0.51805 | -167.4013 | 167.4246 | 0.38641 | -2.2176 | -0.26433 | 2.2728 | 167.4013 | 2.2176 | 31.8334
0.3442 | 1.9618 | 0.62951 | -167.6667 | 167.689 | 0.38535 | -2.2123 | -0.27282 | 2.2686 | 167.6667 | 2.2123 | 32.0007
0.346 | 1.965 | 0.74628 | -167.949 | 167.9713 | 0.38535 | -2.207 | -0.28025 | 2.2643 | 167.949 | 2.207 | 32.1681
0.3478 | 1.9607 | 0.85881 | -168.2505 | 168.2728 | 0.38429 | -2.2006 | -0.28662 | 2.259 | 168.2505 | 2.2006 | 32.3354
0.3496 | 1.9544 | 0.96815 | -168.5711 | 168.5945 | 0.38323 | -2.1943 | -0.29406 | 2.2537 | 168.5711 | 2.1943 | 32.5028
0.3514 | 1.949 | 1.0732 | -168.9406 | 168.9639 | 0.38217 | -2.1879 | -0.30255 | 2.2484 | 168.9406 | 2.1879 | 32.6701
0.3532 | 1.9459 | 1.1826 | -169.3822 | 169.4055 | 0.38323 | -2.1805 | -0.31104 | 2.242 | 169.3822 | 2.1805 | 32.8375
0.355 | 1.9459 | 1.2866 | -169.8429 | 169.8673 | 0.38429 | -2.1741 | -0.32059 | 2.2367 | 169.8429 | 2.1741 | 33.0048
0.3567 | 1.948 | 1.3822 | -170.3206 | 170.345 | 0.38535 | -2.1677 | -0.32909 | 2.2325 | 170.3206 | 2.1677 | 33.1629
0.3585 | 1.948 | 1.4713 | -170.7962 | 170.8217 | 0.38641 | -2.1614 | -0.33758 | 2.2282 | 170.7962 | 2.1614 | 33.3302
0.3603 | 1.9427 | 1.552 | -171.3493 | 171.3758 | 0.38854 | -2.1561 | -0.34713 | 2.2251 | 171.3493 | 2.1561 | 33.4976
0.3621 | 1.9331 | 1.6338 | -171.9904 | 172.017 | 0.39384 | -2.1507 | -0.35775 | 2.2219 | 171.9904 | 2.1507 | 33.6649
0.3639 | 1.9257 | 1.7261 | -172.5265 | 172.5552 | 0.39597 | -2.1454 | -0.36624 | 2.2197 | 172.5265 | 2.1454 | 33.8323
0.3657 | 1.914 | 1.8185 | -173.0531 | 173.0807 | 0.39809 | -2.1412 | -0.37473 | 2.2166 | 173.0531 | 2.1412 | 33.9996
0.3675 | 1.897 | 1.9236 | -173.5892 | 173.6178 | 0.40021 | -2.1348 | -0.38323 | 2.2134 | 173.5892 | 2.1348 | 34.167
0.3693 | 1.8811 | 2.0287 | -174.1274 | 174.1571 | 0.4034 | -2.1295 | -0.39066 | 2.2102 | 174.1274 | 2.1295 | 34.3343
0.3711 | 1.8726 | 2.1168 | -174.7091 | 174.7399 | 0.40552 | -2.1253 | -0.39915 | 2.2081 | 174.7091 | 2.1253 | 34.5017
0.3729 | 1.8684 | 2.1837 | -175.2983 | 175.3301 | 0.4087 | -2.12 | -0.40764 | 2.2049 | 175.2983 | 2.12 | 34.669
0.3747 | 1.8599 | 2.2463 | -175.9395 | 175.9713 | 0.41189 | -2.1168 | -0.41614 | 2.2038 | 175.9395 | 2.1168 | 34.8364
0.3765 | 1.8535 | 2.3089 | -176.5669 | 176.5998 | 0.41401 | -2.1125 | -0.42569 | 2.2028 | 176.5669 | 2.1125 | 35.0037
0.3783 | 1.8355 | 2.3471 | -177.1507 | 177.1837 | 0.41614 | -2.1083 | -0.43312 | 2.1996 | 177.1507 | 2.1083 | 35.1711
0.3801 | 1.8312 | 2.4055 | -177.6667 | 177.7006 | 0.41826 | -2.104 | -0.44161 | 2.1975 | 177.6667 | 2.104 | 35.3384
0.3818 | 1.8163 | 2.4682 | -178.1221 | 178.155 | 0.42038 | -2.0998 | -0.44904 | 2.1953 | 178.1221 | 2.0998 | 35.4965
0.3836 | 1.7962 | 2.5223 | -178.4851 | 178.5191 | 0.42251 | -2.0955 | -0.45648 | 2.1932 | 178.4851 | 2.0955 | 35.6638
0.3854 | 1.7781 | 2.5679 | -178.8142 | 178.8471 | 0.42463 | -2.0913 | -0.46497 | 2.19 | 178.8142 | 2.0913 | 35.8312
0.3872 | 1.759 | 2.5977 | -179.1115 | 179.1444 | 0.42569 | -2.0881 | -0.4724 | 2.189 | 179.1115 | 2.0881 | 35.9985
0.389 | 1.7399 | 2.621 | -179.4098 | 179.4427 | 0.42569 | -2.086 | -0.48089 | 2.1879 | 179.4098 | 2.086 | 36.1659
0.3908 | 1.7176 | 2.655 | -179.8588 | 179.8917 | 0.42569 | -2.0839 | -0.48832 | 2.189 | 179.8588 | 2.0839 | 36.3332
0.3926 | 1.7123 | 2.7081 | -180.3832 | 180.4172 | 0.42675 | -2.0817 | -0.49682 | 2.1879 | 180.3832 | 2.0817 | 36.5006
0.3944 | 1.7113 | 2.758 | -180.9395 | 180.9745 | 0.42994 | -2.0796 | -0.50531 | 2.189 | 180.9395 | 2.0796 | 36.6679
0.3962 | 1.7113 | 2.8121 | -181.5446 | 181.5796 | 0.43312 | -2.0775 | -0.51486 | 2.189 | 181.5446 | 2.0775 | 36.8353
0.398 | 1.7049 | 2.8726 | -182.2049 | 182.241 | 0.43631 | -2.0732 | -0.52442 | 2.189 | 182.2049 | 2.0732 | 37.0026
0.3998 | 1.7049 | 2.9522 | -182.8992 | 182.9374 | 0.44055 | -2.0711 | -0.53397 | 2.19 | 182.8992 | 2.0711 | 37.17
0.4016 | 1.7049 | 3.051 | -183.6008 | 183.6401 | 0.4448 | -2.069 | -0.54352 | 2.1911 | 183.6008 | 2.069 | 37.3373
0.4034 | 1.7006 | 3.1433 | -184.3248 | 184.3662 | 0.44904 | -2.0658 | -0.55202 | 2.1921 | 184.3248 | 2.0658 | 37.5046
0.4051 | 1.6964 | 3.2197 | -184.9788 | 185.0212 | 0.45329 | -2.0637 | -0.56157 | 2.1932 | 184.9788 | 2.0637 | 37.6627
0.4069 | 1.6953 | 3.2951 | -185.6932 | 185.7367 | 0.45754 | -2.0616 | -0.57006 | 2.1943 | 185.6932 | 2.0616 | 37.83
0.4087 | 1.6996 | 3.3726 | -186.5191 | 186.5637 | 0.46391 | -2.0594 | -0.57962 | 2.1964 | 186.5191 | 2.0594 | 37.9974
0.4105 | 1.7113 | 3.4607 | -187.4183 | 187.4639 | 0.47028 | -2.0573 | -0.59023 | 2.1996 | 187.4183 | 2.0573 | 38.1647
0.4123 | 1.7282 | 3.5711 | -188.4204 | 188.4692 | 0.47771 | -2.0563 | -0.59979 | 2.2028 | 188.4204 | 2.0563 | 38.3321
0.4141 | 1.7473 | 3.6688 | -189.3875 | 189.4384 | 0.4862 | -2.0552 | -0.60828 | 2.207 | 189.3875 | 2.0552 | 38.4994
0.4159 | 1.759 | 3.7611 | -190.3015 | 190.3535 | 0.49363 | -2.0552 | -0.61783 | 2.2113 | 190.3015 | 2.0552 | 38.6668
0.4177 | 1.7601 | 3.8323 | -191.1465 | 191.1996 | 0.5 | -2.0541 | -0.62739 | 2.2155 | 191.1465 | 2.0541 | 38.8341
0.4195 | 1.759 | 3.896 | -191.9416 | 191.9958 | 0.50743 | -2.0541 | -0.63588 | 2.2197 | 191.9416 | 2.0541 | 39.0015
0.4213 | 1.7537 | 3.9363 | -192.5924 | 192.6476 | 0.51168 | -2.0563 | -0.64544 | 2.224 | 192.5924 | 2.0563 | 39.1688
0.4231 | 1.7251 | 3.9713 | -193.2442 | 193.2994 | 0.51699 | -2.0552 | -0.65393 | 2.2272 | 193.2442 | 2.0552 | 39.3362
0.4249 | 1.6932 | 4.0053 | -193.9013 | 193.9575 | 0.52229 | -2.0552 | -0.66348 | 2.2304 | 193.9013 | 2.0552 | 39.5035
0.4267 | 1.6783 | 4.0403 | -194.6985 | 194.7558 | 0.52866 | -2.0541 | -0.67304 | 2.2346 | 194.6985 | 2.0541 | 39.6709
0.4285 | 1.6635 | 4.069 | -195.4586 | 195.5149 | 0.53503 | -2.0563 | -0.68259 | 2.2399 | 195.4586 | 2.0563 | 39.8382
0.4302 | 1.6518 | 4.0998 | -196.2378 | 196.2951 | 0.5414 | -2.0573 | -0.69214 | 2.2463 | 196.2378 | 2.0573 | 39.9963
0.432 | 1.6454 | 4.1338 | -197.0924 | 197.1507 | 0.54883 | -2.0573 | -0.70276 | 2.2516 | 197.0924 | 2.0573 | 40.1636
0.4338 | 1.6465 | 4.1762 | -198.0255 | 198.0839 | 0.55839 | -2.0573 | -0.71338 | 2.258 | 198.0255 | 2.0573 | 40.331
0.4356 | 1.6454 | 4.224 | -199.0032 | 199.0626 | 0.569 | -2.0594 | -0.72293 | 2.2643 | 199.0032 | 2.0594 | 40.4983
0.4374 | 1.6401 | 4.2675 | -200.0191 | 200.0786 | 0.57962 | -2.0594 | -0.73248 | 2.2707 | 200.0191 | 2.0594 | 40.6657
0.4392 | 1.6285 | 4.3174 | -201.0924 | 201.1529 | 0.59236 | -2.0573 | -0.7431 | 2.276 | 201.0924 | 2.0573 | 40.833
0.441 | 1.6285 | 4.3758 | -202.2282 | 202.2898 | 0.60616 | -2.0573 | -0.75372 | 2.2834 | 202.2282 | 2.0573 | 41.0004
0.4428 | 1.6316 | 4.4406 | -203.4469 | 203.5106 | 0.62102 | -2.0573 | -0.76433 | 2.2909 | 203.4469 | 2.0573 | 41.1677
0.4446 | 1.6316 | 4.4958 | -204.689 | 204.7537 | 0.638 | -2.0563 | -0.77389 | 2.2994 | 204.689 | 2.0563 | 41.3351
0.4464 | 1.6253 | 4.5382 | -205.8132 | 205.879 | 0.65499 | -2.0552 | -0.78344 | 2.3079 | 205.8132 | 2.0552 | 41.5024
0.4482 | 1.6115 | 4.5817 | -206.7962 | 206.862 | 0.66879 | -2.0541 | -0.79406 | 2.3142 | 206.7962 | 2.0541 | 41.6698
0.45 | 1.5902 | 4.6072 | -207.7006 | 207.7675 | 0.68153 | -2.052 | -0.80361 | 2.3206 | 207.7006 | 2.052 | 41.8371
0.4518 | 1.5626 | 4.6125 | -208.4766 | 208.5425 | 0.69214 | -2.051 | -0.81316 | 2.3248 | 208.4766 | 2.051 | 42.0045
0.4536 | 1.5403 | 4.621 | -209.2686 | 209.3344 | 0.70276 | -2.0488 | -0.82272 | 2.3301 | 209.2686 | 2.0488 | 42.1718
0.4553 | 1.5425 | 4.6423 | -210.1953 | 210.2622 | 0.71656 | -2.0478 | -0.83227 | 2.3365 | 210.1953 | 2.0478 | 42.3299
0.4571 | 1.552 | 4.6667 | -211.3079 | 211.3747 | 0.73355 | -2.0478 | -0.84183 | 2.3461 | 211.3079 | 2.0478 | 42.4972
0.4589 | 1.5658 | 4.6868 | -212.4268 | 212.4947 | 0.75159 | -2.0488 | -0.85138 | 2.3556 | 212.4268 | 2.0488 | 42.6646
0.4607 | 1.5552 | 4.7081 | -213.4225 | 213.4904 | 0.7707 | -2.0478 | -0.86093 | 2.3641 | 213.4225 | 2.0478 | 42.8319
0.4625 | 1.5287 | 4.7134 | -214.2941 | 214.362 | 0.78875 | -2.0456 | -0.87049 | 2.3715 | 214.2941 | 2.0456 | 42.9993
0.4643 | 1.4989 | 4.7166 | -215.1168 | 215.1847 | 0.80786 | -2.0435 | -0.88004 | 2.379 | 215.1168 | 2.0435 | 43.1666
0.4661 | 1.4682 | 4.7261 | -216.0255 | 216.0934 | 0.82696 | -2.0414 | -0.8896 | 2.3875 | 216.0255 | 2.0414 | 43.334
0.4679 | 1.4427 | 4.7452 | -217.1008 | 217.1699 | 0.84713 | -2.0372 | -0.89915 | 2.3949 | 217.1008 | 2.0372 | 43.5013
0.4697 | 1.4204 | 4.7696 | -218.2537 | 218.3227 | 0.86837 | -2.034 | -0.90764 | 2.4034 | 218.2537 | 2.034 | 43.6687
0.4715 | 1.3917 | 4.8068 | -219.4798 | 219.5488 | 0.8896 | -2.0287 | -0.91614 | 2.4108 | 219.4798 | 2.0287 | 43.836
0.4733 | 1.3599 | 4.8344 | -220.7293 | 220.7994 | 0.91189 | -2.0234 | -0.92569 | 2.4193 | 220.7293 | 2.0234 | 44.0033
0.4751 | 1.3323 | 4.8567 | -221.8567 | 221.9268 | 0.93524 | -2.0191 | -0.93418 | 2.4278 | 221.8567 | 2.0191 | 44.1707
0.4769 | 1.3036 | 4.8705 | -222.9406 | 223.0106 | 0.95754 | -2.0138 | -0.94268 | 2.4363 | 222.9406 | 2.0138 | 44.338
0.4786 | 1.2601 | 4.8662 | -223.9958 | 224.0658 | 0.97983 | -2.0074 | -0.95117 | 2.4437 | 223.9958 | 2.0074 | 44.4961
0.4804 | 1.1975 | 4.8652 | -225.0902 | 225.1592 | 1.0021 | -1.9979 | -0.9586 | 2.449 | 225.0902 | 1.9979 | 44.6634
0.4822 | 1.1253 | 4.88 | -226.1985 | 226.2665 | 1.0255 | -1.9873 | -0.96603 | 2.4544 | 226.1985 | 1.9873 | 44.8308
0.484 | 1.0531 | 4.8885 | -227.1911 | 227.259 | 1.0488 | -1.9788 | -0.97452 | 2.4618 | 227.1911 | 1.9788 | 44.9981
0.4858 | 0.98938 | 4.8854 | -228.0892 | 228.1571 | 1.069 | -1.9724 | -0.98195 | 2.4682 | 228.0892 | 1.9724 | 45.1655
0.4876 | 0.92781 | 4.8938 | -228.948 | 229.0159 | 1.0902 | -1.9671 | -0.99045 | 2.4745 | 228.948 | 1.9671 | 45.3328
0.4894 | 0.86412 | 4.9108 | -229.81 | 229.8779 | 1.1115 | -1.9607 | -0.99894 | 2.4809 | 229.81 | 1.9607 | 45.5002
0.4912 | 0.78981 | 4.9321 | -230.7272 | 230.7951 | 1.1327 | -1.9522 | -1.0074 | 2.4873 | 230.7272 | 1.9522 | 45.6675
0.493 | 0.70701 | 4.948 | -231.604 | 231.672 | 1.1518 | -1.9437 | -1.0149 | 2.4926 | 231.604 | 1.9437 | 45.8349
0.4948 | 0.61783 | 4.966 | -232.3429 | 232.4108 | 1.172 | -1.9331 | -1.0223 | 2.4968 | 232.3429 | 1.9331 | 46.0022
0.4966 | 0.52442 | 4.9904 | -232.9873 | 233.0563 | 1.1932 | -1.9214 | -1.0297 | 2.5021 | 232.9873 | 1.9214 | 46.1696
0.4984 | 0.41826 | 5.0456 | -233.5796 | 233.6497 | 1.2134 | -1.9098 | -1.0361 | 2.5053 | 233.5796 | 1.9098 | 46.3369
0.5002 | 0.30573 | 5.1104 | -234.1624 | 234.2346 | 1.2314 | -1.897 | -1.0435 | 2.5085 | 234.1624 | 1.897 | 46.5043
0.502 | 0.19639 | 5.1794 | -234.7335 | 234.8068 | 1.2505 | -1.8854 | -1.0488 | 2.5106 | 234.7335 | 1.8854 | 46.6716
0.5037 | 0.087049 | 5.2505 | -235.3291 | 235.4034 | 1.2686 | -1.8737 | -1.0552 | 2.5138 | 235.3291 | 1.8737 | 46.8297
0.5055 | -0.019108 | 5.3057 | -236.0478 | 236.1242 | 1.2845 | -1.8609 | -1.0616 | 2.5149 | 236.0478 | 1.8609 | 46.997
0.5073 | -0.1327 | 5.3439 | -236.8524 | 236.9289 | 1.3004 | -1.845 | -1.069 | 2.5149 | 236.8524 | 1.845 | 47.1644
0.5091 | -0.24522 | 5.3747 | -237.7144 | 237.7919 | 1.3174 | -1.8291 | -1.0754 | 2.5149 | 237.7144 | 1.8291 | 47.3317
0.5109 | -0.36624 | 5.4151 | -238.6115 | 238.69 | 1.3344 | -1.8121 | -1.0817 | 2.5149 | 238.6115 | 1.8121 | 47.4991
0.5127 | -0.48514 | 5.4416 | -239.3875 | 239.466 | 1.3482 | -1.7951 | -1.0881 | 2.5127 | 239.3875 | 1.7951 | 47.6664
0.5145 | -0.60616 | 5.4628 | -240.0881 | 240.1667 | 1.3609 | -1.7781 | -1.0945 | 2.5106 | 240.0881 | 1.7781 | 47.8338
0.5163 | -0.72399 | 5.4735 | -240.7452 | 240.8248 | 1.3737 | -1.7611 | -1.0998 | 2.5085 | 240.7452 | 1.7611 | 48.0011
0.5181 | -0.8535 | 5.4692 | -241.3206 | 241.4002 | 1.3854 | -1.7442 | -1.1062 | 2.5064 | 241.3206 | 1.7442 | 48.1685
0.5199 | -0.97877 | 5.4682 | -241.8132 | 241.8928 | 1.3981 | -1.7272 | -1.1115 | 2.5032 | 241.8132 | 1.7272 | 48.3358
0.5217 | -1.087 | 5.4713 | -242.2431 | 242.3227 | 1.4098 | -1.7102 | -1.1168 | 2.5011 | 242.2431 | 1.7102 | 48.5032
0.5235 | -1.1815 | 5.4798 | -242.6369 | 242.7176 | 1.4214 | -1.6953 | -1.1221 | 2.4989 | 242.6369 | 1.6953 | 48.6705
0.5253 | -1.2834 | 5.5 | -243.0127 | 243.0934 | 1.431 | -1.6805 | -1.1285 | 2.4958 | 243.0127 | 1.6805 | 48.8379
0.5271 | -1.4204 | 5.4851 | -243.2781 | 243.3609 | 1.4406 | -1.6624 | -1.1327 | 2.4915 | 243.2781 | 1.6624 | 49.0052
0.5288 | -1.5541 | 5.4639 | -243.4724 | 243.5541 | 1.4501 | -1.6444 | -1.1369 | 2.4873 | 243.4724 | 1.6444 | 49.1633
0.5306 | -1.6879 | 5.4798 | -243.6656 | 243.7484 | 1.4628 | -1.6253 | -1.1401 | 2.4841 | 243.6656 | 1.6253 | 49.3306
0.5324 | -1.8217 | 5.4904 | -243.8333 | 243.9161 | 1.4724 | -1.6051 | -1.1444 | 2.4788 | 243.8333 | 1.6051 | 49.498
0.5342 | -1.948 | 5.4904 | -243.9533 | 244.0372 | 1.4798 | -1.5849 | -1.1476 | 2.4724 | 243.9533 | 1.5849 | 49.6653
0.536 | -2.0467 | 5.4936 | -244.0234 | 244.1072 | 1.4851 | -1.5669 | -1.1497 | 2.465 | 244.0234 | 1.5669 | 49.8327
0.5378 | -2.1964 | 5.4968 | -244.0106 | 244.0955 | 1.4894 | -1.5467 | -1.1539 | 2.4565 | 244.0106 | 1.5467 | 50
0.5396 | -2.3524 | 5.5127 | -243.9713 | 244.0584 | 1.4958 | -1.5265 | -1.1571 | 2.449 | 243.9713 | 1.5265 | 50.1673
0.5414 | -2.4841 | 5.4915 | -243.8992 | 243.9873 | 1.4989 | -1.5042 | -1.1592 | 2.4384 | 243.8992 | 1.5042 | 50.3347
0.5432 | -2.6242 | 5.4809 | -243.7516 | 243.8418 | 1.5011 | -1.4809 | -1.1614 | 2.4268 | 243.7516 | 1.4809 | 50.502
0.545 | -2.7622 | 5.4639 | -243.5414 | 243.6316 | 1.5032 | -1.4575 | -1.1624 | 2.414 | 243.5414 | 1.4575 | 50.6694
0.5468 | -2.9023 | 5.4437 | -243.2877 | 243.379 | 1.5042 | -1.4321 | -1.1635 | 2.4013 | 243.2877 | 1.4321 | 50.8367
0.5486 | -3.0626 | 5.4469 | -242.983 | 243.0764 | 1.5053 | -1.4076 | -1.1656 | 2.3875 | 242.983 | 1.4076 | 51.0041
0.5504 | -3.207 | 5.4437 | -242.6125 | 242.7081 | 1.5042 | -1.3832 | -1.1667 | 2.3737 | 242.6125 | 1.3832 | 51.1714
0.5521 | -3.3503 | 5.4448 | -242.155 | 242.2527 | 1.5042 | -1.3577 | -1.1677 | 2.3588 | 242.155 | 1.3577 | 51.3295
0.5539 | -3.4936 | 5.4416 | -241.6423 | 241.742 | 1.5032 | -1.3312 | -1.1688 | 2.3429 | 241.6423 | 1.3312 | 51.4968
0.5557 | -3.6529 | 5.4639 | -241.086 | 241.1879 | 1.5 | -1.3047 | -1.1699 | 2.327 | 241.086 | 1.3047 | 51.6642
0.5575 | -3.8079 | 5.4798 | -240.3822 | 240.4883 | 1.4968 | -1.2739 | -1.1699 | 2.3068 | 240.3822 | 1.2739 | 51.8315
0.5593 | -3.9448 | 5.465 | -239.5669 | 239.6741 | 1.4915 | -1.2442 | -1.1677 | 2.2866 | 239.5669 | 1.2442 | 51.9989
0.5611 | -4.0849 | 5.4416 | -238.7049 | 238.8153 | 1.4862 | -1.2144 | -1.1656 | 2.2654 | 238.7049 | 1.2144 | 52.1662
0.5629 | -4.2335 | 5.4214 | -237.81 | 237.9214 | 1.4809 | -1.1858 | -1.1635 | 2.2452 | 237.81 | 1.1858 | 52.3336
0.5647 | -4.3885 | 5.4055 | -236.8259 | 236.9406 | 1.4745 | -1.155 | -1.1603 | 2.2229 | 236.8259 | 1.155 | 52.5009
0.5665 | -4.5446 | 5.4013 | -235.6858 | 235.8036 | 1.466 | -1.121 | -1.155 | 2.1985 | 235.6858 | 1.121 | 52.6683
0.5683 | -4.6985 | 5.3832 | -234.4459 | 234.5669 | 1.4575 | -1.087 | -1.1486 | 2.1709 | 234.4459 | 1.087 | 52.8356
0.5701 | -4.8546 | 5.3493 | -233.0977 | 233.2219 | 1.4469 | -1.052 | -1.1433 | 2.1433 | 233.0977 | 1.052 | 53.003
0.5719 | -5.0117 | 5.2866 | -231.5297 | 231.6561 | 1.4342 | -1.0159 | -1.1359 | 2.1136 | 231.5297 | 1.0159 | 53.1703
0.5737 | -5.1858 | 5.2272 | -229.845 | 229.9756 | 1.4204 | -0.97771 | -1.1285 | 2.0828 | 229.845 | 0.97771 | 53.3377
0.5755 | -5.3726 | 5.1253 | -227.9798 | 228.1125 | 1.4055 | -0.93206 | -1.12 | 2.0456 | 227.9798 | 0.93206 | 53.505
0.5772 | -5.5616 | 5.0382 | -226.035 | 226.172 | 1.3885 | -0.88641 | -1.1104 | 2.0085 | 226.035 | 0.88641 | 53.6631
0.579 | -5.7601 | 4.9522 | -223.9352 | 224.0764 | 1.3715 | -0.8397 | -1.0998 | 1.9713 | 223.9352 | 0.8397 | 53.8304
0.5808 | -5.965 | 4.8471 | -221.6253 | 221.7707 | 1.3514 | -0.78769 | -1.087 | 1.9289 | 221.6253 | 0.78769 | 53.9978
0.5826 | -6.1677 | 4.758 | -219.1858 | 219.3355 | 1.3312 | -0.73248 | -1.0732 | 1.8875 | 219.1858 | 0.73248 | 54.1651
0.5844 | -6.3662 | 4.6794 | -216.6231 | 216.7781 | 1.31 | -0.6741 | -1.0594 | 1.8461 | 216.6231 | 0.6741 | 54.3325
0.5862 | -6.5552 | 4.6115 | -213.9989 | 214.1603 | 1.2898 | -0.61571 | -1.0456 | 1.8025 | 213.9989 | 0.61571 | 54.4998
0.588 | -6.7293 | 4.5658 | -211.3195 | 211.4873 | 1.2686 | -0.5552 | -1.0297 | 1.759 | 211.3195 | 0.5552 | 54.6672
0.5898 | -6.8949 | 4.5308 | -208.5881 | 208.7622 | 1.2463 | -0.49469 | -1.0149 | 1.7155 | 208.5881 | 0.49469 | 54.8345
0.5916 | -7.0541 | 4.4989 | -205.81 | 205.9904 | 1.224 | -0.43206 | -0.99894 | 1.672 | 205.81 | 0.43206 | 55.0019
0.5934 | -7.2081 | 4.4724 | -202.9851 | 203.173 | 1.2006 | -0.3673 | -0.98195 | 1.6285 | 202.9851 | 0.3673 | 55.1692
0.5952 | -7.3588 | 4.4607 | -200.0616 | 200.2569 | 1.1762 | -0.30042 | -0.96497 | 1.5839 | 200.0616 | 0.30042 | 55.3366
0.597 | -7.5138 | 4.4682 | -197.0032 | 197.2091 | 1.1507 | -0.23248 | -0.94586 | 1.5425 | 197.0032 | 0.23248 | 55.5039
0.5988 | -7.6699 | 4.4851 | -193.8747 | 194.0902 | 1.1263 | -0.16348 | -0.92569 | 1.5 | 193.8747 | 0.16348 | 55.6713
0.6006 | -7.8227 | 4.5127 | -190.6348 | 190.8609 | 1.0998 | -0.095541 | -0.90446 | 1.4597 | 190.6348 | 0.095541 | 55.8386
0.6023 | -7.9692 | 4.5531 | -187.3004 | 187.5372 | 1.0732 | -0.028662 | -0.88217 | 1.4225 | 187.3004 | 0.028662 | 55.9967
0.6041 | -8.104 | 4.6072 | -183.8758 | 184.1253 | 1.0456 | 0.037155 | -0.85987 | 1.3885 | 183.8758 | -0.037155 | 56.164
0.6059 | -8.2272 | 4.672 | -180.4076 | 180.6688 | 1.018 | 0.10297 | -0.83546 | 1.3577 | 180.4076 | -0.10297 | 56.3313
0.6077 | -8.3429 | 4.741 | -176.8758 | 177.1497 | 0.99045 | 0.16879 | -0.81104 | 1.3291 | 176.8758 | -0.16879 | 56.4987
0.6095 | -8.4533 | 4.8079 | -173.258 | 173.5446 | 0.96178 | 0.23355 | -0.78556 | 1.3004 | 173.258 | -0.23355 | 56.666
0.6113 | -8.5446 | 4.8715 | -169.604 | 169.9013 | 0.93312 | 0.29724 | -0.75902 | 1.2739 | 169.604 | -0.29724 | 56.8334
0.6131 | -8.6189 | 4.9352 | -165.9448 | 166.2537 | 0.9034 | 0.35987 | -0.73142 | 1.2505 | 165.9448 | -0.35987 | 57.0007
0.6149 | -8.672 | 4.9926 | -162.2166 | 162.5372 | 0.87367 | 0.41826 | -0.70382 | 1.2304 | 162.2166 | -0.41826 | 57.1681
0.6167 | -8.7028 | 5.0361 | -158.5053 | 158.8365 | 0.84395 | 0.4724 | -0.67516 | 1.2113 | 158.5053 | -0.4724 | 57.3354
0.6185 | -8.7346 | 5.0849 | -154.7654 | 155.1083 | 0.81423 | 0.52654 | -0.6465 | 1.1943 | 154.7654 | -0.52654 | 57.5028
0.6203 | -8.7601 | 5.1423 | -151.0074 | 151.362 | 0.78556 | 0.57962 | -0.61783 | 1.1815 | 151.0074 | -0.57962 | 57.6701
0.6221 | -8.7718 | 5.1964 | -147.2325 | 147.5977 | 0.75584 | 0.62845 | -0.58705 | 1.172 | 147.2325 | -0.62845 | 57.8375
0.6239 | -8.7834 | 5.2399 | -143.5414 | 143.9172 | 0.72824 | 0.67516 | -0.55626 | 1.1624 | 143.5414 | -0.67516 | 58.0048
0.6256 | -8.7866 | 5.2803 | -139.8153 | 140.2028 | 0.70064 | 0.71975 | -0.52548 | 1.1571 | 139.8153 | -0.71975 | 58.1629
0.6274 | -8.7824 | 5.3057 | -136.1083 | 136.5064 | 0.67304 | 0.76221 | -0.49363 | 1.1518 | 136.1083 | -0.76221 | 58.3302
0.6292 | -8.7633 | 5.3132 | -132.4236 | 132.8312 | 0.6465 | 0.80042 | -0.46285 | 1.1465 | 132.4236 | -0.80042 | 58.4976
0.631 | -8.7293 | 5.2962 | -128.7941 | 129.2091 | 0.61996 | 0.83546 | -0.43312 | 1.1433 | 128.7941 | -0.83546 | 58.6649
0.6328 | -8.6783 | 5.2357 | -125.1391 | 125.5594 | 0.59448 | 0.86837 | -0.40234 | 1.1433 | 125.1391 | -0.86837 | 58.8323
0.6346 | -8.6168 | 5.1645 | -121.4533 | 121.879 | 0.57006 | 0.89809 | -0.37155 | 1.1423 | 121.4533 | -0.89809 | 58.9996
0.6364 | -8.5478 | 5.0892 | -117.7909 | 118.2208 | 0.54565 | 0.92569 | -0.34183 | 1.1423 | 117.7909 | -0.92569 | 59.167
0.6382 | -8.4735 | 5.0096 | -114.1614 | 114.5955 | 0.52229 | 0.95117 | -0.31104 | 1.1412 | 114.1614 | -0.95117 | 59.3343
0.64 | -8.3907 | 4.9214 | -110.5786 | 111.0159 | 0.5 | 0.9724 | -0.28132 | 1.1401 | 110.5786 | -0.9724 | 59.5017
0.6418 | -8.2994 | 4.811 | -107.0955 | 107.535 | 0.47983 | 0.98938 | -0.25265 | 1.1369 | 107.0955 | -0.98938 | 59.669
0.6436 | -8.207 | 4.6837 | -103.7006 | 104.1412 | 0.46285 | 1.0032 | -0.22611 | 1.1348 | 103.7006 | -1.0032 | 59.8364
0.6454 | -8.1115 | 4.5541 | -100.3758 | 100.8163 | 0.44586 | 1.0138 | -0.19958 | 1.1316 | 100.3758 | -1.0138 | 60.0037
0.6472 | -8.018 | 4.4098 | -97.1157 | 97.5563 | 0.42994 | 1.0212 | -0.1741 | 1.1274 | 97.1157 | -1.0212 | 60.1711
0.649 | -7.9299 | 4.2431 | -93.8885 | 94.328 | 0.41614 | 1.0244 | -0.14968 | 1.12 | 93.8885 | -1.0244 | 60.3384
0.6507 | -7.8429 | 4.0701 | -90.6943 | 91.1338 | 0.40446 | 1.0234 | -0.12845 | 1.1115 | 90.6943 | -1.0234 | 60.4965
0.6525 | -7.7665 | 3.9098 | -87.5796 | 88.0202 | 0.3949 | 1.0191 | -0.10828 | 1.1019 | 87.5796 | -1.0191 | 60.6638
0.6543 | -7.7113 | 3.759 | -84.5456 | 84.9904 | 0.38641 | 1.0117 | -0.090234 | 1.0892 | 84.5456 | -1.0117 | 60.8312
0.6561 | -7.6624 | 3.6115 | -81.5531 | 82.0032 | 0.37898 | 1.0021 | -0.073248 | 1.0764 | 81.5531 | -1.0021 | 60.9985
0.6579 | -7.6231 | 3.4682 | -78.6465 | 79.1019 | 0.37261 | 0.99151 | -0.057325 | 1.0626 | 78.6465 | -0.99151 | 61.1659
0.6597 | -7.5881 | 3.3291 | -75.8811 | 76.3418 | 0.36624 | 0.97771 | -0.044586 | 1.0467 | 75.8811 | -0.97771 | 61.3332
0.6615 | -7.5276 | 3.1614 | -73.2463 | 73.7091 | 0.35987 | 0.9586 | -0.035032 | 1.0255 | 73.2463 | -0.9586 | 61.5006
0.6633 | -7.4363 | 2.948 | -70.7686 | 71.2282 | 0.3535 | 0.93312 | -0.029724 | 1 | 70.7686 | -0.93312 | 61.6679
0.6651 | -7.345 | 2.7293 | -68.3577 | 68.8142 | 0.34713 | 0.90977 | -0.023355 | 0.97452 | 68.3577 | -0.90977 | 61.8353
0.6669 | -7.2378 | 2.4756 | -66.0308 | 66.482 | 0.34076 | 0.88535 | -0.02017 | 0.94904 | 66.0308 | -0.88535 | 62.0026
0.6687 | -7.1157 | 2.1953 | -63.7484 | 64.19 | 0.33333 | 0.85987 | -0.018047 | 0.92251 | 63.7484 | -0.85987 | 62.17
0.6705 | -6.9915 | 1.9204 | -61.552 | 61.9851 | 0.3259 | 0.83333 | -0.015924 | 0.89597 | 61.552 | -0.83333 | 62.3373
0.6723 | -6.8726 | 1.6518 | -59.4427 | 59.8694 | 0.31847 | 0.80892 | -0.014862 | 0.86943 | 59.4427 | -0.80892 | 62.5046
0.6741 | -6.7431 | 1.3843 | -57.4798 | 57.8992 | 0.3121 | 0.7845 | -0.0138 | 0.84501 | 57.4798 | -0.7845 | 62.672
0.6758 | -6.6104 | 1.1253 | -55.6539 | 56.0669 | 0.30573 | 0.76115 | -0.014862 | 0.82059 | 55.6539 | -0.76115 | 62.83
0.6776 | -6.4904 | 0.86624 | -53.9193 | 54.3259 | 0.30149 | 0.73992 | -0.014862 | 0.79936 | 53.9193 | -0.73992 | 62.9974
0.6794 | -6.3684 | 0.6189 | -52.2144 | 52.6146 | 0.29618 | 0.71868 | -0.015924 | 0.77813 | 52.2144 | -0.71868 | 63.1647
0.6812 | -6.2675 | 0.39703 | -50.6476 | 51.0456 | 0.29087 | 0.69958 | -0.015924 | 0.75796 | 50.6476 | -0.69958 | 63.3321
0.683 | -6.1815 | 0.19321 | -49.1996 | 49.5955 | 0.28662 | 0.68153 | -0.016985 | 0.73992 | 49.1996 | -0.68153 | 63.4994
0.6848 | -6.1093 | 0.016985 | -47.9066 | 48.3025 | 0.28238 | 0.66667 | -0.018047 | 0.72399 | 47.9066 | -0.66667 | 63.6668
0.6866 | -6.0265 | -0.14968 | -46.6688 | 47.0648 | 0.27919 | 0.65074 | -0.02017 | 0.70913 | 46.6688 | -0.65074 | 63.8341
0.6884 | -5.9427 | -0.30573 | -45.4628 | 45.8577 | 0.27601 | 0.63588 | -0.021231 | 0.69321 | 45.4628 | -0.63588 | 64.0015
0.6902 | -5.8747 | -0.44161 | -44.3577 | 44.7558 | 0.27176 | 0.62314 | -0.021231 | 0.68047 | 44.3577 | -0.62314 | 64.1688
0.692 | -5.8163 | -0.56794 | -43.3195 | 43.7197 | 0.26858 | 0.6104 | -0.022293 | 0.66773 | 43.3195 | -0.6104 | 64.3362
0.6938 | -5.7622 | -0.6741 | -42.3705 | 42.7718 | 0.26433 | 0.59873 | -0.023355 | 0.65499 | 42.3705 | -0.59873 | 64.5035
0.6956 | -5.689 | -0.77601 | -41.4352 | 41.8386 | 0.26115 | 0.58705 | -0.023355 | 0.64331 | 41.4352 | -0.58705 | 64.6709
0.6974 | -5.6423 | -0.84183 | -40.5807 | 40.9862 | 0.2569 | 0.57643 | -0.023355 | 0.63163 | 40.5807 | -0.57643 | 64.8382
0.6992 | -5.6062 | -0.89172 | -39.8015 | 40.2113 | 0.25265 | 0.56688 | -0.023355 | 0.62102 | 39.8015 | -0.56688 | 65.0056
0.7009 | -5.5732 | -0.9448 | -39.0817 | 39.4947 | 0.24947 | 0.55839 | -0.023355 | 0.61253 | 39.0817 | -0.55839 | 65.1636
0.7027 | -5.5467 | -0.99469 | -38.4204 | 38.8365 | 0.24522 | 0.54989 | -0.022293 | 0.60297 | 38.4204 | -0.54989 | 65.331
0.7045 | -5.5149 | -1.0244 | -37.81 | 38.2293 | 0.24204 | 0.54246 | -0.022293 | 0.59448 | 37.81 | -0.54246 | 65.4983
0.7063 | -5.4703 | -1.0467 | -37.2548 | 37.673 | 0.23779 | 0.53609 | -0.021231 | 0.58705 | 37.2548 | -0.53609 | 65.6657
0.7081 | -5.4225 | -1.0626 | -36.7558 | 37.173 | 0.23355 | 0.52866 | -0.021231 | 0.57856 | 36.7558 | -0.52866 | 65.833
0.7099 | -5.3875 | -1.0488 | -36.3089 | 36.7261 | 0.2293 | 0.52335 | -0.019108 | 0.57113 | 36.3089 | -0.52335 | 66.0004
0.7117 | -5.345 | -1.035 | -35.8843 | 36.2994 | 0.22505 | 0.51805 | -0.018047 | 0.56476 | 35.8843 | -0.51805 | 66.1677
0.7135 | -5.2813 | -1.0127 | -35.4607 | 35.8705 | 0.22187 | 0.51168 | -0.015924 | 0.55839 | 35.4607 | -0.51168 | 66.3351
0.7153 | -5.2325 | -0.9862 | -35.1019 | 35.5074 | 0.21762 | 0.50637 | -0.014862 | 0.55202 | 35.1019 | -0.50637 | 66.5024
0.7171 | -5.1826 | -0.94268 | -34.7951 | 35.1964 | 0.21444 | 0.50212 | -0.012739 | 0.54565 | 34.7951 | -0.50212 | 66.6698
0.7189 | -5.1221 | -0.89809 | -34.5042 | 34.8992 | 0.21019 | 0.49682 | -0.010616 | 0.54034 | 34.5042 | -0.49682 | 66.8371
0.7207 | -5.0849 | -0.8397 | -34.2304 | 34.6221 | 0.20701 | 0.49257 | -0.0084926 | 0.53503 | 34.2304 | -0.49257 | 67.0045
0.7225 | -5.0435 | -0.76858 | -33.9671 | 34.3535 | 0.20382 | 0.48832 | -0.0063694 | 0.52972 | 33.9671 | -0.48832 | 67.1718
0.7242 | -5.034 | -0.69214 | -33.7601 | 34.1465 | 0.20064 | 0.4862 | -0.0031847 | 0.52654 | 33.7601 | -0.4862 | 67.3299
0.726 | -5.0244 | -0.61253 | -33.5934 | 33.9788 | 0.19851 | 0.48514 | -0.0010616 | 0.52442 | 33.5934 | -0.48514 | 67.4972
0.7278 | -4.9968 | -0.55414 | -33.38 | 33.7633 | 0.19533 | 0.48195 | 0.0010616 | 0.52017 | 33.38 | -0.48195 | 67.6646
0.7296 | -4.9469 | -0.51911 | -33.1083 | 33.4873 | 0.19108 | 0.47771 | 0.0031847 | 0.51486 | 33.1083 | -0.47771 | 67.8319
0.7314 | -4.8938 | -0.48301 | -32.8153 | 33.189 | 0.1879 | 0.4724 | 0.0053079 | 0.50849 | 32.8153 | -0.4724 | 67.9993
0.7332 | -4.8312 | -0.4586 | -32.5149 | 32.8822 | 0.18365 | 0.46709 | 0.007431 | 0.50212 | 32.5149 | -0.46709 | 68.1666
0.735 | -4.741 | -0.44904 | -32.189 | 32.5467 | 0.17941 | 0.45966 | 0.0084926 | 0.49469 | 32.189 | -0.45966 | 68.334
0.7368 | -4.6306 | -0.44904 | -31.8344 | 32.1794 | 0.17516 | 0.45223 | 0.010616 | 0.48514 | 31.8344 | -0.45223 | 68.5013
0.7386 | -4.5 | -0.4586 | -31.4586 | 31.7898 | 0.16985 | 0.44268 | 0.011677 | 0.47558 | 31.4586 | -0.44268 | 68.6687
0.7404 | -4.3609 | -0.47452 | -31.0541 | 31.3694 | 0.16454 | 0.43312 | 0.012739 | 0.46391 | 31.0541 | -0.43312 | 68.836
0.7422 | -4.224 | -0.49045 | -30.6285 | 30.9299 | 0.1603 | 0.42357 | 0.0138 | 0.45329 | 30.6285 | -0.42357 | 69.0033
0.744 | -4.0849 | -0.50955 | -30.1868 | 30.4735 | 0.15499 | 0.41295 | 0.014862 | 0.44161 | 30.1868 | -0.41295 | 69.1707
0.7458 | -3.9416 | -0.53185 | -29.724 | 29.9968 | 0.14968 | 0.40127 | 0.015924 | 0.42994 | 29.724 | -0.40127 | 69.338
0.7476 | -3.7972 | -0.55096 | -29.2389 | 29.4968 | 0.14544 | 0.39066 | 0.015924 | 0.4172 | 29.2389 | -0.39066 | 69.5054
0.7493 | -3.6539 | -0.56051 | -28.7495 | 28.9936 | 0.14119 | 0.37898 | 0.016985 | 0.40552 | 28.7495 | -0.37898 | 69.6634
0.7511 | -3.5202 | -0.57006 | -28.2856 | 28.517 | 0.13694 | 0.36837 | 0.016985 | 0.39384 | 28.2856 | -0.36837 | 69.8308
0.7529 | -3.3907 | -0.56582 | -27.8397 | 28.0594 | 0.13376 | 0.35669 | 0.018047 | 0.38217 | 27.8397 | -0.35669 | 69.9981
0.7547 | -3.2622 | -0.56263 | -27.3949 | 27.6008 | 0.12951 | 0.34607 | 0.018047 | 0.37049 | 27.3949 | -0.34607 | 70.1655
0.7565 | -3.138 | -0.55732 | -26.949 | 27.1433 | 0.12633 | 0.33546 | 0.019108 | 0.35881 | 26.949 | -0.33546 | 70.3328
0.7583 | -3.0212 | -0.54883 | -26.5053 | 26.69 | 0.12314 | 0.32378 | 0.019108 | 0.3482 | 26.5053 | -0.32378 | 70.5002
0.7601 | -2.9204 | -0.54246 | -26.069 | 26.2442 | 0.11996 | 0.31423 | 0.02017 | 0.33758 | 26.069 | -0.31423 | 70.6675
0.7619 | -2.827 | -0.53609 | -25.6433 | 25.811 | 0.11783 | 0.30467 | 0.02017 | 0.32803 | 25.6433 | -0.30467 | 70.8349
0.7637 | -2.7473 | -0.53079 | -25.2155 | 25.3769 | 0.11465 | 0.29512 | 0.02017 | 0.31741 | 25.2155 | -0.29512 | 71.0022
0.7655 | -2.673 | -0.52654 | -24.7898 | 24.9459 | 0.11146 | 0.28662 | 0.021231 | 0.30892 | 24.7898 | -0.28662 | 71.1696
0.7673 | -2.6072 | -0.51592 | -24.3832 | 24.535 | 0.10934 | 0.27707 | 0.021231 | 0.29936 | 24.3832 | -0.27707 | 71.3369
0.7691 | -2.5658 | -0.48514 | -23.9926 | 24.1412 | 0.10722 | 0.2707 | 0.021231 | 0.29299 | 23.9926 | -0.2707 | 71.5043
0.7709 | -2.517 | -0.46072 | -23.6253 | 23.7707 | 0.1051 | 0.26327 | 0.022293 | 0.28556 | 23.6253 | -0.26327 | 71.6716
0.7727 | -2.466 | -0.43843 | -23.2718 | 23.414 | 0.10297 | 0.25584 | 0.022293 | 0.27707 | 23.2718 | -0.25584 | 71.839
0.7744 | -2.4183 | -0.41401 | -22.9427 | 23.0807 | 0.10191 | 0.24841 | 0.023355 | 0.2707 | 22.9427 | -0.24841 | 71.997
0.7762 | -2.3843 | -0.38004 | -22.6635 | 22.7983 | 0.099788 | 0.24204 | 0.023355 | 0.26433 | 22.6635 | -0.24204 | 72.1644
0.778 | -2.3662 | -0.33333 | -22.4448 | 22.5786 | 0.098726 | 0.23673 | 0.024416 | 0.25902 | 22.4448 | -0.23673 | 72.3317
0.7798 | -2.3535 | -0.30786 | -22.2123 | 22.3471 | 0.096603 | 0.23248 | 0.025478 | 0.25372 | 22.2123 | -0.23248 | 72.4991
0.7816 | -2.3386 | -0.28875 | -21.9395 | 22.0743 | 0.095541 | 0.22718 | 0.025478 | 0.24947 | 21.9395 | -0.22718 | 72.6664
0.7834 | -2.3163 | -0.27495 | -21.621 | 21.7548 | 0.093418 | 0.22293 | 0.025478 | 0.24416 | 21.621 | -0.22293 | 72.8338
0.7852 | -2.2941 | -0.25159 | -21.2919 | 21.4257 | 0.091295 | 0.21762 | 0.025478 | 0.23885 | 21.2919 | -0.21762 | 73.0011
0.787 | -2.2845 | -0.21338 | -20.9703 | 21.1051 | 0.090234 | 0.21338 | 0.025478 | 0.23461 | 20.9703 | -0.21338 | 73.1685
0.7888 | -2.2739 | -0.17516 | -20.6614 | 20.7962 | 0.089172 | 0.20913 | 0.025478 | 0.23036 | 20.6614 | -0.20913 | 73.3358
0.7906 | -2.2622 | -0.14225 | -20.3811 | 20.517 | 0.08811 | 0.20594 | 0.025478 | 0.22718 | 20.3811 | -0.20594 | 73.5032
0.7924 | -2.2728 | -0.098726 | -20.1476 | 20.2856 | 0.087049 | 0.20382 | 0.026539 | 0.22399 | 20.1476 | -0.20382 | 73.6705
0.7942 | -2.3057 | -0.050955 | -19.9692 | 20.1136 | 0.085987 | 0.20276 | 0.026539 | 0.22399 | 19.9692 | -0.20276 | 73.8379
0.796 | -2.345 | -0.02017 | -19.7951 | 19.9459 | 0.084926 | 0.20276 | 0.027601 | 0.22293 | 19.7951 | -0.20276 | 74.0052
0.7977 | -2.3471 | -0.0053079 | -19.6316 | 19.7856 | 0.082803 | 0.20276 | 0.027601 | 0.22187 | 19.6316 | -0.20276 | 74.1633
0.7995 | -2.31 | 0.0021231 | -19.4183 | 19.569 | 0.080679 | 0.20064 | 0.026539 | 0.21975 | 19.4183 | -0.20064 | 74.3306
0.8013 | -2.3057 | 0.016985 | -19.1507 | 19.3025 | 0.078556 | 0.19851 | 0.027601 | 0.21656 | 19.1507 | -0.19851 | 74.498
0.8031 | -2.293 | 0.031847 | -18.8769 | 19.0297 | 0.076433 | 0.19533 | 0.027601 | 0.21338 | 18.8769 | -0.19533 | 74.6653
0.8049 | -2.2537 | 0.029724 | -18.5467 | 18.6975 | 0.073248 | 0.19108 | 0.027601 | 0.20807 | 18.5467 | -0.19108 | 74.8327
0.8067 | -2.2325 | 0.03397 | -18.2346 | 18.3864 | 0.070064 | 0.1879 | 0.027601 | 0.20488 | 18.2346 | -0.1879 | 75
0.8085 | -2.2282 | 0.052017 | -17.9427 | 18.0955 | 0.067941 | 0.18684 | 0.027601 | 0.20276 | 17.9427 | -0.18684 | 75.1673
0.8103 | -2.2144 | 0.071125 | -17.7113 | 17.8641 | 0.065817 | 0.18577 | 0.027601 | 0.20064 | 17.7113 | -0.18577 | 75.3347
0.8121 | -2.2219 | 0.10403 | -17.5138 | 17.6699 | 0.063694 | 0.18577 | 0.027601 | 0.19958 | 17.5138 | -0.18577 | 75.502
0.8139 | -2.2272 | 0.14013 | -17.3333 | 17.4926 | 0.062633 | 0.18471 | 0.027601 | 0.19851 | 17.3333 | -0.18471 | 75.6694
0.8157 | -2.2367 | 0.18153 | -17.1688 | 17.3312 | 0.06051 | 0.18471 | 0.027601 | 0.19851 | 17.1688 | -0.18471 | 75.8367
0.8175 | -2.2665 | 0.2293 | -17.0287 | 17.1964 | 0.059448 | 0.18577 | 0.028662 | 0.19851 | 17.0287 | -0.18577 | 76.0041
0.8193 | -2.2834 | 0.25902 | -16.8779 | 17.051 | 0.058386 | 0.18684 | 0.028662 | 0.19958 | 16.8779 | -0.18684 | 76.1714
0.8211 | -2.2792 | 0.27813 | -16.7144 | 16.8885 | 0.056263 | 0.18684 | 0.028662 | 0.19851 | 16.7144 | -0.18684 | 76.3388
0.8228 | -2.2665 | 0.30255 | -16.5669 | 16.7399 | 0.05414 | 0.18684 | 0.028662 | 0.19851 | 16.5669 | -0.18684 | 76.4968
0.8246 | -2.2537 | 0.33652 | -16.4406 | 16.6136 | 0.052017 | 0.18684 | 0.028662 | 0.19745 | 16.4406 | -0.18684 | 76.6642
0.8264 | -2.2399 | 0.37473 | -16.328 | 16.5 | 0.049894 | 0.18684 | 0.029724 | 0.19745 | 16.328 | -0.18684 | 76.8315
0.8282 | -2.2176 | 0.40446 | -16.224 | 16.3917 | 0.047771 | 0.1879 | 0.029724 | 0.19745 | 16.224 | -0.1879 | 76.9989
0.83 | -2.1762 | 0.41614 | -16.1168 | 16.2803 | 0.044586 | 0.18684 | 0.030786 | 0.19745 | 16.1168 | -0.18684 | 77.1662
0.8318 | -2.1614 | 0.45541 | -16.0276 | 16.19 | 0.042463 | 0.1879 | 0.030786 | 0.19745 | 16.0276 | -0.1879 | 77.3336
0.8336 | -2.1571 | 0.50212 | -15.9469 | 16.1115 | 0.04034 | 0.18896 | 0.031847 | 0.19745 | 15.9469 | -0.18896 | 77.5009
0.8354 | -2.1805 | 0.56794 | -15.896 | 16.0669 | 0.039278 | 0.19214 | 0.032909 | 0.20064 | 15.896 | -0.19214 | 77.6683
0.8372 | -2.1953 | 0.6242 | -15.8397 | 16.017 | 0.038217 | 0.19533 | 0.032909 | 0.20276 | 15.8397 | -0.19533 | 77.8356
0.839 | -2.1985 | 0.67091 | -15.7919 | 15.9735 | 0.037155 | 0.19745 | 0.03397 | 0.20594 | 15.7919 | -0.19745 | 78.003
0.8408 | -2.1953 | 0.71444 | -15.7527 | 15.9363 | 0.035032 | 0.20064 | 0.035032 | 0.20807 | 15.7527 | -0.20064 | 78.1703
0.8426 | -2.1932 | 0.76752 | -15.7622 | 15.9469 | 0.032909 | 0.20276 | 0.036093 | 0.21019 | 15.7622 | -0.20276 | 78.3377
0.8444 | -2.2123 | 0.83121 | -15.8132 | 16.0021 | 0.031847 | 0.20594 | 0.036093 | 0.21231 | 15.8132 | -0.20594 | 78.505
0.8462 | -2.2495 | 0.9034 | -15.8758 | 16.0722 | 0.031847 | 0.20913 | 0.037155 | 0.21656 | 15.8758 | -0.20913 | 78.6724
0.8479 | -2.2739 | 0.96285 | -15.9384 | 16.1401 | 0.029724 | 0.21231 | 0.038217 | 0.21975 | 15.9384 | -0.21231 | 78.8304
0.8497 | -2.2771 | 1.0032 | -15.9926 | 16.1964 | 0.027601 | 0.2155 | 0.039278 | 0.22187 | 15.9926 | -0.2155 | 78.9978
0.8515 | -2.2781 | 1.0541 | -16.0541 | 16.2601 | 0.026539 | 0.21762 | 0.04034 | 0.22399 | 16.0541 | -0.21762 | 79.1651
0.8533 | -2.2569 | 1.0839 | -16.1051 | 16.3089 | 0.024416 | 0.21975 | 0.04034 | 0.22611 | 16.1051 | -0.21975 | 79.3325
0.8551 | -2.2558 | 1.1306 | -16.1433 | 16.3503 | 0.022293 | 0.22293 | 0.041401 | 0.22824 | 16.1433 | -0.22293 | 79.4998
0.8569 | -2.2399 | 1.1635 | -16.1624 | 16.3694 | 0.019108 | 0.22505 | 0.042463 | 0.23036 | 16.1624 | -0.22505 | 79.6672
0.8587 | -2.1826 | 1.1741 | -16.1529 | 16.3524 | 0.014862 | 0.22505 | 0.042463 | 0.23036 | 16.1529 | -0.22505 | 79.8345
0.8605 | -2.1178 | 1.1826 | -16.1423 | 16.3344 | 0.011677 | 0.22505 | 0.043524 | 0.23036 | 16.1423 | -0.22505 | 80.0019
0.8623 | -2.0626 | 1.1975 | -16.1783 | 16.3641 | 0.0084926 | 0.22611 | 0.043524 | 0.23036 | 16.1783 | -0.22611 | 80.1692
0.8641 | -2.0021 | 1.2144 | -16.2272 | 16.4055 | 0.0042463 | 0.22611 | 0.044586 | 0.23142 | 16.2272 | -0.22611 | 80.3366
0.8659 | -1.931 | 1.2346 | -16.3025 | 16.4735 | 0.0021231 | 0.22611 | 0.044586 | 0.23142 | 16.3025 | -0.22611 | 80.5039
0.8677 | -1.8673 | 1.2611 | -16.3811 | 16.5456 | -0.0010616 | 0.22505 | 0.045648 | 0.23036 | 16.3811 | -0.22505 | 80.6713
0.8695 | -1.8036 | 1.2877 | -16.4607 | 16.6189 | -0.0031847 | 0.22399 | 0.046709 | 0.2293 | 16.4607 | -0.22399 | 80.8386
0.8712 | -1.7505 | 1.3291 | -16.5711 | 16.7261 | -0.0042463 | 0.22399 | 0.047771 | 0.2293 | 16.5711 | -0.22399 | 80.9967
0.873 | -1.7017 | 1.3705 | -16.7219 | 16.8726 | -0.0053079 | 0.22293 | 0.048832 | 0.2293 | 16.7219 | -0.22293 | 81.164
0.8748 | -1.6773 | 1.4257 | -16.8875 | 17.0382 | -0.0053079 | 0.22399 | 0.048832 | 0.23036 | 16.8875 | -0.22399 | 81.3313
0.8766 | -1.6699 | 1.4947 | -17.0732 | 17.2272 | -0.0042463 | 0.22505 | 0.049894 | 0.23142 | 17.0732 | -0.22505 | 81.4987
0.8784 | -1.6571 | 1.5722 | -17.2834 | 17.4406 | -0.0031847 | 0.22718 | 0.050955 | 0.23355 | 17.2834 | -0.22718 | 81.666
0.8802 | -1.6656 | 1.6529 | -17.5202 | 17.6847 | -0.0031847 | 0.22824 | 0.052017 | 0.23567 | 17.5202 | -0.22824 | 81.8334
0.882 | -1.6741 | 1.7304 | -17.7675 | 17.9384 | -0.0021231 | 0.23036 | 0.052017 | 0.23779 | 17.7675 | -0.23036 | 82.0007
0.8838 | -1.7059 | 1.8153 | -18.0318 | 18.2123 | -0.0010616 | 0.23355 | 0.053079 | 0.24098 | 18.0318 | -0.23355 | 82.1681
0.8856 | -1.7251 | 1.8949 | -18.2898 | 18.4777 | -0.0010616 | 0.23673 | 0.05414 | 0.24416 | 18.2898 | -0.23673 | 82.3354
0.8874 | -1.7357 | 1.9703 | -18.5616 | 18.7558 | 0 | 0.23885 | 0.055202 | 0.24628 | 18.5616 | -0.23885 | 82.5028
0.8892 | -1.759 | 2.0403 | -18.8482 | 19.0499 | 0.0010616 | 0.24098 | 0.055202 | 0.24841 | 18.8482 | -0.24098 | 82.6701
0.891 | -1.7771 | 2.0998 | -19.1178 | 19.3248 | 0.0031847 | 0.24204 | 0.056263 | 0.24947 | 19.1178 | -0.24204 | 82.8375
0.8928 | -1.7856 | 2.1454 | -19.3248 | 19.535 | 0.0031847 | 0.24204 | 0.056263 | 0.24947 | 19.3248 | -0.24204 | 83.0048
0.8946 | -1.8004 | 2.1858 | -19.5149 | 19.7293 | 0.0053079 | 0.24204 | 0.056263 | 0.24947 | 19.5149 | -0.24204 | 83.1722
0.8963 | -1.8142 | 2.2272 | -19.7134 | 19.9321 | 0.0063694 | 0.24204 | 0.055202 | 0.24947 | 19.7134 | -0.24204 | 83.3302
0.8981 | -1.8461 | 2.2665 | -19.8885 | 20.1125 | 0.007431 | 0.2431 | 0.055202 | 0.25053 | 19.8885 | -0.2431 | 83.4976
0.8999 | -1.8652 | 2.2951 | -20.0308 | 20.2569 | 0.0084926 | 0.24204 | 0.05414 | 0.24947 | 20.0308 | -0.24204 | 83.6649
0.9017 | -1.8758 | 2.3057 | -20.1423 | 20.3694 | 0.010616 | 0.23992 | 0.052017 | 0.24628 | 20.1423 | -0.23992 | 83.8323
0.9035 | -1.88 | 2.2962 | -20.2293 | 20.4554 | 0.010616 | 0.23673 | 0.050955 | 0.24416 | 20.2293 | -0.23673 | 83.9996
0.9053 | -1.8609 | 2.2718 | -20.2728 | 20.4947 | 0.011677 | 0.23248 | 0.048832 | 0.23992 | 20.2728 | -0.23248 | 84.167
0.9071 | -1.8291 | 2.2389 | -20.3068 | 20.5223 | 0.011677 | 0.22824 | 0.046709 | 0.23567 | 20.3068 | -0.22824 | 84.3343
0.9089 | -1.7866 | 2.1985 | -20.3376 | 20.5446 | 0.011677 | 0.22187 | 0.044586 | 0.2293 | 20.3376 | -0.22187 | 84.5017
0.9107 | -1.7484 | 2.1529 | -20.3556 | 20.5541 | 0.010616 | 0.21656 | 0.042463 | 0.22399 | 20.3556 | -0.21656 | 84.669
0.9125 | -1.6911 | 2.0913 | -20.3471 | 20.5361 | 0.010616 | 0.20913 | 0.04034 | 0.21656 | 20.3471 | -0.20913 | 84.8364
0.9143 | -1.6444 | 2.017 | -20.3132 | 20.4926 | 0.010616 | 0.2017 | 0.038217 | 0.20807 | 20.3132 | -0.2017 | 85.0037
0.9161 | -1.6115 | 1.9501 | -20.2792 | 20.4512 | 0.0095541 | 0.19427 | 0.036093 | 0.19958 | 20.2792 | -0.19427 | 85.1711
0.9179 | -1.5977 | 1.8907 | -20.2633 | 20.4299 | 0.010616 | 0.18684 | 0.032909 | 0.19321 | 20.2633 | -0.18684 | 85.3384
0.9197 | -1.569 | 1.8121 | -20.2282 | 20.3875 | 0.010616 | 0.18047 | 0.030786 | 0.18684 | 20.2282 | -0.18047 | 85.5058
0.9214 | -1.5096 | 1.7197 | -20.1773 | 20.327 | 0.0095541 | 0.17304 | 0.028662 | 0.18047 | 20.1773 | -0.17304 | 85.6638
0.9232 | -1.4501 | 1.6231 | -20.0892 | 20.2282 | 0.0095541 | 0.16667 | 0.026539 | 0.1741 | 20.0892 | -0.16667 | 85.8312
0.925 | -1.3885 | 1.5425 | -19.9735 | 20.103 | 0.0084926 | 0.15924 | 0.024416 | 0.16879 | 19.9735 | -0.15924 | 85.9985
0.9268 | -1.3376 | 1.4692 | -19.827 | 19.948 | 0.0084926 | 0.15287 | 0.022293 | 0.16454 | 19.827 | -0.15287 | 86.1659
0.9286 | -1.2675 | 1.3949 | -19.7601 | 19.8715 | 0.0084926 | 0.14437 | 0.02017 | 0.15711 | 19.7601 | -0.14437 | 86.3332
0.9304 | -1.1879 | 1.3163 | -19.7314 | 19.8301 | 0.007431 | 0.13482 | 0.018047 | 0.14756 | 19.7314 | -0.13482 | 86.5006
0.9322 | -1.1316 | 1.2537 | -19.7367 | 19.8259 | 0.0084926 | 0.12633 | 0.015924 | 0.13907 | 19.7367 | -0.12633 | 86.6679
0.934 | -1.0987 | 1.2123 | -19.7643 | 19.8471 | 0.010616 | 0.1189 | 0.0138 | 0.13376 | 19.7643 | -0.1189 | 86.8353
0.9358 | -1.0849 | 1.1858 | -19.8142 | 19.8928 | 0.011677 | 0.11253 | 0.012739 | 0.12845 | 19.8142 | -0.11253 | 87.0026
0.9376 | -1.0828 | 1.1645 | -19.9023 | 19.9798 | 0.0138 | 0.10722 | 0.011677 | 0.1242 | 19.9023 | -0.10722 | 87.17
0.9394 | -1.0839 | 1.1444 | -19.9989 | 20.0743 | 0.015924 | 0.10191 | 0.011677 | 0.11996 | 19.9989 | -0.10191 | 87.3373
0.9412 | -1.104 | 1.1295 | -20.1231 | 20.1985 | 0.018047 | 0.099788 | 0.010616 | 0.1189 | 20.1231 | -0.099788 | 87.5046
0.943 | -1.1338 | 1.1125 | -20.2665 | 20.3418 | 0.021231 | 0.097665 | 0.0095541 | 0.11783 | 20.2665 | -0.097665 | 87.672
0.9447 | -1.1529 | 1.0828 | -20.4108 | 20.4851 | 0.023355 | 0.093418 | 0.0095541 | 0.11571 | 20.4108 | -0.093418 | 87.83
0.9465 | -1.1497 | 1.0393 | -20.5212 | 20.5924 | 0.024416 | 0.08811 | 0.0084926 | 0.11253 | 20.5212 | -0.08811 | 87.9974
0.9483 | -1.1253 | 0.98408 | -20.5764 | 20.6423 | 0.025478 | 0.081741 | 0.007431 | 0.11146 | 20.5764 | -0.081741 | 88.1647
0.9501 | -1.0775 | 0.90764 | -20.5839 | 20.6454 | 0.025478 | 0.07431 | 0.0063694 | 0.11146 | 20.5839 | -0.07431 | 88.3321
0.9519 | -1.0265 | 0.83333 | -20.5467 | 20.603 | 0.025478 | 0.067941 | 0.0063694 | 0.11465 | 20.5467 | -0.067941 | 88.4994
0.9537 | -0.99682 | 0.79512 | -20.5446 | 20.5998 | 0.025478 | 0.061571 | 0.0063694 | 0.11677 | 20.5446 | -0.061571 | 88.6668
0.9555 | -1 | 0.76221 | -20.5021 | 20.5563 | 0.026539 | 0.059448 | 0.0063694 | 0.1189 | 20.5021 | -0.059448 | 88.8341
0.9573 | -0.99469 | 0.73673 | -20.518 | 20.5722 | 0.027601 | 0.056263 | 0.0063694 | 0.11783 | 20.518 | -0.056263 | 89.0015
0.9591 | -0.98832 | 0.72718 | -20.551 | 20.604 | 0.028662 | 0.053079 | 0.0053079 | 0.11571 | 20.551 | -0.053079 | 89.1688
0.9609 | -0.99151 | 0.74522 | -20.6263 | 20.6805 | 0.030786 | 0.049894 | 0.0053079 | 0.11571 | 20.6263 | -0.049894 | 89.3362
0.9627 | -1.0032 | 0.78238 | -20.7463 | 20.8004 | 0.032909 | 0.045648 | 0.0053079 | 0.11571 | 20.7463 | -0.045648 | 89.5035
0.9645 | -1.0276 | 0.83333 | -20.9374 | 20.9947 | 0.035032 | 0.042463 | 0.0053079 | 0.11571 | 20.9374 | -0.042463 | 89.6709
0.9663 | -1.0573 | 0.90658 | -21.1752 | 21.2357 | 0.038217 | 0.039278 | 0.0053079 | 0.11677 | 21.1752 | -0.039278 | 89.8382
0.9681 | -1.0849 | 0.97134 | -21.4342 | 21.4968 | 0.04034 | 0.037155 | 0.0053079 | 0.11677 | 21.4342 | -0.037155 | 90.0056
0.9698 | -1.104 | 1.0287 | -21.6985 | 21.7643 | 0.042463 | 0.036093 | 0.0042463 | 0.11677 | 21.6985 | -0.036093 | 90.1636
0.9716 | -1.1295 | 1.087 | -21.9936 | 22.0616 | 0.044586 | 0.032909 | 0.0042463 | 0.11571 | 21.9936 | -0.032909 | 90.331
0.9734 | -1.1401 | 1.1497 | -22.345 | 22.4161 | 0.046709 | 0.027601 | 0.0031847 | 0.11571 | 22.345 | -0.027601 | 90.4983
0.9752 | -1.1359 | 1.2091 | -22.7314 | 22.8047 | 0.047771 | 0.022293 | 0.0021231 | 0.11465 | 22.7314 | -0.022293 | 90.6657
0.977 | -1.1189 | 1.2622 | -23.1242 | 23.1996 | 0.048832 | 0.018047 | 0.0010616 | 0.11465 | 23.1242 | -0.018047 | 90.833
0.9788 | -1.086 | 1.327 | -23.5414 | 23.6168 | 0.049894 | 0.011677 | 0 | 0.11571 | 23.5414 | -0.011677 | 91.0004
0.9806 | -1.0584 | 1.4151 | -24.0042 | 24.0807 | 0.050955 | 0.0042463 | -0.0010616 | 0.11677 | 24.0042 | -0.0042463 | 91.1677
0.9824 | -1.035 | 1.5191 | -24.5106 | 24.5902 | 0.052017 | -0.0031847 | -0.0021231 | 0.1189 | 24.5106 | 0.0031847 | 91.3351
0.9842 | -1.0117 | 1.6348 | -25.0499 | 25.1338 | 0.053079 | -0.011677 | -0.0042463 | 0.12102 | 25.0499 | 0.011677 | 91.5024
0.986 | -0.99894 | 1.7304 | -25.5839 | 25.672 | 0.05414 | -0.021231 | -0.0053079 | 0.11996 | 25.5839 | 0.021231 | 91.6698
0.9878 | -0.96285 | 1.8185 | -26.1231 | 26.2134 | 0.05414 | -0.030786 | -0.007431 | 0.12102 | 26.1231 | 0.030786 | 91.8371
0.9896 | -0.92675 | 1.9246 | -26.7006 | 26.7941 | 0.055202 | -0.041401 | -0.0095541 | 0.12314 | 26.7006 | 0.041401 | 92.0045
0.9914 | -0.88323 | 2.0499 | -27.345 | 27.4437 | 0.055202 | -0.053079 | -0.011677 | 0.12633 | 27.345 | 0.053079 | 92.1718
0.9932 | -0.8259 | 2.1773 | -28.0584 | 28.1624 | 0.056263 | -0.065817 | -0.012739 | 0.12845 | 28.0584 | 0.065817 | 92.3392
0.9949 | -0.7707 | 2.3217 | -28.8227 | 28.9331 | 0.056263 | -0.078556 | -0.014862 | 0.13376 | 28.8227 | 0.078556 | 92.4972
0.9967 | -0.73992 | 2.4873 | -29.6136 | 29.7346 | 0.058386 | -0.090234 | -0.016985 | 0.14013 | 29.6136 | 0.090234 | 92.6646
0.9985 | -0.73779 | 2.673 | -30.4406 | 30.5732 | 0.059448 | -0.10191 | -0.019108 | 0.14756 | 30.4406 | 0.10191 | 92.8319
1.0003 | -0.74416 | 2.8747 | -31.3036 | 31.4512 | 0.061571 | -0.11253 | -0.021231 | 0.15605 | 31.3036 | 0.11253 | 92.9993
1.0021 | -0.73355 | 3.0786 | -32.2187 | 32.3811 | 0.061571 | -0.12527 | -0.024416 | 0.16561 | 32.2187 | 0.12527 | 93.1666
1.0039 | -0.69745 | 3.2728 | -33.1837 | 33.3599 | 0.062633 | -0.14119 | -0.027601 | 0.17728 | 33.1837 | 0.14119 | 93.334
1.0057 | -0.65074 | 3.465 | -34.1762 | 34.3652 | 0.062633 | -0.15711 | -0.030786 | 0.18896 | 34.1762 | 0.15711 | 93.5013
1.0075 | -0.59236 | 3.6454 | -35.1953 | 35.397 | 0.062633 | -0.17304 | -0.035032 | 0.2017 | 35.1953 | 0.17304 | 93.6687
1.0093 | -0.52335 | 3.8174 | -36.2113 | 36.4236 | 0.06051 | -0.19108 | -0.039278 | 0.21656 | 36.2113 | 0.19108 | 93.836
1.0111 | -0.46178 | 4.0096 | -37.2219 | 37.448 | 0.059448 | -0.20913 | -0.043524 | 0.23248 | 37.2219 | 0.20913 | 94.0033
1.0129 | -0.39172 | 4.2038 | -38.3291 | 38.5679 | 0.057325 | -0.2293 | -0.047771 | 0.25053 | 38.3291 | 0.2293 | 94.1707
1.0147 | -0.32484 | 4.4066 | -39.4841 | 39.7378 | 0.056263 | -0.24841 | -0.053079 | 0.26858 | 39.4841 | 0.24841 | 94.338
1.0165 | -0.25053 | 4.6062 | -40.6847 | 40.9533 | 0.05414 | -0.26964 | -0.057325 | 0.28769 | 40.6847 | 0.26964 | 94.5054
1.0182 | -0.16348 | 4.8015 | -41.9151 | 42.1975 | 0.050955 | -0.28981 | -0.063694 | 0.30786 | 41.9151 | 0.28981 | 94.6634
1.02 | -0.084926 | 5.0011 | -43.1295 | 43.4278 | 0.047771 | -0.30998 | -0.069002 | 0.32696 | 43.1295 | 0.30998 | 94.8308
1.0218 | -0.014862 | 5.1932 | -44.3291 | 44.6412 | 0.044586 | -0.33015 | -0.07431 | 0.34713 | 44.3291 | 0.33015 | 94.9981
1.0236 | 0.069002 | 5.3715 | -45.5265 | 45.8524 | 0.04034 | -0.35138 | -0.080679 | 0.3673 | 45.5265 | 0.35138 | 95.1655
1.0254 | 0.16561 | 5.5425 | -46.7304 | 47.0679 | 0.036093 | -0.37261 | -0.085987 | 0.3896 | 46.7304 | 0.37261 | 95.3328
1.0272 | 0.27176 | 5.6985 | -47.949 | 48.2972 | 0.030786 | -0.3949 | -0.092357 | 0.41083 | 47.949 | 0.3949 | 95.5002
1.029 | 0.40552 | 5.8482 | -49.2251 | 49.5849 | 0.025478 | -0.41614 | -0.099788 | 0.43312 | 49.2251 | 0.41614 | 95.6675
1.0308 | 0.55626 | 5.9873 | -50.5372 | 50.9066 | 0.021231 | -0.43631 | -0.10616 | 0.45435 | 50.5372 | 0.43631 | 95.8349
1.0326 | 0.70913 | 6.1306 | -51.9119 | 52.2919 | 0.015924 | -0.45754 | -0.11359 | 0.47665 | 51.9119 | 0.45754 | 96.0022
1.0344 | 0.85456 | 6.3089 | -53.3885 | 53.7813 | 0.011677 | -0.48195 | -0.11996 | 0.5 | 53.3885 | 0.48195 | 96.1696
1.0362 | 1.0096 | 6.4459 | -54.7675 | 55.1688 | 0.0063694 | -0.50425 | -0.12739 | 0.52442 | 54.7675 | 0.50425 | 96.3369
1.038 | 1.1592 | 6.5552 | -56.0605 | 56.4682 | 0.0021231 | -0.52442 | -0.13482 | 0.54565 | 56.0605 | 0.52442 | 96.5043
1.0398 | 1.3142 | 6.6327 | -57.2527 | 57.6656 | -0.0010616 | -0.54459 | -0.14013 | 0.56582 | 57.2527 | 0.54459 | 96.6716
1.0416 | 1.4958 | 6.6783 | -58.3896 | 58.8047 | -0.0031847 | -0.56476 | -0.14544 | 0.58599 | 58.3896 | 0.56476 | 96.839
1.0433 | 1.6858 | 6.7091 | -59.5096 | 59.9268 | -0.0053079 | -0.5828 | -0.14968 | 0.6051 | 59.5096 | 0.5828 | 96.997
1.0451 | 1.8588 | 6.7633 | -60.6157 | 61.0372 | -0.0084926 | -0.59979 | -0.15499 | 0.62314 | 60.6157 | 0.59979 | 97.1644
1.0469 | 2.0478 | 6.7718 | -61.6624 | 62.0828 | -0.010616 | -0.61465 | -0.16136 | 0.63907 | 61.6624 | 0.61465 | 97.3317
1.0487 | 2.2155 | 6.7643 | -62.5998 | 63.018 | -0.012739 | -0.62951 | -0.16667 | 0.65393 | 62.5998 | 0.62951 | 97.4991
1.0505 | 2.38 | 6.7304 | -63.4735 | 63.8875 | -0.0138 | -0.64437 | -0.17091 | 0.66985 | 63.4735 | 0.64437 | 97.6664
1.0523 | 2.5393 | 6.6603 | -64.2569 | 64.6624 | -0.011677 | -0.65605 | -0.17516 | 0.68259 | 64.2569 | 0.65605 | 97.8338
1.0541 | 2.7113 | 6.5701 | -65.0085 | 65.4066 | -0.0084926 | -0.66667 | -0.17834 | 0.69321 | 65.0085 | 0.66667 | 98.0011
1.0559 | 2.8822 | 6.4724 | -65.724 | 66.1146 | -0.0021231 | -0.6741 | -0.18153 | 0.7017 | 65.724 | 0.6741 | 98.1685
1.0577 | 3.0318 | 6.3769 | -66.3546 | 66.7378 | 0.0042463 | -0.67941 | -0.18471 | 0.70701 | 66.3546 | 0.67941 | 98.3358
1.0595 | 3.0955 | 6.2378 | -66.5711 | 66.9437 | 0.015924 | -0.67834 | -0.18471 | 0.70701 | 66.5711 | 0.67834 | 98.5032
1.0613 | 3.0658 | 6.1083 | -66.5403 | 66.9002 | 0.027601 | -0.6741 | -0.18259 | 0.70276 | 66.5403 | 0.6741 | 98.6705
1.0631 | 2.9713 | 5.965 | -66.2548 | 66.5998 | 0.039278 | -0.66242 | -0.18047 | 0.69214 | 66.2548 | 0.66242 | 98.8379
1.0649 | 2.879 | 5.8248 | -65.9597 | 66.2898 | 0.050955 | -0.6518 | -0.17834 | 0.68259 | 65.9597 | 0.6518 | 99.0052
1.0667 | 2.793 | 5.673 | -65.6476 | 65.9628 | 0.061571 | -0.64013 | -0.17622 | 0.67197 | 65.6476 | 0.64013 | 99.1726
1.0684 | 2.6645 | 5.4947 | -65.121 | 65.4172 | 0.071125 | -0.62314 | -0.17197 | 0.65605 | 65.121 | 0.62314 | 99.3306
1.0702 | 2.5361 | 5.3163 | -64.5977 | 64.8758 | 0.080679 | -0.60616 | -0.16879 | 0.63907 | 64.5977 | 0.60616 | 99.498
1.072 | 2.4268 | 5.1391 | -64.1752 | 64.4374 | 0.090234 | -0.58917 | -0.16454 | 0.62314 | 64.1752 | 0.58917 | 99.6653
1.0738 | 2.3408 | 4.9671 | -63.9289 | 64.1752 | 0.10085 | -0.57537 | -0.1603 | 0.6104 | 63.9289 | 0.57537 | 99.8327
1.0756 | 2.258 | 4.7887 | -63.7803 | 64.0117 | 0.1104 | -0.56263 | -0.15605 | 0.59873 | 63.7803 | 0.56263 | 100
| Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My
Minimum | -8.7866 | -17.2781 | -244.0234 | 15.9363 | -0.0138 | -2.6072 | -1.1699 | 0.11146 | 15.7527 | -1.0244
Maximum | 3.0955 | 6.7718 | -15.7527 | 244.1072 | 1.7686 | 1.0244 | 0.1518 | 3.1253 | 244.0234 | 2.6072

## Diagr_BW
Loads in %BW / %BWm
Contact forces and moments, measured telemetrically in knee implants of 8 subjects.
Averages of one individual subject.
Fx  lateral, Fy  anterior, -Fz  distal, Fres = Resultant Force
Mx around x, My around y, Mz around z, Mres = Resultant Moment
z-axis = axis of implant stem!
From Bergmann et al, PlosOne (2014) "Standardized Loads Acting in Knee Implants"
Subject K5R
Subject K5R
G.Bergmann et al., PlosOne, 2014
G.Bergmann et al., PlosOne, 2014

## Dat_AVER75
Data File: | k5r_2242_0_13_i1.akf
Curve Alignment Program: | 1.5.7
------------------------MEASURING INFOS------------------------
Diagram Title #1 | Forces and Moments at Knee Joint
Diagram Title #2 | Measured Relative to Implant
Comment #1 | Knee Joint; Gaitanalysis; Level Walking;
Comment #2 | K5R, 30 Months PO *
BodyWeight [N]: | 942
Patient(en): | k5r | [7 empty] | AVER75 loads in N (forces) and Nm (moments)
Implant Type: | Knee Joint
Varus Angle Leg [°] | 1 | [7 empty] | Subject | k5r
Axes Directions X, Y, Z: | Lateral, Anterior, Superior
Axes Origin: | Tibial Plateau at Stem Axis | [7 empty] | The individual peak value of Fres was adapted to the average peak of Fres from all subjects!
Transformation X Y Z (°): | 7 | -2 | -1 | [5 empty] | The AVER75 forces Fres of all subjects therefore have the same peak values.
Displacement X, Y, Z (mm): | 0 | 0 | 0 | [5 empty] | All 6 load components were then adapted by the same Multiplication Factor.
Activity Code: | KneeActivities | 1P_1_15_2
[9 empty] | Multiplication Factor = |  | Max(Fres, AVER75 of 8 Subjects) / |  |  | Max(Fres, Individual Subject)
[9 empty] | Max(Fres, AVER75 of 8 Subjects) = |  |  | 1960.71 | N
[9 empty] | Max(Fres, Individual Subject) = |  |  | 2299.49 | N
---------------------------DATA--------------------------------
Number of Data Sets = | 601
Number of analog channels = | 0
Time-Offset(sec) = | 0
Time | Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My | Time
[s] | [N] | [N] | [N] | [N] | [Nm] | [Nm] | [Nm] | [Nm] | [N] | (Nm] | [%Cycle]
0 | 9.908 | 24.8639 | -514.3827 | 515.2013 | 1.603 | -3.2743 | -0.89531 | 3.8797 | 514.3827 | 3.2743 | 0
0.0018 | 8.9872 | 23.1756 | -513.1463 | 513.837 | 1.6798 | -3.1464 | -0.85267 | 3.7859 | 513.1463 | 3.1464 | 0.16735
0.0036 | 7.981 | 21.6579 | -512.6774 | 513.2742 | 1.7565 | -3.0014 | -0.81004 | 3.675 | 512.6774 | 3.0014 | 0.3347
0.0054 | 7.1795 | 20.2595 | -513.2998 | 513.82 | 1.8247 | -2.882 | -0.7674 | 3.6068 | 513.2998 | 2.882 | 0.50205
0.0072 | 6.4633 | 18.9378 | -515.3036 | 515.764 | 1.9015 | -2.7797 | -0.72477 | 3.5642 | 515.3036 | 2.7797 | 0.66939
0.009 | 5.7811 | 17.6929 | -519.567 | 519.9848 | 1.9782 | -2.6859 | -0.68214 | 3.5556 | 519.567 | 2.6859 | 0.83674
0.0108 | 5.3633 | 16.3542 | -527.0705 | 527.4371 | 2.0805 | -2.6262 | -0.6395 | 3.5812 | 527.0705 | 2.6262 | 1.0041
0.0125 | 4.9625 | 14.9218 | -536.6716 | 536.987 | 2.1828 | -2.5751 | -0.59687 | 3.5983 | 536.6716 | 2.5751 | 1.1621
0.0143 | 4.6044 | 13.6257 | -547.9098 | 548.1911 | 2.2937 | -2.541 | -0.55424 | 3.6409 | 547.9098 | 2.541 | 1.3295
0.0161 | 4.3572 | 12.2103 | -562.84 | 563.0873 | 2.4216 | -2.541 | -0.49455 | 3.7262 | 562.84 | 2.541 | 1.4968
0.0179 | 4.1696 | 10.7437 | -580.0299 | 580.2516 | 2.5495 | -2.541 | -0.45192 | 3.82 | 580.0299 | 2.541 | 1.6642
0.0197 | 4.0587 | 9.2259 | -598.9166 | 599.1127 | 2.6859 | -2.5495 | -0.40928 | 3.9308 | 598.9166 | 2.5495 | 1.8315
0.0215 | 4.0928 | 7.6996 | -619.483 | 619.6536 | 2.8309 | -2.5836 | -0.35812 | 4.0587 | 619.483 | 2.5836 | 1.9989
0.0233 | 4.1696 | 6.3695 | -640.5184 | 640.6804 | 3.0014 | -2.6262 | -0.31549 | 4.2037 | 640.5184 | 2.6262 | 2.1662
0.0251 | 4.2122 | 5.1842 | -660.9058 | 661.0508 | 3.1719 | -2.6774 | -0.26433 | 4.3657 | 660.9058 | 2.6774 | 2.3336
0.0269 | 4.2889 | 3.9905 | -679.4088 | 679.5452 | 3.3339 | -2.7285 | -0.22169 | 4.5277 | 679.4088 | 2.7285 | 2.5009
0.0287 | 4.4851 | 2.4898 | -696.2746 | 696.4111 | 3.4874 | -2.7797 | -0.18759 | 4.6812 | 696.2746 | 2.7797 | 2.6683
0.0305 | 4.7323 | 0.48602 | -710.4716 | 710.6166 | 3.6324 | -2.8309 | -0.16201 | 4.8091 | 710.4716 | 2.8309 | 2.8356
0.0323 | 4.8943 | -2.0123 | -722.878 | 723.04 | 3.7688 | -2.8565 | -0.13643 | 4.9284 | 722.878 | 2.8565 | 3.003
0.0341 | 5.0563 | -4.5874 | -734.3805 | 734.5767 | 3.9052 | -2.8735 | -0.10232 | 5.0308 | 734.3805 | 2.8735 | 3.1703
0.0359 | 5.278 | -7.1624 | -745.3459 | 745.5676 | 4.0331 | -2.8735 | -0.07674 | 5.1246 | 745.3459 | 2.8735 | 3.3377
0.0376 | 5.8664 | -9.4305 | -758.6987 | 758.946 | 4.1951 | -2.9076 | -0.05116 | 5.261 | 758.6987 | 2.9076 | 3.4957
0.0394 | 6.5059 | -11.6901 | -772.5461 | 772.819 | 4.3572 | -2.9417 | -0.034107 | 5.4145 | 772.5461 | 2.9417 | 3.6631
0.0412 | 7.3244 | -14.0435 | -788.3547 | 788.6872 | 4.5533 | -3.027 | -0.0085267 | 5.6106 | 788.3547 | 3.027 | 3.8304
0.043 | 8.2624 | -16.431 | -805.05 | 805.4422 | 4.7579 | -3.1123 | 0 | 5.8237 | 805.05 | 3.1123 | 3.9978
0.0448 | 9.1492 | -18.7332 | -822.2739 | 822.7258 | 4.9711 | -3.2231 | 0.017053 | 6.0625 | 822.2739 | 3.2231 | 4.1651
0.0466 | 9.8569 | -21.0695 | -839.4382 | 839.9498 | 5.2013 | -3.3425 | 0.034107 | 6.3098 | 839.4382 | 3.3425 | 4.3325
0.0484 | 10.4708 | -23.4826 | -856.3893 | 856.9691 | 5.4315 | -3.496 | 0.059687 | 6.5826 | 856.3893 | 3.496 | 4.4998
0.0502 | 11.1274 | -26.0747 | -874.1334 | 874.807 | 5.6703 | -3.6665 | 0.085267 | 6.8725 | 874.1334 | 3.6665 | 4.6672
0.052 | 11.6304 | -28.6157 | -891.1613 | 891.9202 | 5.9005 | -3.8285 | 0.11085 | 7.1539 | 891.1613 | 3.8285 | 4.8345
0.0538 | 12.0738 | -31.1566 | -907.183 | 908.0271 | 6.1222 | -4.0246 | 0.14495 | 7.4353 | 907.183 | 4.0246 | 5.0019
0.0556 | 12.2955 | -33.6464 | -921.6613 | 922.5908 | 6.3183 | -4.2463 | 0.17053 | 7.7082 | 921.6613 | 4.2463 | 5.1692
0.0574 | 12.4064 | -36.1618 | -935.5002 | 936.5064 | 6.4974 | -4.4765 | 0.18759 | 7.981 | 935.5002 | 4.4765 | 5.3366
0.0592 | 12.449 | -38.6772 | -948.8445 | 949.9189 | 6.6764 | -4.7067 | 0.21317 | 8.2539 | 948.8445 | 4.7067 | 5.5039
0.061 | 12.5684 | -41.3375 | -962.6834 | 963.843 | 6.8555 | -4.9711 | 0.23022 | 8.5523 | 962.6834 | 4.9711 | 5.6713
0.0627 | 12.8157 | -44.1769 | -978.057 | 979.3275 | 7.0516 | -5.2695 | 0.2558 | 8.8848 | 978.057 | 5.2695 | 5.8293
0.0645 | 13.1397 | -47.1186 | -995.102 | 996.4918 | 7.2648 | -5.5765 | 0.28138 | 9.2344 | 995.102 | 5.5765 | 5.9967
0.0663 | 13.566 | -50.2394 | -1012.9228 | 1014.4491 | 7.4779 | -5.909 | 0.30696 | 9.6096 | 1012.9228 | 5.909 | 6.164
0.0681 | 14.0776 | -53.5393 | -1031.3917 | 1033.0714 | 7.6911 | -6.2501 | 0.34107 | 10.0018 | 1031.3917 | 6.2501 | 6.3313
0.0699 | 14.521 | -56.7794 | -1050.6535 | 1052.4953 | 7.9213 | -6.6082 | 0.36665 | 10.4026 | 1050.6535 | 6.6082 | 6.4987
0.0717 | 14.99 | -60.0963 | -1070.4696 | 1072.4819 | 8.143 | -6.9919 | 0.39223 | 10.8204 | 1070.4696 | 6.9919 | 6.666
0.0735 | 15.4334 | -63.3706 | -1090.8229 | 1092.9972 | 8.3647 | -7.3841 | 0.40076 | 11.2382 | 1090.8229 | 7.3841 | 6.8334
0.0753 | 15.9279 | -66.6278 | -1111.5257 | 1113.8621 | 8.5864 | -7.7764 | 0.42634 | 11.6731 | 1111.5257 | 7.7764 | 7.0007
0.0771 | 16.5248 | -70.0385 | -1132.4674 | 1134.9913 | 8.8166 | -8.1686 | 0.44339 | 12.0994 | 1132.4674 | 8.1686 | 7.1681
0.0789 | 16.8658 | -73.2616 | -1154.0826 | 1156.7685 | 9.0554 | -8.5608 | 0.46044 | 12.5343 | 1154.0826 | 8.5608 | 7.3354
0.0807 | 17.1984 | -76.4164 | -1176.6869 | 1179.5434 | 9.3026 | -8.9786 | 0.46897 | 13.0032 | 1176.6869 | 8.9786 | 7.5028
0.0825 | 17.3263 | -79.5287 | -1200.161 | 1203.1709 | 9.5499 | -9.422 | 0.48602 | 13.4893 | 1200.161 | 9.422 | 7.6701
0.0843 | 17.4968 | -82.7006 | -1224.0614 | 1227.2248 | 9.7972 | -9.891 | 0.48602 | 14.0009 | 1224.0614 | 9.891 | 7.8375
0.086 | 17.71 | -85.8726 | -1248.4733 | 1251.7817 | 10.0615 | -10.36 | 0.49455 | 14.521 | 1248.4733 | 10.36 | 7.9955
0.0878 | 17.8976 | -89.036 | -1271.7172 | 1275.1961 | 10.3088 | -10.8119 | 0.50308 | 15.0155 | 1271.7172 | 10.8119 | 8.1629
0.0896 | 18.1619 | -91.9692 | -1294.1595 | 1297.7833 | 10.5475 | -11.2808 | 0.50308 | 15.5186 | 1294.1595 | 11.2808 | 8.3302
0.0914 | 18.2898 | -94.8938 | -1315.6468 | 1319.4242 | 10.7692 | -11.7157 | 0.50308 | 15.9876 | 1315.6468 | 11.7157 | 8.4976
0.0932 | 18.4262 | -97.8185 | -1336.6993 | 1340.6301 | 10.9824 | -12.1676 | 0.50308 | 16.4651 | 1336.6993 | 12.1676 | 8.6649
0.095 | 18.52 | -100.7773 | -1357.4363 | 1361.5291 | 11.1956 | -12.611 | 0.5116 | 16.9341 | 1357.4363 | 12.611 | 8.8323
0.0968 | 18.5456 | -103.719 | -1377.5423 | 1381.8056 | 11.3917 | -13.0373 | 0.5116 | 17.386 | 1377.5423 | 13.0373 | 8.9996
0.0986 | 18.5115 | -106.6351 | -1397.2134 | 1401.6388 | 11.5878 | -13.4637 | 0.5116 | 17.8379 | 1397.2134 | 13.4637 | 9.167
0.1004 | 18.4689 | -109.4745 | -1416.4667 | 1421.0541 | 11.7754 | -13.873 | 0.5116 | 18.2728 | 1416.4667 | 13.873 | 9.3343
0.1022 | 18.4092 | -112.2372 | -1435.4898 | 1440.2392 | 11.9545 | -14.2823 | 0.52013 | 18.6991 | 1435.4898 | 14.2823 | 9.5017
0.104 | 18.3495 | -114.8975 | -1454.2571 | 1459.16 | 12.125 | -14.6745 | 0.52866 | 19.1169 | 1454.2571 | 14.6745 | 9.669
0.1058 | 18.3154 | -117.3276 | -1472.4531 | 1477.4924 | 12.2955 | -15.0326 | 0.54571 | 19.5006 | 1472.4531 | 15.0326 | 9.8364
0.1076 | 18.2642 | -119.4593 | -1489.8732 | 1495.0319 | 12.449 | -15.3651 | 0.55424 | 19.8587 | 1489.8732 | 15.3651 | 10.0037
0.1094 | 18.1875 | -121.4887 | -1507.0034 | 1512.2559 | 12.6025 | -15.6721 | 0.57129 | 20.1998 | 1507.0034 | 15.6721 | 10.1711
0.1111 | 18.1022 | -123.4242 | -1523.801 | 1529.1643 | 12.7645 | -15.9705 | 0.58834 | 20.5238 | 1523.801 | 15.9705 | 10.3291
0.1129 | 17.9317 | -125.2916 | -1540.0871 | 1545.5527 | 12.918 | -16.2775 | 0.6054 | 20.8649 | 1540.0871 | 16.2775 | 10.4965
0.1147 | 17.71 | -126.9713 | -1555.6824 | 1561.2248 | 13.0629 | -16.5759 | 0.62245 | 21.1889 | 1555.6824 | 16.5759 | 10.6638
0.1165 | 17.4542 | -128.3441 | -1570.1949 | 1575.797 | 13.2079 | -16.8573 | 0.64803 | 21.4959 | 1570.1949 | 16.8573 | 10.8312
0.1183 | 17.1728 | -129.4697 | -1583.9314 | 1589.5761 | 13.3358 | -17.1387 | 0.67361 | 21.8028 | 1583.9314 | 17.1387 | 10.9985
0.1201 | 16.9085 | -130.5355 | -1597.199 | 1602.8863 | 13.4637 | -17.4116 | 0.69066 | 22.1012 | 1597.199 | 17.4116 | 11.1659
0.1219 | 16.6527 | -131.5928 | -1610.3302 | 1616.0601 | 13.5831 | -17.6844 | 0.71624 | 22.3912 | 1610.3302 | 17.6844 | 11.3332
0.1237 | 16.3883 | -132.5222 | -1622.7365 | 1628.5006 | 13.7024 | -17.9573 | 0.74182 | 22.6725 | 1622.7365 | 17.9573 | 11.5006
0.1255 | 16.1496 | -133.3152 | -1634.4267 | 1640.2078 | 13.8048 | -18.2216 | 0.7674 | 22.9454 | 1634.4267 | 18.2216 | 11.6679
0.1273 | 15.8767 | -134.057 | -1645.2556 | 1651.0623 | 13.8985 | -18.4944 | 0.79298 | 23.2268 | 1645.2556 | 18.4944 | 11.8353
0.1291 | 15.6465 | -134.6539 | -1654.5838 | 1660.4076 | 13.9668 | -18.7503 | 0.81856 | 23.4741 | 1654.5838 | 18.7503 | 12.0026
0.1309 | 15.3737 | -135.2678 | -1662.7183 | 1668.5591 | 14.0264 | -18.9805 | 0.84414 | 23.6957 | 1662.7183 | 18.9805 | 12.17
0.1327 | 14.9985 | -136.0097 | -1669.8722 | 1675.7557 | 14.0776 | -19.2022 | 0.86973 | 23.9089 | 1669.8722 | 19.2022 | 12.3373
0.1345 | 14.5892 | -136.7259 | -1676.3355 | 1682.2445 | 14.1288 | -19.4068 | 0.88678 | 24.105 | 1676.3355 | 19.4068 | 12.5046
0.1362 | 14.1799 | -137.2972 | -1682.1507 | 1688.0938 | 14.1714 | -19.5773 | 0.90383 | 24.267 | 1682.1507 | 19.5773 | 12.6627
0.138 | 13.8048 | -137.6042 | -1687.5225 | 1693.4656 | 14.197 | -19.7138 | 0.92941 | 24.4035 | 1687.5225 | 19.7138 | 12.83
0.1398 | 13.3443 | -137.9964 | -1691.0526 | 1697.0128 | 14.2055 | -19.8673 | 0.94647 | 24.5314 | 1691.0526 | 19.8673 | 12.9974
0.1416 | 12.8157 | -138.2522 | -1693.6532 | 1699.6219 | 14.197 | -20.0207 | 0.97205 | 24.6507 | 1693.6532 | 20.0207 | 13.1647
0.1434 | 12.4234 | -138.3971 | -1696.2795 | 1702.2482 | 14.1543 | -20.1742 | 0.99763 | 24.7616 | 1696.2795 | 20.1742 | 13.3321
0.1452 | 11.9886 | -138.6018 | -1698.7948 | 1704.7635 | 14.1288 | -20.3192 | 1.0232 | 24.8724 | 1698.7948 | 20.3192 | 13.4994
0.147 | 11.5366 | -138.7808 | -1700.9521 | 1706.9208 | 14.1032 | -20.4471 | 1.0488 | 24.9662 | 1700.9521 | 20.4471 | 13.6668
0.1488 | 10.9824 | -138.687 | -1702.3675 | 1708.3192 | 14.0691 | -20.5068 | 1.0658 | 24.9918 | 1702.3675 | 20.5068 | 13.8341
0.1506 | 10.3344 | -138.4568 | -1702.1203 | 1708.0634 | 14.0009 | -20.592 | 1.0914 | 25.0344 | 1702.1203 | 20.592 | 14.0015
0.1524 | 9.8398 | -138.0817 | -1701.6513 | 1707.5688 | 13.9071 | -20.6858 | 1.0999 | 25.0685 | 1701.6513 | 20.6858 | 14.1688
0.1542 | 9.2174 | -137.5615 | -1701.0203 | 1706.9037 | 13.8218 | -20.754 | 1.117 | 25.0856 | 1701.0203 | 20.754 | 14.3362
0.156 | 8.4926 | -136.9561 | -1700.1506 | 1706.0084 | 13.7451 | -20.8137 | 1.1255 | 25.0941 | 1700.1506 | 20.8137 | 14.5035
0.1578 | 7.8531 | -136.5554 | -1698.684 | 1704.5163 | 13.6598 | -20.8734 | 1.1426 | 25.1027 | 1698.684 | 20.8734 | 14.6709
0.1595 | 7.4012 | -136.129 | -1696.7484 | 1702.5551 | 13.549 | -20.9246 | 1.1511 | 25.0856 | 1696.7484 | 20.9246 | 14.8289
0.1613 | 7.0345 | -135.5663 | -1694.4036 | 1700.1847 | 13.4296 | -20.9416 | 1.1596 | 25.0515 | 1694.4036 | 20.9416 | 14.9963
0.1631 | 6.4974 | -134.7818 | -1691.6495 | 1697.3879 | 13.3102 | -20.9416 | 1.1767 | 24.9833 | 1691.6495 | 20.9416 | 15.1636
0.1649 | 5.9687 | -133.9291 | -1688.7845 | 1694.4718 | 13.1823 | -20.9416 | 1.1852 | 24.9151 | 1688.7845 | 20.9416 | 15.331
0.1667 | 5.3974 | -133.102 | -1685.6637 | 1691.3084 | 13.0715 | -20.9416 | 1.1937 | 24.8554 | 1685.6637 | 20.9416 | 15.4983
0.1685 | 4.7409 | -132.13 | -1681.5368 | 1687.1218 | 12.935 | -20.9246 | 1.2108 | 24.7786 | 1681.5368 | 20.9246 | 15.6657
0.1703 | 3.9052 | -130.6719 | -1676.5572 | 1682.0569 | 12.773 | -20.8649 | 1.2193 | 24.6422 | 1676.5572 | 20.8649 | 15.833
0.1721 | 3.2316 | -129.1115 | -1671.5093 | 1676.9068 | 12.5854 | -20.7967 | 1.2193 | 24.4887 | 1671.5093 | 20.7967 | 16.0004
0.1739 | 2.9417 | -127.7302 | -1666.2739 | 1671.5861 | 12.3637 | -20.7455 | 1.2108 | 24.3352 | 1666.2739 | 20.7455 | 16.1677
0.1757 | 2.5836 | -126.4683 | -1660.058 | 1665.2934 | 12.1506 | -20.6858 | 1.2108 | 24.1818 | 1660.058 | 20.6858 | 16.3351
0.1775 | 2.4045 | -124.9335 | -1653.3815 | 1658.5232 | 11.9289 | -20.6347 | 1.1937 | 24.0283 | 1653.3815 | 20.6347 | 16.5024
0.1793 | 2.4131 | -123.3304 | -1646.6284 | 1651.6762 | 11.7072 | -20.5835 | 1.1767 | 23.8919 | 1646.6284 | 20.5835 | 16.6698
0.1811 | 2.4898 | -121.8212 | -1639.8496 | 1644.8122 | 11.494 | -20.5409 | 1.1682 | 23.7554 | 1639.8496 | 20.5409 | 16.8371
0.1829 | 2.4813 | -120.4825 | -1632.4485 | 1637.3257 | 11.2808 | -20.4471 | 1.1596 | 23.5678 | 1632.4485 | 20.4471 | 17.0045
0.1846 | 2.2852 | -119.2461 | -1624.6039 | 1629.4215 | 11.0506 | -20.3874 | 1.1596 | 23.3973 | 1624.6039 | 20.3874 | 17.1625
0.1864 | 2.0976 | -117.6857 | -1615.4632 | 1620.187 | 10.7948 | -20.3618 | 1.1511 | 23.2524 | 1615.4632 | 20.3618 | 17.3299
0.1882 | 2.1317 | -116.1254 | -1606.5443 | 1611.1743 | 10.5561 | -20.3106 | 1.1426 | 23.0818 | 1606.5443 | 20.3106 | 17.4972
0.19 | 2.1573 | -114.7696 | -1599.2881 | 1603.8413 | 10.3429 | -20.2851 | 1.1341 | 22.9624 | 1599.2881 | 20.2851 | 17.6646
0.1918 | 2.2169 | -113.0216 | -1592.0574 | 1596.5083 | 10.1212 | -20.2595 | 1.117 | 22.826 | 1592.0574 | 20.2595 | 17.8319
0.1936 | 2.396 | -111.3248 | -1585.5174 | 1589.866 | 9.891 | -20.251 | 1.0999 | 22.7152 | 1585.5174 | 20.251 | 17.9993
0.1954 | 2.6518 | -109.6365 | -1579.327 | 1583.5733 | 9.6608 | -20.2424 | 1.0829 | 22.6043 | 1579.327 | 20.2424 | 18.1666
0.1972 | 3.0014 | -108.0591 | -1573.2219 | 1577.3829 | 9.4561 | -20.251 | 1.0573 | 22.5191 | 1573.2219 | 20.251 | 18.334
0.199 | 3.3339 | -106.5925 | -1566.7586 | 1570.8429 | 9.26 | -20.2595 | 1.0317 | 22.4508 | 1566.7586 | 20.2595 | 18.5013
0.2008 | 3.658 | -105.0747 | -1560.4148 | 1564.4223 | 9.0724 | -20.268 | 1.0147 | 22.3826 | 1560.4148 | 20.268 | 18.6687
0.2026 | 3.9564 | -103.3268 | -1554.7104 | 1558.6327 | 8.8934 | -20.2765 | 0.9891 | 22.3144 | 1554.7104 | 20.2765 | 18.836
0.2044 | 4.2719 | -101.647 | -1549.5688 | 1553.3887 | 8.6973 | -20.268 | 0.98057 | 22.2206 | 1549.5688 | 20.268 | 19.0033
0.2062 | 4.7664 | -100.3253 | -1545.8426 | 1549.5943 | 8.5182 | -20.2424 | 0.95499 | 22.1268 | 1545.8426 | 20.2424 | 19.1707
0.208 | 5.4059 | -99.2425 | -1543.1567 | 1546.8573 | 8.3732 | -20.1998 | 0.93794 | 22.033 | 1543.1567 | 20.1998 | 19.338
0.2097 | 6.0966 | -97.9293 | -1541.2979 | 1544.9302 | 8.2198 | -20.1742 | 0.92089 | 21.9478 | 1541.2979 | 20.1742 | 19.4961
0.2115 | 6.7873 | -96.5054 | -1539.3964 | 1542.952 | 8.0577 | -20.1572 | 0.89531 | 21.8625 | 1539.3964 | 20.1572 | 19.6634
0.2133 | 7.5547 | -95.0729 | -1535.5253 | 1539.0042 | 7.9043 | -20.1998 | 0.87825 | 21.854 | 1535.5253 | 20.1998 | 19.8308
0.2151 | 8.1942 | -93.683 | -1530.8953 | 1534.323 | 7.7508 | -20.1913 | 0.85267 | 21.7943 | 1530.8953 | 20.1913 | 19.9981
0.2169 | 8.8081 | -92.0544 | -1528.3031 | 1531.6627 | 7.6399 | -20.2339 | 0.82709 | 21.8028 | 1528.3031 | 20.2339 | 20.1655
0.2187 | 9.405 | -90.5196 | -1525.9071 | 1529.224 | 7.5461 | -20.268 | 0.80151 | 21.8113 | 1525.9071 | 20.268 | 20.3328
0.2205 | 10.0189 | -89.1127 | -1524.2615 | 1527.5187 | 7.4694 | -20.3021 | 0.77593 | 21.8113 | 1524.2615 | 20.3021 | 20.5002
0.2223 | 10.5305 | -87.6888 | -1522.5561 | 1525.7451 | 7.4012 | -20.3362 | 0.75888 | 21.8199 | 1522.5561 | 20.3362 | 20.6675
0.2241 | 10.7692 | -86.2818 | -1520.2113 | 1523.3491 | 7.333 | -20.3789 | 0.7333 | 21.8369 | 1520.2113 | 20.3789 | 20.8349
0.2259 | 10.9142 | -84.9346 | -1516.9285 | 1519.9981 | 7.2477 | -20.4215 | 0.70772 | 21.8454 | 1516.9285 | 20.4215 | 21.0022
0.2277 | 11.1359 | -83.775 | -1513.3814 | 1516.3913 | 7.1539 | -20.4726 | 0.67361 | 21.8625 | 1513.3814 | 20.4726 | 21.1696
0.2295 | 11.2979 | -82.4789 | -1509.6637 | 1512.614 | 7.0601 | -20.5068 | 0.6395 | 21.871 | 1509.6637 | 20.5068 | 21.3369
0.2313 | 11.4855 | -81.0123 | -1504.974 | 1507.839 | 6.9407 | -20.5323 | 0.6054 | 21.8454 | 1504.974 | 20.5323 | 21.5043
0.2331 | 11.7754 | -79.3155 | -1498.8519 | 1501.6145 | 6.7873 | -20.5664 | 0.57982 | 21.8113 | 1498.8519 | 20.5664 | 21.6716
0.2348 | 11.9886 | -77.4908 | -1492.0646 | 1494.7249 | 6.6338 | -20.6005 | 0.54571 | 21.7858 | 1492.0646 | 20.6005 | 21.8297
0.2366 | 12.0227 | -75.6917 | -1484.8595 | 1487.409 | 6.4718 | -20.6176 | 0.52866 | 21.7431 | 1484.8595 | 20.6176 | 21.997
0.2384 | 12.1079 | -74.046 | -1478.8908 | 1481.3465 | 6.3439 | -20.6261 | 0.49455 | 21.709 | 1478.8908 | 20.6261 | 22.1644
0.2402 | 12.125 | -72.2639 | -1474.0476 | 1476.4266 | 6.2416 | -20.6176 | 0.46044 | 21.6664 | 1474.0476 | 20.6176 | 22.3317
0.242 | 12.125 | -70.5586 | -1469.9974 | 1472.3167 | 6.1563 | -20.592 | 0.43486 | 21.6238 | 1469.9974 | 20.592 | 22.4991
0.2438 | 12.2017 | -69.1005 | -1466.3992 | 1468.6673 | 6.0881 | -20.5835 | 0.39223 | 21.5982 | 1466.3992 | 20.5835 | 22.6664
0.2456 | 12.4916 | -68.0773 | -1462.4428 | 1464.6683 | 6.0028 | -20.5835 | 0.3496 | 21.5726 | 1462.4428 | 20.5835 | 22.8338
0.2474 | 12.5599 | -66.7045 | -1458.0345 | 1460.1917 | 5.9175 | -20.5579 | 0.31549 | 21.5214 | 1458.0345 | 20.5579 | 23.0011
0.2492 | 12.6451 | -64.9565 | -1452.3216 | 1454.385 | 5.8067 | -20.5494 | 0.28138 | 21.4703 | 1452.3216 | 20.5494 | 23.1685
0.251 | 12.6451 | -63.1233 | -1445.8413 | 1447.8109 | 5.6873 | -20.5409 | 0.2558 | 21.4276 | 1445.8413 | 20.5409 | 23.3358
0.2528 | 12.5343 | -61.3071 | -1439.2416 | 1441.126 | 5.5679 | -20.5238 | 0.22169 | 21.3765 | 1439.2416 | 20.5238 | 23.5032
0.2546 | 12.4746 | -59.5165 | -1432.6334 | 1434.441 | 5.4486 | -20.5153 | 0.18759 | 21.3253 | 1432.6334 | 20.5153 | 23.6705
0.2564 | 12.5172 | -57.4274 | -1425.7353 | 1427.4406 | 5.3292 | -20.5323 | 0.16201 | 21.3083 | 1425.7353 | 20.5323 | 23.8379
0.2581 | 12.6451 | -55.3725 | -1419.0418 | 1420.6448 | 5.2098 | -20.5409 | 0.1279 | 21.2827 | 1419.0418 | 20.5409 | 23.9959
0.2599 | 12.8839 | -53.7268 | -1413.0049 | 1414.5312 | 5.1075 | -20.5323 | 0.093794 | 21.2401 | 1413.0049 | 20.5323 | 24.1633
0.2617 | 13.097 | -52.1238 | -1406.9935 | 1408.4431 | 4.9881 | -20.5323 | 0.059687 | 21.2059 | 1406.9935 | 20.5323 | 24.3306
0.2635 | 13.1908 | -50.3673 | -1401.545 | 1402.9092 | 4.8773 | -20.5323 | 0.02558 | 21.1718 | 1401.545 | 20.5323 | 24.498
0.2653 | 13.3784 | -48.6108 | -1396.9064 | 1398.2025 | 4.7835 | -20.5068 | -0.017053 | 21.1207 | 1396.9064 | 20.5068 | 24.6653
0.2671 | 13.5319 | -46.7861 | -1392.2764 | 1393.4872 | 4.6982 | -20.4726 | -0.05116 | 21.0695 | 1392.2764 | 20.4726 | 24.8327
0.2689 | 13.6001 | -44.9699 | -1387.5867 | 1388.7293 | 4.613 | -20.413 | -0.093794 | 20.9928 | 1387.5867 | 20.413 | 25
0.2707 | 13.873 | -43.418 | -1383.4854 | 1384.5683 | 4.5277 | -20.3703 | -0.13643 | 20.9331 | 1383.4854 | 20.3703 | 25.1673
0.2725 | 14.1714 | -41.9856 | -1379.8274 | 1380.8677 | 4.4595 | -20.3106 | -0.17053 | 20.8563 | 1379.8274 | 20.3106 | 25.3347
0.2743 | 14.4187 | -40.5701 | -1376.3229 | 1377.312 | 4.3913 | -20.2595 | -0.21317 | 20.7881 | 1376.3229 | 20.2595 | 25.502
0.2761 | 14.5295 | -39.1376 | -1373.3556 | 1374.3021 | 4.3316 | -20.1913 | -0.24727 | 20.7199 | 1373.3556 | 20.1913 | 25.6694
0.2779 | 14.6148 | -37.7392 | -1370.3457 | 1371.2495 | 4.2634 | -20.1401 | -0.28991 | 20.6517 | 1370.3457 | 20.1401 | 25.8367
0.2797 | 14.7171 | -36.3494 | -1367.242 | 1368.1032 | 4.1866 | -20.0889 | -0.34107 | 20.5835 | 1367.242 | 20.0889 | 26.0041
0.2815 | 14.7256 | -34.9851 | -1363.6693 | 1364.4708 | 4.1099 | -20.0207 | -0.3837 | 20.5068 | 1363.6693 | 20.0207 | 26.1714
0.2832 | 14.7939 | -33.7658 | -1360.4291 | 1361.2051 | 4.0417 | -19.9525 | -0.42634 | 20.4215 | 1360.4291 | 19.9525 | 26.3295
0.285 | 14.9132 | -32.4697 | -1357.5215 | 1358.2634 | 3.9735 | -19.9014 | -0.4775 | 20.3533 | 1357.5215 | 19.9014 | 26.4968
0.2868 | 15.1946 | -31.1737 | -1355.1085 | 1355.8162 | 3.9138 | -19.8417 | -0.52013 | 20.2936 | 1355.1085 | 19.8417 | 26.6642
0.2886 | 15.5016 | -29.9288 | -1352.9768 | 1353.6589 | 3.8541 | -19.7905 | -0.56276 | 20.2254 | 1352.9768 | 19.7905 | 26.8315
0.2904 | 15.7147 | -28.7009 | -1350.811 | 1351.4761 | 3.8029 | -19.7308 | -0.61392 | 20.1657 | 1350.811 | 19.7308 | 26.9989
0.2922 | 15.9109 | -27.4475 | -1348.4917 | 1349.1398 | 3.7432 | -19.6711 | -0.65656 | 20.0975 | 1348.4917 | 19.6711 | 27.1662
0.294 | 16.0046 | -26.2964 | -1346.1298 | 1346.7523 | 3.7006 | -19.6029 | -0.70772 | 20.0207 | 1346.1298 | 19.6029 | 27.3336
0.2958 | 16.1667 | -25.1027 | -1344.1261 | 1344.7229 | 3.658 | -19.5347 | -0.75035 | 19.9525 | 1344.1261 | 19.5347 | 27.5009
0.2976 | 16.3457 | -23.8748 | -1342.3866 | 1342.9494 | 3.6068 | -19.4665 | -0.79298 | 19.8758 | 1342.3866 | 19.4665 | 27.6683
0.2994 | 16.4566 | -22.5787 | -1341.082 | 1341.6277 | 3.5642 | -19.4153 | -0.84414 | 19.8161 | 1341.082 | 19.4153 | 27.8356
0.3012 | 16.5248 | -21.3168 | -1339.9309 | 1340.4425 | 3.5215 | -19.3556 | -0.88678 | 19.7479 | 1339.9309 | 19.3556 | 28.003
0.303 | 16.6186 | -20.1145 | -1338.8992 | 1339.3852 | 3.4874 | -19.2789 | -0.92941 | 19.6711 | 1338.8992 | 19.2789 | 28.1703
0.3048 | 16.6868 | -18.9378 | -1337.9186 | 1338.3876 | 3.4533 | -19.1936 | -0.98057 | 19.5773 | 1337.9186 | 19.1936 | 28.3377
0.3066 | 16.7124 | -17.7612 | -1336.9977 | 1337.4496 | 3.4192 | -19.0743 | -1.0232 | 19.4665 | 1336.9977 | 19.0743 | 28.505
0.3083 | 16.7379 | -16.6783 | -1336.1109 | 1336.5373 | 3.4022 | -18.9805 | -1.0658 | 19.3642 | 1336.1109 | 18.9805 | 28.6631
0.3101 | 16.7038 | -15.6551 | -1335.4118 | 1335.8296 | 3.3681 | -18.8867 | -1.117 | 19.2789 | 1335.4118 | 18.8867 | 28.8304
0.3119 | 16.6186 | -14.6745 | -1334.7467 | 1335.1559 | 3.3425 | -18.8099 | -1.1682 | 19.1936 | 1334.7467 | 18.8099 | 28.9978
0.3137 | 16.5077 | -13.6513 | -1334.1242 | 1334.5164 | 3.3254 | -18.7417 | -1.2193 | 19.1254 | 1334.1242 | 18.7417 | 29.1651
0.3155 | 16.4907 | -12.4149 | -1333.6894 | 1334.056 | 3.2998 | -18.6735 | -1.2705 | 19.0572 | 1333.6894 | 18.6735 | 29.3325
0.3173 | 16.4821 | -11.1359 | -1333.2886 | 1333.6297 | 3.2743 | -18.5968 | -1.3216 | 18.989 | 1333.2886 | 18.5968 | 29.4998
0.3191 | 16.4139 | -9.9592 | -1333.2374 | 1333.5615 | 3.2572 | -18.5371 | -1.3728 | 18.9208 | 1333.2374 | 18.5371 | 29.6672
0.3209 | 16.3116 | -8.8166 | -1333.2033 | 1333.5103 | 3.2402 | -18.4689 | -1.424 | 18.8611 | 1333.2033 | 18.4689 | 29.8345
0.3227 | 16.2263 | -7.5803 | -1333.3739 | 1333.6723 | 3.2231 | -18.4092 | -1.4751 | 18.8014 | 1333.3739 | 18.4092 | 30.0019
0.3245 | 16.2008 | -6.3098 | -1333.6297 | 1333.894 | 3.2146 | -18.3495 | -1.5348 | 18.7417 | 1333.6297 | 18.3495 | 30.1692
0.3263 | 16.1411 | -5.0308 | -1333.8769 | 1334.1242 | 3.1975 | -18.2898 | -1.603 | 18.682 | 1333.8769 | 18.2898 | 30.3366
0.3281 | 16.0302 | -3.8455 | -1334.3971 | 1334.6358 | 3.1805 | -18.2386 | -1.6627 | 18.6394 | 1334.3971 | 18.2386 | 30.5039
0.3299 | 15.8938 | -2.7456 | -1335.1304 | 1335.3606 | 3.1634 | -18.196 | -1.7224 | 18.5968 | 1335.1304 | 18.196 | 30.6713
0.3316 | 15.8256 | -1.7224 | -1335.9148 | 1336.128 | 3.1464 | -18.1449 | -1.7821 | 18.5541 | 1335.9148 | 18.1449 | 30.8293
0.3334 | 15.7915 | -0.67361 | -1336.9636 | 1337.1683 | 3.1378 | -18.0937 | -1.8332 | 18.4944 | 1336.9636 | 18.0937 | 30.9967
0.3352 | 15.7403 | 0.34107 | -1338.1915 | 1338.3876 | 3.1293 | -18.034 | -1.9015 | 18.4433 | 1338.1915 | 18.034 | 31.164
0.337 | 15.7147 | 1.2961 | -1339.5728 | 1339.7604 | 3.1208 | -17.9658 | -1.9526 | 18.3921 | 1339.5728 | 17.9658 | 31.3313
0.3388 | 15.6892 | 2.2425 | -1340.9968 | 1341.1843 | 3.1208 | -17.9061 | -2.0123 | 18.341 | 1340.9968 | 17.9061 | 31.4987
0.3406 | 15.6636 | 3.2316 | -1342.6765 | 1342.8641 | 3.1123 | -17.8549 | -2.0635 | 18.2983 | 1342.6765 | 17.8549 | 31.666
0.3424 | 15.6977 | 4.161 | -1344.595 | 1344.7826 | 3.1037 | -17.8123 | -2.1232 | 18.2557 | 1344.595 | 17.8123 | 31.8334
0.3442 | 15.7574 | 5.0563 | -1346.7267 | 1346.9058 | 3.0952 | -17.7697 | -2.1914 | 18.2216 | 1346.7267 | 17.7697 | 32.0007
0.346 | 15.783 | 5.9943 | -1348.9948 | 1349.1739 | 3.0952 | -17.727 | -2.2511 | 18.1875 | 1348.9948 | 17.727 | 32.1681
0.3478 | 15.7488 | 6.8981 | -1351.4164 | 1351.5955 | 3.0867 | -17.6759 | -2.3022 | 18.1449 | 1351.4164 | 17.6759 | 32.3354
0.3496 | 15.6977 | 7.7764 | -1353.9915 | 1354.1791 | 3.0781 | -17.6247 | -2.3619 | 18.1022 | 1353.9915 | 17.6247 | 32.5028
0.3514 | 15.6551 | 8.6205 | -1356.9588 | 1357.1464 | 3.0696 | -17.5736 | -2.4301 | 18.0596 | 1356.9588 | 17.5736 | 32.6701
0.3532 | 15.6295 | 9.4988 | -1360.5059 | 1360.6935 | 3.0781 | -17.5139 | -2.4983 | 18.0084 | 1360.5059 | 17.5139 | 32.8375
0.355 | 15.6295 | 10.3344 | -1364.2065 | 1364.4026 | 3.0867 | -17.4627 | -2.5751 | 17.9658 | 1364.2065 | 17.4627 | 33.0048
0.3567 | 15.6465 | 11.1018 | -1368.0435 | 1368.2396 | 3.0952 | -17.4116 | -2.6433 | 17.9317 | 1368.0435 | 17.4116 | 33.1629
0.3585 | 15.6465 | 11.818 | -1371.8635 | 1372.0681 | 3.1037 | -17.3604 | -2.7115 | 17.8976 | 1371.8635 | 17.3604 | 33.3302
0.3603 | 15.6039 | 12.4661 | -1376.3059 | 1376.5191 | 3.1208 | -17.3178 | -2.7882 | 17.872 | 1376.3059 | 17.3178 | 33.4976
0.3621 | 15.5272 | 13.1226 | -1381.456 | 1381.6692 | 3.1634 | -17.2751 | -2.8735 | 17.8464 | 1381.456 | 17.2751 | 33.6649
0.3639 | 15.4675 | 13.8644 | -1385.762 | 1385.9922 | 3.1805 | -17.2325 | -2.9417 | 17.8294 | 1385.762 | 17.2325 | 33.8323
0.3657 | 15.3737 | 14.6063 | -1389.9913 | 1390.213 | 3.1975 | -17.1984 | -3.0099 | 17.8038 | 1389.9913 | 17.1984 | 33.9996
0.3675 | 15.2372 | 15.4504 | -1394.2973 | 1394.5275 | 3.2146 | -17.1472 | -3.0781 | 17.7782 | 1394.2973 | 17.1472 | 34.167
0.3693 | 15.1093 | 16.2946 | -1398.6203 | 1398.8591 | 3.2402 | -17.1046 | -3.1378 | 17.7526 | 1398.6203 | 17.1046 | 34.3343
0.3711 | 15.0411 | 17.0023 | -1403.2929 | 1403.5402 | 3.2572 | -17.0705 | -3.206 | 17.7356 | 1403.2929 | 17.0705 | 34.5017
0.3729 | 15.007 | 17.5395 | -1408.0253 | 1408.2811 | 3.2828 | -17.0279 | -3.2743 | 17.71 | 1408.0253 | 17.0279 | 34.669
0.3747 | 14.9388 | 18.0425 | -1413.1754 | 1413.4312 | 3.3084 | -17.0023 | -3.3425 | 17.7015 | 1413.1754 | 17.0023 | 34.8364
0.3765 | 14.8876 | 18.5456 | -1418.2147 | 1418.479 | 3.3254 | -16.9682 | -3.4192 | 17.6929 | 1418.2147 | 16.9682 | 35.0037
0.3783 | 14.7427 | 18.8526 | -1422.9044 | 1423.1687 | 3.3425 | -16.9341 | -3.4789 | 17.6674 | 1422.9044 | 16.9341 | 35.1711
0.3801 | 14.7086 | 19.3215 | -1427.0484 | 1427.3212 | 3.3595 | -16.9 | -3.5471 | 17.6503 | 1427.0484 | 16.9 | 35.3384
0.3818 | 14.5892 | 19.8246 | -1430.7063 | 1430.9707 | 3.3766 | -16.8658 | -3.6068 | 17.6333 | 1430.7063 | 16.8658 | 35.4965
0.3836 | 14.4272 | 20.2595 | -1433.6225 | 1433.8953 | 3.3936 | -16.8317 | -3.6665 | 17.6162 | 1433.6225 | 16.8317 | 35.6638
0.3854 | 14.2823 | 20.6261 | -1436.2658 | 1436.5301 | 3.4107 | -16.7976 | -3.7347 | 17.5906 | 1436.2658 | 16.7976 | 35.8312
0.3872 | 14.1288 | 20.8649 | -1438.6532 | 1438.9176 | 3.4192 | -16.7721 | -3.7944 | 17.5821 | 1438.6532 | 16.7721 | 35.9985
0.389 | 13.9753 | 21.0525 | -1441.0492 | 1441.3136 | 3.4192 | -16.755 | -3.8626 | 17.5736 | 1441.0492 | 16.755 | 36.1659
0.3908 | 13.7962 | 21.3253 | -1444.656 | 1444.9204 | 3.4192 | -16.7379 | -3.9223 | 17.5821 | 1444.656 | 16.7379 | 36.3332
0.3926 | 13.7536 | 21.7517 | -1448.8682 | 1449.1411 | 3.4277 | -16.7209 | -3.9905 | 17.5736 | 1448.8682 | 16.7209 | 36.5006
0.3944 | 13.7451 | 22.1524 | -1453.3362 | 1453.6176 | 3.4533 | -16.7038 | -4.0587 | 17.5821 | 1453.3362 | 16.7038 | 36.6679
0.3962 | 13.7451 | 22.5873 | -1458.1965 | 1458.4779 | 3.4789 | -16.6868 | -4.1355 | 17.5821 | 1458.1965 | 16.6868 | 36.8353
0.398 | 13.6939 | 23.0733 | -1463.5001 | 1463.79 | 3.5045 | -16.6527 | -4.2122 | 17.5821 | 1463.5001 | 16.6527 | 37.0026
0.3998 | 13.6939 | 23.7128 | -1469.0766 | 1469.3835 | 3.5386 | -16.6356 | -4.2889 | 17.5906 | 1469.0766 | 16.6356 | 37.17
0.4016 | 13.6939 | 24.5058 | -1474.7127 | 1475.0282 | 3.5727 | -16.6186 | -4.3657 | 17.5991 | 1474.7127 | 16.6186 | 37.3373
0.4034 | 13.6598 | 25.2476 | -1480.5279 | 1480.8605 | 3.6068 | -16.593 | -4.4339 | 17.6077 | 1480.5279 | 16.593 | 37.5046
0.4051 | 13.6257 | 25.8615 | -1485.7804 | 1486.1215 | 3.6409 | -16.5759 | -4.5106 | 17.6162 | 1485.7804 | 16.5759 | 37.6627
0.4069 | 13.6172 | 26.4669 | -1491.5189 | 1491.8685 | 3.675 | -16.5589 | -4.5788 | 17.6247 | 1491.5189 | 16.5589 | 37.83
0.4087 | 13.6513 | 27.0894 | -1498.1527 | 1498.5108 | 3.7262 | -16.5418 | -4.6556 | 17.6418 | 1498.1527 | 16.5418 | 37.9974
0.4105 | 13.7451 | 27.7971 | -1505.3748 | 1505.7414 | 3.7773 | -16.5248 | -4.7409 | 17.6674 | 1505.3748 | 16.5248 | 38.1647
0.4123 | 13.8815 | 28.6839 | -1513.424 | 1513.8162 | 3.837 | -16.5163 | -4.8176 | 17.6929 | 1513.424 | 16.5163 | 38.3321
0.4141 | 14.035 | 29.4683 | -1521.1919 | 1521.6011 | 3.9052 | -16.5077 | -4.8858 | 17.727 | 1521.1919 | 16.5077 | 38.4994
0.4159 | 14.1288 | 30.2102 | -1528.5334 | 1528.9512 | 3.9649 | -16.5077 | -4.9625 | 17.7612 | 1528.5334 | 16.5077 | 38.6668
0.4177 | 14.1373 | 30.7814 | -1535.3206 | 1535.747 | 4.0161 | -16.4992 | -5.0393 | 17.7953 | 1535.3206 | 16.4992 | 38.8341
0.4195 | 14.1288 | 31.2931 | -1541.7071 | 1542.142 | 4.0758 | -16.4992 | -5.1075 | 17.8294 | 1541.7071 | 16.4992 | 39.0015
0.4213 | 14.0861 | 31.6171 | -1546.934 | 1547.3774 | 4.1099 | -16.5163 | -5.1842 | 17.8635 | 1546.934 | 16.5163 | 39.1688
0.4231 | 13.8559 | 31.8984 | -1552.1694 | 1552.6128 | 4.1525 | -16.5077 | -5.2525 | 17.8891 | 1552.1694 | 16.5077 | 39.3362
0.4249 | 13.6001 | 32.1713 | -1557.4475 | 1557.8994 | 4.1951 | -16.5077 | -5.3292 | 17.9146 | 1557.4475 | 16.5077 | 39.5035
0.4267 | 13.4807 | 32.4527 | -1563.851 | 1564.3115 | 4.2463 | -16.4992 | -5.4059 | 17.9487 | 1563.851 | 16.4992 | 39.6709
0.4285 | 13.3614 | 32.6829 | -1569.9561 | 1570.4081 | 4.2975 | -16.5163 | -5.4827 | 17.9914 | 1569.9561 | 16.5163 | 39.8382
0.4302 | 13.2676 | 32.9302 | -1576.2148 | 1576.6752 | 4.3486 | -16.5248 | -5.5594 | 18.0425 | 1576.2148 | 16.5248 | 39.9963
0.432 | 13.2164 | 33.203 | -1583.0788 | 1583.5477 | 4.4083 | -16.5248 | -5.6447 | 18.0852 | 1583.0788 | 16.5248 | 40.1636
0.4338 | 13.2249 | 33.5441 | -1590.5738 | 1591.0427 | 4.4851 | -16.5248 | -5.73 | 18.1363 | 1590.5738 | 16.5248 | 40.331
0.4356 | 13.2164 | 33.9278 | -1598.4269 | 1598.9044 | 4.5703 | -16.5418 | -5.8067 | 18.1875 | 1598.4269 | 16.5418 | 40.4983
0.4374 | 13.1738 | 34.2774 | -1606.5869 | 1607.0644 | 4.6556 | -16.5418 | -5.8834 | 18.2386 | 1606.5869 | 16.5418 | 40.6657
0.4392 | 13.08 | 34.6782 | -1615.2074 | 1615.6935 | 4.7579 | -16.5248 | -5.9687 | 18.2813 | 1615.2074 | 16.5248 | 40.833
0.441 | 13.08 | 35.1471 | -1624.331 | 1624.8256 | 4.8688 | -16.5248 | -6.054 | 18.341 | 1624.331 | 16.5248 | 41.0004
0.4428 | 13.1056 | 35.6673 | -1634.1197 | 1634.6313 | 4.9881 | -16.5248 | -6.1392 | 18.4007 | 1634.1197 | 16.5248 | 41.1677
0.4446 | 13.1056 | 36.1106 | -1644.096 | 1644.6161 | 5.1246 | -16.5163 | -6.216 | 18.4689 | 1644.096 | 16.5163 | 41.3351
0.4464 | 13.0544 | 36.4517 | -1653.1257 | 1653.6544 | 5.261 | -16.5077 | -6.2927 | 18.5371 | 1653.1257 | 16.5077 | 41.5024
0.4482 | 12.9436 | 36.8013 | -1661.0215 | 1661.5501 | 5.3718 | -16.4992 | -6.378 | 18.5882 | 1661.0215 | 16.4992 | 41.6698
0.45 | 12.773 | 37.006 | -1668.2862 | 1668.8234 | 5.4742 | -16.4821 | -6.4547 | 18.6394 | 1668.2862 | 16.4821 | 41.8371
0.4518 | 12.5513 | 37.0486 | -1674.5193 | 1675.0479 | 5.5594 | -16.4736 | -6.5315 | 18.6735 | 1674.5193 | 16.4736 | 42.0045
0.4536 | 12.3723 | 37.1168 | -1680.8802 | 1681.4089 | 5.6447 | -16.4566 | -6.6082 | 18.7161 | 1680.8802 | 16.4566 | 42.1718
0.4553 | 12.3893 | 37.2873 | -1688.324 | 1688.8612 | 5.7555 | -16.448 | -6.6849 | 18.7673 | 1688.324 | 16.448 | 42.3299
0.4571 | 12.4661 | 37.4834 | -1697.26 | 1697.7972 | 5.892 | -16.448 | -6.7617 | 18.844 | 1697.26 | 16.448 | 42.4972
0.4589 | 12.5769 | 37.6455 | -1706.2472 | 1706.7929 | 6.0369 | -16.4566 | -6.8384 | 18.9208 | 1706.2472 | 16.4566 | 42.6646
0.4607 | 12.4916 | 37.816 | -1714.2453 | 1714.791 | 6.1904 | -16.448 | -6.9152 | 18.989 | 1714.2453 | 16.448 | 42.8319
0.4625 | 12.2785 | 37.8586 | -1721.2457 | 1721.7914 | 6.3354 | -16.431 | -6.9919 | 19.0487 | 1721.2457 | 16.431 | 42.9993
0.4643 | 12.0397 | 37.8842 | -1727.8539 | 1728.3996 | 6.4888 | -16.4139 | -7.0686 | 19.1084 | 1727.8539 | 16.4139 | 43.1666
0.4661 | 11.7924 | 37.9609 | -1735.1528 | 1735.6985 | 6.6423 | -16.3969 | -7.1454 | 19.1766 | 1735.1528 | 16.3969 | 43.334
0.4679 | 11.5878 | 38.1144 | -1743.7903 | 1744.3446 | 6.8043 | -16.3628 | -7.2221 | 19.2363 | 1743.7903 | 16.3628 | 43.5013
0.4697 | 11.4087 | 38.3105 | -1753.0503 | 1753.6046 | 6.9749 | -16.3372 | -7.2903 | 19.3045 | 1753.0503 | 16.3372 | 43.6687
0.4715 | 11.1785 | 38.609 | -1762.8987 | 1763.4529 | 7.1454 | -16.2946 | -7.3586 | 19.3642 | 1762.8987 | 16.2946 | 43.836
0.4733 | 10.9227 | 38.8307 | -1772.9346 | 1773.4974 | 7.3244 | -16.2519 | -7.4353 | 19.4324 | 1772.9346 | 16.2519 | 44.0033
0.4751 | 10.701 | 39.0097 | -1781.99 | 1782.5528 | 7.512 | -16.2178 | -7.5035 | 19.5006 | 1781.99 | 16.2178 | 44.1707
0.4769 | 10.4708 | 39.1206 | -1790.6958 | 1791.2586 | 7.6911 | -16.1752 | -7.5717 | 19.5688 | 1790.6958 | 16.1752 | 44.338
0.4786 | 10.1212 | 39.0865 | -1799.1714 | 1799.7341 | 7.8702 | -16.124 | -7.6399 | 19.6285 | 1799.1714 | 16.124 | 44.4961
0.4804 | 9.6181 | 39.0779 | -1807.9624 | 1808.5166 | 8.0492 | -16.0473 | -7.6996 | 19.6711 | 1807.9624 | 16.0473 | 44.6634
0.4822 | 9.0383 | 39.1973 | -1816.8643 | 1817.41 | 8.2368 | -15.962 | -7.7593 | 19.7138 | 1816.8643 | 15.962 | 44.8308
0.484 | 8.4585 | 39.2655 | -1824.8368 | 1825.3825 | 8.4244 | -15.8938 | -7.8275 | 19.7735 | 1824.8368 | 15.8938 | 44.9981
0.4858 | 7.9469 | 39.24 | -1832.0504 | 1832.5961 | 8.5864 | -15.8426 | -7.8872 | 19.8246 | 1832.0504 | 15.8426 | 45.1655
0.4876 | 7.4524 | 39.3082 | -1838.9485 | 1839.4942 | 8.7569 | -15.8 | -7.9554 | 19.8758 | 1838.9485 | 15.8 | 45.3328
0.4894 | 6.9407 | 39.4446 | -1845.8722 | 1846.4179 | 8.9275 | -15.7488 | -8.0236 | 19.9269 | 1845.8722 | 15.7488 | 45.5002
0.4912 | 6.3439 | 39.6151 | -1853.2393 | 1853.785 | 9.098 | -15.6806 | -8.0919 | 19.9781 | 1853.2393 | 15.6806 | 45.6675
0.493 | 5.6788 | 39.743 | -1860.2823 | 1860.828 | 9.2515 | -15.6124 | -8.1515 | 20.0207 | 1860.2823 | 15.6124 | 45.8349
0.4948 | 4.9625 | 39.888 | -1866.2169 | 1866.7626 | 9.4135 | -15.5272 | -8.2112 | 20.0548 | 1866.2169 | 15.5272 | 46.0022
0.4966 | 4.2122 | 40.0841 | -1871.3926 | 1871.9469 | 9.584 | -15.4334 | -8.2709 | 20.0975 | 1871.3926 | 15.4334 | 46.1696
0.4984 | 3.3595 | 40.5275 | -1876.1505 | 1876.7133 | 9.746 | -15.3396 | -8.3221 | 20.1231 | 1876.1505 | 15.3396 | 46.3369
0.5002 | 2.4557 | 41.0476 | -1880.8317 | 1881.4115 | 9.891 | -15.2372 | -8.3818 | 20.1486 | 1880.8317 | 15.2372 | 46.5043
0.502 | 1.5774 | 41.6019 | -1885.4191 | 1886.0074 | 10.0445 | -15.1434 | -8.4244 | 20.1657 | 1885.4191 | 15.1434 | 46.6716
0.5037 | 0.69919 | 42.1731 | -1890.2026 | 1890.7994 | 10.1894 | -15.0497 | -8.4756 | 20.1913 | 1890.2026 | 15.0497 | 46.8297
0.5055 | -0.15348 | 42.6165 | -1895.9752 | 1896.5891 | 10.3173 | -14.9473 | -8.5267 | 20.1998 | 1895.9752 | 14.9473 | 46.997
0.5073 | -1.0658 | 42.9235 | -1902.4384 | 1903.0523 | 10.4452 | -14.8194 | -8.5864 | 20.1998 | 1902.4384 | 14.8194 | 47.1644
0.5091 | -1.9697 | 43.1708 | -1909.3621 | 1909.9846 | 10.5817 | -14.6915 | -8.6376 | 20.1998 | 1909.3621 | 14.6915 | 47.3317
0.5109 | -2.9417 | 43.4948 | -1916.5672 | 1917.1982 | 10.7181 | -14.5551 | -8.6887 | 20.1998 | 1916.5672 | 14.5551 | 47.4991
0.5127 | -3.8967 | 43.708 | -1922.8002 | 1923.4312 | 10.8289 | -14.4187 | -8.7399 | 20.1827 | 1922.8002 | 14.4187 | 47.6664
0.5145 | -4.8688 | 43.8785 | -1928.4279 | 1929.0588 | 10.9313 | -14.2823 | -8.791 | 20.1657 | 1928.4279 | 14.2823 | 47.8338
0.5163 | -5.8152 | 43.9638 | -1933.7059 | 1934.3454 | 11.0336 | -14.1458 | -8.8337 | 20.1486 | 1933.7059 | 14.1458 | 48.0011
0.5181 | -6.8555 | 43.9296 | -1938.3274 | 1938.9669 | 11.1274 | -14.0094 | -8.8848 | 20.1316 | 1938.3274 | 14.0094 | 48.1685
0.5199 | -7.8616 | 43.9211 | -1942.2838 | 1942.9233 | 11.2297 | -13.873 | -8.9275 | 20.106 | 1942.2838 | 13.873 | 48.3358
0.5217 | -8.7314 | 43.9467 | -1945.7371 | 1946.3766 | 11.3235 | -13.7365 | -8.9701 | 20.0889 | 1945.7371 | 13.7365 | 48.5032
0.5235 | -9.4902 | 44.0149 | -1948.9005 | 1949.5485 | 11.4173 | -13.6172 | -9.0127 | 20.0719 | 1948.9005 | 13.6172 | 48.6705
0.5253 | -10.3088 | 44.1769 | -1951.919 | 1952.567 | 11.494 | -13.4978 | -9.0639 | 20.0463 | 1951.919 | 13.4978 | 48.8379
0.5271 | -11.4087 | 44.0575 | -1954.0506 | 1954.7157 | 11.5708 | -13.3528 | -9.098 | 20.0122 | 1954.0506 | 13.3528 | 49.0052
0.5288 | -12.4831 | 43.887 | -1955.611 | 1956.2676 | 11.6475 | -13.2079 | -9.1321 | 19.9781 | 1955.611 | 13.2079 | 49.1633
0.5306 | -13.5575 | 44.0149 | -1957.1629 | 1957.828 | 11.7498 | -13.0544 | -9.1577 | 19.9525 | 1957.1629 | 13.0544 | 49.3306
0.5324 | -14.6318 | 44.1002 | -1958.5101 | 1959.1752 | 11.8266 | -12.8924 | -9.1918 | 19.9099 | 1958.5101 | 12.8924 | 49.498
0.5342 | -15.6465 | 44.1002 | -1959.4736 | 1960.1472 | 11.8862 | -12.7304 | -9.2174 | 19.8587 | 1959.4736 | 12.7304 | 49.6653
0.536 | -16.4395 | 44.1258 | -1960.0364 | 1960.71 | 11.9289 | -12.5854 | -9.2344 | 19.799 | 1960.0364 | 12.5854 | 49.8327
0.5378 | -17.6418 | 44.1513 | -1959.9341 | 1960.6162 | 11.963 | -12.4234 | -9.2685 | 19.7308 | 1959.9341 | 12.4234 | 50
0.5396 | -18.8952 | 44.2792 | -1959.6186 | 1960.3178 | 12.0141 | -12.2614 | -9.2941 | 19.6711 | 1959.6186 | 12.2614 | 50.1673
0.5414 | -19.9525 | 44.1087 | -1959.0388 | 1959.7465 | 12.0397 | -12.0824 | -9.3112 | 19.5859 | 1959.0388 | 12.0824 | 50.3347
0.5432 | -21.078 | 44.0234 | -1957.8535 | 1958.5783 | 12.0568 | -11.8948 | -9.3282 | 19.4921 | 1957.8535 | 11.8948 | 50.502
0.545 | -22.1865 | 43.887 | -1956.1653 | 1956.89 | 12.0738 | -11.7072 | -9.3368 | 19.3898 | 1956.1653 | 11.7072 | 50.6694
0.5468 | -23.312 | 43.725 | -1954.1274 | 1954.8607 | 12.0824 | -11.5025 | -9.3453 | 19.2874 | 1954.1274 | 11.5025 | 50.8367
0.5486 | -24.5996 | 43.7506 | -1951.6802 | 1952.4306 | 12.0909 | -11.3064 | -9.3623 | 19.1766 | 1951.6802 | 11.3064 | 51.0041
0.5504 | -25.7592 | 43.725 | -1948.7044 | 1949.4718 | 12.0824 | -11.1103 | -9.3709 | 19.0657 | 1948.7044 | 11.1103 | 51.1714
0.5521 | -26.9103 | 43.7335 | -1945.0294 | 1945.8138 | 12.0824 | -10.9057 | -9.3794 | 18.9464 | 1945.0294 | 10.9057 | 51.3295
0.5539 | -28.0614 | 43.708 | -1940.911 | 1941.7125 | 12.0738 | -10.6925 | -9.3879 | 18.8185 | 1940.911 | 10.6925 | 51.4968
0.5557 | -29.3404 | 43.887 | -1936.443 | 1937.2615 | 12.0483 | -10.4793 | -9.3964 | 18.6906 | 1936.443 | 10.4793 | 51.6642
0.5575 | -30.5853 | 44.0149 | -1930.7898 | 1931.6424 | 12.0227 | -10.2321 | -9.3964 | 18.5286 | 1930.7898 | 10.2321 | 51.8315
0.5593 | -31.6853 | 43.8955 | -1924.2412 | 1925.1024 | 11.98 | -9.9933 | -9.3794 | 18.3665 | 1924.2412 | 9.9933 | 51.9989
0.5611 | -32.8108 | 43.708 | -1917.3175 | 1918.2043 | 11.9374 | -9.7546 | -9.3623 | 18.196 | 1917.3175 | 9.7546 | 52.1662
0.5629 | -34.0045 | 43.5459 | -1910.1295 | 1911.0248 | 11.8948 | -9.5243 | -9.3453 | 18.034 | 1910.1295 | 9.5243 | 52.3336
0.5647 | -35.2494 | 43.418 | -1902.2252 | 1903.1461 | 11.8436 | -9.2771 | -9.3197 | 17.8549 | 1902.2252 | 9.2771 | 52.5009
0.5665 | -36.5029 | 43.3839 | -1893.0676 | 1894.014 | 11.7754 | -9.0042 | -9.2771 | 17.6588 | 1893.0676 | 9.0042 | 52.6683
0.5683 | -37.7392 | 43.239 | -1883.1084 | 1884.0804 | 11.7072 | -8.7314 | -9.2259 | 17.4371 | 1883.1084 | 8.7314 | 52.8356
0.5701 | -38.9927 | 42.9661 | -1872.2794 | 1873.277 | 11.6219 | -8.45 | -9.1833 | 17.2154 | 1872.2794 | 8.45 | 53.003
0.5719 | -40.2546 | 42.463 | -1859.6855 | 1860.7001 | 11.5196 | -8.1601 | -9.1236 | 16.9767 | 1859.6855 | 8.1601 | 53.1703
0.5737 | -41.653 | 41.9856 | -1846.1536 | 1847.2023 | 11.4087 | -7.8531 | -9.0639 | 16.7294 | 1846.1536 | 7.8531 | 53.3377
0.5755 | -43.1537 | 41.167 | -1831.1721 | 1832.238 | 11.2894 | -7.4865 | -8.9957 | 16.431 | 1831.1721 | 7.4865 | 53.505
0.5772 | -44.6715 | 40.4678 | -1815.5512 | 1816.6511 | 11.1529 | -7.1198 | -8.9189 | 16.1325 | 1815.5512 | 7.1198 | 53.6631
0.579 | -46.266 | 39.7771 | -1798.6853 | 1799.8194 | 11.0165 | -6.7446 | -8.8337 | 15.8341 | 1798.6853 | 6.7446 | 53.8304
0.5808 | -47.9116 | 38.933 | -1780.1312 | 1781.2994 | 10.8545 | -6.3268 | -8.7314 | 15.493 | 1780.1312 | 6.3268 | 53.9978
0.5826 | -49.5402 | 38.2167 | -1760.5368 | 1761.7391 | 10.6925 | -5.8834 | -8.6205 | 15.1605 | 1760.5368 | 5.8834 | 54.1651
0.5844 | -51.1347 | 37.5858 | -1739.9533 | 1741.1982 | 10.522 | -5.4145 | -8.5097 | 14.828 | 1739.9533 | 5.4145 | 54.3325
0.5862 | -52.6525 | 37.0401 | -1718.8753 | 1720.1713 | 10.36 | -4.9455 | -8.3988 | 14.4784 | 1718.8753 | 4.9455 | 54.4998
0.588 | -54.0509 | 36.6734 | -1697.3538 | 1698.701 | 10.1894 | -4.4595 | -8.2709 | 14.1288 | 1697.3538 | 4.4595 | 54.6672
0.5898 | -55.381 | 36.392 | -1675.4146 | 1676.813 | 10.0104 | -3.9735 | -8.1515 | 13.7792 | 1675.4146 | 3.9735 | 54.8345
0.5916 | -56.66 | 36.1362 | -1653.1002 | 1654.5497 | 9.8313 | -3.4704 | -8.0236 | 13.4296 | 1653.1002 | 3.4704 | 55.0019
0.5934 | -57.8964 | 35.9231 | -1630.4106 | 1631.9198 | 9.6437 | -2.9502 | -7.8872 | 13.08 | 1630.4106 | 2.9502 | 55.1692
0.5952 | -59.1072 | 35.8293 | -1606.928 | 1608.4969 | 9.4476 | -2.4131 | -7.7508 | 12.7219 | 1606.928 | 2.4131 | 55.3366
0.597 | -60.3521 | 35.889 | -1582.3625 | 1584.0167 | 9.243 | -1.8674 | -7.5973 | 12.3893 | 1582.3625 | 1.8674 | 55.5039
0.5988 | -61.6055 | 36.0254 | -1557.2343 | 1558.9652 | 9.0468 | -1.3131 | -7.4353 | 12.0483 | 1557.2343 | 1.3131 | 55.6713
0.6006 | -62.8334 | 36.2471 | -1531.2107 | 1533.0269 | 8.8337 | -0.7674 | -7.2648 | 11.7242 | 1531.2107 | 0.7674 | 55.8386
0.6023 | -64.0101 | 36.5711 | -1504.4283 | 1506.3298 | 8.6205 | -0.23022 | -7.0857 | 11.4258 | 1504.4283 | 0.23022 | 55.9967
0.6041 | -65.093 | 37.006 | -1476.9211 | 1478.9249 | 8.3988 | 0.29844 | -6.9066 | 11.1529 | 1476.9211 | -0.29844 | 56.164
0.6059 | -66.0821 | 37.5261 | -1449.0644 | 1451.1619 | 8.1771 | 0.82709 | -6.7105 | 10.9057 | 1449.0644 | -0.82709 | 56.3313
0.6077 | -67.0115 | 38.0803 | -1420.696 | 1422.8959 | 7.9554 | 1.3557 | -6.5144 | 10.6754 | 1420.696 | -1.3557 | 56.4987
0.6095 | -67.8982 | 38.6175 | -1391.6369 | 1393.9391 | 7.7252 | 1.8759 | -6.3098 | 10.4452 | 1391.6369 | -1.8759 | 56.666
0.6113 | -68.6315 | 39.1291 | -1362.288 | 1364.6754 | 7.495 | 2.3875 | -6.0966 | 10.2321 | 1362.288 | -2.3875 | 56.8334
0.6131 | -69.2284 | 39.6407 | -1332.8964 | 1335.3776 | 7.2562 | 2.8906 | -5.8749 | 10.0445 | 1332.8964 | -2.8906 | 57.0007
0.6149 | -69.6547 | 40.1011 | -1302.9505 | 1305.5256 | 7.0175 | 3.3595 | -5.6532 | 9.8825 | 1302.9505 | -3.3595 | 57.1681
0.6167 | -69.902 | 40.4507 | -1273.1411 | 1275.8015 | 6.7787 | 3.7944 | -5.423 | 9.729 | 1273.1411 | -3.7944 | 57.3354
0.6185 | -70.1578 | 40.843 | -1243.1015 | 1245.8556 | 6.54 | 4.2293 | -5.1928 | 9.5926 | 1243.1015 | -4.2293 | 57.5028
0.6203 | -70.3625 | 41.3034 | -1212.9169 | 1215.7649 | 6.3098 | 4.6556 | -4.9625 | 9.4902 | 1212.9169 | -4.6556 | 57.6701
0.6221 | -70.4563 | 41.7383 | -1182.5959 | 1185.5291 | 6.071 | 5.0478 | -4.7153 | 9.4135 | 1182.5959 | -5.0478 | 57.8375
0.6239 | -70.5501 | 42.0879 | -1152.9485 | 1155.967 | 5.8493 | 5.423 | -4.468 | 9.3368 | 1152.9485 | -5.423 | 58.0048
0.6256 | -70.5756 | 42.4119 | -1123.0198 | 1126.132 | 5.6276 | 5.7811 | -4.2207 | 9.2941 | 1123.0198 | -5.7811 | 58.1629
0.6274 | -70.5415 | 42.6165 | -1093.2445 | 1096.442 | 5.4059 | 6.1222 | -3.9649 | 9.2515 | 1093.2445 | -6.1222 | 58.3302
0.6292 | -70.388 | 42.6762 | -1063.6482 | 1066.9225 | 5.1928 | 6.4291 | -3.7176 | 9.2089 | 1063.6482 | -6.4291 | 58.4976
0.631 | -70.1152 | 42.5398 | -1034.4954 | 1037.8293 | 4.9796 | 6.7105 | -3.4789 | 9.1833 | 1034.4954 | -6.7105 | 58.6649
0.6328 | -69.7059 | 42.0538 | -1005.1379 | 1008.5145 | 4.775 | 6.9749 | -3.2316 | 9.1833 | 1005.1379 | -6.9749 | 58.8323
0.6346 | -69.2114 | 41.4825 | -975.5331 | 978.9524 | 4.5788 | 7.2136 | -2.9844 | 9.1747 | 975.5331 | -7.2136 | 58.9996
0.6364 | -68.6571 | 40.8771 | -946.116 | 949.5693 | 4.3827 | 7.4353 | -2.7456 | 9.1747 | 946.116 | -7.4353 | 59.167
0.6382 | -68.0603 | 40.2376 | -916.9631 | 920.4506 | 4.1951 | 7.6399 | -2.4983 | 9.1662 | 916.9631 | -7.6399 | 59.3343
0.64 | -67.3952 | 39.5299 | -888.1855 | 891.6985 | 4.0161 | 7.8105 | -2.2596 | 9.1577 | 888.1855 | -7.8105 | 59.5017
0.6418 | -66.6619 | 38.6431 | -860.2093 | 863.7394 | 3.8541 | 7.9469 | -2.0294 | 9.1321 | 860.2093 | -7.9469 | 59.669
0.6436 | -65.92 | 37.6199 | -832.9409 | 836.4794 | 3.7176 | 8.0577 | -1.8162 | 9.1151 | 832.9409 | -8.0577 | 59.8364
0.6454 | -65.1526 | 36.5796 | -806.2352 | 809.7738 | 3.5812 | 8.143 | -1.603 | 9.0895 | 806.2352 | -8.143 | 60.0037
0.6472 | -64.4023 | 35.42 | -780.0496 | 783.5882 | 3.4533 | 8.2027 | -1.3984 | 9.0554 | 780.0496 | -8.2027 | 60.1711
0.649 | -63.6946 | 34.0813 | -754.1284 | 757.6585 | 3.3425 | 8.2283 | -1.2023 | 8.9957 | 754.1284 | -8.2283 | 60.3384
0.6507 | -62.9954 | 32.6914 | -728.4715 | 732.0016 | 3.2487 | 8.2198 | -1.0317 | 8.9275 | 728.4715 | -8.2198 | 60.4965
0.6525 | -62.3815 | 31.4039 | -703.4541 | 706.9927 | 3.1719 | 8.1856 | -0.86973 | 8.8507 | 703.4541 | -8.1856 | 60.6638
0.6543 | -61.9381 | 30.1931 | -679.0848 | 682.6575 | 3.1037 | 8.126 | -0.72477 | 8.7484 | 679.0848 | -8.126 | 60.8312
0.6561 | -61.5458 | 29.0079 | -655.048 | 658.6633 | 3.044 | 8.0492 | -0.58834 | 8.6461 | 655.048 | -8.0492 | 60.9985
0.6579 | -61.2304 | 27.8568 | -631.7018 | 635.3598 | 2.9929 | 7.964 | -0.46044 | 8.5352 | 631.7018 | -7.964 | 61.1659
0.6597 | -60.949 | 26.7398 | -609.4897 | 613.1903 | 2.9417 | 7.8531 | -0.35812 | 8.4073 | 609.4897 | -7.8531 | 61.3332
0.6615 | -60.4629 | 25.3926 | -588.3264 | 592.0441 | 2.8906 | 7.6996 | -0.28138 | 8.2368 | 588.3264 | -7.6996 | 61.5006
0.6633 | -59.7297 | 23.6787 | -568.425 | 572.1171 | 2.8394 | 7.495 | -0.23875 | 8.0322 | 568.425 | -7.495 | 61.6679
0.6651 | -58.9964 | 21.9222 | -549.0609 | 552.7274 | 2.7882 | 7.3074 | -0.18759 | 7.8275 | 549.0609 | -7.3074 | 61.8353
0.6669 | -58.1352 | 19.8843 | -530.3703 | 533.9942 | 2.7371 | 7.1113 | -0.16201 | 7.6229 | 530.3703 | -7.1113 | 62.0026
0.6687 | -57.1546 | 17.6333 | -512.0379 | 515.585 | 2.6774 | 6.9066 | -0.14495 | 7.4097 | 512.0379 | -6.9066 | 62.17
0.6705 | -56.157 | 15.4248 | -494.3961 | 497.875 | 2.6177 | 6.6935 | -0.1279 | 7.1965 | 494.3961 | -6.6935 | 62.3373
0.6723 | -55.202 | 13.2676 | -477.4535 | 480.8812 | 2.558 | 6.4974 | -0.11937 | 6.9834 | 477.4535 | -6.4974 | 62.5046
0.6741 | -54.1617 | 11.1188 | -461.6876 | 465.0557 | 2.5069 | 6.3012 | -0.11085 | 6.7873 | 461.6876 | -6.3012 | 62.672
0.6758 | -53.0959 | 9.0383 | -447.0217 | 450.3385 | 2.4557 | 6.1137 | -0.11937 | 6.5912 | 447.0217 | -6.1137 | 62.83
0.6776 | -52.1323 | 6.9578 | -433.089 | 436.3547 | 2.4216 | 5.9431 | -0.11937 | 6.4206 | 433.089 | -5.9431 | 62.9974
0.6794 | -51.1518 | 4.9711 | -419.3951 | 422.6097 | 2.379 | 5.7726 | -0.1279 | 6.2501 | 419.3951 | -5.7726 | 63.1647
0.6812 | -50.3417 | 3.189 | -406.8097 | 410.0072 | 2.3363 | 5.6191 | -0.1279 | 6.0881 | 406.8097 | -5.6191 | 63.3321
0.683 | -49.6511 | 1.5519 | -395.1792 | 398.3597 | 2.3022 | 5.4742 | -0.13643 | 5.9431 | 395.1792 | -5.4742 | 63.4994
0.6848 | -49.0713 | 0.13643 | -384.7937 | 387.9741 | 2.2681 | 5.3548 | -0.14495 | 5.8152 | 384.7937 | -5.3548 | 63.6668
0.6866 | -48.4062 | -1.2023 | -374.8515 | 378.032 | 2.2425 | 5.2269 | -0.16201 | 5.6958 | 374.8515 | -5.2269 | 63.8341
0.6884 | -47.7326 | -2.4557 | -365.1652 | 368.3371 | 2.2169 | 5.1075 | -0.17053 | 5.5679 | 365.1652 | -5.1075 | 64.0015
0.6902 | -47.1869 | -3.5471 | -356.2889 | 359.4864 | 2.1828 | 5.0052 | -0.17053 | 5.4656 | 356.2889 | -5.0052 | 64.1688
0.692 | -46.7179 | -4.5618 | -347.9497 | 351.1643 | 2.1573 | 4.9029 | -0.17906 | 5.3633 | 347.9497 | -4.9029 | 64.3362
0.6938 | -46.283 | -5.4145 | -340.3268 | 343.5499 | 2.1232 | 4.8091 | -0.18759 | 5.261 | 340.3268 | -4.8091 | 64.5035
0.6956 | -45.6947 | -6.233 | -332.8148 | 336.055 | 2.0976 | 4.7153 | -0.18759 | 5.1672 | 332.8148 | -4.7153 | 64.6709
0.6974 | -45.3195 | -6.7617 | -325.9508 | 329.208 | 2.0635 | 4.63 | -0.18759 | 5.0734 | 325.9508 | -4.63 | 64.8382
0.6992 | -45.0296 | -7.1624 | -319.6922 | 322.9835 | 2.0294 | 4.5533 | -0.18759 | 4.9881 | 319.6922 | -4.5533 | 65.0056
0.7009 | -44.7653 | -7.5888 | -313.9111 | 317.228 | 2.0038 | 4.4851 | -0.18759 | 4.9199 | 313.9111 | -4.4851 | 65.1636
0.7027 | -44.5521 | -7.9895 | -308.5989 | 311.9414 | 1.9697 | 4.4168 | -0.17906 | 4.8432 | 308.5989 | -4.4168 | 65.331
0.7045 | -44.2963 | -8.2283 | -303.6961 | 307.0641 | 1.9441 | 4.3572 | -0.17906 | 4.775 | 303.6961 | -4.3572 | 65.4983
0.7063 | -43.9382 | -8.4073 | -299.2366 | 302.5961 | 1.91 | 4.306 | -0.17053 | 4.7153 | 299.2366 | -4.306 | 65.6657
0.7081 | -43.5545 | -8.5352 | -295.229 | 298.58 | 1.8759 | 4.2463 | -0.17053 | 4.6471 | 295.229 | -4.2463 | 65.833
0.7099 | -43.2731 | -8.4244 | -291.6393 | 294.9903 | 1.8418 | 4.2037 | -0.15348 | 4.5874 | 291.6393 | -4.2037 | 66.0004
0.7117 | -42.932 | -8.3135 | -288.2286 | 291.5626 | 1.8077 | 4.161 | -0.14495 | 4.5362 | 288.2286 | -4.161 | 66.1677
0.7135 | -42.4204 | -8.1345 | -284.8264 | 288.1178 | 1.7821 | 4.1099 | -0.1279 | 4.4851 | 284.8264 | -4.1099 | 66.3351
0.7153 | -42.0282 | -7.9213 | -281.9444 | 285.2016 | 1.748 | 4.0672 | -0.11937 | 4.4339 | 281.9444 | -4.0672 | 66.5024
0.7171 | -41.6274 | -7.5717 | -279.4802 | 282.7033 | 1.7224 | 4.0331 | -0.10232 | 4.3827 | 279.4802 | -4.0331 | 66.6698
0.7189 | -41.1414 | -7.2136 | -277.1439 | 280.3158 | 1.6883 | 3.9905 | -0.085267 | 4.3401 | 277.1439 | -3.9905 | 66.8371
0.7207 | -40.843 | -6.7446 | -274.944 | 278.0903 | 1.6627 | 3.9564 | -0.068214 | 4.2975 | 274.944 | -3.9564 | 67.0045
0.7225 | -40.5104 | -6.1733 | -272.8294 | 275.9331 | 1.6371 | 3.9223 | -0.05116 | 4.2548 | 272.8294 | -3.9223 | 67.1718
0.7242 | -40.4337 | -5.5594 | -271.1666 | 274.2704 | 1.6115 | 3.9052 | -0.02558 | 4.2293 | 271.1666 | -3.9052 | 67.3299
0.726 | -40.357 | -4.9199 | -269.828 | 272.9232 | 1.5945 | 3.8967 | -0.0085267 | 4.2122 | 269.828 | -3.8967 | 67.4972
0.7278 | -40.1353 | -4.4509 | -268.1141 | 271.1922 | 1.5689 | 3.8711 | 0.0085267 | 4.1781 | 268.1141 | -3.8711 | 67.6646
0.7296 | -39.7345 | -4.1696 | -265.9312 | 268.9753 | 1.5348 | 3.837 | 0.02558 | 4.1355 | 265.9312 | -3.837 | 67.8319
0.7314 | -39.3082 | -3.8797 | -263.5779 | 266.5793 | 1.5092 | 3.7944 | 0.042634 | 4.0843 | 263.5779 | -3.7944 | 67.9993
0.7332 | -38.8051 | -3.6835 | -261.1648 | 264.1151 | 1.4751 | 3.7518 | 0.059687 | 4.0331 | 261.1648 | -3.7518 | 68.1666
0.735 | -38.0803 | -3.6068 | -258.5471 | 261.4206 | 1.441 | 3.6921 | 0.068214 | 3.9735 | 258.5471 | -3.6921 | 68.334
0.7368 | -37.1935 | -3.6068 | -255.6992 | 258.4704 | 1.4069 | 3.6324 | 0.085267 | 3.8967 | 255.6992 | -3.6324 | 68.5013
0.7386 | -36.1448 | -3.6835 | -252.6807 | 255.3411 | 1.3643 | 3.5556 | 0.093794 | 3.82 | 252.6807 | -3.5556 | 68.6687
0.7404 | -35.0278 | -3.8114 | -249.432 | 251.9645 | 1.3216 | 3.4789 | 0.10232 | 3.7262 | 249.432 | -3.4789 | 68.836
0.7422 | -33.9278 | -3.9393 | -246.0128 | 248.4344 | 1.2875 | 3.4022 | 0.11085 | 3.6409 | 246.0128 | -3.4022 | 69.0033
0.744 | -32.8108 | -4.0928 | -242.4657 | 244.7679 | 1.2449 | 3.3169 | 0.11937 | 3.5471 | 242.4657 | -3.3169 | 69.1707
0.7458 | -31.6597 | -4.2719 | -238.7481 | 240.9394 | 1.2023 | 3.2231 | 0.1279 | 3.4533 | 238.7481 | -3.2231 | 69.338
0.7476 | -30.5001 | -4.4254 | -234.8514 | 236.9234 | 1.1682 | 3.1378 | 0.1279 | 3.351 | 234.8514 | -3.1378 | 69.5054
0.7493 | -29.349 | -4.5021 | -230.9205 | 232.8817 | 1.1341 | 3.044 | 0.13643 | 3.2572 | 230.9205 | -3.044 | 69.6634
0.7511 | -28.2746 | -4.5788 | -227.1944 | 229.0532 | 1.0999 | 2.9588 | 0.13643 | 3.1634 | 227.1944 | -2.9588 | 69.8308
0.7529 | -27.2343 | -4.5447 | -223.6131 | 225.3782 | 1.0744 | 2.865 | 0.14495 | 3.0696 | 223.6131 | -2.865 | 69.9981
0.7547 | -26.2026 | -4.5192 | -220.0405 | 221.6946 | 1.0403 | 2.7797 | 0.14495 | 2.9758 | 220.0405 | -2.7797 | 70.1655
0.7565 | -25.205 | -4.4765 | -216.4592 | 218.0196 | 1.0147 | 2.6944 | 0.15348 | 2.882 | 216.4592 | -2.6944 | 70.3328
0.7583 | -24.267 | -4.4083 | -212.8951 | 214.3787 | 0.9891 | 2.6006 | 0.15348 | 2.7968 | 212.8951 | -2.6006 | 70.5002
0.7601 | -23.457 | -4.3572 | -209.3906 | 210.7975 | 0.96352 | 2.5239 | 0.16201 | 2.7115 | 209.3906 | -2.5239 | 70.6675
0.7619 | -22.7066 | -4.306 | -205.9714 | 207.3186 | 0.94647 | 2.4472 | 0.16201 | 2.6348 | 205.9714 | -2.4472 | 70.8349
0.7637 | -22.0671 | -4.2634 | -202.5351 | 203.8312 | 0.92089 | 2.3704 | 0.16201 | 2.5495 | 202.5351 | -2.3704 | 71.0022
0.7655 | -21.4703 | -4.2293 | -199.1159 | 200.3693 | 0.89531 | 2.3022 | 0.17053 | 2.4813 | 199.1159 | -2.3022 | 71.1696
0.7673 | -20.9416 | -4.144 | -195.8502 | 197.0695 | 0.87825 | 2.2255 | 0.17053 | 2.4045 | 195.8502 | -2.2255 | 71.3369
0.7691 | -20.6091 | -3.8967 | -192.7123 | 193.9061 | 0.8612 | 2.1743 | 0.17053 | 2.3534 | 192.7123 | -2.1743 | 71.5043
0.7709 | -20.2168 | -3.7006 | -189.7621 | 190.9302 | 0.84414 | 2.1146 | 0.17906 | 2.2937 | 189.7621 | -2.1146 | 71.6716
0.7727 | -19.8076 | -3.5215 | -186.9227 | 188.0653 | 0.82709 | 2.0549 | 0.17906 | 2.2255 | 186.9227 | -2.0549 | 71.839
0.7744 | -19.4239 | -3.3254 | -184.2794 | 185.3879 | 0.81856 | 1.9953 | 0.18759 | 2.1743 | 184.2794 | -1.9953 | 71.997
0.7762 | -19.151 | -3.0526 | -182.0369 | 183.1198 | 0.80151 | 1.9441 | 0.18759 | 2.1232 | 182.0369 | -1.9441 | 72.1644
0.778 | -19.0061 | -2.6774 | -180.2804 | 181.3547 | 0.79298 | 1.9015 | 0.19611 | 2.0805 | 180.2804 | -1.9015 | 72.3317
0.7798 | -18.9037 | -2.4727 | -178.413 | 179.4959 | 0.77593 | 1.8674 | 0.20464 | 2.0379 | 178.413 | -1.8674 | 72.4991
0.7816 | -18.7844 | -2.3193 | -176.2217 | 177.3045 | 0.7674 | 1.8247 | 0.20464 | 2.0038 | 176.2217 | -1.8247 | 72.6664
0.7834 | -18.6053 | -2.2084 | -173.6636 | 174.738 | 0.75035 | 1.7906 | 0.20464 | 1.9611 | 173.6636 | -1.7906 | 72.8338
0.7852 | -18.4262 | -2.0208 | -171.0204 | 172.0947 | 0.7333 | 1.748 | 0.20464 | 1.9185 | 171.0204 | -1.748 | 73.0011
0.787 | -18.3495 | -1.7139 | -168.4368 | 169.5197 | 0.72477 | 1.7139 | 0.20464 | 1.8844 | 168.4368 | -1.7139 | 73.1685
0.7888 | -18.2642 | -1.4069 | -165.9555 | 167.0384 | 0.71624 | 1.6798 | 0.20464 | 1.8503 | 165.9555 | -1.6798 | 73.3358
0.7906 | -18.1704 | -1.1426 | -163.7044 | 164.7959 | 0.70772 | 1.6542 | 0.20464 | 1.8247 | 163.7044 | -1.6542 | 73.5032
0.7924 | -18.2557 | -0.79298 | -161.8286 | 162.937 | 0.69919 | 1.6371 | 0.21317 | 1.7991 | 161.8286 | -1.6371 | 73.6705
0.7942 | -18.52 | -0.40928 | -160.3961 | 161.5557 | 0.69066 | 1.6286 | 0.21317 | 1.7991 | 160.3961 | -1.6286 | 73.8379
0.796 | -18.8355 | -0.16201 | -158.9977 | 160.2085 | 0.68214 | 1.6286 | 0.22169 | 1.7906 | 158.9977 | -1.6286 | 74.0052
0.7977 | -18.8526 | -0.042634 | -157.6846 | 158.9209 | 0.66508 | 1.6286 | 0.22169 | 1.7821 | 157.6846 | -1.6286 | 74.1633
0.7995 | -18.5541 | 0.017053 | -155.9707 | 157.1815 | 0.64803 | 1.6115 | 0.21317 | 1.765 | 155.9707 | -1.6115 | 74.3306
0.8013 | -18.52 | 0.13643 | -153.822 | 155.0413 | 0.63098 | 1.5945 | 0.22169 | 1.7395 | 153.822 | -1.5945 | 74.498
0.8031 | -18.4177 | 0.2558 | -151.6221 | 152.8499 | 0.61392 | 1.5689 | 0.22169 | 1.7139 | 151.6221 | -1.5689 | 74.6653
0.8049 | -18.1022 | 0.23875 | -148.9703 | 150.1811 | 0.58834 | 1.5348 | 0.22169 | 1.6712 | 148.9703 | -1.5348 | 74.8327
0.8067 | -17.9317 | 0.27285 | -146.4634 | 147.6827 | 0.56276 | 1.5092 | 0.22169 | 1.6457 | 146.4634 | -1.5092 | 75
0.8085 | -17.8976 | 0.41781 | -144.1186 | 145.3464 | 0.54571 | 1.5007 | 0.22169 | 1.6286 | 144.1186 | -1.5007 | 75.1673
0.8103 | -17.7867 | 0.57129 | -142.2597 | 143.4876 | 0.52866 | 1.4922 | 0.22169 | 1.6115 | 142.2597 | -1.4922 | 75.3347
0.8121 | -17.8464 | 0.83562 | -140.6738 | 141.9272 | 0.5116 | 1.4922 | 0.22169 | 1.603 | 140.6738 | -1.4922 | 75.502
0.8139 | -17.8891 | 1.1255 | -139.2242 | 140.5032 | 0.50308 | 1.4836 | 0.22169 | 1.5945 | 139.2242 | -1.4836 | 75.6694
0.8157 | -17.9658 | 1.4581 | -137.9026 | 139.2072 | 0.48602 | 1.4836 | 0.22169 | 1.5945 | 137.9026 | -1.4836 | 75.8367
0.8175 | -18.2045 | 1.8418 | -136.7771 | 138.1243 | 0.4775 | 1.4922 | 0.23022 | 1.5945 | 136.7771 | -1.4922 | 76.0041
0.8193 | -18.341 | 2.0805 | -135.5663 | 136.9561 | 0.46897 | 1.5007 | 0.23022 | 1.603 | 135.5663 | -1.5007 | 76.1714
0.8211 | -18.3069 | 2.234 | -134.2532 | 135.6515 | 0.45192 | 1.5007 | 0.23022 | 1.5945 | 134.2532 | -1.5007 | 76.3388
0.8228 | -18.2045 | 2.4301 | -133.0679 | 134.4578 | 0.43486 | 1.5007 | 0.23022 | 1.5945 | 133.0679 | -1.5007 | 76.4968
0.8246 | -18.1022 | 2.703 | -132.0533 | 133.4431 | 0.41781 | 1.5007 | 0.23022 | 1.586 | 132.0533 | -1.5007 | 76.6642
0.8264 | -17.9914 | 3.0099 | -131.1494 | 132.5308 | 0.40076 | 1.5007 | 0.23875 | 1.586 | 131.1494 | -1.5007 | 76.8315
0.8282 | -17.8123 | 3.2487 | -130.3138 | 131.661 | 0.3837 | 1.5092 | 0.23875 | 1.586 | 130.3138 | -1.5092 | 76.9989
0.83 | -17.4798 | 3.3425 | -129.4526 | 130.7657 | 0.35812 | 1.5007 | 0.24727 | 1.586 | 129.4526 | -1.5007 | 77.1662
0.8318 | -17.3604 | 3.658 | -128.7364 | 130.041 | 0.34107 | 1.5092 | 0.24727 | 1.586 | 128.7364 | -1.5092 | 77.3336
0.8336 | -17.3263 | 4.0331 | -128.0883 | 129.41 | 0.32402 | 1.5178 | 0.2558 | 1.586 | 128.0883 | -1.5178 | 77.5009
0.8354 | -17.5139 | 4.5618 | -127.6791 | 129.0519 | 0.31549 | 1.5433 | 0.26433 | 1.6115 | 127.6791 | -1.5433 | 77.6683
0.8372 | -17.6333 | 5.0137 | -127.2271 | 128.6511 | 0.30696 | 1.5689 | 0.26433 | 1.6286 | 127.2271 | -1.5689 | 77.8356
0.839 | -17.6588 | 5.3889 | -126.8434 | 128.3015 | 0.29844 | 1.586 | 0.27285 | 1.6542 | 126.8434 | -1.586 | 78.003
0.8408 | -17.6333 | 5.7385 | -126.528 | 128.0031 | 0.28138 | 1.6115 | 0.28138 | 1.6712 | 126.528 | -1.6115 | 78.1703
0.8426 | -17.6162 | 6.1648 | -126.6047 | 128.0883 | 0.26433 | 1.6286 | 0.28991 | 1.6883 | 126.6047 | -1.6286 | 78.3377
0.8444 | -17.7697 | 6.6764 | -127.014 | 128.5317 | 0.2558 | 1.6542 | 0.28991 | 1.7053 | 127.014 | -1.6542 | 78.505
0.8462 | -18.0681 | 7.2562 | -127.517 | 129.0945 | 0.2558 | 1.6798 | 0.29844 | 1.7395 | 127.517 | -1.6798 | 78.6724
0.8479 | -18.2642 | 7.7337 | -128.0201 | 129.6402 | 0.23875 | 1.7053 | 0.30696 | 1.765 | 128.0201 | -1.7053 | 78.8304
0.8497 | -18.2898 | 8.0577 | -128.455 | 130.0921 | 0.22169 | 1.7309 | 0.31549 | 1.7821 | 128.455 | -1.7309 | 78.9978
0.8515 | -18.2983 | 8.467 | -128.9495 | 130.6037 | 0.21317 | 1.748 | 0.32402 | 1.7991 | 128.9495 | -1.748 | 79.1651
0.8533 | -18.1278 | 8.7058 | -129.3588 | 130.996 | 0.19611 | 1.765 | 0.32402 | 1.8162 | 129.3588 | -1.765 | 79.3325
0.8551 | -18.1193 | 9.081 | -129.6658 | 131.3285 | 0.17906 | 1.7906 | 0.33254 | 1.8332 | 129.6658 | -1.7906 | 79.4998
0.8569 | -17.9914 | 9.3453 | -129.8193 | 131.482 | 0.15348 | 1.8077 | 0.34107 | 1.8503 | 129.8193 | -1.8077 | 79.6672
0.8587 | -17.5309 | 9.4305 | -129.7425 | 131.3455 | 0.11937 | 1.8077 | 0.34107 | 1.8503 | 129.7425 | -1.8077 | 79.8345
0.8605 | -17.0108 | 9.4988 | -129.6573 | 131.2006 | 0.093794 | 1.8077 | 0.3496 | 1.8503 | 129.6573 | -1.8077 | 80.0019
0.8623 | -16.5674 | 9.6181 | -129.9472 | 131.4393 | 0.068214 | 1.8162 | 0.3496 | 1.8503 | 129.9472 | -1.8162 | 80.1692
0.8641 | -16.0814 | 9.7546 | -130.3394 | 131.7719 | 0.034107 | 1.8162 | 0.35812 | 1.8588 | 130.3394 | -1.8162 | 80.3366
0.8659 | -15.5101 | 9.9166 | -130.9448 | 132.3176 | 0.017053 | 1.8162 | 0.35812 | 1.8588 | 130.9448 | -1.8162 | 80.5039
0.8677 | -14.9985 | 10.1297 | -131.5758 | 132.8974 | -0.0085267 | 1.8077 | 0.36665 | 1.8503 | 131.5758 | -1.8077 | 80.6713
0.8695 | -14.4869 | 10.3429 | -132.2153 | 133.4858 | -0.02558 | 1.7991 | 0.37518 | 1.8418 | 132.2153 | -1.7991 | 80.8386
0.8712 | -14.0606 | 10.6754 | -133.102 | 134.3469 | -0.034107 | 1.7991 | 0.3837 | 1.8418 | 133.102 | -1.7991 | 80.9967
0.873 | -13.6683 | 11.008 | -134.3128 | 135.5236 | -0.042634 | 1.7906 | 0.39223 | 1.8418 | 134.3128 | -1.7906 | 81.164
0.8748 | -13.4722 | 11.4514 | -135.643 | 136.8538 | -0.042634 | 1.7991 | 0.39223 | 1.8503 | 135.643 | -1.7991 | 81.3313
0.8766 | -13.4125 | 12.0056 | -137.1352 | 138.3716 | -0.034107 | 1.8077 | 0.40076 | 1.8588 | 137.1352 | -1.8077 | 81.4987
0.8784 | -13.3102 | 12.6281 | -138.8235 | 140.0854 | -0.02558 | 1.8247 | 0.40928 | 1.8759 | 138.8235 | -1.8247 | 81.666
0.8802 | -13.3784 | 13.2761 | -140.7249 | 142.0466 | -0.02558 | 1.8332 | 0.41781 | 1.8929 | 140.7249 | -1.8332 | 81.8334
0.882 | -13.4466 | 13.8985 | -142.7117 | 144.0845 | -0.017053 | 1.8503 | 0.41781 | 1.91 | 142.7117 | -1.8503 | 82.0007
0.8838 | -13.7024 | 14.5807 | -144.8348 | 146.2844 | -0.0085267 | 1.8759 | 0.42634 | 1.9356 | 144.8348 | -1.8759 | 82.1681
0.8856 | -13.8559 | 15.2202 | -146.9068 | 148.416 | -0.0085267 | 1.9015 | 0.43486 | 1.9611 | 146.9068 | -1.9015 | 82.3354
0.8874 | -13.9412 | 15.8256 | -149.0896 | 150.65 | 0 | 1.9185 | 0.44339 | 1.9782 | 149.0896 | -1.9185 | 82.5028
0.8892 | -14.1288 | 16.3883 | -151.3919 | 153.0119 | 0.0085267 | 1.9356 | 0.44339 | 1.9953 | 151.3919 | -1.9356 | 82.6701
0.891 | -14.2737 | 16.8658 | -153.5576 | 155.2204 | 0.02558 | 1.9441 | 0.45192 | 2.0038 | 153.5576 | -1.9441 | 82.8375
0.8928 | -14.3419 | 17.2325 | -155.2204 | 156.9086 | 0.02558 | 1.9441 | 0.45192 | 2.0038 | 155.2204 | -1.9441 | 83.0048
0.8946 | -14.4613 | 17.5565 | -156.7466 | 158.469 | 0.042634 | 1.9441 | 0.45192 | 2.0038 | 156.7466 | -1.9441 | 83.1722
0.8963 | -14.5722 | 17.8891 | -158.3411 | 160.0976 | 0.05116 | 1.9441 | 0.44339 | 2.0038 | 158.3411 | -1.9441 | 83.3302
0.8981 | -14.828 | 18.2045 | -159.748 | 161.5472 | 0.059687 | 1.9526 | 0.44339 | 2.0123 | 159.748 | -1.9526 | 83.4976
0.8999 | -14.9814 | 18.4348 | -160.8906 | 162.7068 | 0.068214 | 1.9441 | 0.43486 | 2.0038 | 160.8906 | -1.9441 | 83.6649
0.9017 | -15.0667 | 18.52 | -161.7859 | 163.6106 | 0.085267 | 1.927 | 0.41781 | 1.9782 | 161.7859 | -1.927 | 83.8323
0.9035 | -15.1008 | 18.4433 | -162.4851 | 164.3013 | 0.085267 | 1.9015 | 0.40928 | 1.9611 | 162.4851 | -1.9015 | 83.9996
0.9053 | -14.9473 | 18.2472 | -162.8347 | 164.6168 | 0.093794 | 1.8674 | 0.39223 | 1.927 | 162.8347 | -1.8674 | 84.167
0.9071 | -14.6915 | 17.9828 | -163.1076 | 164.8385 | 0.093794 | 1.8332 | 0.37518 | 1.8929 | 163.1076 | -1.8332 | 84.3343
0.9089 | -14.3505 | 17.6588 | -163.3548 | 165.0176 | 0.093794 | 1.7821 | 0.35812 | 1.8418 | 163.3548 | -1.7821 | 84.5017
0.9107 | -14.0435 | 17.2922 | -163.4998 | 165.0943 | 0.085267 | 1.7395 | 0.34107 | 1.7991 | 163.4998 | -1.7395 | 84.669
0.9125 | -13.5831 | 16.7976 | -163.4316 | 164.9493 | 0.085267 | 1.6798 | 0.32402 | 1.7395 | 163.4316 | -1.6798 | 84.8364
0.9143 | -13.2079 | 16.2008 | -163.1587 | 164.5997 | 0.085267 | 1.6201 | 0.30696 | 1.6712 | 163.1587 | -1.6201 | 85.0037
0.9161 | -12.9436 | 15.6636 | -162.8859 | 164.2672 | 0.07674 | 1.5604 | 0.28991 | 1.603 | 162.8859 | -1.5604 | 85.1711
0.9179 | -12.8327 | 15.1861 | -162.758 | 164.0967 | 0.085267 | 1.5007 | 0.26433 | 1.5519 | 162.758 | -1.5007 | 85.3384
0.9197 | -12.6025 | 14.5551 | -162.4766 | 163.7556 | 0.085267 | 1.4495 | 0.24727 | 1.5007 | 162.4766 | -1.4495 | 85.5058
0.9214 | -12.125 | 13.8133 | -162.0673 | 163.2696 | 0.07674 | 1.3899 | 0.23022 | 1.4495 | 162.0673 | -1.3899 | 85.6638
0.9232 | -11.6475 | 13.0373 | -161.3596 | 162.4766 | 0.07674 | 1.3387 | 0.21317 | 1.3984 | 161.3596 | -1.3387 | 85.8312
0.925 | -11.1529 | 12.3893 | -160.4302 | 161.4704 | 0.068214 | 1.279 | 0.19611 | 1.3557 | 160.4302 | -1.279 | 85.9985
0.9268 | -10.7437 | 11.801 | -159.2535 | 160.2255 | 0.068214 | 1.2278 | 0.17906 | 1.3216 | 159.2535 | -1.2278 | 86.1659
0.9286 | -10.1809 | 11.2041 | -158.7163 | 159.6116 | 0.068214 | 1.1596 | 0.16201 | 1.262 | 158.7163 | -1.1596 | 86.3332
0.9304 | -9.5414 | 10.5731 | -158.4861 | 159.2791 | 0.059687 | 1.0829 | 0.14495 | 1.1852 | 158.4861 | -1.0829 | 86.5006
0.9322 | -9.0895 | 10.0701 | -158.5287 | 159.245 | 0.068214 | 1.0147 | 0.1279 | 1.117 | 158.5287 | -1.0147 | 86.6679
0.934 | -8.8252 | 9.7375 | -158.7504 | 159.4155 | 0.085267 | 0.95499 | 0.11085 | 1.0744 | 158.7504 | -0.95499 | 86.8353
0.9358 | -8.7143 | 9.5243 | -159.1512 | 159.7821 | 0.093794 | 0.90383 | 0.10232 | 1.0317 | 159.1512 | -0.90383 | 87.0026
0.9376 | -8.6973 | 9.3538 | -159.8589 | 160.4813 | 0.11085 | 0.8612 | 0.093794 | 0.99763 | 159.8589 | -0.8612 | 87.17
0.9394 | -8.7058 | 9.1918 | -160.6348 | 161.2402 | 0.1279 | 0.81856 | 0.093794 | 0.96352 | 160.6348 | -0.81856 | 87.3373
0.9412 | -8.8678 | 9.0724 | -161.6324 | 162.2378 | 0.14495 | 0.80151 | 0.085267 | 0.95499 | 161.6324 | -0.80151 | 87.5046
0.943 | -9.1065 | 8.936 | -162.7836 | 163.3889 | 0.17053 | 0.78446 | 0.07674 | 0.94647 | 162.7836 | -0.78446 | 87.672
0.9447 | -9.26 | 8.6973 | -163.9432 | 164.5401 | 0.18759 | 0.75035 | 0.07674 | 0.92941 | 163.9432 | -0.75035 | 87.83
0.9465 | -9.2344 | 8.3477 | -164.83 | 165.4013 | 0.19611 | 0.70772 | 0.068214 | 0.90383 | 164.83 | -0.70772 | 87.9974
0.9483 | -9.0383 | 7.9043 | -165.2734 | 165.802 | 0.20464 | 0.65656 | 0.059687 | 0.89531 | 165.2734 | -0.65656 | 88.1647
0.9501 | -8.6546 | 7.2903 | -165.333 | 165.8276 | 0.20464 | 0.59687 | 0.05116 | 0.89531 | 165.333 | -0.59687 | 88.3321
0.9519 | -8.2453 | 6.6935 | -165.0346 | 165.4865 | 0.20464 | 0.54571 | 0.05116 | 0.92089 | 165.0346 | -0.54571 | 88.4994
0.9537 | -8.0066 | 6.3865 | -165.0176 | 165.4609 | 0.20464 | 0.49455 | 0.05116 | 0.93794 | 165.0176 | -0.49455 | 88.6668
0.9555 | -8.0322 | 6.1222 | -164.6765 | 165.1113 | 0.21317 | 0.4775 | 0.05116 | 0.95499 | 164.6765 | -0.4775 | 88.8341
0.9573 | -7.9895 | 5.9175 | -164.8044 | 165.2392 | 0.22169 | 0.45192 | 0.05116 | 0.94647 | 164.8044 | -0.45192 | 89.0015
0.9591 | -7.9384 | 5.8408 | -165.0687 | 165.495 | 0.23022 | 0.42634 | 0.042634 | 0.92941 | 165.0687 | -0.42634 | 89.1688
0.9609 | -7.964 | 5.9858 | -165.6741 | 166.109 | 0.24727 | 0.40076 | 0.042634 | 0.92941 | 165.6741 | -0.40076 | 89.3362
0.9627 | -8.0577 | 6.2842 | -166.6376 | 167.0725 | 0.26433 | 0.36665 | 0.042634 | 0.92941 | 166.6376 | -0.36665 | 89.5035
0.9645 | -8.2539 | 6.6935 | -168.1724 | 168.6329 | 0.28138 | 0.34107 | 0.042634 | 0.92941 | 168.1724 | -0.34107 | 89.6709
0.9663 | -8.4926 | 7.2818 | -170.0824 | 170.5684 | 0.30696 | 0.31549 | 0.042634 | 0.93794 | 170.0824 | -0.31549 | 89.8382
0.9681 | -8.7143 | 7.8019 | -172.1629 | 172.666 | 0.32402 | 0.29844 | 0.042634 | 0.93794 | 172.1629 | -0.29844 | 90.0056
0.9698 | -8.8678 | 8.2624 | -174.2861 | 174.8147 | 0.34107 | 0.28991 | 0.034107 | 0.93794 | 174.2861 | -0.28991 | 90.1636
0.9716 | -9.0724 | 8.7314 | -176.6565 | 177.2022 | 0.35812 | 0.26433 | 0.034107 | 0.92941 | 176.6565 | -0.26433 | 90.331
0.9734 | -9.1577 | 9.2344 | -179.4789 | 180.0502 | 0.37518 | 0.22169 | 0.02558 | 0.92941 | 179.4789 | -0.22169 | 90.4983
0.9752 | -9.1236 | 9.7119 | -182.5826 | 183.1709 | 0.3837 | 0.17906 | 0.017053 | 0.92089 | 182.5826 | -0.17906 | 90.6657
0.977 | -8.9872 | 10.1383 | -185.7375 | 186.3429 | 0.39223 | 0.14495 | 0.0085267 | 0.92089 | 185.7375 | -0.14495 | 90.833
0.9788 | -8.7228 | 10.6584 | -189.0885 | 189.6939 | 0.40076 | 0.093794 | 0 | 0.92941 | 189.0885 | -0.093794 | 91.0004
0.9806 | -8.5011 | 11.3661 | -192.8061 | 193.42 | 0.40928 | 0.034107 | -0.0085267 | 0.93794 | 192.8061 | -0.034107 | 91.1677
0.9824 | -8.3135 | 12.2017 | -196.8734 | 197.5129 | 0.41781 | -0.02558 | -0.017053 | 0.95499 | 196.8734 | 0.02558 | 91.3351
0.9842 | -8.126 | 13.1311 | -201.2049 | 201.8785 | 0.42634 | -0.093794 | -0.034107 | 0.97205 | 201.2049 | 0.093794 | 91.5024
0.986 | -8.0236 | 13.8985 | -205.4939 | 206.2016 | 0.43486 | -0.17053 | -0.042634 | 0.96352 | 205.4939 | 0.17053 | 91.6698
0.9878 | -7.7337 | 14.6063 | -209.8254 | 210.5502 | 0.43486 | -0.24727 | -0.059687 | 0.97205 | 209.8254 | 0.24727 | 91.8371
0.9896 | -7.4438 | 15.4589 | -214.464 | 215.2143 | 0.44339 | -0.33254 | -0.07674 | 0.9891 | 214.464 | 0.33254 | 92.0045
0.9914 | -7.0942 | 16.4651 | -219.6397 | 220.4327 | 0.44339 | -0.42634 | -0.093794 | 1.0147 | 219.6397 | 0.42634 | 92.1718
0.9932 | -6.6338 | 17.4883 | -225.3697 | 226.2053 | 0.45192 | -0.52866 | -0.10232 | 1.0317 | 225.3697 | 0.52866 | 92.3392
0.9949 | -6.1904 | 18.6479 | -231.5089 | 232.3957 | 0.45192 | -0.63098 | -0.11937 | 1.0744 | 231.5089 | 0.63098 | 92.4972
0.9967 | -5.9431 | 19.9781 | -237.8613 | 238.8333 | 0.46897 | -0.72477 | -0.13643 | 1.1255 | 237.8613 | 0.72477 | 92.6646
0.9985 | -5.9261 | 21.4703 | -244.5036 | 245.5694 | 0.4775 | -0.81856 | -0.15348 | 1.1852 | 244.5036 | 0.81856 | 92.8319
1.0003 | -5.9772 | 23.0903 | -251.4358 | 252.621 | 0.49455 | -0.90383 | -0.17053 | 1.2534 | 251.4358 | 0.90383 | 92.9993
1.0021 | -5.892 | 24.7275 | -258.7859 | 260.0904 | 0.49455 | -1.0062 | -0.19611 | 1.3302 | 258.7859 | 1.0062 | 93.1666
1.0039 | -5.6021 | 26.2879 | -266.5366 | 267.9521 | 0.50308 | -1.1341 | -0.22169 | 1.424 | 266.5366 | 1.1341 | 93.334
1.0057 | -5.2269 | 27.8312 | -274.5091 | 276.0269 | 0.50308 | -1.262 | -0.24727 | 1.5178 | 274.5091 | 1.262 | 93.5013
1.0075 | -4.7579 | 29.2807 | -282.6948 | 284.3148 | 0.50308 | -1.3899 | -0.28138 | 1.6201 | 282.6948 | 1.3899 | 93.6687
1.0093 | -4.2037 | 30.6621 | -290.8548 | 292.5602 | 0.48602 | -1.5348 | -0.31549 | 1.7395 | 290.8548 | 1.5348 | 93.836
1.0111 | -3.7091 | 32.2054 | -298.9723 | 300.7885 | 0.4775 | -1.6798 | -0.3496 | 1.8674 | 298.9723 | 1.6798 | 94.0033
1.0129 | -3.1464 | 33.7658 | -307.8656 | 309.7841 | 0.46044 | -1.8418 | -0.3837 | 2.0123 | 307.8656 | 1.8418 | 94.1707
1.0147 | -2.6092 | 35.3944 | -317.1427 | 319.1806 | 0.45192 | -1.9953 | -0.42634 | 2.1573 | 317.1427 | 1.9953 | 94.338
1.0165 | -2.0123 | 36.9974 | -326.7864 | 328.9437 | 0.43486 | -2.1658 | -0.46044 | 2.3107 | 326.7864 | 2.1658 | 94.5054
1.0182 | -1.3131 | 38.5663 | -336.6689 | 338.937 | 0.40928 | -2.3278 | -0.5116 | 2.4727 | 336.6689 | 2.3278 | 94.6634
1.02 | -0.68214 | 40.1694 | -346.4234 | 348.8195 | 0.3837 | -2.4898 | -0.55424 | 2.6262 | 346.4234 | 2.4898 | 94.8308
1.0218 | -0.11937 | 41.7127 | -356.0586 | 358.5655 | 0.35812 | -2.6518 | -0.59687 | 2.7882 | 356.0586 | 2.6518 | 94.9981
1.0236 | 0.55424 | 43.1452 | -365.6768 | 368.2945 | 0.32402 | -2.8223 | -0.64803 | 2.9502 | 365.6768 | 2.8223 | 95.1655
1.0254 | 1.3302 | 44.518 | -375.3461 | 378.0576 | 0.28991 | -2.9929 | -0.69066 | 3.1293 | 375.3461 | 2.9929 | 95.3328
1.0272 | 2.1828 | 45.7714 | -385.1347 | 387.9315 | 0.24727 | -3.1719 | -0.74182 | 3.2998 | 385.1347 | 3.1719 | 95.5002
1.029 | 3.2572 | 46.9737 | -395.3839 | 398.2744 | 0.20464 | -3.3425 | -0.80151 | 3.4789 | 395.3839 | 3.3425 | 95.6675
1.0308 | 4.468 | 48.0907 | -405.9229 | 408.8902 | 0.17053 | -3.5045 | -0.85267 | 3.6494 | 405.9229 | 3.5045 | 95.8349
1.0326 | 5.6958 | 49.2418 | -416.965 | 420.0175 | 0.1279 | -3.675 | -0.91236 | 3.8285 | 416.965 | 3.675 | 96.0022
1.0344 | 6.864 | 50.6743 | -428.8256 | 431.9805 | 0.093794 | -3.8711 | -0.96352 | 4.0161 | 428.8256 | 3.8711 | 96.1696
1.0362 | 8.1089 | 51.7742 | -439.9018 | 443.1249 | 0.05116 | -4.0502 | -1.0232 | 4.2122 | 439.9018 | 4.0502 | 96.3369
1.038 | 9.3112 | 52.6525 | -450.2874 | 453.5616 | 0.017053 | -4.2122 | -1.0829 | 4.3827 | 450.2874 | 4.2122 | 96.5043
1.0398 | 10.5561 | 53.2749 | -459.8629 | 463.1798 | -0.0085267 | -4.3742 | -1.1255 | 4.5447 | 459.8629 | 4.3742 | 96.6716
1.0416 | 12.0141 | 53.6416 | -468.995 | 472.329 | -0.02558 | -4.5362 | -1.1682 | 4.7067 | 468.995 | 4.5362 | 96.839
1.0433 | 13.5404 | 53.8889 | -477.9907 | 481.3417 | -0.042634 | -4.6812 | -1.2023 | 4.8602 | 477.9907 | 4.6812 | 96.997
1.0451 | 14.9303 | 54.3237 | -486.8755 | 490.2606 | -0.068214 | -4.8176 | -1.2449 | 5.0052 | 486.8755 | 4.8176 | 97.1644
1.0469 | 16.448 | 54.3919 | -495.2829 | 498.6595 | -0.085267 | -4.937 | -1.2961 | 5.1331 | 495.2829 | 4.937 | 97.3317
1.0487 | 17.7953 | 54.3322 | -502.812 | 506.1715 | -0.10232 | -5.0563 | -1.3387 | 5.2525 | 502.812 | 5.0563 | 97.4991
1.0505 | 19.1169 | 54.0594 | -509.8295 | 513.1549 | -0.11085 | -5.1757 | -1.3728 | 5.3804 | 509.8295 | 5.1757 | 97.6664
1.0523 | 20.3959 | 53.4966 | -516.1222 | 519.3794 | -0.093794 | -5.2695 | -1.4069 | 5.4827 | 516.1222 | 5.2695 | 97.8338
1.0541 | 21.7772 | 52.7719 | -522.1591 | 525.3566 | -0.068214 | -5.3548 | -1.4325 | 5.5679 | 522.1591 | 5.3548 | 98.0011
1.0559 | 23.15 | 51.9874 | -527.9061 | 531.0439 | -0.017053 | -5.4145 | -1.4581 | 5.6362 | 527.9061 | 5.4145 | 98.1685
1.0577 | 24.3523 | 51.22 | -532.971 | 536.0491 | 0.034107 | -5.4571 | -1.4836 | 5.6788 | 532.971 | 5.4571 | 98.3358
1.0595 | 24.8639 | 50.103 | -534.7104 | 537.7033 | 0.1279 | -5.4486 | -1.4836 | 5.6788 | 534.7104 | 5.4486 | 98.5032
1.0613 | 24.6252 | 49.0627 | -534.4631 | 537.3537 | 0.22169 | -5.4145 | -1.4666 | 5.6447 | 534.4631 | 5.4145 | 98.6705
1.0631 | 23.8663 | 47.9116 | -532.1694 | 534.9406 | 0.31549 | -5.3207 | -1.4495 | 5.5594 | 532.1694 | 5.3207 | 98.8379
1.0649 | 23.1245 | 46.7861 | -529.799 | 532.4508 | 0.40928 | -5.2354 | -1.4325 | 5.4827 | 529.799 | 5.2354 | 99.0052
1.0667 | 22.4338 | 45.5668 | -527.2922 | 529.8246 | 0.49455 | -5.1416 | -1.4154 | 5.3974 | 527.2922 | 5.1416 | 99.1726
1.0684 | 21.4021 | 44.1343 | -523.0629 | 525.4419 | 0.57129 | -5.0052 | -1.3813 | 5.2695 | 523.0629 | 5.0052 | 99.3306
1.0702 | 20.3703 | 42.7018 | -518.8592 | 521.0932 | 0.64803 | -4.8688 | -1.3557 | 5.1331 | 518.8592 | 4.8688 | 99.498
1.072 | 19.4921 | 41.2778 | -515.4656 | 517.5717 | 0.72477 | -4.7323 | -1.3216 | 5.0052 | 515.4656 | 4.7323 | 99.6653
1.0738 | 18.8014 | 39.8965 | -513.4874 | 515.4656 | 0.81004 | -4.6215 | -1.2875 | 4.9029 | 513.4874 | 4.6215 | 99.8327
1.0756 | 18.1363 | 38.464 | -512.2937 | 514.1525 | 0.88678 | -4.5192 | -1.2534 | 4.8091 | 512.2937 | 4.5192 | 100
| Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My
Minimum | -70.5756 | -138.7808 | -1960.0364 | 128.0031 | -0.11085 | -20.9416 | -9.3964 | 0.89531 | 126.528 | -8.2283
Maximum | 24.8639 | 54.3919 | -126.528 | 1960.71 | 14.2055 | 8.2283 | 1.2193 | 25.1027 | 1960.0364 | 20.9416

## Diagr_AVER75
AVER75 = Loads at Average Level in Subject with 75 kg Body Weight.
Contact forces and moments, measured telemetrically in knee implants of 8 subjects.
Averages of one individual subject.
Fx  lateral, Fy  anterior, -Fz  distal, Fres = Resultant Force
Mx around x, My around y, Mz around z, Mres = Resultant Moment
z-axis = axis of implant stem!
From Bergmann et al, PlosOne (2014) "Standardized Loads Acting in Knee Implants"
Subject K5R
Subject K5R
G.Bergmann et al., PlosOne, 2014

## Dat_HIGH100
Data File: | k5r_2242_0_13_i1.akf
Curve Alignment Program: | 1.5.7
------------------------MEASURING INFOS------------------------
Diagram Title #1 | Forces and Moments at Knee Joint
Diagram Title #2 | Measured Relative to Implant
Comment #1 | Knee Joint; Gaitanalysis; Level Walking;
Comment #2 | K5R, 30 Months PO *
BodyWeight [N]: | 942
Patient(en): | k5r | [7 empty] | Subject | k5r
Implant Type: | Knee Joint
Varus Angle Leg [°] | 1 | [7 empty] | HIGH100 Loads in N (forces) and Nm (moments)
Axes Directions X, Y, Z: | Lateral, Anterior, Superior
Axes Origin: | Tibial Plateau at Stem Axis | [7 empty] | = AVER75 Loads * 1/Caver
Transformation X Y Z (°): | 7 | -2 | -1
Displacement X, Y, Z (mm): | 0 | 0 | 0 | [5 empty] | Caver = | 0.58
Activity Code: | KneeActivities | 1P_1_15_2
---------------------------DATA--------------------------------
Number of Data Sets = | 601
Number of analog channels = | 0
Time-Offset(sec) = | 0
Time | Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My | Time
[s] | [N] | [N] | [N] | [N] | [Nm] | [Nm] | [Nm] | [Nm] | [N] | (Nm] | [%Cycle]
0 | 17.0828 | 42.8688 | -886.8668 | 888.2781 | 2.7638 | -5.6453 | -1.5436 | 6.6891 | 886.8668 | 5.6453 | 0
0.0018 | 15.4951 | 39.958 | -884.7351 | 885.9259 | 2.8961 | -5.4248 | -1.4701 | 6.5273 | 884.7351 | 5.4248 | 0.16735
0.0036 | 13.7604 | 37.3411 | -883.9265 | 884.9556 | 3.0285 | -5.1748 | -1.3966 | 6.3362 | 883.9265 | 5.1748 | 0.3347
0.0054 | 12.3784 | 34.9301 | -884.9997 | 885.8965 | 3.1461 | -4.969 | -1.3231 | 6.2186 | 884.9997 | 4.969 | 0.50205
0.0072 | 11.1435 | 32.6514 | -888.4545 | 889.2484 | 3.2784 | -4.7926 | -1.2496 | 6.1451 | 888.4545 | 4.7926 | 0.66939
0.009 | 9.9674 | 30.5051 | -895.8051 | 896.5255 | 3.4107 | -4.6309 | -1.1761 | 6.1304 | 895.8051 | 4.6309 | 0.83674
0.0108 | 9.2471 | 28.197 | -908.7422 | 909.3743 | 3.5871 | -4.528 | -1.1026 | 6.1745 | 908.7422 | 4.528 | 1.0041
0.0125 | 8.5561 | 25.7272 | -925.2958 | 925.8397 | 3.7635 | -4.4398 | -1.0291 | 6.2039 | 925.2958 | 4.4398 | 1.1621
0.0143 | 7.9387 | 23.4926 | -944.672 | 945.1572 | 3.9546 | -4.381 | -0.95558 | 6.2774 | 944.672 | 4.381 | 1.3295
0.0161 | 7.5123 | 21.0522 | -970.4139 | 970.8402 | 4.1752 | -4.381 | -0.85267 | 6.4244 | 970.4139 | 4.381 | 1.4968
0.0179 | 7.1889 | 18.5236 | -1000.0516 | 1000.4338 | 4.3957 | -4.381 | -0.77917 | 6.5862 | 1000.0516 | 4.381 | 1.6642
0.0197 | 6.9978 | 15.9067 | -1032.6148 | 1032.9529 | 4.6309 | -4.3957 | -0.70566 | 6.7773 | 1032.6148 | 4.3957 | 1.8315
0.0215 | 7.0566 | 13.2752 | -1068.0742 | 1068.3682 | 4.8808 | -4.4545 | -0.61745 | 6.9978 | 1068.0742 | 4.4545 | 1.9989
0.0233 | 7.1889 | 10.9818 | -1104.3421 | 1104.6215 | 5.1748 | -4.528 | -0.54395 | 7.2477 | 1104.3421 | 4.528 | 2.1662
0.0251 | 7.2624 | 8.9384 | -1139.4928 | 1139.7427 | 5.4689 | -4.6162 | -0.45574 | 7.527 | 1139.4928 | 4.6162 | 2.3336
0.0269 | 7.3947 | 6.8802 | -1171.3945 | 1171.6297 | 5.7482 | -4.7044 | -0.38223 | 7.8064 | 1171.3945 | 4.7044 | 2.5009
0.0287 | 7.7329 | 4.2928 | -1200.4735 | 1200.7087 | 6.0128 | -4.7926 | -0.32343 | 8.071 | 1200.4735 | 4.7926 | 2.6683
0.0305 | 8.1592 | 0.83797 | -1224.9511 | 1225.201 | 6.2627 | -4.8808 | -0.27932 | 8.2915 | 1224.9511 | 4.8808 | 2.8356
0.0323 | 8.4385 | -3.4695 | -1246.3414 | 1246.6207 | 6.4979 | -4.9249 | -0.23522 | 8.4973 | 1246.3414 | 4.9249 | 3.003
0.0341 | 8.7178 | -7.9093 | -1266.1733 | 1266.5115 | 6.7332 | -4.9543 | -0.17641 | 8.6737 | 1266.1733 | 4.9543 | 3.1703
0.0359 | 9.1001 | -12.349 | -1285.0791 | 1285.4614 | 6.9537 | -4.9543 | -0.13231 | 8.8354 | 1285.0791 | 4.9543 | 3.3377
0.0376 | 10.1145 | -16.2596 | -1308.1013 | 1308.5276 | 7.233 | -5.0131 | -0.088207 | 9.0707 | 1308.1013 | 5.0131 | 3.4957
0.0394 | 11.217 | -20.1554 | -1331.9761 | 1332.4465 | 7.5123 | -5.0719 | -0.058805 | 9.3353 | 1331.9761 | 5.0719 | 3.6631
0.0412 | 12.6284 | -24.2129 | -1359.2322 | 1359.8055 | 7.8505 | -5.2189 | -0.014701 | 9.6734 | 1359.2322 | 5.2189 | 3.8304
0.043 | 14.2455 | -28.3293 | -1388.0172 | 1388.6934 | 8.2033 | -5.366 | 0 | 10.0409 | 1388.0172 | 5.366 | 3.9978
0.0448 | 15.7744 | -32.2986 | -1417.7137 | 1418.4928 | 8.5708 | -5.5571 | 0.029402 | 10.4526 | 1417.7137 | 5.5571 | 4.1651
0.0466 | 16.9946 | -36.3268 | -1447.3073 | 1448.1893 | 8.9678 | -5.7629 | 0.058805 | 10.8789 | 1447.3073 | 5.7629 | 4.3325
0.0484 | 18.0531 | -40.4872 | -1476.5333 | 1477.533 | 9.3647 | -6.0275 | 0.10291 | 11.3494 | 1476.5333 | 6.0275 | 4.4998
0.0502 | 19.1851 | -44.9564 | -1507.1266 | 1508.288 | 9.7763 | -6.3215 | 0.14701 | 11.8492 | 1507.1266 | 6.3215 | 4.6672
0.052 | 20.0525 | -49.3373 | -1536.485 | 1537.7934 | 10.1733 | -6.6009 | 0.19112 | 12.3343 | 1536.485 | 6.6009 | 4.8345
0.0538 | 20.8169 | -53.7183 | -1564.1086 | 1565.564 | 10.5555 | -6.939 | 0.24992 | 12.8195 | 1564.1086 | 6.939 | 5.0019
0.0556 | 21.1992 | -58.0111 | -1589.0713 | 1590.6737 | 10.8936 | -7.3212 | 0.29402 | 13.2899 | 1589.0713 | 7.3212 | 5.1692
0.0574 | 21.3903 | -62.3479 | -1612.9314 | 1614.6661 | 11.2023 | -7.7181 | 0.32343 | 13.7604 | 1612.9314 | 7.7181 | 5.3366
0.0592 | 21.4638 | -66.6848 | -1635.9388 | 1637.7912 | 11.5111 | -8.1151 | 0.36753 | 14.2308 | 1635.9388 | 8.1151 | 5.5039
0.061 | 21.6696 | -71.2716 | -1659.7989 | 1661.7983 | 11.8198 | -8.5708 | 0.39693 | 14.7453 | 1659.7989 | 8.5708 | 5.6713
0.0627 | 22.096 | -76.1671 | -1686.3053 | 1688.4957 | 12.1579 | -9.0854 | 0.44104 | 15.3187 | 1686.3053 | 9.0854 | 5.8293
0.0645 | 22.6546 | -81.239 | -1715.693 | 1718.0893 | 12.5255 | -9.6146 | 0.48514 | 15.9214 | 1715.693 | 9.6146 | 5.9967
0.0663 | 23.3897 | -86.6197 | -1746.4186 | 1749.0501 | 12.893 | -10.188 | 0.52924 | 16.5683 | 1746.4186 | 10.188 | 6.164
0.0681 | 24.2717 | -92.3091 | -1778.2615 | 1781.1576 | 13.2605 | -10.776 | 0.58805 | 17.2445 | 1778.2615 | 10.776 | 6.3313
0.0699 | 25.0362 | -97.8955 | -1811.4716 | 1814.647 | 13.6574 | -11.3935 | 0.63215 | 17.9355 | 1811.4716 | 11.3935 | 6.4987
0.0717 | 25.8448 | -103.6143 | -1845.6373 | 1849.1067 | 14.0397 | -12.055 | 0.67626 | 18.6559 | 1845.6373 | 12.055 | 6.666
0.0735 | 26.6092 | -109.2596 | -1880.7291 | 1884.4779 | 14.4219 | -12.7313 | 0.69096 | 19.3762 | 1880.7291 | 12.7313 | 6.8334
0.0753 | 27.4619 | -114.8755 | -1916.4237 | 1920.4518 | 14.8041 | -13.4075 | 0.73506 | 20.126 | 1916.4237 | 13.4075 | 7.0007
0.0771 | 28.491 | -120.756 | -1952.5299 | 1956.8815 | 15.2011 | -14.0838 | 0.76446 | 20.8611 | 1952.5299 | 14.0838 | 7.1681
0.0789 | 29.079 | -126.313 | -1989.7976 | 1994.4285 | 15.6127 | -14.76 | 0.79387 | 21.6108 | 1989.7976 | 14.76 | 7.3354
0.0807 | 29.6524 | -131.7525 | -2028.7705 | 2033.6955 | 16.039 | -15.4804 | 0.80857 | 22.4194 | 2028.7705 | 15.4804 | 7.5028
0.0825 | 29.8729 | -137.1184 | -2069.243 | 2074.4326 | 16.4654 | -16.2449 | 0.83797 | 23.2574 | 2069.243 | 16.2449 | 7.6701
0.0843 | 30.1669 | -142.5873 | -2110.4506 | 2115.9048 | 16.8917 | -17.0534 | 0.83797 | 24.1394 | 2110.4506 | 17.0534 | 7.8375
0.086 | 30.5345 | -148.0561 | -2152.5403 | 2158.2443 | 17.3475 | -17.862 | 0.85267 | 25.0362 | 2152.5403 | 17.862 | 7.9955
0.0878 | 30.8579 | -153.5103 | -2192.6158 | 2198.6139 | 17.7738 | -18.6412 | 0.86737 | 25.8889 | 2192.6158 | 18.6412 | 8.1629
0.0896 | 31.3136 | -158.5675 | -2231.3095 | 2237.5575 | 18.1854 | -19.4497 | 0.86737 | 26.7562 | 2231.3095 | 19.4497 | 8.3302
0.0914 | 31.5342 | -163.6101 | -2268.3566 | 2274.8692 | 18.5677 | -20.1995 | 0.86737 | 27.5648 | 2268.3566 | 20.1995 | 8.4976
0.0932 | 31.7694 | -168.6526 | -2304.6539 | 2311.4312 | 18.9352 | -20.9787 | 0.86737 | 28.3881 | 2304.6539 | 20.9787 | 8.6649
0.095 | 31.9311 | -173.7539 | -2340.4073 | 2347.4639 | 19.3027 | -21.7431 | 0.88207 | 29.1967 | 2340.4073 | 21.7431 | 8.8323
0.0968 | 31.9752 | -178.8258 | -2375.0729 | 2382.4235 | 19.6409 | -22.4782 | 0.88207 | 29.9758 | 2375.0729 | 22.4782 | 8.9996
0.0986 | 31.9164 | -183.8537 | -2408.9886 | 2416.6186 | 19.979 | -23.2133 | 0.88207 | 30.755 | 2408.9886 | 23.2133 | 9.167
0.1004 | 31.8429 | -188.7492 | -2442.184 | 2450.0933 | 20.3024 | -23.9189 | 0.88207 | 31.5047 | 2442.184 | 23.9189 | 9.3343
0.1022 | 31.74 | -193.5124 | -2474.9825 | 2483.171 | 20.6111 | -24.6246 | 0.89678 | 32.2398 | 2474.9825 | 24.6246 | 9.5017
0.104 | 31.6371 | -198.0992 | -2507.3399 | 2515.7931 | 20.9052 | -25.3008 | 0.91148 | 32.9602 | 2507.3399 | 25.3008 | 9.669
0.1058 | 31.5783 | -202.289 | -2538.7123 | 2547.4007 | 21.1992 | -25.9183 | 0.94088 | 33.6217 | 2538.7123 | 25.9183 | 9.8364
0.1076 | 31.49 | -205.9643 | -2568.7469 | 2577.6412 | 21.4638 | -26.4916 | 0.95558 | 34.2392 | 2568.7469 | 26.4916 | 10.0037
0.1094 | 31.3577 | -209.4632 | -2598.2817 | 2607.3377 | 21.7284 | -27.0209 | 0.98498 | 34.8272 | 2598.2817 | 27.0209 | 10.1711
0.1111 | 31.2107 | -212.8004 | -2627.2432 | 2636.4902 | 22.0078 | -27.5354 | 1.0144 | 35.3859 | 2627.2432 | 27.5354 | 10.3291
0.1129 | 30.9167 | -216.02 | -2655.3225 | 2664.746 | 22.2724 | -28.0647 | 1.0438 | 35.9739 | 2655.3225 | 28.0647 | 10.4965
0.1147 | 30.5345 | -218.9161 | -2682.2111 | 2691.7669 | 22.5223 | -28.5792 | 1.0732 | 36.5326 | 2682.2111 | 28.5792 | 10.6638
0.1165 | 30.0934 | -221.283 | -2707.2326 | 2716.8913 | 22.7722 | -29.0643 | 1.1173 | 37.0618 | 2707.2326 | 29.0643 | 10.8312
0.1183 | 29.6083 | -223.2236 | -2730.9163 | 2740.6485 | 22.9927 | -29.5495 | 1.1614 | 37.5911 | 2730.9163 | 29.5495 | 10.9985
0.1201 | 29.1526 | -225.0612 | -2753.7914 | 2763.5971 | 23.2133 | -30.0199 | 1.1908 | 38.1056 | 2753.7914 | 30.0199 | 11.1659
0.1219 | 28.7115 | -226.8842 | -2776.4313 | 2786.3105 | 23.4191 | -30.4904 | 1.2349 | 38.6054 | 2776.4313 | 30.4904 | 11.3332
0.1237 | 28.2558 | -228.4866 | -2797.8216 | 2807.7596 | 23.6249 | -30.9608 | 1.279 | 39.0906 | 2797.8216 | 30.9608 | 11.5006
0.1255 | 27.8441 | -229.8538 | -2817.977 | 2827.9444 | 23.8013 | -31.4165 | 1.3231 | 39.561 | 2817.977 | 31.4165 | 11.6679
0.1273 | 27.3737 | -231.1328 | -2836.6476 | 2846.6591 | 23.963 | -31.887 | 1.3672 | 40.0462 | 2836.6476 | 31.887 | 11.8353
0.1291 | 26.9768 | -232.1619 | -2852.7307 | 2862.7717 | 24.0806 | -32.328 | 1.4113 | 40.4725 | 2852.7307 | 32.328 | 12.0026
0.1309 | 26.5063 | -233.2204 | -2866.7557 | 2876.826 | 24.1835 | -32.725 | 1.4554 | 40.8547 | 2866.7557 | 32.725 | 12.17
0.1327 | 25.8595 | -234.4994 | -2879.09 | 2889.2339 | 24.2717 | -33.1072 | 1.4995 | 41.2223 | 2879.09 | 33.1072 | 12.3373
0.1345 | 25.1538 | -235.7343 | -2890.2336 | 2900.4215 | 24.3599 | -33.46 | 1.5289 | 41.5604 | 2890.2336 | 33.46 | 12.5046
0.1362 | 24.4482 | -236.7193 | -2900.2598 | 2910.5066 | 24.4335 | -33.754 | 1.5583 | 41.8397 | 2900.2598 | 33.754 | 12.6627
0.138 | 23.8013 | -237.2485 | -2909.5216 | 2919.7683 | 24.4776 | -33.9893 | 1.6024 | 42.0749 | 2909.5216 | 33.9893 | 12.83
0.1398 | 23.0074 | -237.9248 | -2915.6079 | 2925.8841 | 24.4923 | -34.2539 | 1.6318 | 42.2955 | 2915.6079 | 34.2539 | 12.9974
0.1416 | 22.096 | -238.3658 | -2920.0918 | 2930.3826 | 24.4776 | -34.5185 | 1.6759 | 42.5013 | 2920.0918 | 34.5185 | 13.1647
0.1434 | 21.4197 | -238.6158 | -2924.6198 | 2934.9106 | 24.4041 | -34.7831 | 1.72 | 42.6924 | 2924.6198 | 34.7831 | 13.3321
0.1452 | 20.6699 | -238.9686 | -2928.9566 | 2939.2475 | 24.3599 | -35.033 | 1.7641 | 42.8835 | 2928.9566 | 35.033 | 13.4994
0.147 | 19.8908 | -239.2773 | -2932.676 | 2942.9669 | 24.3158 | -35.2536 | 1.8083 | 43.0452 | 2932.676 | 35.2536 | 13.6668
0.1488 | 18.9352 | -239.1156 | -2935.1164 | 2945.3779 | 24.257 | -35.3565 | 1.8377 | 43.0893 | 2935.1164 | 35.3565 | 13.8341
0.1506 | 17.8179 | -238.7187 | -2934.6901 | 2944.9369 | 24.1394 | -35.5035 | 1.8818 | 43.1628 | 2934.6901 | 35.5035 | 14.0015
0.1524 | 16.9652 | -238.0718 | -2933.8815 | 2944.0842 | 23.9777 | -35.6652 | 1.8965 | 43.2216 | 2933.8815 | 35.6652 | 14.1688
0.1542 | 15.892 | -237.175 | -2932.7936 | 2942.9375 | 23.8307 | -35.7828 | 1.9259 | 43.251 | 2932.7936 | 35.7828 | 14.3362
0.156 | 14.6424 | -236.1313 | -2931.2941 | 2941.3939 | 23.6984 | -35.8857 | 1.9406 | 43.2657 | 2931.2941 | 35.8857 | 14.5035
0.1578 | 13.5398 | -235.4403 | -2928.7655 | 2938.8212 | 23.5514 | -35.9886 | 1.97 | 43.2804 | 2928.7655 | 35.9886 | 14.6709
0.1595 | 12.7607 | -234.7052 | -2925.4283 | 2935.4399 | 23.3603 | -36.0768 | 1.9847 | 43.251 | 2925.4283 | 36.0768 | 14.8289
0.1613 | 12.1285 | -233.7349 | -2921.3855 | 2931.3529 | 23.1544 | -36.1062 | 1.9994 | 43.1922 | 2921.3855 | 36.1062 | 14.9963
0.1631 | 11.2023 | -232.3824 | -2916.637 | 2926.5309 | 22.9486 | -36.1062 | 2.0288 | 43.0746 | 2916.637 | 36.1062 | 15.1636
0.1649 | 10.2909 | -230.9123 | -2911.6974 | 2921.5031 | 22.7281 | -36.1062 | 2.0435 | 42.957 | 2911.6974 | 36.1062 | 15.331
0.1667 | 9.3059 | -229.4863 | -2906.3167 | 2916.0489 | 22.537 | -36.1062 | 2.0582 | 42.8541 | 2906.3167 | 36.1062 | 15.4983
0.1685 | 8.1739 | -227.8104 | -2899.2013 | 2908.8306 | 22.3018 | -36.0768 | 2.0876 | 42.7218 | 2899.2013 | 36.0768 | 15.6657
0.1703 | 6.7332 | -225.2964 | -2890.6158 | 2900.0981 | 22.0225 | -35.9739 | 2.1023 | 42.4866 | 2890.6158 | 35.9739 | 15.833
0.1721 | 5.5718 | -222.6061 | -2881.9127 | 2891.2185 | 21.699 | -35.8563 | 2.1023 | 42.2219 | 2881.9127 | 35.8563 | 16.0004
0.1739 | 5.0719 | -220.2245 | -2872.8861 | 2882.045 | 21.3168 | -35.7681 | 2.0876 | 41.9573 | 2872.8861 | 35.7681 | 16.1677
0.1757 | 4.4545 | -218.0487 | -2862.1689 | 2871.1955 | 20.9493 | -35.6652 | 2.0876 | 41.6927 | 2862.1689 | 35.6652 | 16.3351
0.1775 | 4.1457 | -215.4025 | -2850.6578 | 2859.5227 | 20.567 | -35.577 | 2.0582 | 41.4281 | 2850.6578 | 35.577 | 16.5024
0.1793 | 4.1604 | -212.6387 | -2839.0145 | 2847.7176 | 20.1848 | -35.4888 | 2.0288 | 41.1929 | 2839.0145 | 35.4888 | 16.6698
0.1811 | 4.2928 | -210.0366 | -2827.327 | 2835.8831 | 19.8173 | -35.4153 | 2.0141 | 40.9576 | 2827.327 | 35.4153 | 16.8371
0.1829 | 4.2781 | -207.7285 | -2814.5663 | 2822.9754 | 19.4497 | -35.2536 | 1.9994 | 40.6342 | 2814.5663 | 35.2536 | 17.0045
0.1846 | 3.9399 | -205.5968 | -2801.0412 | 2809.3474 | 19.0528 | -35.1507 | 1.9994 | 40.3402 | 2801.0412 | 35.1507 | 17.1625
0.1864 | 3.6165 | -202.9065 | -2785.2814 | 2793.4259 | 18.6118 | -35.1066 | 1.9847 | 40.0903 | 2785.2814 | 35.1066 | 17.3299
0.1882 | 3.6753 | -200.2161 | -2769.904 | 2777.8867 | 18.2001 | -35.0183 | 1.97 | 39.7962 | 2769.904 | 35.0183 | 17.4972
0.19 | 3.7194 | -197.8786 | -2757.3932 | 2765.2437 | 17.8326 | -34.9742 | 1.9553 | 39.5904 | 2757.3932 | 34.9742 | 17.6646
0.1918 | 3.8223 | -194.8649 | -2744.9266 | 2752.6006 | 17.4504 | -34.9301 | 1.9259 | 39.3552 | 2744.9266 | 34.9301 | 17.8319
0.1936 | 4.131 | -191.9393 | -2733.6507 | 2741.1483 | 17.0534 | -34.9154 | 1.8965 | 39.1641 | 2733.6507 | 34.9154 | 17.9993
0.1954 | 4.5721 | -189.0285 | -2722.9776 | 2730.2988 | 16.6565 | -34.9007 | 1.8671 | 38.973 | 2722.9776 | 34.9007 | 18.1666
0.1972 | 5.1748 | -186.3088 | -2712.4515 | 2719.6257 | 16.3037 | -34.9154 | 1.823 | 38.826 | 2712.4515 | 34.9154 | 18.334
0.199 | 5.7482 | -183.7801 | -2701.308 | 2708.3499 | 15.9655 | -34.9301 | 1.7788 | 38.7084 | 2701.308 | 34.9301 | 18.5013
0.2008 | 6.3068 | -181.1633 | -2690.3703 | 2697.2798 | 15.6421 | -34.9448 | 1.7494 | 38.5907 | 2690.3703 | 34.9448 | 18.6687
0.2026 | 6.8214 | -178.1496 | -2680.5351 | 2687.2977 | 15.3334 | -34.9595 | 1.7053 | 38.4731 | 2680.5351 | 34.9595 | 18.836
0.2044 | 7.3653 | -175.2534 | -2671.6703 | 2678.2564 | 14.9953 | -34.9448 | 1.6906 | 38.3114 | 2671.6703 | 34.9448 | 19.0033
0.2062 | 8.218 | -172.9747 | -2665.2459 | 2671.7144 | 14.6865 | -34.9007 | 1.6465 | 38.1497 | 2665.2459 | 34.9007 | 19.1707
0.208 | 9.3206 | -171.1077 | -2660.615 | 2666.9953 | 14.4366 | -34.8272 | 1.6171 | 37.988 | 2660.615 | 34.8272 | 19.338
0.2097 | 10.5114 | -168.8437 | -2657.4101 | 2663.6728 | 14.172 | -34.7831 | 1.5877 | 37.841 | 2657.4101 | 34.7831 | 19.4961
0.2115 | 11.7022 | -166.3886 | -2654.1317 | 2660.2621 | 13.8927 | -34.7537 | 1.5436 | 37.694 | 2654.1317 | 34.7537 | 19.6634
0.2133 | 13.0253 | -163.9188 | -2647.4574 | 2653.4555 | 13.628 | -34.8272 | 1.5142 | 37.6793 | 2647.4574 | 34.8272 | 19.8308
0.2151 | 14.1279 | -161.5225 | -2639.4746 | 2645.3845 | 13.3634 | -34.8125 | 1.4701 | 37.5764 | 2639.4746 | 34.8125 | 19.9981
0.2169 | 15.1864 | -158.7145 | -2635.0054 | 2640.7977 | 13.1723 | -34.886 | 1.426 | 37.5911 | 2635.0054 | 34.886 | 20.1655
0.2187 | 16.2155 | -156.0683 | -2630.8744 | 2636.5931 | 13.0106 | -34.9448 | 1.3819 | 37.6058 | 2630.8744 | 34.9448 | 20.3328
0.2205 | 17.274 | -153.6426 | -2628.037 | 2633.6529 | 12.8783 | -35.0036 | 1.3378 | 37.6058 | 2628.037 | 35.0036 | 20.5002
0.2223 | 18.156 | -151.1875 | -2625.0968 | 2630.595 | 12.7607 | -35.0624 | 1.3084 | 37.6205 | 2625.0968 | 35.0624 | 20.6675
0.2241 | 18.5677 | -148.7618 | -2621.0539 | 2626.464 | 12.6431 | -35.136 | 1.2643 | 37.6499 | 2621.0539 | 35.136 | 20.8349
0.2259 | 18.8176 | -146.439 | -2615.394 | 2620.6864 | 12.4961 | -35.2095 | 1.2202 | 37.6646 | 2615.394 | 35.2095 | 21.0022
0.2277 | 19.1998 | -144.4396 | -2609.2782 | 2614.4678 | 12.3343 | -35.2977 | 1.1614 | 37.694 | 2609.2782 | 35.2977 | 21.1696
0.2295 | 19.4791 | -142.2051 | -2602.8685 | 2607.9551 | 12.1726 | -35.3565 | 1.1026 | 37.7087 | 2602.8685 | 35.3565 | 21.3369
0.2313 | 19.8026 | -139.6764 | -2594.7828 | 2599.7224 | 11.9668 | -35.4006 | 1.0438 | 37.6646 | 2594.7828 | 35.4006 | 21.5043
0.2331 | 20.3024 | -136.7509 | -2584.2273 | 2588.9905 | 11.7022 | -35.4594 | 0.99968 | 37.6058 | 2584.2273 | 35.4594 | 21.6716
0.2348 | 20.6699 | -133.6048 | -2572.5252 | 2577.1119 | 11.4376 | -35.5182 | 0.94088 | 37.5617 | 2572.5252 | 35.5182 | 21.8297
0.2366 | 20.7287 | -130.5029 | -2560.1026 | 2564.4983 | 11.1582 | -35.5476 | 0.91148 | 37.4882 | 2560.1026 | 35.5476 | 21.997
0.2384 | 20.8758 | -127.6655 | -2549.8117 | 2554.0457 | 10.9377 | -35.5623 | 0.85267 | 37.4293 | 2549.8117 | 35.5623 | 22.1644
0.2402 | 20.9052 | -124.593 | -2541.4614 | 2545.5631 | 10.7613 | -35.5476 | 0.79387 | 37.3558 | 2541.4614 | 35.5476 | 22.3317
0.242 | 20.9052 | -121.6527 | -2534.4784 | 2538.4771 | 10.6143 | -35.5035 | 0.74976 | 37.2823 | 2534.4784 | 35.5035 | 22.4991
0.2438 | 21.0375 | -119.1388 | -2528.2744 | 2532.185 | 10.4967 | -35.4888 | 0.67626 | 37.2382 | 2528.2744 | 35.4888 | 22.6664
0.2456 | 21.5373 | -117.3747 | -2521.4531 | 2525.2901 | 10.3497 | -35.4888 | 0.60275 | 37.1941 | 2521.4531 | 35.4888 | 22.8338
0.2474 | 21.6549 | -115.0078 | -2513.8525 | 2517.5719 | 10.2027 | -35.4447 | 0.54395 | 37.1059 | 2513.8525 | 35.4447 | 23.0011
0.2492 | 21.8019 | -111.994 | -2504.0027 | 2507.5604 | 10.0115 | -35.43 | 0.48514 | 37.0177 | 2504.0027 | 35.43 | 23.1685
0.251 | 21.8019 | -108.8332 | -2492.8298 | 2496.2257 | 9.8057 | -35.4153 | 0.44104 | 36.9442 | 2492.8298 | 35.4153 | 23.3358
0.2528 | 21.6108 | -105.7019 | -2481.451 | 2484.7 | 9.5999 | -35.3859 | 0.38223 | 36.856 | 2481.451 | 35.3859 | 23.5032
0.2546 | 21.5079 | -102.6146 | -2470.0575 | 2473.1742 | 9.3941 | -35.3712 | 0.32343 | 36.7678 | 2470.0575 | 35.3712 | 23.6705
0.2564 | 21.5814 | -99.0128 | -2458.1642 | 2461.1045 | 9.1883 | -35.4006 | 0.27932 | 36.7384 | 2458.1642 | 35.4006 | 23.8379
0.2581 | 21.8019 | -95.4698 | -2446.6238 | 2449.3876 | 8.9825 | -35.4153 | 0.22052 | 36.6943 | 2446.6238 | 35.4153 | 23.9959
0.2599 | 22.2136 | -92.6325 | -2436.2153 | 2438.8468 | 8.806 | -35.4006 | 0.16171 | 36.6208 | 2436.2153 | 35.4006 | 24.1633
0.2617 | 22.5811 | -89.8687 | -2425.8509 | 2428.3501 | 8.6002 | -35.4006 | 0.10291 | 36.562 | 2425.8509 | 35.4006 | 24.3306
0.2635 | 22.7428 | -86.8402 | -2416.4568 | 2418.809 | 8.4091 | -35.4006 | 0.044104 | 36.5032 | 2416.4568 | 35.4006 | 24.498
0.2653 | 23.0662 | -83.8117 | -2408.4594 | 2410.694 | 8.2474 | -35.3565 | -0.029402 | 36.415 | 2408.4594 | 35.3565 | 24.6653
0.2671 | 23.3309 | -80.6657 | -2400.4766 | 2402.5642 | 8.1004 | -35.2977 | -0.088207 | 36.3268 | 2400.4766 | 35.2977 | 24.8327
0.2689 | 23.4485 | -77.5343 | -2392.3909 | 2394.3609 | 7.9534 | -35.1948 | -0.16171 | 36.1944 | 2392.3909 | 35.1948 | 25
0.2707 | 23.9189 | -74.8587 | -2385.3196 | 2387.1867 | 7.8064 | -35.1213 | -0.23522 | 36.0915 | 2385.3196 | 35.1213 | 25.1673
0.2725 | 24.4335 | -72.3889 | -2379.0128 | 2380.8063 | 7.6887 | -35.0183 | -0.29402 | 35.9592 | 2379.0128 | 35.0183 | 25.3347
0.2743 | 24.8598 | -69.9485 | -2372.9706 | 2374.6759 | 7.5711 | -34.9301 | -0.36753 | 35.8416 | 2372.9706 | 34.9301 | 25.502
0.2761 | 25.0509 | -67.4787 | -2367.8546 | 2369.4864 | 7.4682 | -34.8125 | -0.42634 | 35.724 | 2367.8546 | 34.8125 | 25.6694
0.2779 | 25.1979 | -65.0677 | -2362.665 | 2364.2233 | 7.3506 | -34.7243 | -0.49984 | 35.6064 | 2362.665 | 34.7243 | 25.8367
0.2797 | 25.3743 | -62.6714 | -2357.3138 | 2358.7986 | 7.2183 | -34.6361 | -0.58805 | 35.4888 | 2357.3138 | 34.6361 | 26.0041
0.2815 | 25.389 | -60.3192 | -2351.1539 | 2352.5359 | 7.086 | -34.5185 | -0.66156 | 35.3565 | 2351.1539 | 34.5185 | 26.1714
0.2832 | 25.5066 | -58.2169 | -2345.5675 | 2346.9053 | 6.9684 | -34.4009 | -0.73506 | 35.2095 | 2345.5675 | 34.4009 | 26.3295
0.285 | 25.7125 | -55.9823 | -2340.5544 | 2341.8334 | 6.8508 | -34.3127 | -0.82327 | 35.0918 | 2340.5544 | 34.3127 | 26.4968
0.2868 | 26.1976 | -53.7477 | -2336.3939 | 2337.6141 | 6.7479 | -34.2098 | -0.89678 | 34.9889 | 2336.3939 | 34.2098 | 26.6642
0.2886 | 26.7268 | -51.6013 | -2332.7186 | 2333.8947 | 6.645 | -34.1216 | -0.97028 | 34.8713 | 2332.7186 | 34.1216 | 26.8315
0.2904 | 27.0944 | -49.4844 | -2328.9845 | 2330.1312 | 6.5568 | -34.0187 | -1.0585 | 34.7684 | 2328.9845 | 34.0187 | 26.9989
0.2922 | 27.4325 | -47.3233 | -2324.9857 | 2326.103 | 6.4538 | -33.9158 | -1.132 | 34.6508 | 2324.9857 | 33.9158 | 27.1662
0.294 | 27.5942 | -45.3386 | -2320.9135 | 2321.9867 | 6.3803 | -33.7981 | -1.2202 | 34.5185 | 2320.9135 | 33.7981 | 27.3336
0.2958 | 27.8735 | -43.2804 | -2317.4587 | 2318.4878 | 6.3068 | -33.6805 | -1.2937 | 34.4009 | 2317.4587 | 33.6805 | 27.5009
0.2976 | 28.1823 | -41.1635 | -2314.4597 | 2315.4299 | 6.2186 | -33.5629 | -1.3672 | 34.2686 | 2314.4597 | 33.5629 | 27.6683
0.2994 | 28.3734 | -38.9289 | -2312.2104 | 2313.1513 | 6.1451 | -33.4747 | -1.4554 | 34.1657 | 2312.2104 | 33.4747 | 27.8356
0.3012 | 28.491 | -36.7531 | -2310.2257 | 2311.1078 | 6.0716 | -33.3718 | -1.5289 | 34.0481 | 2310.2257 | 33.3718 | 28.003
0.303 | 28.6527 | -34.6802 | -2308.4469 | 2309.2848 | 6.0128 | -33.2395 | -1.6024 | 33.9158 | 2308.4469 | 33.2395 | 28.1703
0.3048 | 28.7703 | -32.6514 | -2306.7562 | 2307.5648 | 5.954 | -33.0925 | -1.6906 | 33.754 | 2306.7562 | 33.0925 | 28.3377
0.3066 | 28.8144 | -30.6227 | -2305.1685 | 2305.9476 | 5.8952 | -32.8867 | -1.7641 | 33.5629 | 2305.1685 | 32.8867 | 28.505
0.3083 | 28.8585 | -28.7556 | -2303.6396 | 2304.3746 | 5.8658 | -32.725 | -1.8377 | 33.3865 | 2303.6396 | 32.725 | 28.6631
0.3101 | 28.7997 | -26.9915 | -2302.4341 | 2303.1544 | 5.807 | -32.5632 | -1.9259 | 33.2395 | 2302.4341 | 32.5632 | 28.8304
0.3119 | 28.6527 | -25.3008 | -2301.2874 | 2301.993 | 5.7629 | -32.4309 | -2.0141 | 33.0925 | 2301.2874 | 32.4309 | 28.9978
0.3137 | 28.4616 | -23.5367 | -2300.2142 | 2300.8904 | 5.7335 | -32.3133 | -2.1023 | 32.9749 | 2300.2142 | 32.3133 | 29.1651
0.3155 | 28.4322 | -21.405 | -2299.4644 | 2300.0966 | 5.6894 | -32.1957 | -2.1905 | 32.8573 | 2299.4644 | 32.1957 | 29.3325
0.3173 | 28.4175 | -19.1998 | -2298.7734 | 2299.3615 | 5.6453 | -32.0634 | -2.2787 | 32.7397 | 2298.7734 | 32.0634 | 29.4998
0.3191 | 28.2999 | -17.171 | -2298.6852 | 2299.2439 | 5.6159 | -31.9605 | -2.3669 | 32.622 | 2298.6852 | 31.9605 | 29.6672
0.3209 | 28.1235 | -15.2011 | -2298.6264 | 2299.1557 | 5.5865 | -31.8429 | -2.4551 | 32.5191 | 2298.6264 | 31.8429 | 29.8345
0.3227 | 27.9765 | -13.0694 | -2298.9205 | 2299.435 | 5.5571 | -31.74 | -2.5433 | 32.4162 | 2298.9205 | 31.74 | 30.0019
0.3245 | 27.9323 | -10.8789 | -2299.3615 | 2299.8172 | 5.5424 | -31.6371 | -2.6462 | 32.3133 | 2299.3615 | 31.6371 | 30.1692
0.3263 | 27.8294 | -8.6737 | -2299.7878 | 2300.2142 | 5.513 | -31.5342 | -2.7638 | 32.2104 | 2299.7878 | 31.5342 | 30.3366
0.3281 | 27.6383 | -6.6303 | -2300.6846 | 2301.0962 | 5.4836 | -31.4459 | -2.8667 | 32.1369 | 2300.6846 | 31.4459 | 30.5039
0.3299 | 27.4031 | -4.7338 | -2301.9489 | 2302.3458 | 5.4542 | -31.3724 | -2.9696 | 32.0634 | 2301.9489 | 31.3724 | 30.6713
0.3316 | 27.2855 | -2.9696 | -2303.3014 | 2303.669 | 5.4248 | -31.2842 | -3.0726 | 31.9899 | 2303.3014 | 31.2842 | 30.8293
0.3334 | 27.2267 | -1.1614 | -2305.1097 | 2305.4625 | 5.4101 | -31.196 | -3.1608 | 31.887 | 2305.1097 | 31.196 | 30.9967
0.3352 | 27.1385 | 0.58805 | -2307.2267 | 2307.5648 | 5.3954 | -31.0931 | -3.2784 | 31.7988 | 2307.2267 | 31.0931 | 31.164
0.337 | 27.0944 | 2.2346 | -2309.6083 | 2309.9317 | 5.3807 | -30.9755 | -3.3666 | 31.7106 | 2309.6083 | 30.9755 | 31.3313
0.3388 | 27.0503 | 3.8664 | -2312.0634 | 2312.3868 | 5.3807 | -30.8726 | -3.4695 | 31.6224 | 2312.0634 | 30.8726 | 31.4987
0.3406 | 27.0062 | 5.5718 | -2314.9595 | 2315.2829 | 5.366 | -30.7844 | -3.5577 | 31.5489 | 2314.9595 | 30.7844 | 31.666
0.3424 | 27.065 | 7.1742 | -2318.2673 | 2318.5907 | 5.3512 | -30.7109 | -3.6606 | 31.4753 | 2318.2673 | 30.7109 | 31.8334
0.3442 | 27.1679 | 8.7178 | -2321.9426 | 2322.2513 | 5.3365 | -30.6374 | -3.7782 | 31.4165 | 2321.9426 | 30.6374 | 32.0007
0.346 | 27.212 | 10.335 | -2325.8531 | 2326.1618 | 5.3365 | -30.5639 | -3.8811 | 31.3577 | 2325.8531 | 30.5639 | 32.1681
0.3478 | 27.1532 | 11.8933 | -2330.0283 | 2330.337 | 5.3218 | -30.4757 | -3.9693 | 31.2842 | 2330.0283 | 30.4757 | 32.3354
0.3496 | 27.065 | 13.4075 | -2334.468 | 2334.7915 | 5.3071 | -30.3875 | -4.0722 | 31.2107 | 2334.468 | 30.3875 | 32.5028
0.3514 | 26.9915 | 14.8629 | -2339.5841 | 2339.9075 | 5.2924 | -30.2992 | -4.1899 | 31.1372 | 2339.5841 | 30.2992 | 32.6701
0.3532 | 26.9474 | 16.3772 | -2345.6998 | 2346.0232 | 5.3071 | -30.1963 | -4.3075 | 31.049 | 2345.6998 | 30.1963 | 32.8375
0.355 | 26.9474 | 17.8179 | -2352.0801 | 2352.4183 | 5.3218 | -30.1081 | -4.4398 | 30.9755 | 2352.0801 | 30.1081 | 33.0048
0.3567 | 26.9768 | 19.141 | -2358.6957 | 2359.0338 | 5.3365 | -30.0199 | -4.5574 | 30.9167 | 2358.6957 | 30.0199 | 33.1629
0.3585 | 26.9768 | 20.3759 | -2365.2818 | 2365.6347 | 5.3512 | -29.9317 | -4.675 | 30.8579 | 2365.2818 | 29.9317 | 33.3302
0.3603 | 26.9033 | 21.4932 | -2372.9412 | 2373.3087 | 5.3807 | -29.8582 | -4.8073 | 30.8138 | 2372.9412 | 29.8582 | 33.4976
0.3621 | 26.771 | 22.6252 | -2381.8207 | 2382.1883 | 5.4542 | -29.7847 | -4.9543 | 30.7697 | 2381.8207 | 29.7847 | 33.6649
0.3639 | 26.668 | 23.9042 | -2389.2449 | 2389.6418 | 5.4836 | -29.7112 | -5.0719 | 30.7403 | 2389.2449 | 29.7112 | 33.8323
0.3657 | 26.5063 | 25.1832 | -2396.5367 | 2396.9189 | 5.513 | -29.6524 | -5.1895 | 30.6962 | 2396.5367 | 29.6524 | 33.9996
0.3675 | 26.2711 | 26.6386 | -2403.9608 | 2404.3577 | 5.5424 | -29.5642 | -5.3071 | 30.6521 | 2403.9608 | 29.5642 | 34.167
0.3693 | 26.0506 | 28.0941 | -2411.4143 | 2411.8259 | 5.5865 | -29.4907 | -5.4101 | 30.608 | 2411.4143 | 29.4907 | 34.3343
0.3711 | 25.933 | 29.3143 | -2419.4706 | 2419.8969 | 5.6159 | -29.4319 | -5.5277 | 30.5786 | 2419.4706 | 29.4319 | 34.5017
0.3729 | 25.8742 | 30.2404 | -2427.6298 | 2428.0708 | 5.66 | -29.3584 | -5.6453 | 30.5345 | 2427.6298 | 29.3584 | 34.669
0.3747 | 25.7566 | 31.1078 | -2436.5093 | 2436.9504 | 5.7041 | -29.3143 | -5.7629 | 30.5198 | 2436.5093 | 29.3143 | 34.8364
0.3765 | 25.6684 | 31.9752 | -2445.1978 | 2445.6535 | 5.7335 | -29.2555 | -5.8952 | 30.5051 | 2445.1978 | 29.2555 | 35.0037
0.3783 | 25.4184 | 32.5044 | -2453.2834 | 2453.7392 | 5.7629 | -29.1967 | -5.9981 | 30.461 | 2453.2834 | 29.1967 | 35.1711
0.3801 | 25.3596 | 33.313 | -2460.4282 | 2460.8987 | 5.7923 | -29.1378 | -6.1157 | 30.4316 | 2460.4282 | 29.1378 | 35.3384
0.3818 | 25.1538 | 34.1804 | -2466.7351 | 2467.1908 | 5.8217 | -29.079 | -6.2186 | 30.4022 | 2466.7351 | 29.079 | 35.4965
0.3836 | 24.8745 | 34.9301 | -2471.7629 | 2472.2333 | 5.8511 | -29.0202 | -6.3215 | 30.3728 | 2471.7629 | 29.0202 | 35.6638
0.3854 | 24.6246 | 35.5623 | -2476.3203 | 2476.776 | 5.8805 | -28.9614 | -6.4391 | 30.3286 | 2476.3203 | 28.9614 | 35.8312
0.3872 | 24.3599 | 35.9739 | -2480.4366 | 2480.8924 | 5.8952 | -28.9173 | -6.542 | 30.3139 | 2480.4366 | 28.9173 | 35.9985
0.389 | 24.0953 | 36.2974 | -2484.5677 | 2485.0234 | 5.8952 | -28.8879 | -6.6597 | 30.2992 | 2484.5677 | 28.8879 | 36.1659
0.3908 | 23.7866 | 36.7678 | -2490.7863 | 2491.242 | 5.8952 | -28.8585 | -6.7626 | 30.3139 | 2490.7863 | 28.8585 | 36.3332
0.3926 | 23.7131 | 37.5029 | -2498.0487 | 2498.5191 | 5.9099 | -28.8291 | -6.8802 | 30.2992 | 2498.0487 | 28.8291 | 36.5006
0.3944 | 23.6984 | 38.1938 | -2505.7521 | 2506.2373 | 5.954 | -28.7997 | -6.9978 | 30.3139 | 2505.7521 | 28.7997 | 36.6679
0.3962 | 23.6984 | 38.9436 | -2514.1318 | 2514.617 | 5.9981 | -28.7703 | -7.1301 | 30.3139 | 2514.1318 | 28.7703 | 36.8353
0.398 | 23.6102 | 39.7815 | -2523.276 | 2523.7759 | 6.0422 | -28.7115 | -7.2624 | 30.3139 | 2523.276 | 28.7115 | 37.0026
0.3998 | 23.6102 | 40.8841 | -2532.8906 | 2533.4199 | 6.101 | -28.6821 | -7.3947 | 30.3286 | 2532.8906 | 28.6821 | 37.17
0.4016 | 23.6102 | 42.2514 | -2542.6081 | 2543.1521 | 6.1598 | -28.6527 | -7.527 | 30.3434 | 2542.6081 | 28.6527 | 37.3373
0.4034 | 23.5514 | 43.5304 | -2552.6344 | 2553.2077 | 6.2186 | -28.6086 | -7.6446 | 30.3581 | 2552.6344 | 28.6086 | 37.5046
0.4051 | 23.4926 | 44.5888 | -2561.6903 | 2562.2784 | 6.2774 | -28.5792 | -7.777 | 30.3728 | 2561.6903 | 28.5792 | 37.6627
0.4069 | 23.4779 | 45.6326 | -2571.5843 | 2572.187 | 6.3362 | -28.5498 | -7.8946 | 30.3875 | 2571.5843 | 28.5498 | 37.83
0.4087 | 23.5367 | 46.7058 | -2583.0218 | 2583.6393 | 6.4244 | -28.5204 | -8.0269 | 30.4169 | 2583.0218 | 28.5204 | 37.9974
0.4105 | 23.6984 | 47.926 | -2595.4738 | 2596.1059 | 6.5126 | -28.491 | -8.1739 | 30.461 | 2595.4738 | 28.491 | 38.1647
0.4123 | 23.9336 | 49.455 | -2609.3518 | 2610.028 | 6.6156 | -28.4763 | -8.3062 | 30.5051 | 2609.3518 | 28.4763 | 38.3321
0.4141 | 24.1982 | 50.8075 | -2622.7446 | 2623.4502 | 6.7332 | -28.4616 | -8.4238 | 30.5639 | 2622.7446 | 28.4616 | 38.4994
0.4159 | 24.3599 | 52.0865 | -2635.4023 | 2636.1227 | 6.8361 | -28.4616 | -8.5561 | 30.6227 | 2635.4023 | 28.4616 | 38.6668
0.4177 | 24.3746 | 53.0715 | -2647.1045 | 2647.8396 | 6.9243 | -28.4469 | -8.6884 | 30.6815 | 2647.1045 | 28.4469 | 38.8341
0.4195 | 24.3599 | 53.9535 | -2658.1158 | 2658.8655 | 7.0272 | -28.4469 | -8.806 | 30.7403 | 2658.1158 | 28.4469 | 39.0015
0.4213 | 24.2864 | 54.5122 | -2667.1276 | 2667.8921 | 7.086 | -28.4763 | -8.9384 | 30.7991 | 2667.1276 | 28.4763 | 39.1688
0.4231 | 23.8895 | 54.9973 | -2676.1542 | 2676.9186 | 7.1595 | -28.4616 | -9.056 | 30.8432 | 2676.1542 | 28.4616 | 39.3362
0.4249 | 23.4485 | 55.4678 | -2685.2542 | 2686.0334 | 7.233 | -28.4616 | -9.1883 | 30.8873 | 2685.2542 | 28.4616 | 39.5035
0.4267 | 23.2427 | 55.9529 | -2696.2949 | 2697.0887 | 7.3212 | -28.4469 | -9.3206 | 30.9461 | 2696.2949 | 28.4469 | 39.6709
0.4285 | 23.0368 | 56.3498 | -2706.8209 | 2707.6001 | 7.4094 | -28.4763 | -9.4529 | 31.0196 | 2706.8209 | 28.4763 | 39.8382
0.4302 | 22.8751 | 56.7762 | -2717.6117 | 2718.4055 | 7.4976 | -28.491 | -9.5852 | 31.1078 | 2717.6117 | 28.491 | 39.9963
0.432 | 22.7869 | 57.2466 | -2729.4461 | 2730.2547 | 7.6005 | -28.491 | -9.7322 | 31.1813 | 2729.4461 | 28.491 | 40.1636
0.4338 | 22.8016 | 57.8347 | -2742.3685 | 2743.1771 | 7.7329 | -28.491 | -9.8792 | 31.2695 | 2742.3685 | 28.491 | 40.331
0.4356 | 22.7869 | 58.4962 | -2755.9084 | 2756.7316 | 7.8799 | -28.5204 | -10.0115 | 31.3577 | 2755.9084 | 28.5204 | 40.4983
0.4374 | 22.7134 | 59.099 | -2769.9775 | 2770.8007 | 8.0269 | -28.5204 | -10.1439 | 31.4459 | 2769.9775 | 28.5204 | 40.6657
0.4392 | 22.5517 | 59.7899 | -2784.8404 | 2785.6784 | 8.2033 | -28.491 | -10.2909 | 31.5194 | 2784.8404 | 28.491 | 40.833
0.441 | 22.5517 | 60.5985 | -2800.5707 | 2801.4234 | 8.3944 | -28.491 | -10.4379 | 31.6224 | 2800.5707 | 28.491 | 41.0004
0.4428 | 22.5958 | 61.4953 | -2817.4477 | 2818.3298 | 8.6002 | -28.491 | -10.5849 | 31.7253 | 2817.4477 | 28.491 | 41.1677
0.4446 | 22.5958 | 62.2597 | -2834.6482 | 2835.545 | 8.8354 | -28.4763 | -10.7172 | 31.8429 | 2834.6482 | 28.4763 | 41.3351
0.4464 | 22.5076 | 62.8478 | -2850.2168 | 2851.1283 | 9.0707 | -28.4616 | -10.8495 | 31.9605 | 2850.2168 | 28.4616 | 41.5024
0.4482 | 22.3165 | 63.4505 | -2863.8301 | 2864.7416 | 9.2618 | -28.4469 | -10.9965 | 32.0487 | 2863.8301 | 28.4469 | 41.6698
0.45 | 22.0225 | 63.8034 | -2876.3556 | 2877.2818 | 9.4382 | -28.4175 | -11.1288 | 32.1369 | 2876.3556 | 28.4175 | 41.8371
0.4518 | 21.6402 | 63.8769 | -2887.1022 | 2888.0137 | 9.5852 | -28.4028 | -11.2611 | 32.1957 | 2887.1022 | 28.4028 | 42.0045
0.4536 | 21.3315 | 63.9945 | -2898.0693 | 2898.9808 | 9.7322 | -28.3734 | -11.3935 | 32.2692 | 2898.0693 | 28.3734 | 42.1718
0.4553 | 21.3609 | 64.2885 | -2910.9035 | 2911.8297 | 9.9233 | -28.3587 | -11.5258 | 32.3574 | 2910.9035 | 28.3587 | 42.3299
0.4571 | 21.4932 | 64.6266 | -2926.3104 | 2927.2366 | 10.1586 | -28.3587 | -11.6581 | 32.4897 | 2926.3104 | 28.3587 | 42.4972
0.4589 | 21.6843 | 64.906 | -2941.8055 | 2942.7464 | 10.4085 | -28.3734 | -11.7904 | 32.622 | 2941.8055 | 28.3734 | 42.6646
0.4607 | 21.5373 | 65.2 | -2955.5953 | 2956.5361 | 10.6731 | -28.3587 | -11.9227 | 32.7397 | 2955.5953 | 28.3587 | 42.8319
0.4625 | 21.1698 | 65.2735 | -2967.665 | 2968.6059 | 10.923 | -28.3293 | -12.055 | 32.8426 | 2967.665 | 28.3293 | 42.9993
0.4643 | 20.7581 | 65.3176 | -2979.0584 | 2979.9993 | 11.1876 | -28.2999 | -12.1873 | 32.9455 | 2979.0584 | 28.2999 | 43.1666
0.4661 | 20.3318 | 65.4499 | -2991.6427 | 2992.5836 | 11.4523 | -28.2705 | -12.3196 | 33.0631 | 2991.6427 | 28.2705 | 43.334
0.4679 | 19.979 | 65.7145 | -3006.535 | 3007.4906 | 11.7316 | -28.2117 | -12.4519 | 33.166 | 3006.535 | 28.2117 | 43.5013
0.4697 | 19.6703 | 66.0527 | -3022.5006 | 3023.4562 | 12.0256 | -28.1676 | -12.5696 | 33.2836 | 3022.5006 | 28.1676 | 43.6687
0.4715 | 19.2733 | 66.5672 | -3039.4805 | 3040.4361 | 12.3196 | -28.0941 | -12.6872 | 33.3865 | 3039.4805 | 28.0941 | 43.836
0.4733 | 18.8323 | 66.9494 | -3056.7839 | 3057.7541 | 12.6284 | -28.0206 | -12.8195 | 33.5041 | 3056.7839 | 28.0206 | 44.0033
0.4751 | 18.4501 | 67.2582 | -3072.3966 | 3073.3669 | 12.9518 | -27.9618 | -12.9371 | 33.6217 | 3072.3966 | 27.9618 | 44.1707
0.4769 | 18.0531 | 67.4493 | -3087.4065 | 3088.3768 | 13.2605 | -27.8882 | -13.0547 | 33.7393 | 3087.4065 | 27.8882 | 44.338
0.4786 | 17.4504 | 67.3905 | -3102.0196 | 3102.9899 | 13.5692 | -27.8 | -13.1723 | 33.8422 | 3102.0196 | 27.8 | 44.4961
0.4804 | 16.583 | 67.3758 | -3117.1765 | 3118.1321 | 13.878 | -27.6677 | -13.2752 | 33.9158 | 3117.1765 | 27.6677 | 44.6634
0.4822 | 15.5833 | 67.5816 | -3132.5246 | 3133.4655 | 14.2014 | -27.5207 | -13.3781 | 33.9893 | 3132.5246 | 27.5207 | 44.8308
0.484 | 14.5836 | 67.6992 | -3146.2703 | 3147.2112 | 14.5248 | -27.4031 | -13.4957 | 34.0922 | 3146.2703 | 27.4031 | 44.9981
0.4858 | 13.7016 | 67.6551 | -3158.7075 | 3159.6484 | 14.8041 | -27.3149 | -13.5986 | 34.1804 | 3158.7075 | 27.3149 | 45.1655
0.4876 | 12.8489 | 67.7727 | -3170.6008 | 3171.5417 | 15.0982 | -27.2414 | -13.7163 | 34.2686 | 3170.6008 | 27.2414 | 45.3328
0.4894 | 11.9668 | 68.0079 | -3182.5382 | 3183.4791 | 15.3922 | -27.1532 | -13.8339 | 34.3568 | 3182.5382 | 27.1532 | 45.5002
0.4912 | 10.9377 | 68.3019 | -3195.2401 | 3196.181 | 15.6862 | -27.0356 | -13.9515 | 34.445 | 3195.2401 | 27.0356 | 45.6675
0.493 | 9.791 | 68.5225 | -3207.3833 | 3208.3242 | 15.9508 | -26.918 | -14.0544 | 34.5185 | 3207.3833 | 26.918 | 45.8349
0.4948 | 8.5561 | 68.7724 | -3217.6154 | 3218.5563 | 16.2302 | -26.771 | -14.1573 | 34.5773 | 3217.6154 | 26.771 | 46.0022
0.4966 | 7.2624 | 69.1105 | -3226.539 | 3227.4946 | 16.5242 | -26.6092 | -14.2602 | 34.6508 | 3226.539 | 26.6092 | 46.1696
0.4984 | 5.7923 | 69.875 | -3234.7423 | 3235.7126 | 16.8035 | -26.4475 | -14.3484 | 34.6949 | 3234.7423 | 26.4475 | 46.3369
0.5002 | 4.234 | 70.7717 | -3242.8133 | 3243.813 | 17.0534 | -26.2711 | -14.4513 | 34.739 | 3242.8133 | 26.2711 | 46.5043
0.502 | 2.7197 | 71.7273 | -3250.7226 | 3251.737 | 17.3181 | -26.1094 | -14.5248 | 34.7684 | 3250.7226 | 26.1094 | 46.6716
0.5037 | 1.2055 | 72.7123 | -3258.97 | 3259.9991 | 17.568 | -25.9477 | -14.613 | 34.8125 | 3258.97 | 25.9477 | 46.8297
0.5055 | -0.26462 | 73.4768 | -3268.9227 | 3269.9812 | 17.7885 | -25.7713 | -14.7012 | 34.8272 | 3268.9227 | 25.7713 | 46.997
0.5073 | -1.8377 | 74.006 | -3280.0662 | 3281.1247 | 18.009 | -25.5507 | -14.8041 | 34.8272 | 3280.0662 | 25.5507 | 47.1644
0.5091 | -3.396 | 74.4324 | -3292.0036 | 3293.0768 | 18.2442 | -25.3302 | -14.8924 | 34.8272 | 3292.0036 | 25.3302 | 47.3317
0.5109 | -5.0719 | 74.991 | -3304.4262 | 3305.5141 | 18.4795 | -25.095 | -14.9806 | 34.8272 | 3304.4262 | 25.095 | 47.4991
0.5127 | -6.7185 | 75.3585 | -3315.1728 | 3316.2607 | 18.6706 | -24.8598 | -15.0688 | 34.7978 | 3315.1728 | 24.8598 | 47.6664
0.5145 | -8.3944 | 75.6526 | -3324.8756 | 3325.9635 | 18.847 | -24.6246 | -15.157 | 34.7684 | 3324.8756 | 24.6246 | 47.8338
0.5163 | -10.0262 | 75.7996 | -3333.9757 | 3335.0783 | 19.0234 | -24.3894 | -15.2305 | 34.739 | 3333.9757 | 24.3894 | 48.0011
0.5181 | -11.8198 | 75.7408 | -3341.9437 | 3343.0463 | 19.1851 | -24.1541 | -15.3187 | 34.7096 | 3341.9437 | 24.1541 | 48.1685
0.5199 | -13.5545 | 75.7261 | -3348.7651 | 3349.8677 | 19.3615 | -23.9189 | -15.3922 | 34.6655 | 3348.7651 | 23.9189 | 48.3358
0.5217 | -15.0541 | 75.7702 | -3354.7191 | 3355.8217 | 19.5232 | -23.6837 | -15.4657 | 34.6361 | 3354.7191 | 23.6837 | 48.5032
0.5235 | -16.3625 | 75.8878 | -3360.1733 | 3361.2906 | 19.685 | -23.4779 | -15.5392 | 34.6067 | 3360.1733 | 23.4779 | 48.6705
0.5253 | -17.7738 | 76.1671 | -3365.3775 | 3366.4948 | 19.8173 | -23.2721 | -15.6274 | 34.5626 | 3365.3775 | 23.2721 | 48.8379
0.5271 | -19.6703 | 75.9613 | -3369.0528 | 3370.1995 | 19.9496 | -23.0221 | -15.6862 | 34.5038 | 3369.0528 | 23.0221 | 49.0052
0.5288 | -21.5226 | 75.6673 | -3371.7431 | 3372.8751 | 20.0819 | -22.7722 | -15.745 | 34.445 | 3371.7431 | 22.7722 | 49.1633
0.5306 | -23.375 | 75.8878 | -3374.4188 | 3375.5655 | 20.2583 | -22.5076 | -15.7891 | 34.4009 | 3374.4188 | 22.5076 | 49.3306
0.5324 | -25.2273 | 76.0348 | -3376.7416 | 3377.8883 | 20.3906 | -22.2283 | -15.8479 | 34.3274 | 3376.7416 | 22.2283 | 49.498
0.5342 | -26.9768 | 76.0348 | -3378.4028 | 3379.5642 | 20.4935 | -21.9489 | -15.892 | 34.2392 | 3378.4028 | 21.9489 | 49.6653
0.536 | -28.344 | 76.0789 | -3379.3731 | 3380.5345 | 20.567 | -21.699 | -15.9214 | 34.1363 | 3379.3731 | 21.699 | 49.8327
0.5378 | -30.4169 | 76.123 | -3379.1967 | 3380.3728 | 20.6258 | -21.4197 | -15.9802 | 34.0187 | 3379.1967 | 21.4197 | 50
0.5396 | -32.5779 | 76.3435 | -3378.6527 | 3379.8582 | 20.714 | -21.1404 | -16.0243 | 33.9158 | 3378.6527 | 21.1404 | 50.1673
0.5414 | -34.4009 | 76.0495 | -3377.653 | 3378.8732 | 20.7581 | -20.8317 | -16.0537 | 33.7687 | 3377.653 | 20.8317 | 50.3347
0.5432 | -36.3415 | 75.9025 | -3375.6096 | 3376.8592 | 20.7875 | -20.5082 | -16.0832 | 33.607 | 3375.6096 | 20.5082 | 50.502
0.545 | -38.2526 | 75.6673 | -3372.6987 | 3373.9483 | 20.8169 | -20.1848 | -16.0979 | 33.4306 | 3372.6987 | 20.1848 | 50.6694
0.5468 | -40.1932 | 75.3879 | -3369.1851 | 3370.4494 | 20.8317 | -19.832 | -16.1126 | 33.2542 | 3369.1851 | 19.832 | 50.8367
0.5486 | -42.4131 | 75.432 | -3364.9659 | 3366.2596 | 20.8464 | -19.4938 | -16.142 | 33.0631 | 3364.9659 | 19.4938 | 51.0041
0.5504 | -44.4124 | 75.3879 | -3359.8351 | 3361.1583 | 20.8317 | -19.1557 | -16.1567 | 32.872 | 3359.8351 | 19.1557 | 51.1714
0.5521 | -46.3971 | 75.4026 | -3353.4989 | 3354.8514 | 20.8317 | -18.8029 | -16.1714 | 32.6661 | 3353.4989 | 18.8029 | 51.3295
0.5539 | -48.3818 | 75.3585 | -3346.3982 | 3347.7801 | 20.8169 | -18.4353 | -16.1861 | 32.4456 | 3346.3982 | 18.4353 | 51.4968
0.5557 | -50.587 | 75.6673 | -3338.6948 | 3340.1061 | 20.7728 | -18.0678 | -16.2008 | 32.2251 | 3338.6948 | 18.0678 | 51.6642
0.5575 | -52.7333 | 75.8878 | -3328.9478 | 3330.418 | 20.7287 | -17.6415 | -16.2008 | 31.9458 | 3328.9478 | 17.6415 | 51.8315
0.5593 | -54.6298 | 75.682 | -3317.6573 | 3319.1421 | 20.6552 | -17.2298 | -16.1714 | 31.6665 | 3317.6573 | 17.2298 | 51.9989
0.5611 | -56.5704 | 75.3585 | -3305.7199 | 3307.2488 | 20.5817 | -16.8182 | -16.142 | 31.3724 | 3305.7199 | 16.8182 | 52.1662
0.5629 | -58.6285 | 75.0792 | -3293.3268 | 3294.8704 | 20.5082 | -16.4213 | -16.1126 | 31.0931 | 3293.3268 | 16.4213 | 52.3336
0.5647 | -60.7749 | 74.8587 | -3279.6987 | 3281.2864 | 20.42 | -15.9949 | -16.0685 | 30.7844 | 3279.6987 | 15.9949 | 52.5009
0.5665 | -62.936 | 74.7999 | -3263.9096 | 3265.5414 | 20.3024 | -15.5245 | -15.9949 | 30.4463 | 3263.9096 | 15.5245 | 52.6683
0.5683 | -65.0677 | 74.55 | -3246.7385 | 3248.4145 | 20.1848 | -15.0541 | -15.9067 | 30.064 | 3246.7385 | 15.0541 | 52.8356
0.5701 | -67.2288 | 74.0795 | -3228.068 | 3229.788 | 20.0378 | -14.5689 | -15.8332 | 29.6818 | 3228.068 | 14.5689 | 53.003
0.5719 | -69.4045 | 73.2122 | -3206.3542 | 3208.1037 | 19.8614 | -14.0691 | -15.7303 | 29.2702 | 3206.3542 | 14.0691 | 53.1703
0.5737 | -71.8155 | 72.3889 | -3183.0234 | 3184.8316 | 19.6703 | -13.5398 | -15.6274 | 28.8438 | 3183.0234 | 13.5398 | 53.3377
0.5755 | -74.403 | 70.9776 | -3157.1933 | 3159.031 | 19.4644 | -12.9077 | -15.5098 | 28.3293 | 3157.1933 | 12.9077 | 53.505
0.5772 | -77.0198 | 69.7721 | -3130.2606 | 3132.1571 | 19.2292 | -12.2755 | -15.3775 | 27.8147 | 3130.2606 | 12.2755 | 53.6631
0.579 | -79.7689 | 68.5813 | -3101.1816 | 3103.1369 | 18.994 | -11.6287 | -15.2305 | 27.3002 | 3101.1816 | 11.6287 | 53.8304
0.5808 | -82.6062 | 67.1258 | -3069.1917 | 3071.2058 | 18.7147 | -10.9083 | -15.0541 | 26.7121 | 3069.1917 | 10.9083 | 53.9978
0.5826 | -85.4142 | 65.8909 | -3035.4083 | 3037.4811 | 18.4353 | -10.1439 | -14.8629 | 26.1388 | 3035.4083 | 10.1439 | 54.1651
0.5844 | -88.1633 | 64.803 | -2999.9195 | 3002.0659 | 18.1413 | -9.3353 | -14.6718 | 25.5654 | 2999.9195 | 9.3353 | 54.3325
0.5862 | -90.7801 | 63.8622 | -2963.578 | 2965.8126 | 17.862 | -8.5267 | -14.4807 | 24.9627 | 2963.578 | 8.5267 | 54.4998
0.588 | -93.1911 | 63.23 | -2926.4721 | 2928.7949 | 17.568 | -7.6887 | -14.2602 | 24.3599 | 2926.4721 | 7.6887 | 54.6672
0.5898 | -95.4845 | 62.7449 | -2888.6458 | 2891.0568 | 17.2593 | -6.8508 | -14.0544 | 23.7572 | 2888.6458 | 6.8508 | 54.8345
0.5916 | -97.6897 | 62.3038 | -2850.1727 | 2852.6719 | 16.9505 | -5.9834 | -13.8339 | 23.1544 | 2850.1727 | 5.9834 | 55.0019
0.5934 | -99.8214 | 61.9363 | -2811.0527 | 2813.6548 | 16.6271 | -5.0866 | -13.5986 | 22.5517 | 2811.0527 | 5.0866 | 55.1692
0.5952 | -101.909 | 61.7746 | -2770.5655 | 2773.2705 | 16.289 | -4.1604 | -13.3634 | 21.9342 | 2770.5655 | 4.1604 | 55.3366
0.597 | -104.0553 | 61.8775 | -2728.2112 | 2731.0633 | 15.9361 | -3.2196 | -13.0988 | 21.3609 | 2728.2112 | 3.2196 | 55.5039
0.5988 | -106.2164 | 62.1127 | -2684.8867 | 2687.8711 | 15.598 | -2.264 | -12.8195 | 20.7728 | 2684.8867 | 2.264 | 55.6713
0.6006 | -108.3334 | 62.495 | -2640.0185 | 2643.1499 | 15.2305 | -1.3231 | -12.5255 | 20.2142 | 2640.0185 | 1.3231 | 55.8386
0.6023 | -110.3622 | 63.0536 | -2593.8419 | 2597.1203 | 14.8629 | -0.39693 | -12.2167 | 19.6997 | 2593.8419 | 0.39693 | 55.9967
0.6041 | -112.2292 | 63.8034 | -2546.4158 | 2549.8706 | 14.4807 | 0.51454 | -11.908 | 19.2292 | 2546.4158 | -0.51454 | 56.164
0.6059 | -113.9346 | 64.7001 | -2498.3868 | 2502.0033 | 14.0985 | 1.426 | -11.5699 | 18.8029 | 2498.3868 | -1.426 | 56.3313
0.6077 | -115.537 | 65.6557 | -2449.4758 | 2453.2687 | 13.7163 | 2.3375 | -11.2317 | 18.4059 | 2449.4758 | -2.3375 | 56.4987
0.6095 | -117.0659 | 66.5819 | -2399.374 | 2403.3433 | 13.3193 | 3.2343 | -10.8789 | 18.009 | 2399.374 | -3.2343 | 56.666
0.6113 | -118.3302 | 67.464 | -2348.7723 | 2352.8887 | 12.9224 | 4.1163 | -10.5114 | 17.6415 | 2348.7723 | -4.1163 | 56.8334
0.6131 | -119.3593 | 68.346 | -2298.0972 | 2302.3752 | 12.5108 | 4.9837 | -10.1292 | 17.3181 | 2298.0972 | -4.9837 | 57.0007
0.6149 | -120.0944 | 69.1399 | -2246.4664 | 2250.9062 | 12.0991 | 5.7923 | -9.7469 | 17.0387 | 2246.4664 | -5.7923 | 57.1681
0.6167 | -120.5207 | 69.7427 | -2195.0709 | 2199.6577 | 11.6875 | 6.542 | -9.35 | 16.7741 | 2195.0709 | -6.542 | 57.3354
0.6185 | -120.9618 | 70.4189 | -2143.2785 | 2148.027 | 11.2758 | 7.2918 | -8.9531 | 16.5389 | 2143.2785 | -7.2918 | 57.5028
0.6203 | -121.3146 | 71.2128 | -2091.2361 | 2096.1463 | 10.8789 | 8.0269 | -8.5561 | 16.3625 | 2091.2361 | -8.0269 | 57.6701
0.6221 | -121.4763 | 71.9625 | -2038.9585 | 2044.0157 | 10.4673 | 8.7031 | -8.1298 | 16.2302 | 2038.9585 | -8.7031 | 57.8375
0.6239 | -121.638 | 72.5653 | -1987.8423 | 1993.0465 | 10.085 | 9.35 | -7.7034 | 16.0979 | 1987.8423 | -9.35 | 58.0048
0.6256 | -121.6821 | 73.1239 | -1936.241 | 1941.6069 | 9.7028 | 9.9674 | -7.2771 | 16.0243 | 1936.241 | -9.9674 | 58.1629
0.6274 | -121.6233 | 73.4768 | -1884.9043 | 1890.4172 | 9.3206 | 10.5555 | -6.8361 | 15.9508 | 1884.9043 | -10.5555 | 58.3302
0.6292 | -121.3587 | 73.5797 | -1833.8763 | 1839.5215 | 8.9531 | 11.0847 | -6.4097 | 15.8773 | 1833.8763 | -11.0847 | 58.4976
0.631 | -120.8883 | 73.3445 | -1783.6127 | 1789.3609 | 8.5855 | 11.5699 | -5.9981 | 15.8332 | 1783.6127 | -11.5699 | 58.6649
0.6328 | -120.1826 | 72.5065 | -1732.9964 | 1738.8181 | 8.2327 | 12.0256 | -5.5718 | 15.8332 | 1732.9964 | -12.0256 | 58.8323
0.6346 | -119.3299 | 71.5215 | -1681.9537 | 1687.8489 | 7.8946 | 12.4372 | -5.1454 | 15.8185 | 1681.9537 | -12.4372 | 58.9996
0.6364 | -118.3744 | 70.4777 | -1631.2344 | 1637.1884 | 7.5564 | 12.8195 | -4.7338 | 15.8185 | 1631.2344 | -12.8195 | 59.167
0.6382 | -117.3453 | 69.3751 | -1580.9709 | 1586.9837 | 7.233 | 13.1723 | -4.3075 | 15.8038 | 1580.9709 | -13.1723 | 59.3343
0.64 | -116.1986 | 68.1549 | -1531.3542 | 1537.4111 | 6.9243 | 13.4663 | -3.8958 | 15.7891 | 1531.3542 | -13.4663 | 59.5017
0.6418 | -114.9343 | 66.626 | -1483.1195 | 1489.2058 | 6.645 | 13.7016 | -3.4989 | 15.745 | 1483.1195 | -13.7016 | 59.669
0.6436 | -113.6553 | 64.8619 | -1436.1049 | 1442.2059 | 6.4097 | 13.8927 | -3.1314 | 15.7156 | 1436.1049 | -13.8927 | 59.8364
0.6454 | -112.3321 | 63.0683 | -1390.0607 | 1396.1617 | 6.1745 | 14.0397 | -2.7638 | 15.6715 | 1390.0607 | -14.0397 | 60.0037
0.6472 | -111.0384 | 61.0689 | -1344.9132 | 1351.0142 | 5.954 | 14.1426 | -2.411 | 15.6127 | 1344.9132 | -14.1426 | 60.1711
0.649 | -109.8182 | 58.7608 | -1300.2214 | 1306.3077 | 5.7629 | 14.1867 | -2.0729 | 15.5098 | 1300.2214 | -14.1867 | 60.3384
0.6507 | -108.6127 | 56.3645 | -1255.9854 | 1262.0717 | 5.6012 | 14.172 | -1.7788 | 15.3922 | 1255.9854 | -14.172 | 60.4965
0.6525 | -107.5542 | 54.1447 | -1212.852 | 1218.953 | 5.4689 | 14.1132 | -1.4995 | 15.2599 | 1212.852 | -14.1132 | 60.6638
0.6543 | -106.7898 | 52.0571 | -1170.8358 | 1176.9956 | 5.3512 | 14.0103 | -1.2496 | 15.0835 | 1170.8358 | -14.0103 | 60.8312
0.6561 | -106.1135 | 50.0136 | -1129.393 | 1135.6264 | 5.2483 | 13.878 | -1.0144 | 14.9071 | 1129.393 | -13.878 | 60.9985
0.6579 | -105.5696 | 48.0289 | -1089.1411 | 1095.4479 | 5.1601 | 13.731 | -0.79387 | 14.7159 | 1089.1411 | -13.731 | 61.1659
0.6597 | -105.0844 | 46.1031 | -1050.8443 | 1057.2247 | 5.0719 | 13.5398 | -0.61745 | 14.4954 | 1050.8443 | -13.5398 | 61.3332
0.6615 | -104.2465 | 43.7803 | -1014.3559 | 1020.7656 | 4.9837 | 13.2752 | -0.48514 | 14.2014 | 1014.3559 | -13.2752 | 61.5006
0.6633 | -102.9822 | 40.8253 | -980.0432 | 986.4088 | 4.8955 | 12.9224 | -0.41163 | 13.8486 | 980.0432 | -12.9224 | 61.6679
0.6651 | -101.7179 | 37.7969 | -946.6567 | 952.9782 | 4.8073 | 12.599 | -0.32343 | 13.4957 | 946.6567 | -12.599 | 61.8353
0.6669 | -100.233 | 34.2833 | -914.4316 | 920.6796 | 4.7191 | 12.2608 | -0.27932 | 13.1429 | 914.4316 | -12.2608 | 62.0026
0.6687 | -98.5424 | 30.4022 | -882.8239 | 888.9396 | 4.6162 | 11.908 | -0.24992 | 12.7754 | 882.8239 | -11.908 | 62.17
0.6705 | -96.8223 | 26.5945 | -852.4071 | 858.4052 | 4.5133 | 11.5405 | -0.22052 | 12.4078 | 852.4071 | -11.5405 | 62.3373
0.6723 | -95.1758 | 22.8751 | -823.1957 | 829.1056 | 4.4104 | 11.2023 | -0.20582 | 12.0403 | 823.1957 | -11.2023 | 62.5046
0.6741 | -93.3823 | 19.1704 | -796.0131 | 801.8201 | 4.3222 | 10.8642 | -0.19112 | 11.7022 | 796.0131 | -10.8642 | 62.672
0.6758 | -91.5446 | 15.5833 | -770.727 | 776.4458 | 4.234 | 10.5408 | -0.20582 | 11.3641 | 770.727 | -10.5408 | 62.83
0.6776 | -89.8834 | 11.9962 | -746.7052 | 752.3357 | 4.1752 | 10.2468 | -0.20582 | 11.07 | 746.7052 | -10.2468 | 62.9974
0.6794 | -88.1927 | 8.5708 | -723.095 | 728.6374 | 4.1016 | 9.9527 | -0.22052 | 10.776 | 723.095 | -9.9527 | 63.1647
0.6812 | -86.7961 | 5.4983 | -701.396 | 706.9089 | 4.0281 | 9.6881 | -0.22052 | 10.4967 | 701.396 | -9.6881 | 63.3321
0.683 | -85.6053 | 2.6756 | -681.3435 | 686.827 | 3.9693 | 9.4382 | -0.23522 | 10.2468 | 681.3435 | -9.4382 | 63.4994
0.6848 | -84.6056 | 0.23522 | -663.4374 | 668.9209 | 3.9105 | 9.2324 | -0.24992 | 10.0262 | 663.4374 | -9.2324 | 63.6668
0.6866 | -83.4589 | -2.0729 | -646.2957 | 651.7793 | 3.8664 | 9.0119 | -0.27932 | 9.8204 | 646.2957 | -9.0119 | 63.8341
0.6884 | -82.2975 | -4.234 | -629.5951 | 635.064 | 3.8223 | 8.806 | -0.29402 | 9.5999 | 629.5951 | -8.806 | 64.0015
0.6902 | -81.3566 | -6.1157 | -614.2911 | 619.8041 | 3.7635 | 8.6296 | -0.29402 | 9.4235 | 614.2911 | -8.6296 | 64.1688
0.692 | -80.5481 | -7.8652 | -599.9133 | 605.4557 | 3.7194 | 8.4532 | -0.30873 | 9.2471 | 599.9133 | -8.4532 | 64.3362
0.6938 | -79.7983 | -9.3353 | -586.7704 | 592.3275 | 3.6606 | 8.2915 | -0.32343 | 9.0707 | 586.7704 | -8.2915 | 64.5035
0.6956 | -78.7839 | -10.7466 | -573.8186 | 579.4051 | 3.6165 | 8.1298 | -0.32343 | 8.9089 | 573.8186 | -8.1298 | 64.6709
0.6974 | -78.1371 | -11.6581 | -561.9841 | 567.6 | 3.5577 | 7.9828 | -0.32343 | 8.7472 | 561.9841 | -7.9828 | 64.8382
0.6992 | -77.6372 | -12.349 | -551.1934 | 556.8681 | 3.4989 | 7.8505 | -0.32343 | 8.6002 | 551.1934 | -7.8505 | 65.0056
0.7009 | -77.1815 | -13.0841 | -541.226 | 546.9448 | 3.4548 | 7.7329 | -0.32343 | 8.4826 | 541.226 | -7.7329 | 65.1636
0.7027 | -76.814 | -13.7751 | -532.0671 | 537.83 | 3.396 | 7.6152 | -0.30873 | 8.3503 | 532.0671 | -7.6152 | 65.331
0.7045 | -76.3729 | -14.1867 | -523.6139 | 529.4209 | 3.3519 | 7.5123 | -0.30873 | 8.2327 | 523.6139 | -7.5123 | 65.4983
0.7063 | -75.7555 | -14.4954 | -515.9252 | 521.7175 | 3.2931 | 7.4241 | -0.29402 | 8.1298 | 515.9252 | -7.4241 | 65.6657
0.7081 | -75.0939 | -14.7159 | -509.0156 | 514.7932 | 3.2343 | 7.3212 | -0.29402 | 8.0122 | 509.0156 | -7.3212 | 65.833
0.7099 | -74.6088 | -14.5248 | -502.8264 | 508.604 | 3.1755 | 7.2477 | -0.26462 | 7.9093 | 502.8264 | -7.2477 | 66.0004
0.7117 | -74.0207 | -14.3337 | -496.9459 | 502.6941 | 3.1167 | 7.1742 | -0.24992 | 7.8211 | 496.9459 | -7.1742 | 66.1677
0.7135 | -73.1386 | -14.025 | -491.0801 | 496.7548 | 3.0726 | 7.086 | -0.22052 | 7.7329 | 491.0801 | -7.086 | 66.3351
0.7153 | -72.4624 | -13.6574 | -486.1111 | 491.7269 | 3.0138 | 7.0125 | -0.20582 | 7.6446 | 486.1111 | -7.0125 | 66.5024
0.7171 | -71.7714 | -13.0547 | -481.8624 | 487.4195 | 2.9696 | 6.9537 | -0.17641 | 7.5564 | 481.8624 | -6.9537 | 66.6698
0.7189 | -70.9335 | -12.4372 | -477.8343 | 483.3031 | 2.9108 | 6.8802 | -0.14701 | 7.4829 | 477.8343 | -6.8802 | 66.8371
0.7207 | -70.4189 | -11.6287 | -474.0414 | 479.4661 | 2.8667 | 6.8214 | -0.11761 | 7.4094 | 474.0414 | -6.8214 | 67.0045
0.7225 | -69.8456 | -10.6437 | -470.3954 | 475.7467 | 2.8226 | 6.7626 | -0.088207 | 7.3359 | 470.3954 | -6.7626 | 67.1718
0.7242 | -69.7133 | -9.5852 | -467.5287 | 472.88 | 2.7785 | 6.7332 | -0.044104 | 7.2918 | 467.5287 | -6.7332 | 67.3299
0.726 | -69.5809 | -8.4826 | -465.2206 | 470.5572 | 2.7491 | 6.7185 | -0.014701 | 7.2624 | 465.2206 | -6.7185 | 67.4972
0.7278 | -69.1987 | -7.674 | -462.2657 | 467.5728 | 2.705 | 6.6744 | 0.014701 | 7.2036 | 462.2657 | -6.6744 | 67.6646
0.7296 | -68.5078 | -7.1889 | -458.5021 | 463.7505 | 2.6462 | 6.6156 | 0.044104 | 7.1301 | 458.5021 | -6.6156 | 67.8319
0.7314 | -67.7727 | -6.6891 | -454.4446 | 459.6194 | 2.6021 | 6.542 | 0.073506 | 7.0419 | 454.4446 | -6.542 | 67.9993
0.7332 | -66.9053 | -6.3509 | -450.2842 | 455.3708 | 2.5433 | 6.4685 | 0.10291 | 6.9537 | 450.2842 | -6.4685 | 68.1666
0.735 | -65.6557 | -6.2186 | -445.7709 | 450.7252 | 2.4845 | 6.3656 | 0.11761 | 6.8508 | 445.7709 | -6.3656 | 68.334
0.7368 | -64.1268 | -6.2186 | -440.8607 | 445.6386 | 2.4257 | 6.2627 | 0.14701 | 6.7185 | 440.8607 | -6.2627 | 68.5013
0.7386 | -62.3185 | -6.3509 | -435.6564 | 440.2432 | 2.3522 | 6.1304 | 0.16171 | 6.5862 | 435.6564 | -6.1304 | 68.6687
0.7404 | -60.3927 | -6.5715 | -430.0553 | 434.4215 | 2.2787 | 5.9981 | 0.17641 | 6.4244 | 430.0553 | -5.9981 | 68.836
0.7422 | -58.4962 | -6.792 | -424.1601 | 428.3352 | 2.2199 | 5.8658 | 0.19112 | 6.2774 | 424.1601 | -5.8658 | 69.0033
0.744 | -56.5704 | -7.0566 | -418.0443 | 422.0137 | 2.1464 | 5.7188 | 0.20582 | 6.1157 | 418.0443 | -5.7188 | 69.1707
0.7458 | -54.5857 | -7.3653 | -411.6346 | 415.4128 | 2.0729 | 5.5571 | 0.22052 | 5.954 | 411.6346 | -5.5571 | 69.338
0.7476 | -52.5863 | -7.6299 | -404.9161 | 408.4885 | 2.0141 | 5.4101 | 0.22052 | 5.7776 | 404.9161 | -5.4101 | 69.5054
0.7493 | -50.6017 | -7.7623 | -398.1389 | 401.5202 | 1.9553 | 5.2483 | 0.23522 | 5.6159 | 398.1389 | -5.2483 | 69.6634
0.7511 | -48.7493 | -7.8946 | -391.7144 | 394.9193 | 1.8965 | 5.1013 | 0.23522 | 5.4542 | 391.7144 | -5.1013 | 69.8308
0.7529 | -46.9557 | -7.8358 | -385.5399 | 388.5831 | 1.8524 | 4.9396 | 0.24992 | 5.2924 | 385.5399 | -4.9396 | 69.9981
0.7547 | -45.1769 | -7.7917 | -379.3801 | 382.2321 | 1.7936 | 4.7926 | 0.24992 | 5.1307 | 379.3801 | -4.7926 | 70.1655
0.7565 | -43.4569 | -7.7181 | -373.2056 | 375.8959 | 1.7494 | 4.6456 | 0.26462 | 4.969 | 373.2056 | -4.6456 | 70.3328
0.7583 | -41.8397 | -7.6005 | -367.0605 | 369.6185 | 1.7053 | 4.4839 | 0.26462 | 4.822 | 367.0605 | -4.4839 | 70.5002
0.7601 | -40.4431 | -7.5123 | -361.0182 | 363.444 | 1.6612 | 4.3516 | 0.27932 | 4.675 | 361.0182 | -4.3516 | 70.6675
0.7619 | -39.1494 | -7.4241 | -355.1231 | 357.4458 | 1.6318 | 4.2193 | 0.27932 | 4.5427 | 355.1231 | -4.2193 | 70.8349
0.7637 | -38.0468 | -7.3506 | -349.1985 | 351.433 | 1.5877 | 4.0869 | 0.27932 | 4.3957 | 349.1985 | -4.0869 | 71.0022
0.7655 | -37.0177 | -7.2918 | -343.3033 | 345.4643 | 1.5436 | 3.9693 | 0.29402 | 4.2781 | 343.3033 | -3.9693 | 71.1696
0.7673 | -36.1062 | -7.1448 | -337.6727 | 339.775 | 1.5142 | 3.837 | 0.29402 | 4.1457 | 337.6727 | -3.837 | 71.3369
0.7691 | -35.5329 | -6.7185 | -332.2626 | 334.3208 | 1.4848 | 3.7488 | 0.29402 | 4.0575 | 332.2626 | -3.7488 | 71.5043
0.7709 | -34.8566 | -6.3803 | -327.176 | 329.1901 | 1.4554 | 3.6459 | 0.30873 | 3.9546 | 327.176 | -3.6459 | 71.6716
0.7727 | -34.151 | -6.0716 | -322.2805 | 324.2505 | 1.426 | 3.543 | 0.30873 | 3.837 | 322.2805 | -3.543 | 71.839
0.7744 | -33.4894 | -5.7335 | -317.7231 | 319.6343 | 1.4113 | 3.4401 | 0.32343 | 3.7488 | 317.7231 | -3.4401 | 71.997
0.7762 | -33.019 | -5.263 | -313.8567 | 315.7237 | 1.3819 | 3.3519 | 0.32343 | 3.6606 | 313.8567 | -3.3519 | 72.1644
0.778 | -32.7691 | -4.6162 | -310.8282 | 312.6806 | 1.3672 | 3.2784 | 0.33813 | 3.5871 | 310.8282 | -3.2784 | 72.3317
0.7798 | -32.5926 | -4.2634 | -307.6087 | 309.4757 | 1.3378 | 3.2196 | 0.35283 | 3.5136 | 307.6087 | -3.2196 | 72.4991
0.7816 | -32.3868 | -3.9987 | -303.8304 | 305.6975 | 1.3231 | 3.1461 | 0.35283 | 3.4548 | 303.8304 | -3.1461 | 72.6664
0.7834 | -32.0781 | -3.8076 | -299.4201 | 301.2724 | 1.2937 | 3.0873 | 0.35283 | 3.3813 | 299.4201 | -3.0873 | 72.8338
0.7852 | -31.7694 | -3.4842 | -294.8627 | 296.715 | 1.2643 | 3.0138 | 0.35283 | 3.3078 | 294.8627 | -3.0138 | 73.0011
0.787 | -31.6371 | -2.9549 | -290.4082 | 292.2753 | 1.2496 | 2.9549 | 0.35283 | 3.249 | 290.4082 | -2.9549 | 73.1685
0.7888 | -31.49 | -2.4257 | -286.1302 | 287.9972 | 1.2349 | 2.8961 | 0.35283 | 3.1902 | 286.1302 | -2.8961 | 73.3358
0.7906 | -31.3283 | -1.97 | -282.249 | 284.1308 | 1.2202 | 2.852 | 0.35283 | 3.1461 | 282.249 | -2.852 | 73.5032
0.7924 | -31.4753 | -1.3672 | -279.0148 | 280.9259 | 1.2055 | 2.8226 | 0.36753 | 3.102 | 279.0148 | -2.8226 | 73.6705
0.7942 | -31.9311 | -0.70566 | -276.5449 | 278.5443 | 1.1908 | 2.8079 | 0.36753 | 3.102 | 276.5449 | -2.8079 | 73.8379
0.796 | -32.475 | -0.27932 | -274.1339 | 276.2215 | 1.1761 | 2.8079 | 0.38223 | 3.0873 | 274.1339 | -2.8079 | 74.0052
0.7977 | -32.5044 | -0.073506 | -271.87 | 274.0016 | 1.1467 | 2.8079 | 0.38223 | 3.0726 | 271.87 | -2.8079 | 74.1633
0.7995 | -31.9899 | 0.029402 | -268.915 | 271.0026 | 1.1173 | 2.7785 | 0.36753 | 3.0432 | 268.915 | -2.7785 | 74.3306
0.8013 | -31.9311 | 0.23522 | -265.2103 | 267.3126 | 1.0879 | 2.7491 | 0.38223 | 2.9991 | 265.2103 | -2.7491 | 74.498
0.8031 | -31.7547 | 0.44104 | -261.4174 | 263.5344 | 1.0585 | 2.705 | 0.38223 | 2.9549 | 261.4174 | -2.705 | 74.6653
0.8049 | -31.2107 | 0.41163 | -256.8453 | 258.9329 | 1.0144 | 2.6462 | 0.38223 | 2.8814 | 256.8453 | -2.6462 | 74.8327
0.8067 | -30.9167 | 0.47044 | -252.5231 | 254.6254 | 0.97028 | 2.6021 | 0.38223 | 2.8373 | 252.5231 | -2.6021 | 75
0.8085 | -30.8579 | 0.72036 | -248.4803 | 250.5973 | 0.94088 | 2.5874 | 0.38223 | 2.8079 | 248.4803 | -2.5874 | 75.1673
0.8103 | -30.6668 | 0.98498 | -245.2754 | 247.3924 | 0.91148 | 2.5727 | 0.38223 | 2.7785 | 245.2754 | -2.5727 | 75.3347
0.8121 | -30.7697 | 1.4407 | -242.541 | 244.7021 | 0.88207 | 2.5727 | 0.38223 | 2.7638 | 242.541 | -2.5727 | 75.502
0.8139 | -30.8432 | 1.9406 | -240.0418 | 242.247 | 0.86737 | 2.558 | 0.38223 | 2.7491 | 240.0418 | -2.558 | 75.6694
0.8157 | -30.9755 | 2.5139 | -237.7631 | 240.0124 | 0.83797 | 2.558 | 0.38223 | 2.7491 | 237.7631 | -2.558 | 75.8367
0.8175 | -31.3871 | 3.1755 | -235.8225 | 238.1453 | 0.82327 | 2.5727 | 0.39693 | 2.7491 | 235.8225 | -2.5727 | 76.0041
0.8193 | -31.6224 | 3.5871 | -233.7349 | 236.1313 | 0.80857 | 2.5874 | 0.39693 | 2.7638 | 233.7349 | -2.5874 | 76.1714
0.8211 | -31.5636 | 3.8517 | -231.471 | 233.882 | 0.77917 | 2.5874 | 0.39693 | 2.7491 | 231.471 | -2.5874 | 76.3388
0.8228 | -31.3871 | 4.1899 | -229.4275 | 231.8238 | 0.74976 | 2.5874 | 0.39693 | 2.7491 | 229.4275 | -2.5874 | 76.4968
0.8246 | -31.2107 | 4.6603 | -227.678 | 230.0743 | 0.72036 | 2.5874 | 0.39693 | 2.7344 | 227.678 | -2.5874 | 76.6642
0.8264 | -31.0196 | 5.1895 | -226.1197 | 228.5013 | 0.69096 | 2.5874 | 0.41163 | 2.7344 | 226.1197 | -2.5874 | 76.8315
0.8282 | -30.7109 | 5.6012 | -224.679 | 227.0018 | 0.66156 | 2.6021 | 0.41163 | 2.7344 | 224.679 | -2.6021 | 76.9989
0.83 | -30.1375 | 5.7629 | -223.1942 | 225.4582 | 0.61745 | 2.5874 | 0.42634 | 2.7344 | 223.1942 | -2.5874 | 77.1662
0.8318 | -29.9317 | 6.3068 | -221.9593 | 224.2085 | 0.58805 | 2.6021 | 0.42634 | 2.7344 | 221.9593 | -2.6021 | 77.3336
0.8336 | -29.8729 | 6.9537 | -220.842 | 223.1207 | 0.55865 | 2.6168 | 0.44104 | 2.7344 | 220.842 | -2.6168 | 77.5009
0.8354 | -30.1963 | 7.8652 | -220.1363 | 222.5032 | 0.54395 | 2.6609 | 0.45574 | 2.7785 | 220.1363 | -2.6609 | 77.6683
0.8372 | -30.4022 | 8.6443 | -219.3571 | 221.8122 | 0.52924 | 2.705 | 0.45574 | 2.8079 | 219.3571 | -2.705 | 77.8356
0.839 | -30.4463 | 9.2912 | -218.6956 | 221.2095 | 0.51454 | 2.7344 | 0.47044 | 2.852 | 218.6956 | -2.7344 | 78.003
0.8408 | -30.4022 | 9.8939 | -218.1516 | 220.695 | 0.48514 | 2.7785 | 0.48514 | 2.8814 | 218.1516 | -2.7785 | 78.1703
0.8426 | -30.3728 | 10.629 | -218.2839 | 220.842 | 0.45574 | 2.8079 | 0.49984 | 2.9108 | 218.2839 | -2.8079 | 78.3377
0.8444 | -30.6374 | 11.5111 | -218.9896 | 221.6064 | 0.44104 | 2.852 | 0.49984 | 2.9402 | 218.9896 | -2.852 | 78.505
0.8462 | -31.1519 | 12.5108 | -219.857 | 222.5767 | 0.44104 | 2.8961 | 0.51454 | 2.9991 | 219.857 | -2.8961 | 78.6724
0.8479 | -31.49 | 13.334 | -220.7244 | 223.5176 | 0.41163 | 2.9402 | 0.52924 | 3.0432 | 220.7244 | -2.9402 | 78.8304
0.8497 | -31.5342 | 13.8927 | -221.4741 | 224.2968 | 0.38223 | 2.9844 | 0.54395 | 3.0726 | 221.4741 | -2.9844 | 78.9978
0.8515 | -31.5489 | 14.5983 | -222.3268 | 225.1788 | 0.36753 | 3.0138 | 0.55865 | 3.102 | 222.3268 | -3.0138 | 79.1651
0.8533 | -31.2548 | 15.01 | -223.0324 | 225.8551 | 0.33813 | 3.0432 | 0.55865 | 3.1314 | 223.0324 | -3.0432 | 79.3325
0.8551 | -31.2401 | 15.6568 | -223.5617 | 226.4284 | 0.30873 | 3.0873 | 0.57335 | 3.1608 | 223.5617 | -3.0873 | 79.4998
0.8569 | -31.0196 | 16.1126 | -223.8263 | 226.6931 | 0.26462 | 3.1167 | 0.58805 | 3.1902 | 223.8263 | -3.1167 | 79.6672
0.8587 | -30.2257 | 16.2596 | -223.694 | 226.4578 | 0.20582 | 3.1167 | 0.58805 | 3.1902 | 223.694 | -3.1167 | 79.8345
0.8605 | -29.329 | 16.3772 | -223.547 | 226.2079 | 0.16171 | 3.1167 | 0.60275 | 3.1902 | 223.547 | -3.1167 | 80.0019
0.8623 | -28.5645 | 16.583 | -224.0468 | 226.6196 | 0.11761 | 3.1314 | 0.60275 | 3.1902 | 224.0468 | -3.1314 | 80.1692
0.8641 | -27.7265 | 16.8182 | -224.7231 | 227.1929 | 0.058805 | 3.1314 | 0.61745 | 3.2049 | 224.7231 | -3.1314 | 80.3366
0.8659 | -26.7415 | 17.0975 | -225.7669 | 228.1338 | 0.029402 | 3.1314 | 0.61745 | 3.2049 | 225.7669 | -3.1314 | 80.5039
0.8677 | -25.8595 | 17.4651 | -226.8548 | 229.1335 | -0.014701 | 3.1167 | 0.63215 | 3.1902 | 226.8548 | -3.1167 | 80.6713
0.8695 | -24.9774 | 17.8326 | -227.9574 | 230.1478 | -0.044104 | 3.102 | 0.64685 | 3.1755 | 227.9574 | -3.102 | 80.8386
0.8712 | -24.2423 | 18.4059 | -229.4863 | 231.6327 | -0.058805 | 3.102 | 0.66156 | 3.1755 | 229.4863 | -3.102 | 80.9967
0.873 | -23.5661 | 18.9793 | -231.5739 | 233.6614 | -0.073506 | 3.0873 | 0.67626 | 3.1755 | 231.5739 | -3.0873 | 81.164
0.8748 | -23.228 | 19.7438 | -233.8673 | 235.9548 | -0.073506 | 3.102 | 0.67626 | 3.1902 | 233.8673 | -3.102 | 81.3313
0.8766 | -23.125 | 20.6993 | -236.44 | 238.5717 | -0.058805 | 3.1167 | 0.69096 | 3.2049 | 236.44 | -3.1167 | 81.4987
0.8784 | -22.9486 | 21.7725 | -239.3508 | 241.5266 | -0.044104 | 3.1461 | 0.70566 | 3.2343 | 239.3508 | -3.1461 | 81.666
0.8802 | -23.0662 | 22.8898 | -242.6292 | 244.9079 | -0.044104 | 3.1608 | 0.72036 | 3.2637 | 242.6292 | -3.1608 | 81.8334
0.882 | -23.1838 | 23.963 | -246.0546 | 248.4215 | -0.029402 | 3.1902 | 0.72036 | 3.2931 | 246.0546 | -3.1902 | 82.0007
0.8838 | -23.6249 | 25.1391 | -249.7152 | 252.2144 | -0.014701 | 3.2343 | 0.73506 | 3.3372 | 249.7152 | -3.2343 | 82.1681
0.8856 | -23.8895 | 26.2417 | -253.2876 | 255.8897 | -0.014701 | 3.2784 | 0.74976 | 3.3813 | 253.2876 | -3.2784 | 82.3354
0.8874 | -24.0365 | 27.2855 | -257.0511 | 259.7414 | 0 | 3.3078 | 0.76446 | 3.4107 | 257.0511 | -3.3078 | 82.5028
0.8892 | -24.3599 | 28.2558 | -261.0204 | 263.8137 | 0.014701 | 3.3372 | 0.76446 | 3.4401 | 261.0204 | -3.3372 | 82.6701
0.891 | -24.6099 | 29.079 | -264.7546 | 267.6213 | 0.044104 | 3.3519 | 0.77917 | 3.4548 | 264.7546 | -3.3519 | 82.8375
0.8928 | -24.7275 | 29.7112 | -267.6213 | 270.5321 | 0.044104 | 3.3519 | 0.77917 | 3.4548 | 267.6213 | -3.3519 | 83.0048
0.8946 | -24.9333 | 30.2698 | -270.2528 | 273.2225 | 0.073506 | 3.3519 | 0.77917 | 3.4548 | 270.2528 | -3.3519 | 83.1722
0.8963 | -25.1244 | 30.8432 | -273.0019 | 276.0304 | 0.088207 | 3.3519 | 0.76446 | 3.4548 | 273.0019 | -3.3519 | 83.3302
0.8981 | -25.5654 | 31.3871 | -275.4277 | 278.5296 | 0.10291 | 3.3666 | 0.76446 | 3.4695 | 275.4277 | -3.3666 | 83.4976
0.8999 | -25.8301 | 31.7841 | -277.3976 | 280.529 | 0.11761 | 3.3519 | 0.74976 | 3.4548 | 277.3976 | -3.3519 | 83.6649
0.9017 | -25.9771 | 31.9311 | -278.9412 | 282.0873 | 0.14701 | 3.3225 | 0.72036 | 3.4107 | 278.9412 | -3.3225 | 83.8323
0.9035 | -26.0359 | 31.7988 | -280.1468 | 283.2781 | 0.14701 | 3.2784 | 0.70566 | 3.3813 | 280.1468 | -3.2784 | 83.9996
0.9053 | -25.7713 | 31.4606 | -280.7495 | 283.8221 | 0.16171 | 3.2196 | 0.67626 | 3.3225 | 280.7495 | -3.2196 | 84.167
0.9071 | -25.3302 | 31.0049 | -281.2199 | 284.2043 | 0.16171 | 3.1608 | 0.64685 | 3.2637 | 281.2199 | -3.1608 | 84.3343
0.9089 | -24.7422 | 30.4463 | -281.6463 | 284.513 | 0.16171 | 3.0726 | 0.61745 | 3.1755 | 281.6463 | -3.0726 | 84.5017
0.9107 | -24.2129 | 29.8141 | -281.8962 | 284.6453 | 0.14701 | 2.9991 | 0.58805 | 3.102 | 281.8962 | -2.9991 | 84.669
0.9125 | -23.4191 | 28.9614 | -281.7786 | 284.3954 | 0.14701 | 2.8961 | 0.55865 | 2.9991 | 281.7786 | -2.8961 | 84.8364
0.9143 | -22.7722 | 27.9323 | -281.3081 | 283.7927 | 0.14701 | 2.7932 | 0.52924 | 2.8814 | 281.3081 | -2.7932 | 85.0037
0.9161 | -22.3165 | 27.0062 | -280.8377 | 283.2193 | 0.13231 | 2.6903 | 0.49984 | 2.7638 | 280.8377 | -2.6903 | 85.1711
0.9179 | -22.1254 | 26.1829 | -280.6172 | 282.9253 | 0.14701 | 2.5874 | 0.45574 | 2.6756 | 280.6172 | -2.5874 | 85.3384
0.9197 | -21.7284 | 25.095 | -280.132 | 282.3372 | 0.14701 | 2.4992 | 0.42634 | 2.5874 | 280.132 | -2.4992 | 85.5058
0.9214 | -20.9052 | 23.816 | -279.4264 | 281.4993 | 0.13231 | 2.3963 | 0.39693 | 2.4992 | 279.4264 | -2.3963 | 85.6638
0.9232 | -20.0819 | 22.4782 | -278.2062 | 280.132 | 0.13231 | 2.3081 | 0.36753 | 2.411 | 278.2062 | -2.3081 | 85.8312
0.925 | -19.2292 | 21.3609 | -276.6038 | 278.3973 | 0.11761 | 2.2052 | 0.33813 | 2.3375 | 276.6038 | -2.2052 | 85.9985
0.9268 | -18.5236 | 20.3465 | -274.575 | 276.2509 | 0.11761 | 2.117 | 0.30873 | 2.2787 | 274.575 | -2.117 | 86.1659
0.9286 | -17.5533 | 19.3174 | -273.6488 | 275.1924 | 0.11761 | 1.9994 | 0.27932 | 2.1758 | 273.6488 | -1.9994 | 86.3332
0.9304 | -16.4507 | 18.2295 | -273.2519 | 274.6191 | 0.10291 | 1.8671 | 0.24992 | 2.0435 | 273.2519 | -1.8671 | 86.5006
0.9322 | -15.6715 | 17.3622 | -273.3254 | 274.5603 | 0.11761 | 1.7494 | 0.22052 | 1.9259 | 273.3254 | -1.7494 | 86.6679
0.934 | -15.2158 | 16.7888 | -273.7076 | 274.8543 | 0.14701 | 1.6465 | 0.19112 | 1.8524 | 273.7076 | -1.6465 | 86.8353
0.9358 | -15.0247 | 16.4213 | -274.3986 | 275.4865 | 0.16171 | 1.5583 | 0.17641 | 1.7788 | 274.3986 | -1.5583 | 87.0026
0.9376 | -14.9953 | 16.1273 | -275.6188 | 276.692 | 0.19112 | 1.4848 | 0.16171 | 1.72 | 275.6188 | -1.4848 | 87.17
0.9394 | -15.01 | 15.8479 | -276.9566 | 278.0004 | 0.22052 | 1.4113 | 0.16171 | 1.6612 | 276.9566 | -1.4113 | 87.3373
0.9412 | -15.2893 | 15.6421 | -278.6766 | 279.7204 | 0.24992 | 1.3819 | 0.14701 | 1.6465 | 278.6766 | -1.3819 | 87.5046
0.943 | -15.7009 | 15.4069 | -280.6613 | 281.7051 | 0.29402 | 1.3525 | 0.13231 | 1.6318 | 280.6613 | -1.3525 | 87.672
0.9447 | -15.9655 | 14.9953 | -282.6607 | 283.6897 | 0.32343 | 1.2937 | 0.13231 | 1.6024 | 282.6607 | -1.2937 | 87.83
0.9465 | -15.9214 | 14.3925 | -284.1896 | 285.1746 | 0.33813 | 1.2202 | 0.11761 | 1.5583 | 284.1896 | -1.2202 | 87.9974
0.9483 | -15.5833 | 13.628 | -284.9541 | 285.8655 | 0.35283 | 1.132 | 0.10291 | 1.5436 | 284.9541 | -1.132 | 88.1647
0.9501 | -14.9218 | 12.5696 | -285.057 | 285.9096 | 0.35283 | 1.0291 | 0.088207 | 1.5436 | 285.057 | -1.0291 | 88.3321
0.9519 | -14.2161 | 11.5405 | -284.5424 | 285.3216 | 0.35283 | 0.94088 | 0.088207 | 1.5877 | 284.5424 | -0.94088 | 88.4994
0.9537 | -13.8045 | 11.0112 | -284.513 | 285.2775 | 0.35283 | 0.85267 | 0.088207 | 1.6171 | 284.513 | -0.85267 | 88.6668
0.9555 | -13.8486 | 10.5555 | -283.925 | 284.6747 | 0.36753 | 0.82327 | 0.088207 | 1.6465 | 283.925 | -0.82327 | 88.8341
0.9573 | -13.7751 | 10.2027 | -284.1455 | 284.8952 | 0.38223 | 0.77917 | 0.088207 | 1.6318 | 284.1455 | -0.77917 | 89.0015
0.9591 | -13.6869 | 10.0703 | -284.6012 | 285.3363 | 0.39693 | 0.73506 | 0.073506 | 1.6024 | 284.6012 | -0.73506 | 89.1688
0.9609 | -13.731 | 10.3203 | -285.645 | 286.3948 | 0.42634 | 0.69096 | 0.073506 | 1.6024 | 285.645 | -0.69096 | 89.3362
0.9627 | -13.8927 | 10.8348 | -287.3063 | 288.056 | 0.45574 | 0.63215 | 0.073506 | 1.6024 | 287.3063 | -0.63215 | 89.5035
0.9645 | -14.2308 | 11.5405 | -289.9525 | 290.7463 | 0.48514 | 0.58805 | 0.073506 | 1.6024 | 289.9525 | -0.58805 | 89.6709
0.9663 | -14.6424 | 12.5549 | -293.2456 | 294.0835 | 0.52924 | 0.54395 | 0.073506 | 1.6171 | 293.2456 | -0.54395 | 89.8382
0.9681 | -15.0247 | 13.4516 | -296.8327 | 297.7 | 0.55865 | 0.51454 | 0.073506 | 1.6171 | 296.8327 | -0.51454 | 90.0056
0.9698 | -15.2893 | 14.2455 | -300.4933 | 301.4047 | 0.58805 | 0.49984 | 0.058805 | 1.6171 | 300.4933 | -0.49984 | 90.1636
0.9716 | -15.6421 | 15.0541 | -304.5802 | 305.5211 | 0.61745 | 0.45574 | 0.058805 | 1.6024 | 304.5802 | -0.45574 | 90.331
0.9734 | -15.7891 | 15.9214 | -309.4463 | 310.4313 | 0.64685 | 0.38223 | 0.044104 | 1.6024 | 309.4463 | -0.38223 | 90.4983
0.9752 | -15.7303 | 16.7447 | -314.7976 | 315.8119 | 0.66156 | 0.30873 | 0.029402 | 1.5877 | 314.7976 | -0.30873 | 90.6657
0.977 | -15.4951 | 17.4798 | -320.237 | 321.2808 | 0.67626 | 0.24992 | 0.014701 | 1.5877 | 320.237 | -0.24992 | 90.833
0.9788 | -15.0394 | 18.3765 | -326.0146 | 327.0584 | 0.69096 | 0.16171 | 0 | 1.6024 | 326.0146 | -0.16171 | 91.0004
0.9806 | -14.6571 | 19.5967 | -332.4243 | 333.4828 | 0.70566 | 0.058805 | -0.014701 | 1.6171 | 332.4243 | -0.058805 | 91.1677
0.9824 | -14.3337 | 21.0375 | -339.4368 | 340.5394 | 0.72036 | -0.044104 | -0.029402 | 1.6465 | 339.4368 | 0.044104 | 91.3351
0.9842 | -14.0103 | 22.6399 | -346.9051 | 348.0665 | 0.73506 | -0.16171 | -0.058805 | 1.6759 | 346.9051 | 0.16171 | 91.5024
0.986 | -13.8339 | 23.963 | -354.2998 | 355.52 | 0.74976 | -0.29402 | -0.073506 | 1.6612 | 354.2998 | 0.29402 | 91.6698
0.9878 | -13.334 | 25.1832 | -361.768 | 363.0176 | 0.74976 | -0.42634 | -0.10291 | 1.6759 | 361.768 | 0.42634 | 91.8371
0.9896 | -12.8342 | 26.6533 | -369.7655 | 371.0592 | 0.76446 | -0.57335 | -0.13231 | 1.7053 | 369.7655 | 0.57335 | 92.0045
0.9914 | -12.2314 | 28.3881 | -378.6891 | 380.0563 | 0.76446 | -0.73506 | -0.16171 | 1.7494 | 378.6891 | 0.73506 | 92.1718
0.9932 | -11.4376 | 30.1522 | -388.5684 | 390.0091 | 0.77917 | -0.91148 | -0.17641 | 1.7788 | 388.5684 | 0.91148 | 92.3392
0.9949 | -10.6731 | 32.1516 | -399.1533 | 400.6822 | 0.77917 | -1.0879 | -0.20582 | 1.8524 | 399.1533 | 1.0879 | 92.4972
0.9967 | -10.2468 | 34.445 | -410.1057 | 411.7816 | 0.80857 | -1.2496 | -0.23522 | 1.9406 | 410.1057 | 1.2496 | 92.6646
0.9985 | -10.2174 | 37.0177 | -421.5579 | 423.3956 | 0.82327 | -1.4113 | -0.26462 | 2.0435 | 421.5579 | 1.4113 | 92.8319
1.0003 | -10.3056 | 39.8109 | -433.51 | 435.5535 | 0.85267 | -1.5583 | -0.29402 | 2.1611 | 433.51 | 1.5583 | 92.9993
1.0021 | -10.1586 | 42.6336 | -446.1825 | 448.4318 | 0.85267 | -1.7347 | -0.33813 | 2.2934 | 446.1825 | 1.7347 | 93.1666
1.0039 | -9.6587 | 45.3239 | -459.5459 | 461.9863 | 0.86737 | -1.9553 | -0.38223 | 2.4551 | 459.5459 | 1.9553 | 93.334
1.0057 | -9.0119 | 47.9848 | -473.2916 | 475.9084 | 0.86737 | -2.1758 | -0.42634 | 2.6168 | 473.2916 | 2.1758 | 93.5013
1.0075 | -8.2033 | 50.484 | -487.4048 | 490.198 | 0.86737 | -2.3963 | -0.48514 | 2.7932 | 487.4048 | 2.3963 | 93.6687
1.0093 | -7.2477 | 52.8656 | -501.4739 | 504.4141 | 0.83797 | -2.6462 | -0.54395 | 2.9991 | 501.4739 | 2.6462 | 93.836
1.0111 | -6.395 | 55.5266 | -515.4694 | 518.6008 | 0.82327 | -2.8961 | -0.60275 | 3.2196 | 515.4694 | 2.8961 | 94.0033
1.0129 | -5.4248 | 58.2169 | -530.8028 | 534.1106 | 0.79387 | -3.1755 | -0.66156 | 3.4695 | 530.8028 | 3.1755 | 94.1707
1.0147 | -4.4986 | 61.0248 | -546.7978 | 550.3114 | 0.77917 | -3.4401 | -0.73506 | 3.7194 | 546.7978 | 3.4401 | 94.338
1.0165 | -3.4695 | 63.7887 | -563.4249 | 567.1443 | 0.74976 | -3.7341 | -0.79387 | 3.984 | 563.4249 | 3.7341 | 94.5054
1.0182 | -2.264 | 66.4937 | -580.4636 | 584.3741 | 0.70566 | -4.0134 | -0.88207 | 4.2634 | 580.4636 | 4.0134 | 94.6634
1.02 | -1.1761 | 69.2575 | -597.2818 | 601.4129 | 0.66156 | -4.2928 | -0.95558 | 4.528 | 597.2818 | 4.2928 | 94.8308
1.0218 | -0.20582 | 71.9184 | -613.8942 | 618.2164 | 0.61745 | -4.5721 | -1.0291 | 4.8073 | 613.8942 | 4.5721 | 94.9981
1.0236 | 0.95558 | 74.3883 | -630.4772 | 634.9905 | 0.55865 | -4.8661 | -1.1173 | 5.0866 | 630.4772 | 4.8661 | 95.1655
1.0254 | 2.2934 | 76.7552 | -647.1484 | 651.8234 | 0.49984 | -5.1601 | -1.1908 | 5.3954 | 647.1484 | 5.1601 | 95.3328
1.0272 | 3.7635 | 78.9162 | -664.0254 | 668.8474 | 0.42634 | -5.4689 | -1.279 | 5.6894 | 664.0254 | 5.4689 | 95.5002
1.029 | 5.6159 | 80.9891 | -681.6963 | 686.68 | 0.35283 | -5.7629 | -1.3819 | 5.9981 | 681.6963 | 5.7629 | 95.6675
1.0308 | 7.7034 | 82.915 | -699.867 | 704.9831 | 0.29402 | -6.0422 | -1.4701 | 6.2921 | 699.867 | 6.0422 | 95.8349
1.0326 | 9.8204 | 84.8996 | -718.9051 | 724.1682 | 0.22052 | -6.3362 | -1.573 | 6.6009 | 718.9051 | 6.3362 | 96.0022
1.0344 | 11.8345 | 87.3694 | -739.3546 | 744.794 | 0.16171 | -6.6744 | -1.6612 | 6.9243 | 739.3546 | 6.6744 | 96.1696
1.0362 | 13.9809 | 89.2659 | -758.4515 | 764.0085 | 0.088207 | -6.9831 | -1.7641 | 7.2624 | 758.4515 | 6.9831 | 96.3369
1.038 | 16.0537 | 90.7801 | -776.3576 | 782.0028 | 0.029402 | -7.2624 | -1.8671 | 7.5564 | 776.3576 | 7.2624 | 96.5043
1.0398 | 18.2001 | 91.8533 | -792.8671 | 798.5858 | -0.014701 | -7.5417 | -1.9406 | 7.8358 | 792.8671 | 7.5417 | 96.6716
1.0416 | 20.714 | 92.4855 | -808.6121 | 814.3603 | -0.044104 | -7.8211 | -2.0141 | 8.1151 | 808.6121 | 7.8211 | 96.839
1.0433 | 23.3456 | 92.9118 | -824.1219 | 829.8995 | -0.073506 | -8.071 | -2.0729 | 8.3797 | 824.1219 | 8.071 | 96.997
1.0451 | 25.7419 | 93.6616 | -839.4406 | 845.277 | -0.11761 | -8.3062 | -2.1464 | 8.6296 | 839.4406 | 8.3062 | 97.1644
1.0469 | 28.3587 | 93.7792 | -853.936 | 859.7577 | -0.14701 | -8.512 | -2.2346 | 8.8501 | 853.936 | 8.512 | 97.3317
1.0487 | 30.6815 | 93.6763 | -866.9172 | 872.7095 | -0.17641 | -8.7178 | -2.3081 | 9.056 | 866.9172 | 8.7178 | 97.4991
1.0505 | 32.9602 | 93.2058 | -879.0163 | 884.7498 | -0.19112 | -8.9237 | -2.3669 | 9.2765 | 879.0163 | 8.9237 | 97.6664
1.0523 | 35.1654 | 92.2356 | -889.8658 | 895.4817 | -0.16171 | -9.0854 | -2.4257 | 9.4529 | 889.8658 | 9.0854 | 97.8338
1.0541 | 37.547 | 90.9859 | -900.2743 | 905.7872 | -0.11761 | -9.2324 | -2.4698 | 9.5999 | 900.2743 | 9.2324 | 98.0011
1.0559 | 39.9139 | 89.6334 | -910.1829 | 915.593 | -0.029402 | -9.3353 | -2.5139 | 9.7175 | 910.1829 | 9.3353 | 98.1685
1.0577 | 41.9867 | 88.3103 | -918.9154 | 924.2226 | 0.058805 | -9.4088 | -2.558 | 9.791 | 918.9154 | 9.4088 | 98.3358
1.0595 | 42.8688 | 86.3845 | -921.9145 | 927.0746 | 0.22052 | -9.3941 | -2.558 | 9.791 | 921.9145 | 9.3941 | 98.5032
1.0613 | 42.4572 | 84.5909 | -921.4882 | 926.4719 | 0.38223 | -9.3353 | -2.5286 | 9.7322 | 921.4882 | 9.3353 | 98.6705
1.0631 | 41.1488 | 82.6062 | -917.5335 | 922.3114 | 0.54395 | -9.1736 | -2.4992 | 9.5852 | 917.5335 | 9.1736 | 98.8379
1.0649 | 39.8698 | 80.6657 | -913.4466 | 918.0187 | 0.70566 | -9.0266 | -2.4698 | 9.4529 | 913.4466 | 9.0266 | 99.0052
1.0667 | 38.679 | 78.5634 | -909.1244 | 913.4907 | 0.85267 | -8.8648 | -2.4404 | 9.3059 | 909.1244 | 8.8648 | 99.1726
1.0684 | 36.9001 | 76.0936 | -901.8326 | 905.9343 | 0.98498 | -8.6296 | -2.3816 | 9.0854 | 901.8326 | 8.6296 | 99.3306
1.0702 | 35.1213 | 73.6238 | -894.5849 | 898.4366 | 1.1173 | -8.3944 | -2.3375 | 8.8501 | 894.5849 | 8.3944 | 99.498
1.072 | 33.607 | 71.1687 | -888.7338 | 892.365 | 1.2496 | -8.1592 | -2.2787 | 8.6296 | 888.7338 | 8.1592 | 99.6653
1.0738 | 32.4162 | 68.7871 | -885.3231 | 888.7338 | 1.3966 | -7.9681 | -2.2199 | 8.4532 | 885.3231 | 7.9681 | 99.8327
1.0756 | 31.2695 | 66.3173 | -883.265 | 886.4698 | 1.5289 | -7.7917 | -2.1611 | 8.2915 | 883.265 | 7.7917 | 100
| Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My
Minimum | -121.6821 | -239.2773 | -3379.3731 | 220.695 | -0.19112 | -36.1062 | -16.2008 | 1.5436 | 218.1516 | -14.1867
Maximum | 42.8688 | 93.7792 | -218.1516 | 3380.5345 | 24.4923 | 14.1867 | 2.1023 | 43.2804 | 3379.3731 | 36.1062

## Diagr_HIGH100
HIGH100 = Loads at High Level in Subject with 100kg Body Weight.
Contact forces and moments, measured telemetrically in knee implants of 8 subjects.
Averages of one individual subject.
Fx  lateral, Fy  anterior, -Fz  distal, Fres = Resultant Force
Mx around x, My around y, Mz around z, Mres = Resultant Moment
z-axis = axis of implant stem!
From Bergmann et al, PlosOne (2014) "Standardized Loads Acting in Knee Implants"
Subject K5R
Subject K5R
G.Bergmann et al., PlosOne, 2014
G.Bergmann et al., PlosOne, 2014