## Dat_Individ
Data File: | k9l_2246_863_6_i1.akf
Curve Alignment Program: | 1.4.11
------------------------MEASURING INFOS------------------------
Diagram Title #1 | Forces and Moments at Knee Joint
Diagram Title #2 | Measured Relative to Implant
Comment #1 | Knee Joint; Gaitanalysis; Stand up; Seat Height: 45 cm;
Comment #2 | K9L, 14,5 Months PO *
BodyWeight [N]: | 1070
Patient(en): | K9L | [7 empty] | Subject | K9L
Implant Type: | Knee Joint
Varus Angle Leg [°] | 7 | [7 empty] | Individual loads in N (forces) and Nm (moments)
Axes Directions X, Y, Z: | Lateral, Anterior, Superior
Axes Origin: | Tibial Plateau at Stem Axis | [7 empty] | BW [kg] = | 109.0724
Transformation X Y Z (°): | 6.5 | -3.5 | -2
Displacement X, Y, Z (mm): | 0 | 0 | 0
Activity Code: | KneeActivities | 1P_1_15_8 | 2P_1_31_2
---------------------------DATA--------------------------------
Number of Data Sets = | 601
Number of analog channels = | 0
Time-Offset(sec) = | 0
Time | Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My | Time
[s] | [N] | [N] | [N] | [N] | [Nm] | [Nm] | [Nm] | [Nm] | [N] | (Nm] | [%Cycle]
0 | -1.14 | -7.75 | -151.53 | 151.81 | 0.11 | -0.36 | -0.29 | 0.52 | 151.53 | 0.36 | 0
0.0018 | -1.22 | -7.31 | -152.63 | 152.87 | 0.12 | -0.36 | -0.29 | 0.52 | 152.63 | 0.36 | 0.16735
0.0036 | -1.54 | -6.94 | -154.61 | 154.83 | 0.13 | -0.35 | -0.3 | 0.52 | 154.61 | 0.35 | 0.3347
0.0054 | -1.83 | -6.47 | -156.13 | 156.32 | 0.15 | -0.34 | -0.3 | 0.52 | 156.13 | 0.34 | 0.50205
0.0072 | -2 | -5.71 | -156.21 | 156.37 | 0.17 | -0.35 | -0.3 | 0.54 | 156.21 | 0.35 | 0.66939
0.009 | -1.78 | -5.68 | -154.5 | 154.66 | 0.17 | -0.35 | -0.3 | 0.54 | 154.5 | 0.35 | 0.83674
0.0108 | -1.65 | -5.7 | -153.33 | 153.48 | 0.16 | -0.35 | -0.31 | 0.54 | 153.33 | 0.35 | 1.0041
0.0125 | -1.95 | -5.07 | -154.49 | 154.63 | 0.17 | -0.34 | -0.31 | 0.54 | 154.49 | 0.34 | 1.1621
0.0143 | -2.13 | -4.37 | -154.83 | 154.95 | 0.19 | -0.34 | -0.32 | 0.55 | 154.83 | 0.34 | 1.3295
0.0161 | -2.11 | -3.64 | -153.77 | 153.88 | 0.21 | -0.34 | -0.32 | 0.56 | 153.77 | 0.34 | 1.4968
0.0179 | -1.92 | -3.08 | -151.73 | 151.82 | 0.22 | -0.35 | -0.32 | 0.58 | 151.73 | 0.35 | 1.6642
0.0197 | -1.57 | -2.78 | -149.12 | 149.2 | 0.22 | -0.35 | -0.33 | 0.58 | 149.12 | 0.35 | 1.8315
0.0215 | -1.07 | -2.37 | -145.46 | 145.53 | 0.22 | -0.37 | -0.33 | 0.6 | 145.46 | 0.37 | 1.9989
0.0233 | -0.43 | -2.55 | -141.72 | 141.79 | 0.19 | -0.39 | -0.34 | 0.61 | 141.72 | 0.39 | 2.1662
0.0251 | 0.19 | -3.08 | -138.63 | 138.71 | 0.15 | -0.41 | -0.34 | 0.62 | 138.63 | 0.41 | 2.3336
0.0269 | 0.67 | -3.88 | -136.62 | 136.73 | 0.1 | -0.42 | -0.35 | 0.62 | 136.62 | 0.42 | 2.5009
0.0287 | 0.94 | -4.36 | -136.14 | 136.27 | 0.06 | -0.42 | -0.35 | 0.61 | 136.14 | 0.42 | 2.6683
0.0305 | 1.17 | -4.71 | -136.98 | 137.12 | 0.02 | -0.43 | -0.36 | 0.62 | 136.98 | 0.43 | 2.8356
0.0323 | 1.22 | -4.54 | -138.75 | 138.9 | 0.01 | -0.44 | -0.37 | 0.62 | 138.75 | 0.44 | 3.003
0.0341 | 1.17 | -4.05 | -141.62 | 141.76 | 0.01 | -0.44 | -0.37 | 0.64 | 141.62 | 0.44 | 3.1703
0.0359 | 0.83 | -2.95 | -145.07 | 145.18 | 0.03 | -0.45 | -0.38 | 0.65 | 145.07 | 0.45 | 3.3377
0.0376 | 0.46 | -1.75 | -148.06 | 148.14 | 0.06 | -0.45 | -0.38 | 0.67 | 148.06 | 0.45 | 3.4957
0.0394 | 0.23 | -0.19 | -149.6 | 149.68 | 0.09 | -0.47 | -0.39 | 0.69 | 149.6 | 0.47 | 3.6631
0.0412 | 0.06 | 1.4 | -150.57 | 150.65 | 0.12 | -0.48 | -0.39 | 0.71 | 150.57 | 0.48 | 3.8304
0.043 | 0.04 | 2.78 | -150.06 | 150.17 | 0.14 | -0.5 | -0.39 | 0.73 | 150.06 | 0.5 | 3.9978
0.0448 | 0.12 | 3.81 | -148.63 | 148.76 | 0.15 | -0.53 | -0.39 | 0.75 | 148.63 | 0.53 | 4.1651
0.0466 | 0.27 | 4.87 | -146.4 | 146.57 | 0.16 | -0.56 | -0.39 | 0.77 | 146.4 | 0.56 | 4.3325
0.0484 | 0.52 | 5.48 | -143.4 | 143.61 | 0.15 | -0.58 | -0.39 | 0.78 | 143.4 | 0.58 | 4.4998
0.0502 | 0.95 | 5.81 | -139.6 | 139.86 | 0.12 | -0.6 | -0.38 | 0.79 | 139.6 | 0.6 | 4.6672
0.052 | 1.38 | 5.93 | -136.31 | 136.6 | 0.09 | -0.63 | -0.38 | 0.8 | 136.31 | 0.63 | 4.8345
0.0538 | 1.78 | 5.54 | -133.96 | 134.25 | 0.04 | -0.65 | -0.39 | 0.82 | 133.96 | 0.65 | 5.0019
0.0556 | 2.15 | 5.05 | -132.36 | 132.65 | -0.01 | -0.68 | -0.39 | 0.83 | 132.36 | 0.68 | 5.1692
0.0574 | 2.38 | 4.66 | -132.5 | 132.8 | -0.06 | -0.7 | -0.4 | 0.86 | 132.5 | 0.7 | 5.3366
0.0592 | 2.47 | 4.47 | -133.37 | 133.68 | -0.1 | -0.71 | -0.4 | 0.87 | 133.37 | 0.71 | 5.5039
0.061 | 2.51 | 4.69 | -135.16 | 135.49 | -0.13 | -0.73 | -0.41 | 0.9 | 135.16 | 0.73 | 5.6713
0.0627 | 2.35 | 5.63 | -137.84 | 138.21 | -0.13 | -0.74 | -0.42 | 0.92 | 137.84 | 0.74 | 5.8293
0.0645 | 2.16 | 7.16 | -140.55 | 141.04 | -0.11 | -0.76 | -0.43 | 0.95 | 140.55 | 0.76 | 5.9967
0.0663 | 1.97 | 8.7 | -141.92 | 142.51 | -0.1 | -0.79 | -0.44 | 0.97 | 141.92 | 0.79 | 6.164
0.0681 | 1.82 | 10.13 | -142.56 | 143.22 | -0.08 | -0.81 | -0.44 | 1 | 142.56 | 0.81 | 6.3313
0.0699 | 1.85 | 11.49 | -141.63 | 142.39 | -0.06 | -0.84 | -0.45 | 1.03 | 141.63 | 0.84 | 6.4987
0.0717 | 1.96 | 12.76 | -140.22 | 141.09 | -0.05 | -0.87 | -0.46 | 1.06 | 140.22 | 0.87 | 6.666
0.0735 | 2.18 | 13.57 | -137.96 | 138.92 | -0.06 | -0.91 | -0.46 | 1.09 | 137.96 | 0.91 | 6.8334
0.0753 | 2.54 | 14.26 | -134.89 | 135.91 | -0.06 | -0.94 | -0.47 | 1.12 | 134.89 | 0.94 | 7.0007
0.0771 | 2.97 | 14.19 | -131.4 | 132.43 | -0.1 | -0.97 | -0.48 | 1.14 | 131.4 | 0.97 | 7.1681
0.0789 | 3.39 | 14.04 | -128.09 | 129.14 | -0.13 | -1 | -0.49 | 1.15 | 128.09 | 1 | 7.3354
0.0807 | 3.61 | 13.7 | -127.03 | 128.06 | -0.19 | -1.03 | -0.49 | 1.18 | 127.03 | 1.03 | 7.5028
0.0825 | 3.71 | 13.26 | -126.9 | 127.9 | -0.24 | -1.05 | -0.5 | 1.21 | 126.9 | 1.05 | 7.6701
0.0843 | 3.78 | 13.41 | -127.96 | 128.93 | -0.28 | -1.09 | -0.51 | 1.25 | 127.96 | 1.09 | 7.8375
0.086 | 3.77 | 13.37 | -130.35 | 131.31 | -0.31 | -1.11 | -0.52 | 1.28 | 130.35 | 1.11 | 7.9955
0.0878 | 3.63 | 13.85 | -133.72 | 134.72 | -0.33 | -1.12 | -0.53 | 1.31 | 133.72 | 1.12 | 8.1629
0.0896 | 3.41 | 14.97 | -137.41 | 138.51 | -0.31 | -1.14 | -0.54 | 1.33 | 137.41 | 1.14 | 8.3302
0.0914 | 3.24 | 16.28 | -140.32 | 141.56 | -0.29 | -1.16 | -0.55 | 1.34 | 140.32 | 1.16 | 8.4976
0.0932 | 3.26 | 17.68 | -141.13 | 142.55 | -0.28 | -1.2 | -0.56 | 1.38 | 141.13 | 1.2 | 8.6649
0.095 | 3.29 | 19.04 | -140.98 | 142.58 | -0.27 | -1.24 | -0.57 | 1.42 | 140.98 | 1.24 | 8.8323
0.0968 | 3.39 | 20.19 | -140.15 | 141.93 | -0.27 | -1.28 | -0.59 | 1.46 | 140.15 | 1.28 | 8.9996
0.0986 | 3.51 | 20.99 | -138.54 | 140.46 | -0.28 | -1.32 | -0.6 | 1.5 | 138.54 | 1.32 | 9.167
0.1004 | 3.7 | 21.41 | -136.23 | 138.23 | -0.28 | -1.35 | -0.61 | 1.52 | 136.23 | 1.35 | 9.3343
0.1022 | 3.85 | 21.23 | -133.74 | 135.75 | -0.31 | -1.36 | -0.62 | 1.55 | 133.74 | 1.36 | 9.5017
0.104 | 3.92 | 20.87 | -132.61 | 134.58 | -0.35 | -1.38 | -0.63 | 1.58 | 132.61 | 1.38 | 9.669
0.1058 | 3.97 | 20.28 | -132.21 | 134.11 | -0.41 | -1.4 | -0.64 | 1.62 | 132.21 | 1.4 | 9.8364
0.1076 | 3.92 | 19.95 | -133.25 | 135.1 | -0.45 | -1.42 | -0.66 | 1.65 | 133.25 | 1.42 | 10.0037
0.1094 | 3.74 | 19.68 | -135.6 | 137.39 | -0.49 | -1.43 | -0.67 | 1.68 | 135.6 | 1.43 | 10.1711
0.1111 | 3.54 | 19.74 | -138.95 | 140.76 | -0.52 | -1.45 | -0.69 | 1.71 | 138.95 | 1.45 | 10.3291
0.1129 | 3.3 | 19.92 | -143.29 | 145.09 | -0.54 | -1.47 | -0.7 | 1.74 | 143.29 | 1.47 | 10.4965
0.1147 | 3 | 20.33 | -147.9 | 149.75 | -0.53 | -1.49 | -0.71 | 1.76 | 147.9 | 1.49 | 10.6638
0.1165 | 2.77 | 21.32 | -151.04 | 152.98 | -0.52 | -1.52 | -0.73 | 1.79 | 151.04 | 1.52 | 10.8312
0.1183 | 2.6 | 22.55 | -152.89 | 154.98 | -0.5 | -1.55 | -0.74 | 1.82 | 152.89 | 1.55 | 10.9985
0.1201 | 2.56 | 23.86 | -153.58 | 155.86 | -0.48 | -1.59 | -0.76 | 1.85 | 153.58 | 1.59 | 11.1659
0.1219 | 2.57 | 24.95 | -153.32 | 155.76 | -0.47 | -1.63 | -0.78 | 1.89 | 153.32 | 1.63 | 11.3332
0.1237 | 2.59 | 25.72 | -152.32 | 154.91 | -0.47 | -1.66 | -0.8 | 1.92 | 152.32 | 1.66 | 11.5006
0.1255 | 2.63 | 26 | -150.35 | 153.03 | -0.49 | -1.69 | -0.82 | 1.96 | 150.35 | 1.69 | 11.6679
0.1273 | 2.57 | 25.86 | -148.42 | 151.11 | -0.51 | -1.71 | -0.83 | 1.99 | 148.42 | 1.71 | 11.8353
0.1291 | 2.41 | 25.46 | -147.12 | 149.76 | -0.55 | -1.72 | -0.85 | 2.01 | 147.12 | 1.72 | 12.0026
0.1309 | 2.13 | 24.75 | -147.2 | 149.68 | -0.6 | -1.72 | -0.86 | 2.04 | 147.2 | 1.72 | 12.17
0.1327 | 1.77 | 24.3 | -148.19 | 150.55 | -0.64 | -1.73 | -0.88 | 2.06 | 148.19 | 1.73 | 12.3373
0.1345 | 1.33 | 24.19 | -150.14 | 152.42 | -0.66 | -1.74 | -0.89 | 2.08 | 150.14 | 1.74 | 12.5046
0.1362 | 0.86 | 24.39 | -153.01 | 155.26 | -0.68 | -1.75 | -0.91 | 2.11 | 153.01 | 1.75 | 12.6627
0.138 | 0.39 | 25.19 | -155.34 | 157.69 | -0.68 | -1.78 | -0.93 | 2.14 | 155.34 | 1.78 | 12.83
0.1398 | -0.06 | 26.12 | -157.23 | 159.71 | -0.67 | -1.81 | -0.94 | 2.16 | 157.23 | 1.81 | 12.9974
0.1416 | -0.38 | 27.17 | -158.12 | 160.79 | -0.66 | -1.84 | -0.95 | 2.19 | 158.12 | 1.84 | 13.1647
0.1434 | -0.53 | 28.14 | -157.79 | 160.65 | -0.64 | -1.88 | -0.97 | 2.22 | 157.79 | 1.88 | 13.3321
0.1452 | -0.64 | 28.94 | -156.69 | 159.74 | -0.63 | -1.92 | -0.98 | 2.26 | 156.69 | 1.92 | 13.4994
0.147 | -0.78 | 29.48 | -154.81 | 158.01 | -0.63 | -1.94 | -1 | 2.28 | 154.81 | 1.94 | 13.6668
0.1488 | -0.72 | 29.77 | -151.98 | 155.27 | -0.63 | -1.97 | -1.01 | 2.32 | 151.98 | 1.97 | 13.8341
0.1506 | -0.55 | 29.55 | -148.6 | 151.89 | -0.64 | -1.99 | -1.02 | 2.34 | 148.6 | 1.99 | 14.0015
0.1524 | -0.71 | 28.74 | -146.38 | 149.57 | -0.69 | -2 | -1.03 | 2.36 | 146.38 | 2 | 14.1688
0.1542 | -1.15 | 28.06 | -145.14 | 148.26 | -0.73 | -1.99 | -1.04 | 2.37 | 145.14 | 1.99 | 14.3362
0.156 | -1.68 | 27.14 | -145.53 | 148.53 | -0.78 | -1.98 | -1.05 | 2.38 | 145.53 | 1.98 | 14.5035
0.1578 | -2.21 | 26.24 | -146.67 | 149.54 | -0.82 | -1.97 | -1.06 | 2.39 | 146.67 | 1.97 | 14.6709
0.1595 | -2.71 | 25.57 | -148.98 | 151.74 | -0.86 | -1.96 | -1.07 | 2.4 | 148.98 | 1.96 | 14.8289
0.1613 | -3.13 | 25.16 | -152 | 154.68 | -0.89 | -1.97 | -1.08 | 2.42 | 152 | 1.97 | 14.9963
0.1631 | -3.39 | 25.29 | -155.72 | 158.41 | -0.91 | -1.99 | -1.08 | 2.45 | 155.72 | 1.99 | 15.1636
0.1649 | -3.81 | 25.83 | -159.74 | 162.49 | -0.9 | -1.99 | -1.09 | 2.45 | 159.74 | 1.99 | 15.331
0.1667 | -4.11 | 26.61 | -163.03 | 165.85 | -0.89 | -2 | -1.1 | 2.46 | 163.03 | 2 | 15.4983
0.1685 | -4.32 | 27.52 | -165.87 | 168.79 | -0.87 | -2.02 | -1.11 | 2.48 | 165.87 | 2.02 | 15.6657
0.1703 | -4.36 | 28.71 | -168.26 | 171.36 | -0.85 | -2.05 | -1.12 | 2.49 | 168.26 | 2.05 | 15.833
0.1721 | -4.41 | 29.83 | -169.63 | 172.87 | -0.82 | -2.08 | -1.13 | 2.51 | 169.63 | 2.08 | 16.0004
0.1739 | -4.33 | 31.03 | -170.01 | 173.47 | -0.79 | -2.1 | -1.13 | 2.53 | 170.01 | 2.1 | 16.1677
0.1757 | -4.19 | 32.05 | -169.75 | 173.39 | -0.77 | -2.13 | -1.14 | 2.54 | 169.75 | 2.13 | 16.3351
0.1775 | -3.94 | 32.99 | -168.71 | 172.54 | -0.75 | -2.15 | -1.15 | 2.56 | 168.71 | 2.15 | 16.5024
0.1793 | -3.68 | 33.78 | -166.87 | 170.89 | -0.73 | -2.17 | -1.15 | 2.58 | 166.87 | 2.17 | 16.6698
0.1811 | -3.41 | 34.15 | -164.41 | 168.55 | -0.72 | -2.19 | -1.16 | 2.59 | 164.41 | 2.19 | 16.8371
0.1829 | -3.1 | 34.28 | -161.61 | 165.8 | -0.72 | -2.2 | -1.16 | 2.6 | 161.61 | 2.2 | 17.0045
0.1846 | -2.83 | 34.07 | -158.91 | 163.1 | -0.72 | -2.2 | -1.16 | 2.6 | 158.91 | 2.2 | 17.1625
0.1864 | -2.75 | 33.65 | -157.63 | 161.82 | -0.74 | -2.18 | -1.16 | 2.59 | 157.63 | 2.18 | 17.3299
0.1882 | -2.82 | 32.65 | -156.99 | 161.04 | -0.78 | -2.15 | -1.16 | 2.58 | 156.99 | 2.15 | 17.4972
0.19 | -3.04 | 31.59 | -157.65 | 161.52 | -0.83 | -2.12 | -1.16 | 2.57 | 157.65 | 2.12 | 17.6646
0.1918 | -3.33 | 30.55 | -159.13 | 162.77 | -0.87 | -2.1 | -1.16 | 2.56 | 159.13 | 2.1 | 17.8319
0.1936 | -3.6 | 29.7 | -161.87 | 165.35 | -0.91 | -2.07 | -1.16 | 2.55 | 161.87 | 2.07 | 17.9993
0.1954 | -3.81 | 29.47 | -165.12 | 168.48 | -0.94 | -2.05 | -1.16 | 2.55 | 165.12 | 2.05 | 18.1666
0.1972 | -3.8 | 30.1 | -169.52 | 172.91 | -0.93 | -2.06 | -1.16 | 2.56 | 169.52 | 2.06 | 18.334
0.199 | -3.68 | 31 | -173.65 | 177.12 | -0.92 | -2.08 | -1.15 | 2.56 | 173.65 | 2.08 | 18.5013
0.2008 | -3.51 | 31.82 | -176.32 | 179.91 | -0.9 | -2.1 | -1.15 | 2.57 | 176.32 | 2.1 | 18.6687
0.2026 | -3.48 | 32.11 | -178.86 | 182.53 | -0.89 | -2.09 | -1.14 | 2.55 | 178.86 | 2.09 | 18.836
0.2044 | -3.42 | 32.53 | -180.73 | 184.5 | -0.87 | -2.09 | -1.13 | 2.54 | 180.73 | 2.09 | 19.0033
0.2062 | -3.21 | 33.13 | -181.73 | 185.6 | -0.84 | -2.1 | -1.12 | 2.53 | 181.73 | 2.1 | 19.1707
0.208 | -3.07 | 33.08 | -182.27 | 186.11 | -0.82 | -2.07 | -1.11 | 2.5 | 182.27 | 2.07 | 19.338
0.2097 | -2.94 | 32.96 | -182.78 | 186.57 | -0.8 | -2.05 | -1.09 | 2.47 | 182.78 | 2.05 | 19.4961
0.2115 | -3.29 | 32.89 | -183.56 | 187.31 | -0.77 | -1.99 | -1.08 | 2.4 | 183.56 | 1.99 | 19.6634
0.2133 | -3.74 | 32.3 | -186.07 | 189.67 | -0.75 | -1.92 | -1.06 | 2.33 | 186.07 | 1.92 | 19.8308
0.2151 | -4.24 | 31.18 | -189.83 | 193.21 | -0.75 | -1.82 | -1.04 | 2.24 | 189.83 | 1.82 | 19.9981
0.2169 | -4.86 | 29.96 | -194.75 | 197.86 | -0.74 | -1.71 | -1.03 | 2.15 | 194.75 | 1.71 | 20.1655
0.2187 | -5.46 | 28.82 | -201.16 | 203.99 | -0.74 | -1.61 | -1.01 | 2.06 | 201.16 | 1.61 | 20.3328
0.2205 | -6.27 | 27.58 | -208.97 | 211.61 | -0.72 | -1.5 | -1 | 1.96 | 208.97 | 1.5 | 20.5002
0.2223 | -7.21 | 26.23 | -217.37 | 219.83 | -0.7 | -1.37 | -1 | 1.86 | 217.37 | 1.37 | 20.6675
0.2241 | -8.19 | 24.68 | -227.38 | 229.65 | -0.67 | -1.22 | -0.99 | 1.75 | 227.38 | 1.22 | 20.8349
0.2259 | -9.36 | 23 | -239.09 | 241.15 | -0.64 | -1.05 | -0.98 | 1.63 | 239.09 | 1.05 | 21.0022
0.2277 | -10.74 | 21.23 | -252.6 | 254.48 | -0.58 | -0.86 | -0.98 | 1.5 | 252.6 | 0.86 | 21.1696
0.2295 | -12.29 | 19.37 | -267.94 | 269.69 | -0.52 | -0.65 | -0.98 | 1.39 | 267.94 | 0.65 | 21.3369
0.2313 | -13.87 | 17.34 | -284.98 | 286.62 | -0.45 | -0.43 | -1 | 1.32 | 284.98 | 0.43 | 21.5043
0.2331 | -15.57 | 15.01 | -303.46 | 305.05 | -0.37 | -0.18 | -1.01 | 1.3 | 303.46 | 0.18 | 21.6716
0.2348 | -17.69 | 12.25 | -324.5 | 326.03 | -0.28 | 0.1 | -1.04 | 1.33 | 324.5 | -0.1 | 21.8297
0.2366 | -19.98 | 9.26 | -347.71 | 349.15 | -0.19 | 0.4 | -1.07 | 1.39 | 347.71 | -0.4 | 21.997
0.2384 | -22.33 | 6.11 | -372.64 | 374.04 | -0.09 | 0.71 | -1.1 | 1.51 | 372.64 | -0.71 | 22.1644
0.2402 | -24.83 | 2.67 | -400.32 | 401.74 | 0.02 | 1.06 | -1.15 | 1.71 | 400.32 | -1.06 | 22.3317
0.242 | -27.4 | -0.92 | -429.86 | 431.32 | 0.13 | 1.42 | -1.21 | 1.97 | 429.86 | -1.42 | 22.4991
0.2438 | -29.97 | -4.82 | -462.35 | 463.88 | 0.26 | 1.82 | -1.27 | 2.31 | 462.35 | -1.82 | 22.6664
0.2456 | -32.55 | -8.98 | -497.33 | 498.98 | 0.4 | 2.23 | -1.35 | 2.69 | 497.33 | -2.23 | 22.8338
0.2474 | -35.06 | -13.34 | -535.42 | 537.22 | 0.55 | 2.67 | -1.44 | 3.12 | 535.42 | -2.67 | 23.0011
0.2492 | -37.5 | -17.77 | -576.33 | 578.27 | 0.71 | 3.12 | -1.54 | 3.59 | 576.33 | -3.12 | 23.1685
0.251 | -39.92 | -22.29 | -620.64 | 622.75 | 0.89 | 3.61 | -1.65 | 4.09 | 620.64 | -3.61 | 23.3358
0.2528 | -42.3 | -26.97 | -667.51 | 669.78 | 1.08 | 4.11 | -1.77 | 4.63 | 667.51 | -4.11 | 23.5032
0.2546 | -44.46 | -31.77 | -720.36 | 722.8 | 1.29 | 4.64 | -1.91 | 5.21 | 720.36 | -4.64 | 23.6705
0.2564 | -46.27 | -36.32 | -773.24 | 775.83 | 1.51 | 5.15 | -2.05 | 5.77 | 773.24 | -5.15 | 23.8379
0.2581 | -47.95 | -41.4 | -827.88 | 830.66 | 1.72 | 5.7 | -2.21 | 6.36 | 827.88 | -5.7 | 23.9959
0.2599 | -49.17 | -46.55 | -882.25 | 885.21 | 1.93 | 6.23 | -2.36 | 6.95 | 882.25 | -6.23 | 24.1633
0.2617 | -49.83 | -52.11 | -938.5 | 941.62 | 2.14 | 6.77 | -2.52 | 7.55 | 938.5 | -6.77 | 24.3306
0.2635 | -50.31 | -57.72 | -992.63 | 995.92 | 2.34 | 7.29 | -2.68 | 8.12 | 992.63 | -7.29 | 24.498
0.2653 | -50.41 | -63.35 | -1046.9 | 1050.35 | 2.54 | 7.8 | -2.84 | 8.69 | 1046.9 | -7.8 | 24.6653
0.2671 | -50.31 | -68.2 | -1102.15 | 1105.71 | 2.76 | 8.32 | -3 | 9.28 | 1102.15 | -8.32 | 24.8327
0.2689 | -49.99 | -73.15 | -1156.39 | 1160.07 | 2.97 | 8.79 | -3.16 | 9.81 | 1156.39 | -8.79 | 25
0.2707 | -49.3 | -77.54 | -1206.86 | 1210.64 | 3.16 | 9.18 | -3.3 | 10.26 | 1206.86 | -9.18 | 25.1673
0.2725 | -48.4 | -81.62 | -1254.84 | 1258.69 | 3.34 | 9.51 | -3.44 | 10.66 | 1254.84 | -9.51 | 25.3347
0.2743 | -47.06 | -85.44 | -1302.56 | 1306.47 | 3.51 | 9.79 | -3.57 | 11 | 1302.56 | -9.79 | 25.502
0.2761 | -45.77 | -88.92 | -1348.71 | 1352.66 | 3.68 | 10.06 | -3.7 | 11.34 | 1348.71 | -10.06 | 25.6694
0.2779 | -44.45 | -92.14 | -1392.54 | 1396.53 | 3.84 | 10.34 | -3.82 | 11.68 | 1392.54 | -10.34 | 25.8367
0.2797 | -43.01 | -94.94 | -1432.99 | 1437.01 | 4 | 10.59 | -3.93 | 11.99 | 1432.99 | -10.59 | 26.0041
0.2815 | -41.34 | -97.34 | -1470.36 | 1474.39 | 4.13 | 10.78 | -4.02 | 12.23 | 1470.36 | -10.78 | 26.1714
0.2832 | -39.5 | -99.45 | -1505.24 | 1509.28 | 4.25 | 10.93 | -4.11 | 12.44 | 1505.24 | -10.93 | 26.3295
0.285 | -37.63 | -101.21 | -1537.45 | 1541.49 | 4.36 | 11.04 | -4.19 | 12.6 | 1537.45 | -11.04 | 26.4968
0.2868 | -35.74 | -102.36 | -1565.24 | 1569.25 | 4.45 | 11.09 | -4.26 | 12.7 | 1565.24 | -11.09 | 26.6642
0.2886 | -33.84 | -103.35 | -1590.89 | 1594.88 | 4.53 | 11.13 | -4.32 | 12.78 | 1590.89 | -11.13 | 26.8315
0.2904 | -31.95 | -104.27 | -1614.35 | 1618.31 | 4.6 | 11.15 | -4.38 | 12.85 | 1614.35 | -11.15 | 26.9989
0.2922 | -30.15 | -104.67 | -1633.11 | 1637.02 | 4.66 | 11.15 | -4.41 | 12.89 | 1633.11 | -11.15 | 27.1662
0.294 | -28.39 | -104.37 | -1649.65 | 1653.49 | 4.71 | 11.09 | -4.45 | 12.86 | 1649.65 | -11.09 | 27.3336
0.2958 | -26.64 | -103.97 | -1665 | 1668.77 | 4.76 | 11.04 | -4.47 | 12.85 | 1665 | -11.04 | 27.5009
0.2976 | -24.89 | -102.82 | -1678.91 | 1682.55 | 4.81 | 10.95 | -4.49 | 12.8 | 1678.91 | -10.95 | 27.6683
0.2994 | -23.25 | -101.32 | -1692.37 | 1695.87 | 4.86 | 10.86 | -4.5 | 12.75 | 1692.37 | -10.86 | 27.8356
0.3012 | -21.26 | -99.15 | -1704.28 | 1707.62 | 4.89 | 10.66 | -4.51 | 12.61 | 1704.28 | -10.66 | 28.003
0.303 | -19.36 | -96.85 | -1714.35 | 1717.54 | 4.91 | 10.47 | -4.52 | 12.46 | 1714.35 | -10.47 | 28.1703
0.3048 | -17.55 | -94.5 | -1722.29 | 1725.35 | 4.91 | 10.26 | -4.52 | 12.29 | 1722.29 | -10.26 | 28.3377
0.3066 | -15.78 | -92 | -1728.5 | 1731.43 | 4.9 | 10.04 | -4.51 | 12.11 | 1728.5 | -10.04 | 28.505
0.3083 | -13.99 | -88.89 | -1733.44 | 1736.25 | 4.88 | 9.76 | -4.5 | 11.88 | 1733.44 | -9.76 | 28.6631
0.3101 | -12.32 | -85.95 | -1736.87 | 1739.53 | 4.87 | 9.53 | -4.48 | 11.69 | 1736.87 | -9.53 | 28.8304
0.3119 | -10.73 | -82.78 | -1739.89 | 1742.44 | 4.83 | 9.17 | -4.47 | 11.41 | 1739.89 | -9.17 | 28.9978
0.3137 | -9.27 | -79.71 | -1741.8 | 1744.24 | 4.81 | 8.91 | -4.45 | 11.21 | 1741.8 | -8.91 | 29.1651
0.3155 | -7.72 | -76.33 | -1743.51 | 1745.85 | 4.78 | 8.59 | -4.43 | 11.02 | 1743.51 | -8.59 | 29.3325
0.3173 | -6.26 | -72.71 | -1745.16 | 1747.45 | 4.74 | 8.23 | -4.41 | 10.86 | 1745.16 | -8.23 | 29.4998
0.3191 | -4.89 | -69.02 | -1746.46 | 1748.73 | 4.7 | 7.92 | -4.39 | 10.76 | 1746.46 | -7.92 | 29.6672
0.3209 | -3.49 | -65.6 | -1747.37 | 1749.6 | 4.68 | 7.66 | -4.36 | 10.71 | 1747.37 | -7.66 | 29.8345
0.3227 | -1.93 | -62.03 | -1747.57 | 1749.7 | 4.65 | 7.4 | -4.33 | 10.58 | 1747.57 | -7.4 | 30.0019
0.3245 | -0.53 | -58.47 | -1747.47 | 1749.53 | 4.63 | 7.14 | -4.29 | 10.5 | 1747.47 | -7.14 | 30.1692
0.3263 | 0.7 | -54.89 | -1746.63 | 1748.66 | 4.59 | 6.87 | -4.26 | 10.39 | 1746.63 | -6.87 | 30.3366
0.3281 | 1.74 | -51.6 | -1745.19 | 1747.2 | 4.55 | 6.61 | -4.23 | 10.26 | 1745.19 | -6.61 | 30.5039
0.3299 | 2.98 | -47.69 | -1742.96 | 1744.88 | 4.5 | 6.33 | -4.19 | 10.07 | 1742.96 | -6.33 | 30.6713
0.3316 | 4.03 | -44.04 | -1740.32 | 1742.17 | 4.46 | 6.06 | -4.15 | 9.88 | 1740.32 | -6.06 | 30.8293
0.3334 | 5.46 | -40.23 | -1736.03 | 1737.81 | 4.41 | 5.78 | -4.1 | 9.7 | 1736.03 | -5.78 | 30.9967
0.3352 | 6.64 | -36.04 | -1732 | 1733.71 | 4.35 | 5.49 | -4.06 | 9.49 | 1732 | -5.49 | 31.164
0.337 | 7.84 | -32.18 | -1727.72 | 1729.44 | 4.29 | 5.2 | -4.01 | 9.36 | 1727.72 | -5.2 | 31.3313
0.3388 | 8.96 | -28.26 | -1723.04 | 1724.78 | 4.23 | 4.91 | -3.97 | 9.23 | 1723.04 | -4.91 | 31.4987
0.3406 | 9.94 | -24.69 | -1717.98 | 1719.76 | 4.18 | 4.66 | -3.92 | 9.12 | 1717.98 | -4.66 | 31.666
0.3424 | 10.78 | -21.59 | -1712.67 | 1714.5 | 4.13 | 4.44 | -3.87 | 9.05 | 1712.67 | -4.44 | 31.8334
0.3442 | 11.84 | -17.87 | -1706.18 | 1707.95 | 4.08 | 4.2 | -3.82 | 8.9 | 1706.18 | -4.2 | 32.0007
0.346 | 12.86 | -13.38 | -1700.22 | 1701.97 | 4 | 3.91 | -3.77 | 8.68 | 1700.22 | -3.91 | 32.1681
0.3478 | 13.9 | -9.3 | -1694.66 | 1696.44 | 3.92 | 3.65 | -3.72 | 8.51 | 1694.66 | -3.65 | 32.3354
0.3496 | 14.7 | -6.12 | -1686.27 | 1688.04 | 3.87 | 3.47 | -3.66 | 8.42 | 1686.27 | -3.47 | 32.5028
0.3514 | 15.4 | -2.26 | -1676.31 | 1678.06 | 3.8 | 3.25 | -3.6 | 8.27 | 1676.31 | -3.25 | 32.6701
0.3532 | 16.5 | 2.03 | -1668.11 | 1669.84 | 3.71 | 3.01 | -3.55 | 8.08 | 1668.11 | -3.01 | 32.8375
0.355 | 17.64 | 6.33 | -1661.3 | 1663.03 | 3.64 | 2.76 | -3.49 | 7.91 | 1661.3 | -2.76 | 33.0048
0.3567 | 18.6 | 10.2 | -1655.38 | 1657.21 | 3.55 | 2.54 | -3.44 | 7.79 | 1655.38 | -2.54 | 33.1629
0.3585 | 19.41 | 13.9 | -1650.6 | 1652.57 | 3.47 | 2.34 | -3.39 | 7.72 | 1650.6 | -2.34 | 33.3302
0.3603 | 20.27 | 17.54 | -1644.8 | 1646.86 | 3.41 | 2.12 | -3.34 | 7.65 | 1644.8 | -2.12 | 33.4976
0.3621 | 21.19 | 21.02 | -1636.87 | 1639 | 3.34 | 1.89 | -3.28 | 7.61 | 1636.87 | -1.89 | 33.6649
0.3639 | 22.09 | 24.42 | -1628.64 | 1630.86 | 3.27 | 1.66 | -3.22 | 7.58 | 1628.64 | -1.66 | 33.8323
0.3657 | 23.01 | 27.68 | -1619.9 | 1622.21 | 3.2 | 1.44 | -3.16 | 7.54 | 1619.9 | -1.44 | 33.9996
0.3675 | 24.46 | 31.18 | -1607.87 | 1610.12 | 3.12 | 1.2 | -3.08 | 7.37 | 1607.87 | -1.2 | 34.167
0.3693 | 25.99 | 34.91 | -1594.84 | 1596.99 | 3.04 | 0.93 | -3 | 7.15 | 1594.84 | -0.93 | 34.3343
0.3711 | 27.57 | 38.75 | -1582.48 | 1584.52 | 2.93 | 0.62 | -2.91 | 6.88 | 1582.48 | -0.62 | 34.5017
0.3729 | 29.04 | 42.79 | -1570.95 | 1572.91 | 2.82 | 0.3 | -2.83 | 6.58 | 1570.95 | -0.3 | 34.669
0.3747 | 30.23 | 46.53 | -1561.75 | 1563.71 | 2.72 | 0.03 | -2.76 | 6.37 | 1561.75 | -0.03 | 34.8364
0.3765 | 31.02 | 48.84 | -1554.8 | 1556.87 | 2.66 | -0.15 | -2.71 | 6.38 | 1554.8 | 0.15 | 35.0037
0.3783 | 31.92 | 50.62 | -1547.71 | 1549.9 | 2.59 | -0.32 | -2.66 | 6.41 | 1547.71 | 0.32 | 35.1711
0.3801 | 33.23 | 53.23 | -1539.59 | 1541.83 | 2.52 | -0.55 | -2.6 | 6.3 | 1539.59 | 0.55 | 35.3384
0.3818 | 34.25 | 54.68 | -1533.13 | 1535.43 | 2.49 | -0.7 | -2.55 | 6.23 | 1533.13 | 0.7 | 35.4965
0.3836 | 35 | 55.84 | -1526.94 | 1529.27 | 2.48 | -0.83 | -2.51 | 6.18 | 1526.94 | 0.83 | 35.6638
0.3854 | 35.66 | 57.39 | -1518.73 | 1521.1 | 2.42 | -0.99 | -2.46 | 6.08 | 1518.73 | 0.99 | 35.8312
0.3872 | 36.33 | 58.95 | -1510.51 | 1512.93 | 2.37 | -1.15 | -2.42 | 5.98 | 1510.51 | 1.15 | 35.9985
0.389 | 37.08 | 59.77 | -1501.04 | 1503.54 | 2.33 | -1.25 | -2.38 | 5.97 | 1501.04 | 1.25 | 36.1659
0.3908 | 37.87 | 60.74 | -1491.2 | 1493.76 | 2.28 | -1.36 | -2.34 | 5.95 | 1491.2 | 1.36 | 36.3332
0.3926 | 38.67 | 62.25 | -1480.91 | 1483.52 | 2.23 | -1.5 | -2.29 | 5.85 | 1480.91 | 1.5 | 36.5006
0.3944 | 39.46 | 63.36 | -1470.4 | 1473.06 | 2.19 | -1.62 | -2.24 | 5.79 | 1470.4 | 1.62 | 36.6679
0.3962 | 40.25 | 64.5 | -1459.55 | 1462.27 | 2.15 | -1.74 | -2.19 | 5.76 | 1459.55 | 1.74 | 36.8353
0.398 | 40.9 | 65.5 | -1448.13 | 1450.92 | 2.12 | -1.84 | -2.14 | 5.72 | 1448.13 | 1.84 | 37.0026
0.3998 | 41.38 | 65.58 | -1435.2 | 1438.04 | 2.07 | -1.84 | -2.1 | 5.66 | 1435.2 | 1.84 | 37.17
0.4016 | 41.93 | 65.6 | -1421.77 | 1424.65 | 2.02 | -1.8 | -2.05 | 5.58 | 1421.77 | 1.8 | 37.3373
0.4034 | 42.67 | 65.49 | -1408.68 | 1411.61 | 1.96 | -1.77 | -2.03 | 5.52 | 1408.68 | 1.77 | 37.5046
0.4051 | 43.62 | 65.25 | -1395.94 | 1398.91 | 1.9 | -1.75 | -2.03 | 5.47 | 1395.94 | 1.75 | 37.6627
0.4069 | 44.51 | 65 | -1383.46 | 1386.49 | 1.84 | -1.73 | -2.03 | 5.43 | 1383.46 | 1.73 | 37.83
0.4087 | 45.25 | 64.68 | -1371.65 | 1374.73 | 1.77 | -1.74 | -2.04 | 5.41 | 1371.65 | 1.74 | 37.9974
0.4105 | 46.15 | 64.33 | -1359.05 | 1362.19 | 1.69 | -1.73 | -2.06 | 5.37 | 1359.05 | 1.73 | 38.1647
0.4123 | 46.63 | 64.02 | -1347.27 | 1350.41 | 1.61 | -1.76 | -2.09 | 5.33 | 1347.27 | 1.76 | 38.3321
0.4141 | 47.12 | 64.07 | -1335.99 | 1339.15 | 1.52 | -1.83 | -2.11 | 5.28 | 1335.99 | 1.83 | 38.4994
0.4159 | 47.62 | 64.29 | -1325.67 | 1328.83 | 1.44 | -1.93 | -2.11 | 5.23 | 1325.67 | 1.93 | 38.6668
0.4177 | 48.06 | 64.48 | -1315.98 | 1319.17 | 1.37 | -2.04 | -2.1 | 5.22 | 1315.98 | 2.04 | 38.8341
0.4195 | 48.56 | 64.25 | -1306.68 | 1309.89 | 1.3 | -2.11 | -2.1 | 5.23 | 1306.68 | 2.11 | 39.0015
0.4213 | 49.15 | 63.49 | -1297.46 | 1300.68 | 1.22 | -2.13 | -2.12 | 5.24 | 1297.46 | 2.13 | 39.1688
0.4231 | 49.53 | 62.48 | -1288.32 | 1291.55 | 1.16 | -2.18 | -2.13 | 5.28 | 1288.32 | 2.18 | 39.3362
0.4249 | 50.34 | 61.99 | -1278.91 | 1282.19 | 1.09 | -2.22 | -2.16 | 5.33 | 1278.91 | 2.22 | 39.5035
0.4267 | 50.98 | 61.87 | -1269.31 | 1272.63 | 1.01 | -2.28 | -2.18 | 5.35 | 1269.31 | 2.28 | 39.6709
0.4285 | 51.28 | 61.87 | -1259.81 | 1263.13 | 0.96 | -2.36 | -2.19 | 5.33 | 1259.81 | 2.36 | 39.8382
0.4302 | 51.46 | 61.85 | -1250.92 | 1254.23 | 0.91 | -2.45 | -2.18 | 5.29 | 1250.92 | 2.45 | 39.9963
0.432 | 51.47 | 61.59 | -1242.65 | 1245.94 | 0.86 | -2.5 | -2.18 | 5.25 | 1242.65 | 2.5 | 40.1636
0.4338 | 51.43 | 61.25 | -1234.6 | 1237.87 | 0.81 | -2.53 | -2.17 | 5.19 | 1234.6 | 2.53 | 40.331
0.4356 | 51.69 | 62.11 | -1226.28 | 1229.56 | 0.77 | -2.71 | -2.16 | 5.05 | 1226.28 | 2.71 | 40.4983
0.4374 | 52.63 | 64.67 | -1218.26 | 1221.61 | 0.68 | -2.94 | -2.11 | 4.81 | 1218.26 | 2.94 | 40.6657
0.4392 | 52.95 | 66.16 | -1210.63 | 1214.04 | 0.6 | -3.08 | -2.07 | 4.64 | 1210.63 | 3.08 | 40.833
0.441 | 52.97 | 66.5 | -1200.96 | 1204.42 | 0.56 | -3.09 | -2.07 | 4.59 | 1200.96 | 3.09 | 41.0004
0.4428 | 52.65 | 66.16 | -1192.12 | 1195.61 | 0.54 | -3.09 | -2.09 | 4.58 | 1192.12 | 3.09 | 41.1677
0.4446 | 52.09 | 65.68 | -1184.28 | 1187.76 | 0.51 | -3.09 | -2.09 | 4.56 | 1184.28 | 3.09 | 41.3351
0.4464 | 51.59 | 65.24 | -1176.41 | 1179.88 | 0.48 | -3.1 | -2.09 | 4.54 | 1176.41 | 3.1 | 41.5024
0.4482 | 50.62 | 63.97 | -1169.92 | 1173.34 | 0.5 | -3.09 | -2.09 | 4.53 | 1169.92 | 3.09 | 41.6698
0.45 | 50.11 | 62.71 | -1163.72 | 1167.08 | 0.52 | -3.08 | -2.1 | 4.55 | 1163.72 | 3.08 | 41.8371
0.4518 | 49.6 | 61.06 | -1159.27 | 1162.58 | 0.57 | -3.09 | -2.1 | 4.56 | 1159.27 | 3.09 | 42.0045
0.4536 | 48.89 | 59.02 | -1155.38 | 1158.62 | 0.63 | -3.07 | -2.1 | 4.57 | 1155.38 | 3.07 | 42.1718
0.4553 | 48.15 | 56.76 | -1152.22 | 1155.38 | 0.7 | -3.03 | -2.11 | 4.57 | 1152.22 | 3.03 | 42.3299
0.4571 | 47.82 | 55.3 | -1148.22 | 1151.31 | 0.72 | -3.02 | -2.11 | 4.57 | 1148.22 | 3.02 | 42.4972
0.4589 | 47.18 | 52.95 | -1147.27 | 1150.22 | 0.79 | -2.98 | -2.12 | 4.55 | 1147.27 | 2.98 | 42.6646
0.4607 | 46.51 | 50.42 | -1146.57 | 1149.39 | 0.86 | -2.94 | -2.12 | 4.54 | 1146.57 | 2.94 | 42.8319
0.4625 | 45.84 | 47.93 | -1145.84 | 1148.52 | 0.93 | -2.89 | -2.12 | 4.52 | 1145.84 | 2.89 | 42.9993
0.4643 | 45 | 45.86 | -1144.73 | 1147.31 | 1.01 | -2.85 | -2.13 | 4.51 | 1144.73 | 2.85 | 43.1666
0.4661 | 44.17 | 43.79 | -1143.62 | 1146.1 | 1.09 | -2.81 | -2.14 | 4.51 | 1143.62 | 2.81 | 43.334
0.4679 | 43.88 | 42.48 | -1139.93 | 1142.36 | 1.16 | -2.83 | -2.15 | 4.54 | 1139.93 | 2.83 | 43.5013
0.4697 | 43.68 | 41.29 | -1135.84 | 1138.24 | 1.22 | -2.86 | -2.15 | 4.57 | 1135.84 | 2.86 | 43.6687
0.4715 | 43.58 | 39.95 | -1131.37 | 1133.73 | 1.28 | -2.88 | -2.16 | 4.61 | 1131.37 | 2.88 | 43.836
0.4733 | 43.49 | 38.59 | -1126.84 | 1129.16 | 1.33 | -2.89 | -2.17 | 4.64 | 1126.84 | 2.89 | 44.0033
0.4751 | 42.73 | 37.42 | -1122.35 | 1124.62 | 1.38 | -2.87 | -2.17 | 4.64 | 1122.35 | 2.87 | 44.1707
0.4769 | 41.87 | 36.47 | -1117.43 | 1119.66 | 1.42 | -2.83 | -2.18 | 4.63 | 1117.43 | 2.83 | 44.338
0.4786 | 41.13 | 35.67 | -1111.52 | 1113.73 | 1.46 | -2.8 | -2.19 | 4.64 | 1111.52 | 2.8 | 44.4961
0.4804 | 40.45 | 34.96 | -1104.9 | 1107.09 | 1.49 | -2.76 | -2.2 | 4.63 | 1104.9 | 2.76 | 44.6634
0.4822 | 39.91 | 34.41 | -1096.91 | 1099.08 | 1.49 | -2.71 | -2.21 | 4.6 | 1096.91 | 2.71 | 44.8308
0.484 | 39.67 | 33.99 | -1087.49 | 1089.69 | 1.5 | -2.67 | -2.23 | 4.57 | 1087.49 | 2.67 | 44.9981
0.4858 | 40 | 33.53 | -1077.18 | 1079.42 | 1.48 | -2.65 | -2.25 | 4.55 | 1077.18 | 2.65 | 45.1655
0.4876 | 40.52 | 32.7 | -1064.99 | 1067.24 | 1.42 | -2.64 | -2.3 | 4.54 | 1064.99 | 2.64 | 45.3328
0.4894 | 40.89 | 31.91 | -1054.89 | 1057.14 | 1.37 | -2.62 | -2.35 | 4.55 | 1054.89 | 2.62 | 45.5002
0.4912 | 40.78 | 31.17 | -1046.19 | 1048.41 | 1.34 | -2.58 | -2.38 | 4.56 | 1046.19 | 2.58 | 45.6675
0.493 | 40.21 | 30.39 | -1038.51 | 1040.66 | 1.31 | -2.52 | -2.39 | 4.55 | 1038.51 | 2.52 | 45.8349
0.4948 | 39.6 | 30.23 | -1032.05 | 1034.16 | 1.31 | -2.47 | -2.4 | 4.51 | 1032.05 | 2.47 | 46.0022
0.4966 | 38.95 | 29.77 | -1026.42 | 1028.46 | 1.31 | -2.44 | -2.39 | 4.49 | 1026.42 | 2.44 | 46.1696
0.4984 | 38.59 | 28.41 | -1022.64 | 1024.56 | 1.36 | -2.4 | -2.38 | 4.44 | 1022.64 | 2.4 | 46.3369
0.5002 | 38.03 | 26.23 | -1018.87 | 1020.61 | 1.43 | -2.36 | -2.37 | 4.38 | 1018.87 | 2.36 | 46.5043
0.502 | 37.51 | 24.43 | -1014.19 | 1015.8 | 1.49 | -2.32 | -2.36 | 4.34 | 1014.19 | 2.32 | 46.6716
0.5037 | 36.76 | 23.42 | -1008.55 | 1010.05 | 1.52 | -2.28 | -2.35 | 4.29 | 1008.55 | 2.28 | 46.8297
0.5055 | 35.65 | 23.9 | -1001.02 | 1002.44 | 1.52 | -2.23 | -2.34 | 4.24 | 1001.02 | 2.23 | 46.997
0.5073 | 34.58 | 25.17 | -992.77 | 994.12 | 1.49 | -2.22 | -2.33 | 4.21 | 992.77 | 2.22 | 47.1644
0.5091 | 33.55 | 25.07 | -987.07 | 988.36 | 1.49 | -2.18 | -2.33 | 4.17 | 987.07 | 2.18 | 47.3317
0.5109 | 32.54 | 24.81 | -982.42 | 983.66 | 1.49 | -2.14 | -2.33 | 4.14 | 982.42 | 2.14 | 47.4991
0.5127 | 31.54 | 24.47 | -978.6 | 979.8 | 1.49 | -2.1 | -2.33 | 4.11 | 978.6 | 2.1 | 47.6664
0.5145 | 30.61 | 23.85 | -974.98 | 976.13 | 1.5 | -2.06 | -2.34 | 4.08 | 974.98 | 2.06 | 47.8338
0.5163 | 29.66 | 23.18 | -971.72 | 972.82 | 1.51 | -2.03 | -2.35 | 4.06 | 971.72 | 2.03 | 48.0011
0.5181 | 28.62 | 22.31 | -969.68 | 970.72 | 1.52 | -2.02 | -2.37 | 4.07 | 969.68 | 2.02 | 48.1685
0.5199 | 27.73 | 22.12 | -968.16 | 969.16 | 1.52 | -2.03 | -2.39 | 4.09 | 968.16 | 2.03 | 48.3358
0.5217 | 27.09 | 21 | -967.18 | 968.12 | 1.56 | -2.05 | -2.42 | 4.13 | 967.18 | 2.05 | 48.5032
0.5235 | 26.52 | 19.91 | -966.58 | 967.48 | 1.6 | -2.09 | -2.46 | 4.18 | 966.58 | 2.09 | 48.6705
0.5253 | 26.08 | 18.87 | -966.57 | 967.44 | 1.65 | -2.17 | -2.49 | 4.26 | 966.57 | 2.17 | 48.8379
0.5271 | 25.61 | 17.92 | -966.92 | 967.77 | 1.7 | -2.25 | -2.53 | 4.34 | 966.92 | 2.25 | 49.0052
0.5288 | 25.05 | 17.27 | -968.35 | 969.17 | 1.75 | -2.33 | -2.56 | 4.43 | 968.35 | 2.33 | 49.1633
0.5306 | 24.44 | 16.64 | -970.41 | 971.2 | 1.78 | -2.36 | -2.6 | 4.48 | 970.41 | 2.36 | 49.3306
0.5324 | 23.96 | 15.92 | -973.01 | 973.76 | 1.81 | -2.38 | -2.65 | 4.52 | 973.01 | 2.38 | 49.498
0.5342 | 23.68 | 15.12 | -975.36 | 976.1 | 1.86 | -2.43 | -2.69 | 4.56 | 975.36 | 2.43 | 49.6653
0.536 | 23.02 | 14.55 | -976.97 | 977.68 | 1.89 | -2.45 | -2.73 | 4.6 | 976.97 | 2.45 | 49.8327
0.5378 | 22.91 | 14.36 | -979.83 | 980.53 | 1.9 | -2.46 | -2.77 | 4.65 | 979.83 | 2.46 | 50
0.5396 | 22.49 | 13.92 | -982.55 | 983.24 | 1.93 | -2.49 | -2.81 | 4.7 | 982.55 | 2.49 | 50.1673
0.5414 | 22.47 | 13.78 | -985.56 | 986.27 | 1.93 | -2.5 | -2.86 | 4.75 | 985.56 | 2.5 | 50.3347
0.5432 | 22.5 | 13.59 | -988.13 | 988.83 | 1.94 | -2.5 | -2.91 | 4.8 | 988.13 | 2.5 | 50.502
0.545 | 22.32 | 13.13 | -989.59 | 990.29 | 1.97 | -2.51 | -2.96 | 4.85 | 989.59 | 2.51 | 50.6694
0.5468 | 21.92 | 12.64 | -991.15 | 991.85 | 2.03 | -2.53 | -3 | 4.89 | 991.15 | 2.53 | 50.8367
0.5486 | 21.57 | 12.35 | -993.48 | 994.18 | 2.07 | -2.52 | -3.04 | 4.93 | 993.48 | 2.52 | 51.0041
0.5504 | 21.25 | 12.08 | -993.38 | 994.08 | 2.11 | -2.58 | -3.07 | 4.97 | 993.38 | 2.58 | 51.1714
0.5521 | 20.74 | 12.2 | -993.19 | 993.88 | 2.15 | -2.63 | -3.1 | 5.02 | 993.19 | 2.63 | 51.3295
0.5539 | 20.71 | 12.1 | -994.97 | 995.65 | 2.26 | -2.58 | -3.14 | 5.08 | 994.97 | 2.58 | 51.4968
0.5557 | 21.01 | 12.09 | -996.27 | 996.95 | 2.36 | -2.58 | -3.17 | 5.16 | 996.27 | 2.58 | 51.6642
0.5575 | 20.56 | 12.16 | -997.03 | 997.7 | 2.4 | -2.6 | -3.2 | 5.2 | 997.03 | 2.6 | 51.8315
0.5593 | 20.28 | 12.38 | -996.08 | 996.73 | 2.43 | -2.63 | -3.22 | 5.24 | 996.08 | 2.63 | 51.9989
0.5611 | 19.96 | 12.62 | -995.09 | 995.7 | 2.44 | -2.67 | -3.23 | 5.28 | 995.09 | 2.67 | 52.1662
0.5629 | 19.63 | 12.95 | -994.55 | 995.14 | 2.45 | -2.69 | -3.25 | 5.31 | 994.55 | 2.69 | 52.3336
0.5647 | 19.3 | 13.33 | -994.27 | 994.81 | 2.45 | -2.72 | -3.26 | 5.34 | 994.27 | 2.72 | 52.5009
0.5665 | 19.03 | 12.96 | -994.23 | 994.73 | 2.44 | -2.71 | -3.27 | 5.36 | 994.23 | 2.71 | 52.6683
0.5683 | 18.56 | 12.15 | -995.84 | 996.31 | 2.43 | -2.69 | -3.29 | 5.37 | 995.84 | 2.69 | 52.8356
0.5701 | 17.88 | 11.71 | -999.22 | 999.67 | 2.43 | -2.69 | -3.32 | 5.38 | 999.22 | 2.69 | 53.003
0.5719 | 17.31 | 11.74 | -1003.28 | 1003.7 | 2.42 | -2.71 | -3.34 | 5.4 | 1003.28 | 2.71 | 53.1703
0.5737 | 17.24 | 12.24 | -1007.04 | 1007.44 | 2.42 | -2.76 | -3.37 | 5.45 | 1007.04 | 2.76 | 53.3377
0.5755 | 17.09 | 11.79 | -1009.47 | 1009.85 | 2.43 | -2.79 | -3.38 | 5.48 | 1009.47 | 2.79 | 53.505
0.5772 | 16.77 | 11.22 | -1011.05 | 1011.41 | 2.44 | -2.8 | -3.4 | 5.5 | 1011.05 | 2.8 | 53.6631
0.579 | 16.39 | 10.85 | -1013.07 | 1013.42 | 2.45 | -2.8 | -3.43 | 5.52 | 1013.07 | 2.8 | 53.8304
0.5808 | 15.79 | 10.74 | -1016.38 | 1016.72 | 2.46 | -2.81 | -3.45 | 5.55 | 1016.38 | 2.81 | 53.9978
0.5826 | 15.07 | 10.67 | -1020.8 | 1021.14 | 2.47 | -2.82 | -3.49 | 5.59 | 1020.8 | 2.82 | 54.1651
0.5844 | 14.42 | 10.94 | -1026.29 | 1026.62 | 2.49 | -2.85 | -3.54 | 5.64 | 1026.29 | 2.85 | 54.3325
0.5862 | 13.92 | 11.18 | -1032.46 | 1032.78 | 2.53 | -2.89 | -3.58 | 5.71 | 1032.46 | 2.89 | 54.4998
0.588 | 13.77 | 10.97 | -1036.82 | 1037.15 | 2.59 | -2.95 | -3.6 | 5.8 | 1036.82 | 2.95 | 54.6672
0.5898 | 13.46 | 10.48 | -1040.65 | 1040.98 | 2.65 | -2.99 | -3.62 | 5.87 | 1040.65 | 2.99 | 54.8345
0.5916 | 13.14 | 10.26 | -1045.24 | 1045.57 | 2.69 | -3 | -3.65 | 5.92 | 1045.24 | 3 | 55.0019
0.5934 | 13.13 | 10.22 | -1049.21 | 1049.54 | 2.74 | -3.03 | -3.67 | 5.97 | 1049.21 | 3.03 | 55.1692
0.5952 | 13.61 | 10.36 | -1051.82 | 1052.16 | 2.79 | -3.1 | -3.69 | 6.06 | 1051.82 | 3.1 | 55.3366
0.597 | 13.54 | 10.38 | -1056.17 | 1056.5 | 2.84 | -3.13 | -3.72 | 6.12 | 1056.17 | 3.13 | 55.5039
0.5988 | 13.26 | 10.63 | -1061.83 | 1062.17 | 2.9 | -3.15 | -3.76 | 6.19 | 1061.83 | 3.15 | 55.6713
0.6006 | 12.44 | 10.69 | -1068.49 | 1068.81 | 2.96 | -3.13 | -3.8 | 6.24 | 1068.49 | 3.13 | 55.8386
0.6023 | 11.55 | 10.66 | -1075.01 | 1075.33 | 3.02 | -3.09 | -3.83 | 6.27 | 1075.01 | 3.09 | 55.9967
0.6041 | 11 | 10.88 | -1083.44 | 1083.76 | 3.1 | -3.08 | -3.88 | 6.34 | 1083.44 | 3.08 | 56.164
0.6059 | 11 | 11.33 | -1091.96 | 1092.29 | 3.18 | -3.13 | -3.93 | 6.44 | 1091.96 | 3.13 | 56.3313
0.6077 | 11.08 | 11.97 | -1100.31 | 1100.65 | 3.28 | -3.16 | -3.97 | 6.55 | 1100.31 | 3.16 | 56.4987
0.6095 | 10.98 | 12.54 | -1108.33 | 1108.68 | 3.38 | -3.17 | -4.01 | 6.64 | 1108.33 | 3.17 | 56.666
0.6113 | 10.66 | 12.66 | -1115.24 | 1115.6 | 3.44 | -3.14 | -4.04 | 6.68 | 1115.24 | 3.14 | 56.8334
0.6131 | 10.56 | 12.79 | -1121.43 | 1121.8 | 3.49 | -3.14 | -4.08 | 6.73 | 1121.43 | 3.14 | 57.0007
0.6149 | 10.26 | 12.84 | -1127.29 | 1127.65 | 3.52 | -3.14 | -4.12 | 6.77 | 1127.29 | 3.14 | 57.1681
0.6167 | 10 | 13.13 | -1132.94 | 1133.3 | 3.53 | -3.13 | -4.17 | 6.82 | 1132.94 | 3.13 | 57.3354
0.6185 | 9.77 | 13.56 | -1138.46 | 1138.82 | 3.53 | -3.13 | -4.24 | 6.86 | 1138.46 | 3.13 | 57.5028
0.6203 | 10.15 | 14.04 | -1142.52 | 1142.89 | 3.57 | -3.14 | -4.28 | 6.92 | 1142.52 | 3.14 | 57.6701
0.6221 | 10.1 | 14.37 | -1146.81 | 1147.18 | 3.59 | -3.11 | -4.32 | 6.94 | 1146.81 | 3.11 | 57.8375
0.6239 | 9.71 | 14.6 | -1152.45 | 1152.84 | 3.61 | -3.09 | -4.37 | 6.96 | 1152.45 | 3.09 | 58.0048
0.6256 | 9.54 | 15.01 | -1159.21 | 1159.62 | 3.64 | -3.09 | -4.4 | 7.01 | 1159.21 | 3.09 | 58.1629
0.6274 | 9.36 | 15.92 | -1168.04 | 1168.47 | 3.73 | -3.09 | -4.42 | 7.08 | 1168.04 | 3.09 | 58.3302
0.6292 | 9.9 | 17.26 | -1176.86 | 1177.35 | 3.85 | -3.12 | -4.47 | 7.2 | 1176.86 | 3.12 | 58.4976
0.631 | 10.61 | 18.71 | -1185.68 | 1186.24 | 3.97 | -3.15 | -4.53 | 7.33 | 1185.68 | 3.15 | 58.6649
0.6328 | 10.82 | 19.79 | -1195.14 | 1195.74 | 4.1 | -3.16 | -4.58 | 7.45 | 1195.14 | 3.16 | 58.8323
0.6346 | 10.81 | 20.72 | -1204.84 | 1205.47 | 4.23 | -3.15 | -4.63 | 7.57 | 1204.84 | 3.15 | 58.9996
0.6364 | 11.21 | 21.99 | -1212.69 | 1213.37 | 4.34 | -3.15 | -4.69 | 7.68 | 1212.69 | 3.15 | 59.167
0.6382 | 11.44 | 23.04 | -1220.66 | 1221.39 | 4.45 | -3.12 | -4.75 | 7.78 | 1220.66 | 3.12 | 59.3343
0.64 | 11.73 | 23.7 | -1226.57 | 1227.32 | 4.5 | -3.1 | -4.8 | 7.83 | 1226.57 | 3.1 | 59.5017
0.6418 | 12.12 | 24.25 | -1232.05 | 1232.85 | 4.55 | -3.09 | -4.85 | 7.88 | 1232.05 | 3.09 | 59.669
0.6436 | 12.61 | 24.94 | -1236.96 | 1237.79 | 4.6 | -3.07 | -4.89 | 7.93 | 1236.96 | 3.07 | 59.8364
0.6454 | 13.1 | 25.63 | -1241.45 | 1242.31 | 4.65 | -3.05 | -4.94 | 7.97 | 1241.45 | 3.05 | 60.0037
0.6472 | 13.65 | 26.25 | -1245.38 | 1246.27 | 4.69 | -3.04 | -4.97 | 8 | 1245.38 | 3.04 | 60.1711
0.649 | 14.25 | 26.84 | -1248.36 | 1249.28 | 4.74 | -3.05 | -5 | 8.04 | 1248.36 | 3.05 | 60.3384
0.6507 | 14.76 | 27.47 | -1251.13 | 1252.09 | 4.81 | -3.05 | -5.03 | 8.07 | 1251.13 | 3.05 | 60.4965
0.6525 | 15.23 | 27.83 | -1253.32 | 1254.3 | 4.84 | -3.06 | -5.05 | 8.1 | 1253.32 | 3.06 | 60.6638
0.6543 | 15.36 | 27.88 | -1255.22 | 1256.21 | 4.86 | -3.07 | -5.07 | 8.12 | 1255.22 | 3.07 | 60.8312
0.6561 | 15.75 | 28.36 | -1257.07 | 1258.07 | 4.87 | -3.08 | -5.09 | 8.14 | 1257.07 | 3.08 | 60.9985
0.6579 | 16.37 | 29.45 | -1258.64 | 1259.68 | 4.89 | -3.07 | -5.12 | 8.16 | 1258.64 | 3.07 | 61.1659
0.6597 | 16.83 | 30.47 | -1260.3 | 1261.37 | 4.9 | -3.06 | -5.16 | 8.18 | 1260.3 | 3.06 | 61.3332
0.6615 | 16.58 | 31.05 | -1262.68 | 1263.76 | 4.89 | -3.01 | -5.19 | 8.16 | 1262.68 | 3.01 | 61.5006
0.6633 | 16.01 | 31.46 | -1266.11 | 1267.19 | 4.89 | -2.93 | -5.22 | 8.15 | 1266.11 | 2.93 | 61.6679
0.6651 | 15.62 | 31.52 | -1269.81 | 1270.89 | 4.9 | -2.9 | -5.25 | 8.17 | 1269.81 | 2.9 | 61.8353
0.6669 | 15.2 | 31.51 | -1274.49 | 1275.58 | 4.91 | -2.9 | -5.28 | 8.2 | 1274.49 | 2.9 | 62.0026
0.6687 | 15.11 | 31.38 | -1279.37 | 1280.47 | 4.96 | -2.92 | -5.32 | 8.27 | 1279.37 | 2.92 | 62.17
0.6705 | 15.16 | 31.35 | -1283.59 | 1284.7 | 4.99 | -2.92 | -5.36 | 8.32 | 1283.59 | 2.92 | 62.3373
0.6723 | 15.22 | 31.44 | -1287.3 | 1288.43 | 5.02 | -2.94 | -5.4 | 8.38 | 1287.3 | 2.94 | 62.5046
0.6741 | 15.64 | 30.99 | -1291.53 | 1292.69 | 5.07 | -2.97 | -5.43 | 8.45 | 1291.53 | 2.97 | 62.672
0.6758 | 16.13 | 30.85 | -1294.87 | 1296.04 | 5.08 | -2.98 | -5.46 | 8.48 | 1294.87 | 2.98 | 62.83
0.6776 | 16.59 | 30.81 | -1297.72 | 1298.91 | 5.09 | -2.98 | -5.49 | 8.51 | 1297.72 | 2.98 | 62.9974
0.6794 | 17.04 | 30.72 | -1300.46 | 1301.65 | 5.11 | -2.99 | -5.52 | 8.55 | 1300.46 | 2.99 | 63.1647
0.6812 | 17.51 | 30.86 | -1302.69 | 1303.89 | 5.13 | -3.03 | -5.55 | 8.59 | 1302.69 | 3.03 | 63.3321
0.683 | 17.65 | 30.47 | -1304.32 | 1305.5 | 5.13 | -3.04 | -5.58 | 8.62 | 1304.32 | 3.04 | 63.4994
0.6848 | 17.25 | 29.65 | -1306.3 | 1307.42 | 5.12 | -3.04 | -5.59 | 8.63 | 1306.3 | 3.04 | 63.6668
0.6866 | 16.85 | 28.7 | -1308.64 | 1309.7 | 5.1 | -3.03 | -5.61 | 8.63 | 1308.64 | 3.03 | 63.8341
0.6884 | 16.57 | 27.79 | -1311.14 | 1312.14 | 5.1 | -3.04 | -5.63 | 8.65 | 1311.14 | 3.04 | 64.0015
0.6902 | 16.63 | 27.26 | -1313.53 | 1314.5 | 5.11 | -3.08 | -5.65 | 8.69 | 1313.53 | 3.08 | 64.1688
0.692 | 17.19 | 27.41 | -1315.24 | 1316.22 | 5.13 | -3.13 | -5.67 | 8.74 | 1315.24 | 3.13 | 64.3362
0.6938 | 17.74 | 27.61 | -1316.22 | 1317.21 | 5.13 | -3.15 | -5.71 | 8.78 | 1316.22 | 3.15 | 64.5035
0.6956 | 18.38 | 27.8 | -1316.77 | 1317.78 | 5.12 | -3.18 | -5.74 | 8.8 | 1316.77 | 3.18 | 64.6709
0.6974 | 19.07 | 27.99 | -1317.22 | 1318.25 | 5.1 | -3.2 | -5.77 | 8.82 | 1317.22 | 3.2 | 64.8382
0.6992 | 19.83 | 28.3 | -1317.41 | 1318.47 | 5.09 | -3.22 | -5.8 | 8.85 | 1317.41 | 3.22 | 65.0056
0.7009 | 20.6 | 28.62 | -1317.55 | 1318.64 | 5.07 | -3.25 | -5.83 | 8.87 | 1317.55 | 3.25 | 65.1636
0.7027 | 21.49 | 28.81 | -1317.43 | 1318.54 | 5.04 | -3.27 | -5.86 | 8.89 | 1317.43 | 3.27 | 65.331
0.7045 | 22.06 | 28.88 | -1317.05 | 1318.18 | 5.02 | -3.28 | -5.88 | 8.9 | 1317.05 | 3.28 | 65.4983
0.7063 | 22.4 | 28.85 | -1316.36 | 1317.5 | 5.01 | -3.28 | -5.89 | 8.89 | 1316.36 | 3.28 | 65.6657
0.7081 | 22.68 | 28.66 | -1315.67 | 1316.81 | 4.97 | -3.27 | -5.89 | 8.88 | 1315.67 | 3.27 | 65.833
0.7099 | 22.96 | 28.48 | -1314.99 | 1316.13 | 4.94 | -3.26 | -5.9 | 8.86 | 1314.99 | 3.26 | 66.0004
0.7117 | 22.71 | 28.12 | -1315.09 | 1316.22 | 4.89 | -3.22 | -5.9 | 8.82 | 1315.09 | 3.22 | 66.1677
0.7135 | 22.5 | 27.63 | -1315.95 | 1317.07 | 4.84 | -3.2 | -5.9 | 8.78 | 1315.95 | 3.2 | 66.3351
0.7153 | 22.16 | 27.01 | -1317.24 | 1318.32 | 4.78 | -3.17 | -5.9 | 8.73 | 1317.24 | 3.17 | 66.5024
0.7171 | 21.76 | 26.4 | -1318.86 | 1319.92 | 4.73 | -3.15 | -5.9 | 8.69 | 1318.86 | 3.15 | 66.6698
0.7189 | 21.05 | 25.55 | -1320.51 | 1321.53 | 4.68 | -3.11 | -5.9 | 8.64 | 1320.51 | 3.11 | 66.8371
0.7207 | 20.5 | 24.91 | -1322.19 | 1323.17 | 4.64 | -3.09 | -5.89 | 8.6 | 1322.19 | 3.09 | 67.0045
0.7225 | 20.16 | 24.57 | -1323.89 | 1324.87 | 4.6 | -3.09 | -5.87 | 8.57 | 1323.89 | 3.09 | 67.1718
0.7242 | 19.74 | 24.34 | -1325.03 | 1325.99 | 4.57 | -3.08 | -5.87 | 8.56 | 1325.03 | 3.08 | 67.3299
0.726 | 19.25 | 24.31 | -1326.15 | 1327.11 | 4.57 | -3.06 | -5.87 | 8.55 | 1326.15 | 3.06 | 67.4972
0.7278 | 18.79 | 24.31 | -1327.19 | 1328.15 | 4.56 | -3.04 | -5.87 | 8.54 | 1327.19 | 3.04 | 67.6646
0.7296 | 18.55 | 24.45 | -1327.65 | 1328.62 | 4.55 | -3.04 | -5.87 | 8.54 | 1327.65 | 3.04 | 67.8319
0.7314 | 18.5 | 24.6 | -1327.67 | 1328.64 | 4.55 | -3.05 | -5.88 | 8.54 | 1327.67 | 3.05 | 67.9993
0.7332 | 18.57 | 24.77 | -1327.36 | 1328.35 | 4.53 | -3.06 | -5.89 | 8.55 | 1327.36 | 3.06 | 68.1666
0.735 | 18.64 | 24.95 | -1327.05 | 1328.04 | 4.51 | -3.07 | -5.89 | 8.54 | 1327.05 | 3.07 | 68.334
0.7368 | 18.71 | 25.13 | -1326.73 | 1327.74 | 4.48 | -3.07 | -5.9 | 8.54 | 1326.73 | 3.07 | 68.5013
0.7386 | 18.72 | 25.3 | -1327.25 | 1328.26 | 4.45 | -3.08 | -5.91 | 8.54 | 1327.25 | 3.08 | 68.6687
0.7404 | 18.6 | 25.29 | -1328.26 | 1329.25 | 4.43 | -3.12 | -5.93 | 8.55 | 1328.26 | 3.12 | 68.836
0.7422 | 18.29 | 25.17 | -1329.76 | 1330.74 | 4.41 | -3.15 | -5.94 | 8.57 | 1329.76 | 3.15 | 69.0033
0.744 | 17.97 | 25.19 | -1331.98 | 1332.95 | 4.4 | -3.14 | -5.94 | 8.56 | 1331.98 | 3.14 | 69.1707
0.7458 | 17.76 | 25.24 | -1334.65 | 1335.63 | 4.38 | -3.13 | -5.94 | 8.54 | 1334.65 | 3.13 | 69.338
0.7476 | 18.19 | 25.47 | -1335.73 | 1336.73 | 4.37 | -3.16 | -5.96 | 8.56 | 1335.73 | 3.16 | 69.5054
0.7493 | 18.73 | 25.82 | -1336.34 | 1337.35 | 4.38 | -3.2 | -5.97 | 8.59 | 1336.34 | 3.2 | 69.6634
0.7511 | 19.18 | 26.31 | -1336.95 | 1337.97 | 4.37 | -3.24 | -5.98 | 8.61 | 1336.95 | 3.24 | 69.8308
0.7529 | 19.63 | 26.94 | -1336.95 | 1337.98 | 4.38 | -3.27 | -5.99 | 8.63 | 1336.95 | 3.27 | 69.9981
0.7547 | 20.08 | 27.66 | -1336.53 | 1337.59 | 4.39 | -3.29 | -6 | 8.66 | 1336.53 | 3.29 | 70.1655
0.7565 | 20.62 | 28.25 | -1335.65 | 1336.73 | 4.4 | -3.31 | -6.02 | 8.68 | 1335.65 | 3.31 | 70.3328
0.7583 | 21.29 | 28.66 | -1334.18 | 1335.28 | 4.4 | -3.33 | -6.02 | 8.7 | 1334.18 | 3.33 | 70.5002
0.7601 | 21.51 | 28.97 | -1332.59 | 1333.7 | 4.42 | -3.32 | -6.03 | 8.71 | 1332.59 | 3.32 | 70.6675
0.7619 | 21.45 | 29.22 | -1330.94 | 1332.03 | 4.43 | -3.31 | -6.03 | 8.72 | 1330.94 | 3.31 | 70.8349
0.7637 | 21.4 | 29.46 | -1329.29 | 1330.37 | 4.45 | -3.29 | -6.02 | 8.72 | 1329.29 | 3.29 | 71.0022
0.7655 | 20.83 | 29.43 | -1328.02 | 1329.06 | 4.46 | -3.24 | -6.03 | 8.7 | 1328.02 | 3.24 | 71.1696
0.7673 | 20.25 | 29.39 | -1326.76 | 1327.77 | 4.48 | -3.2 | -6.03 | 8.69 | 1326.76 | 3.2 | 71.3369
0.7691 | 19.62 | 29.32 | -1325.54 | 1326.51 | 4.5 | -3.14 | -6.03 | 8.67 | 1325.54 | 3.14 | 71.5043
0.7709 | 18.96 | 29.22 | -1324.36 | 1325.3 | 4.52 | -3.08 | -6.02 | 8.65 | 1324.36 | 3.08 | 71.6716
0.7727 | 18.29 | 29.12 | -1323.18 | 1324.08 | 4.54 | -3.02 | -6.02 | 8.63 | 1323.18 | 3.02 | 71.839
0.7744 | 17.78 | 29.15 | -1321.76 | 1322.64 | 4.55 | -2.96 | -6.02 | 8.6 | 1321.76 | 2.96 | 71.997
0.7762 | 17.48 | 29.37 | -1320 | 1320.87 | 4.54 | -2.9 | -6.02 | 8.56 | 1320 | 2.9 | 72.1644
0.778 | 17.6 | 29.54 | -1316.97 | 1317.84 | 4.54 | -2.85 | -6.02 | 8.54 | 1316.97 | 2.85 | 72.3317
0.7798 | 17.93 | 29.89 | -1313.09 | 1313.95 | 4.53 | -2.8 | -6.02 | 8.5 | 1313.09 | 2.8 | 72.4991
0.7816 | 18.61 | 30.23 | -1308.85 | 1309.73 | 4.52 | -2.75 | -6.03 | 8.47 | 1308.85 | 2.75 | 72.6664
0.7834 | 19.18 | 30.43 | -1303.92 | 1304.82 | 4.48 | -2.7 | -6.03 | 8.42 | 1303.92 | 2.7 | 72.8338
0.7852 | 19.69 | 30.73 | -1297.97 | 1298.9 | 4.42 | -2.63 | -6.02 | 8.33 | 1297.97 | 2.63 | 73.0011
0.787 | 20.01 | 30.9 | -1292.31 | 1293.25 | 4.35 | -2.54 | -6.02 | 8.25 | 1292.31 | 2.54 | 73.1685
0.7888 | 19.94 | 30.81 | -1287.24 | 1288.18 | 4.3 | -2.42 | -6.02 | 8.16 | 1287.24 | 2.42 | 73.3358
0.7906 | 19.72 | 31.23 | -1281.99 | 1282.94 | 4.19 | -2.28 | -5.99 | 8.02 | 1281.99 | 2.28 | 73.5032
0.7924 | 19.53 | 31.45 | -1275.83 | 1276.78 | 4.06 | -2.12 | -5.95 | 7.87 | 1275.83 | 2.12 | 73.6705
0.7942 | 19.37 | 31.46 | -1269.39 | 1270.33 | 3.94 | -1.96 | -5.92 | 7.71 | 1269.39 | 1.96 | 73.8379
0.796 | 19.19 | 31.43 | -1265.03 | 1265.96 | 3.88 | -1.85 | -5.92 | 7.64 | 1265.03 | 1.85 | 74.0052
0.7977 | 19.4 | 31.55 | -1261.42 | 1262.35 | 3.85 | -1.77 | -5.92 | 7.61 | 1261.42 | 1.77 | 74.1633
0.7995 | 19.68 | 31.71 | -1258.75 | 1259.68 | 3.83 | -1.71 | -5.92 | 7.59 | 1258.75 | 1.71 | 74.3306
0.8013 | 19.65 | 31.73 | -1256.62 | 1257.52 | 3.8 | -1.64 | -5.93 | 7.55 | 1256.62 | 1.64 | 74.498
0.8031 | 19.62 | 31.75 | -1254.48 | 1255.37 | 3.77 | -1.57 | -5.94 | 7.52 | 1254.48 | 1.57 | 74.6653
0.8049 | 19.48 | 31.58 | -1252.94 | 1253.81 | 3.72 | -1.49 | -5.95 | 7.48 | 1252.94 | 1.49 | 74.8327
0.8067 | 18.78 | 30.83 | -1252.07 | 1252.89 | 3.67 | -1.38 | -5.96 | 7.42 | 1252.07 | 1.38 | 75
0.8085 | 17.77 | 30.2 | -1250.8 | 1251.59 | 3.62 | -1.27 | -5.97 | 7.37 | 1250.8 | 1.27 | 75.1673
0.8103 | 17.21 | 29.64 | -1249.93 | 1250.68 | 3.58 | -1.19 | -5.99 | 7.34 | 1249.93 | 1.19 | 75.3347
0.8121 | 17.04 | 29.56 | -1248.18 | 1248.92 | 3.54 | -1.12 | -6 | 7.31 | 1248.18 | 1.12 | 75.502
0.8139 | 17.05 | 29.58 | -1245.91 | 1246.64 | 3.52 | -1.08 | -6 | 7.29 | 1245.91 | 1.08 | 75.6694
0.8157 | 16.86 | 29.41 | -1243.02 | 1243.73 | 3.49 | -1.02 | -6 | 7.25 | 1243.02 | 1.02 | 75.8367
0.8175 | 16.84 | 29.49 | -1239.28 | 1239.98 | 3.44 | -0.95 | -6 | 7.21 | 1239.28 | 0.95 | 76.0041
0.8193 | 16.85 | 29.42 | -1234.81 | 1235.49 | 3.36 | -0.87 | -6.01 | 7.15 | 1234.81 | 0.87 | 76.1714
0.8211 | 16.22 | 28.19 | -1229.3 | 1229.93 | 3.22 | -0.73 | -6 | 7.07 | 1229.3 | 0.73 | 76.3388
0.8228 | 16.15 | 27.53 | -1225.17 | 1225.76 | 3.16 | -0.64 | -6 | 7.04 | 1225.17 | 0.64 | 76.4968
0.8246 | 16.06 | 27.03 | -1221.71 | 1222.28 | 3.13 | -0.57 | -5.99 | 7.02 | 1221.71 | 0.57 | 76.6642
0.8264 | 15.63 | 26.19 | -1218.56 | 1219.1 | 3.09 | -0.48 | -5.98 | 6.99 | 1218.56 | 0.48 | 76.8315
0.8282 | 14.9 | 25.63 | -1215.94 | 1216.45 | 3.07 | -0.4 | -5.97 | 6.97 | 1215.94 | 0.4 | 76.9989
0.83 | 14.92 | 25.15 | -1214.07 | 1214.58 | 3.08 | -0.37 | -5.97 | 6.97 | 1214.07 | 0.37 | 77.1662
0.8318 | 14.96 | 23.92 | -1211.66 | 1212.14 | 3.07 | -0.33 | -5.96 | 6.96 | 1211.66 | 0.33 | 77.3336
0.8336 | 14.71 | 22.16 | -1208.83 | 1209.26 | 3.02 | -0.29 | -5.94 | 6.91 | 1208.83 | 0.29 | 77.5009
0.8354 | 14.23 | 21.05 | -1206.9 | 1207.3 | 2.99 | -0.23 | -5.94 | 6.87 | 1206.9 | 0.23 | 77.6683
0.8372 | 13.73 | 19.82 | -1205.48 | 1205.85 | 2.96 | -0.2 | -5.93 | 6.83 | 1205.48 | 0.2 | 77.8356
0.839 | 13.05 | 18.93 | -1204.36 | 1204.69 | 2.99 | -0.14 | -5.91 | 6.81 | 1204.36 | 0.14 | 78.003
0.8408 | 12.48 | 17.83 | -1203.46 | 1203.77 | 2.97 | -0.09 | -5.9 | 6.77 | 1203.46 | 0.09 | 78.1703
0.8426 | 11.9 | 16.63 | -1203.72 | 1204.01 | 2.95 | -0.06 | -5.89 | 6.74 | 1203.72 | 0.06 | 78.3377
0.8444 | 11.65 | 15.49 | -1204.78 | 1205.06 | 2.93 | -0.04 | -5.88 | 6.7 | 1204.78 | 0.04 | 78.505
0.8462 | 11.57 | 14.39 | -1205.62 | 1205.91 | 2.92 | -0.02 | -5.86 | 6.68 | 1205.62 | 0.02 | 78.6724
0.8479 | 11.82 | 13.4 | -1205.73 | 1206.02 | 2.92 | 0.03 | -5.83 | 6.67 | 1205.73 | -0.03 | 78.8304
0.8497 | 12.22 | 12.86 | -1205.32 | 1205.63 | 2.93 | 0.05 | -5.81 | 6.67 | 1205.32 | -0.05 | 78.9978
0.8515 | 12.43 | 11.76 | -1204.91 | 1205.23 | 2.94 | 0.05 | -5.79 | 6.64 | 1204.91 | -0.05 | 79.1651
0.8533 | 12.52 | 10.88 | -1204.27 | 1204.6 | 2.94 | 0.05 | -5.76 | 6.62 | 1204.27 | -0.05 | 79.3325
0.8551 | 12.46 | 10.28 | -1203.35 | 1203.67 | 2.96 | 0.04 | -5.75 | 6.6 | 1203.35 | -0.04 | 79.4998
0.8569 | 12.63 | 9.87 | -1202.21 | 1202.53 | 2.97 | 0.03 | -5.74 | 6.59 | 1202.21 | -0.03 | 79.6672
0.8587 | 13.05 | 9.68 | -1200.82 | 1201.14 | 2.98 | 0.02 | -5.74 | 6.59 | 1200.82 | -0.02 | 79.8345
0.8605 | 13.29 | 9.31 | -1198.96 | 1199.28 | 3.01 | 0.01 | -5.73 | 6.59 | 1198.96 | -0.01 | 80.0019
0.8623 | 13.4 | 8.84 | -1196.85 | 1197.16 | 3.05 | 0.01 | -5.72 | 6.59 | 1196.85 | -0.01 | 80.1692
0.8641 | 13.95 | 8.68 | -1194.38 | 1194.7 | 3.07 | -0.04 | -5.72 | 6.58 | 1194.38 | 0.04 | 80.3366
0.8659 | 14.67 | 8.88 | -1191.2 | 1191.52 | 3.08 | -0.09 | -5.71 | 6.57 | 1191.2 | 0.09 | 80.5039
0.8677 | 14.91 | 8.74 | -1188.15 | 1188.47 | 3.1 | -0.12 | -5.71 | 6.57 | 1188.15 | 0.12 | 80.6713
0.8695 | 15.14 | 8.61 | -1185.13 | 1185.45 | 3.13 | -0.15 | -5.7 | 6.57 | 1185.13 | 0.15 | 80.8386
0.8712 | 15.26 | 8.74 | -1183.19 | 1183.5 | 3.13 | -0.19 | -5.7 | 6.57 | 1183.19 | 0.19 | 80.9967
0.873 | 14.97 | 8.74 | -1181.75 | 1182.05 | 3.13 | -0.19 | -5.7 | 6.57 | 1181.75 | 0.19 | 81.164
0.8748 | 14.28 | 8.6 | -1180.92 | 1181.2 | 3.13 | -0.14 | -5.7 | 6.57 | 1180.92 | 0.14 | 81.3313
0.8766 | 13.67 | 8.56 | -1180.46 | 1180.71 | 3.13 | -0.11 | -5.7 | 6.56 | 1180.46 | 0.11 | 81.4987
0.8784 | 13.51 | 9.43 | -1181.77 | 1182 | 3.15 | -0.14 | -5.7 | 6.59 | 1181.77 | 0.14 | 81.666
0.8802 | 12.98 | 10.34 | -1183.75 | 1183.97 | 3.21 | -0.18 | -5.71 | 6.65 | 1183.75 | 0.18 | 81.8334
0.882 | 12.44 | 10.96 | -1184.82 | 1185.04 | 3.28 | -0.18 | -5.72 | 6.71 | 1184.82 | 0.18 | 82.0007
0.8838 | 11.95 | 11.47 | -1185.42 | 1185.63 | 3.34 | -0.17 | -5.73 | 6.76 | 1185.42 | 0.17 | 82.1681
0.8856 | 12.3 | 11.86 | -1184.27 | 1184.49 | 3.38 | -0.18 | -5.73 | 6.78 | 1184.27 | 0.18 | 82.3354
0.8874 | 12.89 | 12.33 | -1182.26 | 1182.48 | 3.41 | -0.19 | -5.74 | 6.81 | 1182.26 | 0.19 | 82.5028
0.8892 | 13.38 | 12.83 | -1179.83 | 1180.06 | 3.44 | -0.2 | -5.74 | 6.83 | 1179.83 | 0.2 | 82.6701
0.891 | 14.02 | 13.53 | -1176.88 | 1177.11 | 3.49 | -0.22 | -5.75 | 6.89 | 1176.88 | 0.22 | 82.8375
0.8928 | 14.43 | 14.08 | -1174.17 | 1174.41 | 3.53 | -0.22 | -5.77 | 6.92 | 1174.17 | 0.22 | 83.0048
0.8946 | 14.16 | 14.24 | -1172.2 | 1172.44 | 3.54 | -0.18 | -5.77 | 6.93 | 1172.2 | 0.18 | 83.1722
0.8963 | 13.43 | 14.13 | -1170.8 | 1171.03 | 3.53 | -0.11 | -5.77 | 6.93 | 1170.8 | 0.11 | 83.3302
0.8981 | 13.01 | 14.24 | -1170.21 | 1170.43 | 3.53 | -0.06 | -5.78 | 6.93 | 1170.21 | 0.06 | 83.4976
0.8999 | 12.77 | 14.39 | -1170.51 | 1170.73 | 3.53 | -0.04 | -5.79 | 6.94 | 1170.51 | 0.04 | 83.6649
0.9017 | 12.53 | 14.76 | -1171.77 | 1171.99 | 3.52 | 0 | -5.82 | 6.94 | 1171.77 | 0 | 83.8323
0.9035 | 12.41 | 15.18 | -1172.78 | 1173 | 3.5 | 0.04 | -5.84 | 6.94 | 1172.78 | -0.04 | 83.9996
0.9053 | 12.59 | 15.57 | -1172.4 | 1172.63 | 3.49 | 0.06 | -5.85 | 6.93 | 1172.4 | -0.06 | 84.167
0.9071 | 12.93 | 16.05 | -1171.23 | 1171.48 | 3.49 | 0.07 | -5.86 | 6.93 | 1171.23 | -0.07 | 84.3343
0.9089 | 12.49 | 16.21 | -1169.86 | 1170.11 | 3.48 | 0.13 | -5.86 | 6.93 | 1169.86 | -0.13 | 84.5017
0.9107 | 12.01 | 16.36 | -1168.49 | 1168.73 | 3.46 | 0.19 | -5.86 | 6.93 | 1168.49 | -0.19 | 84.669
0.9125 | 12.21 | 16.95 | -1167.04 | 1167.29 | 3.45 | 0.22 | -5.86 | 6.92 | 1167.04 | -0.22 | 84.8364
0.9143 | 12.52 | 17.27 | -1164.75 | 1165.01 | 3.43 | 0.24 | -5.87 | 6.91 | 1164.75 | -0.24 | 85.0037
0.9161 | 12.36 | 17.38 | -1161.41 | 1161.68 | 3.36 | 0.3 | -5.88 | 6.87 | 1161.41 | -0.3 | 85.1711
0.9179 | 12.21 | 17.44 | -1158.53 | 1158.8 | 3.3 | 0.36 | -5.89 | 6.84 | 1158.53 | -0.36 | 85.3384
0.9197 | 11.92 | 17.3 | -1156.56 | 1156.82 | 3.27 | 0.4 | -5.89 | 6.83 | 1156.56 | -0.4 | 85.5058
0.9214 | 11.47 | 17.24 | -1154.98 | 1155.24 | 3.24 | 0.45 | -5.9 | 6.82 | 1154.98 | -0.45 | 85.6638
0.9232 | 11.23 | 17.54 | -1154.06 | 1154.31 | 3.23 | 0.49 | -5.9 | 6.82 | 1154.06 | -0.49 | 85.8312
0.925 | 11.52 | 18.13 | -1152.24 | 1152.51 | 3.24 | 0.5 | -5.91 | 6.83 | 1152.24 | -0.5 | 85.9985
0.9268 | 11.9 | 18.69 | -1150.05 | 1150.34 | 3.25 | 0.5 | -5.91 | 6.84 | 1150.05 | -0.5 | 86.1659
0.9286 | 11.84 | 18.89 | -1147.36 | 1147.66 | 3.24 | 0.53 | -5.91 | 6.84 | 1147.36 | -0.53 | 86.3332
0.9304 | 11.62 | 19.06 | -1145.2 | 1145.5 | 3.23 | 0.57 | -5.92 | 6.84 | 1145.2 | -0.57 | 86.5006
0.9322 | 11.37 | 19.23 | -1143.89 | 1144.18 | 3.21 | 0.61 | -5.92 | 6.84 | 1143.89 | -0.61 | 86.6679
0.934 | 11.31 | 19.32 | -1142.62 | 1142.92 | 3.2 | 0.64 | -5.92 | 6.83 | 1142.62 | -0.64 | 86.8353
0.9358 | 11 | 19.38 | -1142 | 1142.3 | 3.17 | 0.68 | -5.92 | 6.82 | 1142 | -0.68 | 87.0026
0.9376 | 10.92 | 19.45 | -1142.29 | 1142.58 | 3.15 | 0.69 | -5.92 | 6.82 | 1142.29 | -0.69 | 87.17
0.9394 | 11.16 | 19.46 | -1142.2 | 1142.49 | 3.14 | 0.69 | -5.92 | 6.81 | 1142.2 | -0.69 | 87.3373
0.9412 | 11.3 | 19.39 | -1141.91 | 1142.21 | 3.13 | 0.7 | -5.92 | 6.8 | 1141.91 | -0.7 | 87.5046
0.943 | 11.24 | 19.43 | -1141.98 | 1142.28 | 3.12 | 0.71 | -5.92 | 6.8 | 1141.98 | -0.71 | 87.672
0.9447 | 11.09 | 19.52 | -1142.28 | 1142.58 | 3.1 | 0.73 | -5.92 | 6.8 | 1142.28 | -0.73 | 87.83
0.9465 | 11.04 | 19.86 | -1142.94 | 1143.24 | 3.1 | 0.73 | -5.93 | 6.8 | 1142.94 | -0.73 | 87.9974
0.9483 | 11.34 | 20.03 | -1143.27 | 1143.58 | 3.1 | 0.71 | -5.94 | 6.81 | 1143.27 | -0.71 | 88.1647
0.9501 | 11.5 | 20.08 | -1143.45 | 1143.76 | 3.09 | 0.71 | -5.94 | 6.8 | 1143.45 | -0.71 | 88.3321
0.9519 | 11.61 | 20.01 | -1143.24 | 1143.55 | 3.07 | 0.71 | -5.95 | 6.8 | 1143.24 | -0.71 | 88.4994
0.9537 | 11.82 | 20.12 | -1142.75 | 1143.06 | 3.06 | 0.71 | -5.95 | 6.8 | 1142.75 | -0.71 | 88.6668
0.9555 | 11.98 | 20.17 | -1141.62 | 1141.94 | 3.04 | 0.71 | -5.96 | 6.81 | 1141.62 | -0.71 | 88.8341
0.9573 | 12.78 | 20.26 | -1140.22 | 1140.53 | 2.99 | 0.71 | -5.99 | 6.81 | 1140.22 | -0.71 | 89.0015
0.9591 | 13.59 | 20.28 | -1138.02 | 1138.33 | 2.97 | 0.69 | -6.01 | 6.81 | 1138.02 | -0.69 | 89.1688
0.9609 | 14.08 | 20.13 | -1135.59 | 1135.91 | 2.95 | 0.68 | -6.02 | 6.8 | 1135.59 | -0.68 | 89.3362
0.9627 | 14.12 | 19.95 | -1134.68 | 1135 | 2.94 | 0.69 | -6.02 | 6.8 | 1134.68 | -0.69 | 89.5035
0.9645 | 14.13 | 19.8 | -1134.44 | 1134.76 | 2.93 | 0.7 | -6.02 | 6.79 | 1134.44 | -0.7 | 89.6709
0.9663 | 14.14 | 19.67 | -1134.67 | 1135 | 2.92 | 0.71 | -6.02 | 6.78 | 1134.67 | -0.71 | 89.8382
0.9681 | 14.17 | 19.55 | -1134.95 | 1135.27 | 2.9 | 0.71 | -6.02 | 6.78 | 1134.95 | -0.71 | 90.0056
0.9698 | 13.83 | 19.18 | -1134.93 | 1135.24 | 2.88 | 0.73 | -6.02 | 6.78 | 1134.93 | -0.73 | 90.1636
0.9716 | 13.59 | 19.06 | -1134.77 | 1135.08 | 2.87 | 0.75 | -6.02 | 6.78 | 1134.77 | -0.75 | 90.331
0.9734 | 13.44 | 19.05 | -1134.56 | 1134.87 | 2.87 | 0.76 | -6.02 | 6.78 | 1134.56 | -0.76 | 90.4983
0.9752 | 13.4 | 19.04 | -1134.68 | 1135 | 2.86 | 0.77 | -6.02 | 6.77 | 1134.68 | -0.77 | 90.6657
0.977 | 13.51 | 19.07 | -1134.9 | 1135.22 | 2.85 | 0.76 | -6.03 | 6.77 | 1134.9 | -0.76 | 90.833
0.9788 | 13.72 | 19.17 | -1134.73 | 1135.05 | 2.84 | 0.75 | -6.03 | 6.77 | 1134.73 | -0.75 | 91.0004
0.9806 | 13.99 | 19.41 | -1135.12 | 1135.46 | 2.84 | 0.74 | -6.03 | 6.77 | 1135.12 | -0.74 | 91.1677
0.9824 | 14.13 | 19.49 | -1135.55 | 1135.89 | 2.83 | 0.74 | -6.04 | 6.77 | 1135.55 | -0.74 | 91.3351
0.9842 | 13.77 | 19.3 | -1136.3 | 1136.64 | 2.82 | 0.76 | -6.04 | 6.77 | 1136.3 | -0.76 | 91.5024
0.986 | 13.21 | 19.08 | -1137.71 | 1138.04 | 2.81 | 0.8 | -6.04 | 6.77 | 1137.71 | -0.8 | 91.6698
0.9878 | 12.64 | 18.88 | -1140.03 | 1140.36 | 2.8 | 0.83 | -6.05 | 6.78 | 1140.03 | -0.83 | 91.8371
0.9896 | 12.26 | 18.75 | -1142.63 | 1142.97 | 2.8 | 0.85 | -6.06 | 6.79 | 1142.63 | -0.85 | 92.0045
0.9914 | 12.32 | 18.89 | -1144.29 | 1144.63 | 2.8 | 0.85 | -6.06 | 6.8 | 1144.29 | -0.85 | 92.1718
0.9932 | 12.84 | 19.15 | -1144.94 | 1145.29 | 2.81 | 0.82 | -6.07 | 6.8 | 1144.94 | -0.82 | 92.3392
0.9949 | 13.54 | 19.36 | -1144.79 | 1145.16 | 2.8 | 0.78 | -6.08 | 6.8 | 1144.79 | -0.78 | 92.4972
0.9967 | 13.9 | 19.42 | -1144.2 | 1144.58 | 2.79 | 0.76 | -6.08 | 6.8 | 1144.2 | -0.76 | 92.6646
0.9985 | 13.7 | 19.51 | -1143.57 | 1143.94 | 2.77 | 0.77 | -6.09 | 6.8 | 1143.57 | -0.77 | 92.8319
1.0003 | 13.52 | 19.61 | -1142.91 | 1143.28 | 2.75 | 0.78 | -6.09 | 6.8 | 1142.91 | -0.78 | 92.9993
1.0021 | 13.44 | 19.73 | -1142.17 | 1142.54 | 2.72 | 0.78 | -6.1 | 6.79 | 1142.17 | -0.78 | 93.1666
1.0039 | 13.37 | 19.91 | -1141.81 | 1142.19 | 2.7 | 0.78 | -6.11 | 6.79 | 1141.81 | -0.78 | 93.334
1.0057 | 13.33 | 20.16 | -1141.88 | 1142.27 | 2.69 | 0.78 | -6.11 | 6.79 | 1141.88 | -0.78 | 93.5013
1.0075 | 13.29 | 20.41 | -1141.96 | 1142.35 | 2.67 | 0.78 | -6.12 | 6.79 | 1141.96 | -0.78 | 93.6687
1.0093 | 13.26 | 20.49 | -1142.38 | 1142.77 | 2.66 | 0.78 | -6.12 | 6.79 | 1142.38 | -0.78 | 93.836
1.0111 | 13.23 | 20.56 | -1142.81 | 1143.2 | 2.66 | 0.78 | -6.13 | 6.79 | 1142.81 | -0.78 | 94.0033
1.0129 | 13.16 | 20.65 | -1143.66 | 1144.05 | 2.65 | 0.78 | -6.13 | 6.79 | 1143.66 | -0.78 | 94.1707
1.0147 | 13.1 | 20.81 | -1144.86 | 1145.26 | 2.65 | 0.77 | -6.14 | 6.8 | 1144.86 | -0.77 | 94.338
1.0165 | 13.1 | 21.04 | -1146.15 | 1146.55 | 2.66 | 0.76 | -6.14 | 6.8 | 1146.15 | -0.76 | 94.5054
1.0182 | 13.58 | 21.27 | -1147.43 | 1147.85 | 2.67 | 0.72 | -6.15 | 6.8 | 1147.43 | -0.72 | 94.6634
1.02 | 13.79 | 21.34 | -1148.31 | 1148.73 | 2.68 | 0.71 | -6.15 | 6.81 | 1148.31 | -0.71 | 94.8308
1.0218 | 13.65 | 21.11 | -1149.21 | 1149.62 | 2.68 | 0.71 | -6.16 | 6.82 | 1149.21 | -0.71 | 94.9981
1.0236 | 13.33 | 21.03 | -1150.61 | 1151.01 | 2.68 | 0.72 | -6.17 | 6.82 | 1150.61 | -0.72 | 95.1655
1.0254 | 13.35 | 21.39 | -1151.17 | 1151.58 | 2.67 | 0.72 | -6.17 | 6.83 | 1151.17 | -0.72 | 95.3328
1.0272 | 13.39 | 21.56 | -1151.77 | 1152.19 | 2.67 | 0.71 | -6.18 | 6.83 | 1151.77 | -0.71 | 95.5002
1.029 | 13.81 | 21.54 | -1152.13 | 1152.55 | 2.67 | 0.69 | -6.18 | 6.83 | 1152.13 | -0.69 | 95.6675
1.0308 | 14.18 | 21.61 | -1152.09 | 1152.52 | 2.67 | 0.66 | -6.19 | 6.83 | 1152.09 | -0.66 | 95.8349
1.0326 | 14.61 | 21.88 | -1151.59 | 1152.03 | 2.68 | 0.64 | -6.19 | 6.84 | 1151.59 | -0.64 | 96.0022
1.0344 | 15.09 | 21.99 | -1150.45 | 1150.91 | 2.68 | 0.62 | -6.19 | 6.84 | 1150.45 | -0.62 | 96.1696
1.0362 | 15.56 | 22.01 | -1148.93 | 1149.39 | 2.67 | 0.6 | -6.2 | 6.84 | 1148.93 | -0.6 | 96.3369
1.038 | 15.49 | 21.8 | -1146.48 | 1146.94 | 2.67 | 0.62 | -6.19 | 6.84 | 1146.48 | -0.62 | 96.5043
1.0398 | 15.32 | 21.49 | -1143.18 | 1143.64 | 2.68 | 0.64 | -6.18 | 6.83 | 1143.18 | -0.64 | 96.6716
1.0416 | 14.86 | 21.48 | -1142.11 | 1142.56 | 2.66 | 0.67 | -6.18 | 6.83 | 1142.11 | -0.67 | 96.839
1.0433 | 14.82 | 21.67 | -1141.06 | 1141.52 | 2.65 | 0.68 | -6.19 | 6.83 | 1141.06 | -0.68 | 96.997
1.0451 | 14.76 | 21.83 | -1140.53 | 1140.99 | 2.64 | 0.69 | -6.2 | 6.83 | 1140.53 | -0.69 | 97.1644
1.0469 | 15.22 | 22.15 | -1140.19 | 1140.67 | 2.65 | 0.65 | -6.2 | 6.84 | 1140.19 | -0.65 | 97.3317
1.0487 | 15.14 | 22.1 | -1139.57 | 1140.05 | 2.64 | 0.66 | -6.21 | 6.83 | 1139.57 | -0.66 | 97.4991
1.0505 | 14.55 | 21.97 | -1139.02 | 1139.49 | 2.64 | 0.7 | -6.2 | 6.84 | 1139.02 | -0.7 | 97.6664
1.0523 | 14.46 | 21.95 | -1139.38 | 1139.85 | 2.64 | 0.69 | -6.21 | 6.84 | 1139.38 | -0.69 | 97.8338
1.0541 | 14.28 | 21.86 | -1139.62 | 1140.08 | 2.63 | 0.69 | -6.21 | 6.84 | 1139.62 | -0.69 | 98.0011
1.0559 | 14.06 | 22.01 | -1139.65 | 1140.12 | 2.63 | 0.71 | -6.21 | 6.84 | 1139.65 | -0.71 | 98.1685
1.0577 | 13.86 | 22.18 | -1139.76 | 1140.23 | 2.62 | 0.72 | -6.21 | 6.84 | 1139.76 | -0.72 | 98.3358
1.0595 | 13.74 | 22.23 | -1139.71 | 1140.19 | 2.62 | 0.72 | -6.21 | 6.84 | 1139.71 | -0.72 | 98.5032
1.0613 | 13.9 | 22.44 | -1140.06 | 1140.54 | 2.62 | 0.7 | -6.22 | 6.85 | 1140.06 | -0.7 | 98.6705
1.0631 | 14.27 | 22.86 | -1140.55 | 1141.04 | 2.63 | 0.67 | -6.23 | 6.85 | 1140.55 | -0.67 | 98.8379
1.0649 | 14.36 | 22.98 | -1140.58 | 1141.09 | 2.63 | 0.67 | -6.24 | 6.86 | 1140.58 | -0.67 | 99.0052
1.0667 | 14.4 | 23.08 | -1140.52 | 1141.04 | 2.63 | 0.67 | -6.24 | 6.87 | 1140.52 | -0.67 | 99.1726
1.0684 | 14.42 | 23.16 | -1140.39 | 1140.91 | 2.65 | 0.68 | -6.24 | 6.88 | 1140.39 | -0.68 | 99.3306
1.0702 | 14.57 | 23.16 | -1140.27 | 1140.79 | 2.65 | 0.67 | -6.25 | 6.88 | 1140.27 | -0.67 | 99.498
1.072 | 14.35 | 22.8 | -1139.88 | 1140.39 | 2.65 | 0.69 | -6.26 | 6.89 | 1139.88 | -0.69 | 99.6653
1.0738 | 13.91 | 22.61 | -1139.85 | 1140.35 | 2.65 | 0.71 | -6.26 | 6.89 | 1139.85 | -0.71 | 99.8327
1.0756 | 13.82 | 22.74 | -1139.85 | 1140.34 | 2.65 | 0.71 | -6.26 | 6.9 | 1139.85 | -0.71 | 100
| Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My
Minimum | -50.41 | -104.67 | -1747.57 | 127.9 | -0.94 | -3.33 | -6.26 | 0.52 | 126.9 | -11.15
Maximum | 52.97 | 66.5 | -126.9 | 1749.7 | 5.13 | 11.15 | -0.29 | 12.89 | 1747.57 | 3.33

## Diagr_Individ
Individual Loads
Contact forces and moments, measured telemetrically in knee implants of 8 subjects.
Averages of one individual subject.
Fx  lateral, Fy  anterior, -Fz  distal, Fres = Resultant Force
Mx around x, My around y, Mz around z, Mres = Resultant Moment
z-axis = axis of implant stem!
From Bergmann et al, PlosOne (2014) "Standardized Loads Acting in Knee Implants"
Subject K9L
Subject K9L

## Dat_BW
Data File: | k9l_2246_863_6_i1.akf
Curve Alignment Program: | 1.4.11
------------------------MEASURING INFOS------------------------
Diagram Title #1 | Forces and Moments at Knee Joint
Diagram Title #2 | Measured Relative to Implant
Comment #1 | Knee Joint; Gaitanalysis; Stand up; Seat Height: 45 cm;
Comment #2 | K9L, 14,5 Months PO *
BodyWeight [N]: | 1070
Patient(en): | K9L | [7 empty] | Subject | K9L
Implant Type: | Knee Joint
Varus Angle Leg [°] | 7 | [7 empty] | Normalized Loads in %BW (Forces) and %BWm (moments)
Axes Directions X, Y, Z: | Lateral, Anterior, Superior
Axes Origin: | Tibial Plateau at Stem Axis | [7 empty] | BW [kg] = | 109.0724
Transformation X Y Z (°): | 6.5 | -3.5 | -2
Displacement X, Y, Z (mm): | 0 | 0 | 0
Activity Code: | KneeActivities | 1P_1_15_8
---------------------------DATA--------------------------------
Number of Data Sets = | 601
Number of analog channels = | 0
Time-Offset(sec) = | 0
Time | Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My | Time
[s] | [%BW] | [%BW] | [%BW] | [%BW] | [%BWm] | [%BWm] | [%BWm] | [%BWm] | [%BW] | [%BWm] | [%Cycle]
0 | -0.10654 | -0.7243 | -14.1617 | 14.1879 | 0.01028 | -0.033645 | -0.027103 | 0.048598 | 14.1617 | 0.033645 | 0
0.0018 | -0.11402 | -0.68318 | -14.2645 | 14.2869 | 0.011215 | -0.033645 | -0.027103 | 0.048598 | 14.2645 | 0.033645 | 0.16735
0.0036 | -0.14393 | -0.6486 | -14.4495 | 14.4701 | 0.01215 | -0.03271 | -0.028037 | 0.048598 | 14.4495 | 0.03271 | 0.3347
0.0054 | -0.17103 | -0.60467 | -14.5916 | 14.6093 | 0.014019 | -0.031776 | -0.028037 | 0.048598 | 14.5916 | 0.031776 | 0.50205
0.0072 | -0.18692 | -0.53364 | -14.5991 | 14.614 | 0.015888 | -0.03271 | -0.028037 | 0.050467 | 14.5991 | 0.03271 | 0.66939
0.009 | -0.16636 | -0.53084 | -14.4393 | 14.4542 | 0.015888 | -0.03271 | -0.028037 | 0.050467 | 14.4393 | 0.03271 | 0.83674
0.0108 | -0.15421 | -0.53271 | -14.3299 | 14.3439 | 0.014953 | -0.03271 | -0.028972 | 0.050467 | 14.3299 | 0.03271 | 1.0041
0.0125 | -0.18224 | -0.47383 | -14.4383 | 14.4514 | 0.015888 | -0.031776 | -0.028972 | 0.050467 | 14.4383 | 0.031776 | 1.1621
0.0143 | -0.19907 | -0.40841 | -14.4701 | 14.4813 | 0.017757 | -0.031776 | -0.029907 | 0.051402 | 14.4701 | 0.031776 | 1.3295
0.0161 | -0.1972 | -0.34019 | -14.371 | 14.3813 | 0.019626 | -0.031776 | -0.029907 | 0.052336 | 14.371 | 0.031776 | 1.4968
0.0179 | -0.17944 | -0.28785 | -14.1804 | 14.1888 | 0.020561 | -0.03271 | -0.029907 | 0.054206 | 14.1804 | 0.03271 | 1.6642
0.0197 | -0.14673 | -0.25981 | -13.9364 | 13.9439 | 0.020561 | -0.03271 | -0.030841 | 0.054206 | 13.9364 | 0.03271 | 1.8315
0.0215 | -0.1 | -0.2215 | -13.5944 | 13.6009 | 0.020561 | -0.034579 | -0.030841 | 0.056075 | 13.5944 | 0.034579 | 1.9989
0.0233 | -0.040187 | -0.23832 | -13.2449 | 13.2514 | 0.017757 | -0.036449 | -0.031776 | 0.057009 | 13.2449 | 0.036449 | 2.1662
0.0251 | 0.017757 | -0.28785 | -12.9561 | 12.9636 | 0.014019 | -0.038318 | -0.031776 | 0.057944 | 12.9561 | 0.038318 | 2.3336
0.0269 | 0.062617 | -0.36262 | -12.7682 | 12.7785 | 0.0093458 | -0.039252 | -0.03271 | 0.057944 | 12.7682 | 0.039252 | 2.5009
0.0287 | 0.08785 | -0.40748 | -12.7234 | 12.7355 | 0.0056075 | -0.039252 | -0.03271 | 0.057009 | 12.7234 | 0.039252 | 2.6683
0.0305 | 0.10935 | -0.44019 | -12.8019 | 12.815 | 0.0018692 | -0.040187 | -0.033645 | 0.057944 | 12.8019 | 0.040187 | 2.8356
0.0323 | 0.11402 | -0.4243 | -12.9673 | 12.9813 | 0.00093458 | -0.041121 | -0.034579 | 0.057944 | 12.9673 | 0.041121 | 3.003
0.0341 | 0.10935 | -0.3785 | -13.2355 | 13.2486 | 0.00093458 | -0.041121 | -0.034579 | 0.059813 | 13.2355 | 0.041121 | 3.1703
0.0359 | 0.07757 | -0.2757 | -13.5579 | 13.5682 | 0.0028037 | -0.042056 | -0.035514 | 0.060748 | 13.5579 | 0.042056 | 3.3377
0.0376 | 0.042991 | -0.16355 | -13.8374 | 13.8449 | 0.0056075 | -0.042056 | -0.035514 | 0.062617 | 13.8374 | 0.042056 | 3.4957
0.0394 | 0.021495 | -0.017757 | -13.9813 | 13.9888 | 0.0084112 | -0.043925 | -0.036449 | 0.064486 | 13.9813 | 0.043925 | 3.6631
0.0412 | 0.0056075 | 0.13084 | -14.072 | 14.0794 | 0.011215 | -0.04486 | -0.036449 | 0.066355 | 14.072 | 0.04486 | 3.8304
0.043 | 0.0037383 | 0.25981 | -14.0243 | 14.0346 | 0.013084 | -0.046729 | -0.036449 | 0.068224 | 14.0243 | 0.046729 | 3.9978
0.0448 | 0.011215 | 0.35607 | -13.8907 | 13.9028 | 0.014019 | -0.049533 | -0.036449 | 0.070093 | 13.8907 | 0.049533 | 4.1651
0.0466 | 0.025234 | 0.45514 | -13.6822 | 13.6981 | 0.014953 | -0.052336 | -0.036449 | 0.071963 | 13.6822 | 0.052336 | 4.3325
0.0484 | 0.048598 | 0.51215 | -13.4019 | 13.4215 | 0.014019 | -0.054206 | -0.036449 | 0.072897 | 13.4019 | 0.054206 | 4.4998
0.0502 | 0.088785 | 0.54299 | -13.0467 | 13.071 | 0.011215 | -0.056075 | -0.035514 | 0.073832 | 13.0467 | 0.056075 | 4.6672
0.052 | 0.12897 | 0.55421 | -12.7393 | 12.7664 | 0.0084112 | -0.058879 | -0.035514 | 0.074766 | 12.7393 | 0.058879 | 4.8345
0.0538 | 0.16636 | 0.51776 | -12.5196 | 12.5467 | 0.0037383 | -0.060748 | -0.036449 | 0.076636 | 12.5196 | 0.060748 | 5.0019
0.0556 | 0.20093 | 0.47196 | -12.3701 | 12.3972 | -0.00093458 | -0.063551 | -0.036449 | 0.07757 | 12.3701 | 0.063551 | 5.1692
0.0574 | 0.22243 | 0.43551 | -12.3832 | 12.4112 | -0.0056075 | -0.065421 | -0.037383 | 0.080374 | 12.3832 | 0.065421 | 5.3366
0.0592 | 0.23084 | 0.41776 | -12.4645 | 12.4935 | -0.0093458 | -0.066355 | -0.037383 | 0.081308 | 12.4645 | 0.066355 | 5.5039
0.061 | 0.23458 | 0.43832 | -12.6318 | 12.6626 | -0.01215 | -0.068224 | -0.038318 | 0.084112 | 12.6318 | 0.068224 | 5.6713
0.0627 | 0.21963 | 0.52617 | -12.8822 | 12.9168 | -0.01215 | -0.069159 | -0.039252 | 0.085981 | 12.8822 | 0.069159 | 5.8293
0.0645 | 0.20187 | 0.66916 | -13.1355 | 13.1813 | -0.01028 | -0.071028 | -0.040187 | 0.088785 | 13.1355 | 0.071028 | 5.9967
0.0663 | 0.18411 | 0.81308 | -13.2636 | 13.3187 | -0.0093458 | -0.073832 | -0.041121 | 0.090654 | 13.2636 | 0.073832 | 6.164
0.0681 | 0.17009 | 0.94673 | -13.3234 | 13.385 | -0.0074766 | -0.075701 | -0.041121 | 0.093458 | 13.3234 | 0.075701 | 6.3313
0.0699 | 0.1729 | 1.0738 | -13.2364 | 13.3075 | -0.0056075 | -0.078505 | -0.042056 | 0.096262 | 13.2364 | 0.078505 | 6.4987
0.0717 | 0.18318 | 1.1925 | -13.1047 | 13.186 | -0.0046729 | -0.081308 | -0.042991 | 0.099065 | 13.1047 | 0.081308 | 6.666
0.0735 | 0.20374 | 1.2682 | -12.8935 | 12.9832 | -0.0056075 | -0.085047 | -0.042991 | 0.10187 | 12.8935 | 0.085047 | 6.8334
0.0753 | 0.23738 | 1.3327 | -12.6065 | 12.7019 | -0.0056075 | -0.08785 | -0.043925 | 0.10467 | 12.6065 | 0.08785 | 7.0007
0.0771 | 0.27757 | 1.3262 | -12.2804 | 12.3766 | -0.0093458 | -0.090654 | -0.04486 | 0.10654 | 12.2804 | 0.090654 | 7.1681
0.0789 | 0.31682 | 1.3121 | -11.971 | 12.0692 | -0.01215 | -0.093458 | -0.045794 | 0.10748 | 11.971 | 0.093458 | 7.3354
0.0807 | 0.33738 | 1.2804 | -11.872 | 11.9682 | -0.017757 | -0.096262 | -0.045794 | 0.11028 | 11.872 | 0.096262 | 7.5028
0.0825 | 0.34673 | 1.2393 | -11.8598 | 11.9533 | -0.02243 | -0.098131 | -0.046729 | 0.11308 | 11.8598 | 0.098131 | 7.6701
0.0843 | 0.35327 | 1.2533 | -11.9589 | 12.0495 | -0.026168 | -0.10187 | -0.047664 | 0.11682 | 11.9589 | 0.10187 | 7.8375
0.086 | 0.35234 | 1.2495 | -12.1822 | 12.272 | -0.028972 | -0.10374 | -0.048598 | 0.11963 | 12.1822 | 0.10374 | 7.9955
0.0878 | 0.33925 | 1.2944 | -12.4972 | 12.5907 | -0.030841 | -0.10467 | -0.049533 | 0.12243 | 12.4972 | 0.10467 | 8.1629
0.0896 | 0.31869 | 1.3991 | -12.8421 | 12.9449 | -0.028972 | -0.10654 | -0.050467 | 0.1243 | 12.8421 | 0.10654 | 8.3302
0.0914 | 0.3028 | 1.5215 | -13.114 | 13.2299 | -0.027103 | -0.10841 | -0.051402 | 0.12523 | 13.114 | 0.10841 | 8.4976
0.0932 | 0.30467 | 1.6523 | -13.1897 | 13.3224 | -0.026168 | -0.11215 | -0.052336 | 0.12897 | 13.1897 | 0.11215 | 8.6649
0.095 | 0.30748 | 1.7794 | -13.1757 | 13.3252 | -0.025234 | -0.11589 | -0.053271 | 0.13271 | 13.1757 | 0.11589 | 8.8323
0.0968 | 0.31682 | 1.8869 | -13.0981 | 13.2645 | -0.025234 | -0.11963 | -0.05514 | 0.13645 | 13.0981 | 0.11963 | 8.9996
0.0986 | 0.32804 | 1.9617 | -12.9477 | 13.1271 | -0.026168 | -0.12336 | -0.056075 | 0.14019 | 12.9477 | 0.12336 | 9.167
0.1004 | 0.34579 | 2.0009 | -12.7318 | 12.9187 | -0.026168 | -0.12617 | -0.057009 | 0.14206 | 12.7318 | 0.12617 | 9.3343
0.1022 | 0.35981 | 1.9841 | -12.4991 | 12.6869 | -0.028972 | -0.1271 | -0.057944 | 0.14486 | 12.4991 | 0.1271 | 9.5017
0.104 | 0.36636 | 1.9505 | -12.3935 | 12.5776 | -0.03271 | -0.12897 | -0.058879 | 0.14766 | 12.3935 | 0.12897 | 9.669
0.1058 | 0.37103 | 1.8953 | -12.3561 | 12.5336 | -0.038318 | -0.13084 | -0.059813 | 0.1514 | 12.3561 | 0.13084 | 9.8364
0.1076 | 0.36636 | 1.8645 | -12.4533 | 12.6262 | -0.042056 | -0.13271 | -0.061682 | 0.15421 | 12.4533 | 0.13271 | 10.0037
0.1094 | 0.34953 | 1.8393 | -12.6729 | 12.8402 | -0.045794 | -0.13364 | -0.062617 | 0.15701 | 12.6729 | 0.13364 | 10.1711
0.1111 | 0.33084 | 1.8449 | -12.986 | 13.1551 | -0.048598 | -0.13551 | -0.064486 | 0.15981 | 12.986 | 0.13551 | 10.3291
0.1129 | 0.30841 | 1.8617 | -13.3916 | 13.5598 | -0.050467 | -0.13738 | -0.065421 | 0.16262 | 13.3916 | 0.13738 | 10.4965
0.1147 | 0.28037 | 1.9 | -13.8224 | 13.9953 | -0.049533 | -0.13925 | -0.066355 | 0.16449 | 13.8224 | 0.13925 | 10.6638
0.1165 | 0.25888 | 1.9925 | -14.1159 | 14.2972 | -0.048598 | -0.14206 | -0.068224 | 0.16729 | 14.1159 | 0.14206 | 10.8312
0.1183 | 0.24299 | 2.1075 | -14.2888 | 14.4841 | -0.046729 | -0.14486 | -0.069159 | 0.17009 | 14.2888 | 0.14486 | 10.9985
0.1201 | 0.23925 | 2.2299 | -14.3533 | 14.5664 | -0.04486 | -0.1486 | -0.071028 | 0.1729 | 14.3533 | 0.1486 | 11.1659
0.1219 | 0.24019 | 2.3318 | -14.329 | 14.557 | -0.043925 | -0.15234 | -0.072897 | 0.17664 | 14.329 | 0.15234 | 11.3332
0.1237 | 0.24206 | 2.4037 | -14.2355 | 14.4776 | -0.043925 | -0.15514 | -0.074766 | 0.17944 | 14.2355 | 0.15514 | 11.5006
0.1255 | 0.24579 | 2.4299 | -14.0514 | 14.3019 | -0.045794 | -0.15794 | -0.076636 | 0.18318 | 14.0514 | 0.15794 | 11.6679
0.1273 | 0.24019 | 2.4168 | -13.871 | 14.1224 | -0.047664 | -0.15981 | -0.07757 | 0.18598 | 13.871 | 0.15981 | 11.8353
0.1291 | 0.22523 | 2.3794 | -13.7495 | 13.9963 | -0.051402 | -0.16075 | -0.079439 | 0.18785 | 13.7495 | 0.16075 | 12.0026
0.1309 | 0.19907 | 2.3131 | -13.757 | 13.9888 | -0.056075 | -0.16075 | -0.080374 | 0.19065 | 13.757 | 0.16075 | 12.17
0.1327 | 0.16542 | 2.271 | -13.8495 | 14.0701 | -0.059813 | -0.16168 | -0.082243 | 0.19252 | 13.8495 | 0.16168 | 12.3373
0.1345 | 0.1243 | 2.2607 | -14.0318 | 14.2449 | -0.061682 | -0.16262 | -0.083178 | 0.19439 | 14.0318 | 0.16262 | 12.5046
0.1362 | 0.080374 | 2.2794 | -14.3 | 14.5103 | -0.063551 | -0.16355 | -0.085047 | 0.1972 | 14.3 | 0.16355 | 12.6627
0.138 | 0.036449 | 2.3542 | -14.5178 | 14.7374 | -0.063551 | -0.16636 | -0.086916 | 0.2 | 14.5178 | 0.16636 | 12.83
0.1398 | -0.0056075 | 2.4411 | -14.6944 | 14.9262 | -0.062617 | -0.16916 | -0.08785 | 0.20187 | 14.6944 | 0.16916 | 12.9974
0.1416 | -0.035514 | 2.5393 | -14.7776 | 15.0271 | -0.061682 | -0.17196 | -0.088785 | 0.20467 | 14.7776 | 0.17196 | 13.1647
0.1434 | -0.049533 | 2.6299 | -14.7467 | 15.014 | -0.059813 | -0.1757 | -0.090654 | 0.20748 | 14.7467 | 0.1757 | 13.3321
0.1452 | -0.059813 | 2.7047 | -14.6439 | 14.929 | -0.058879 | -0.17944 | -0.091589 | 0.21121 | 14.6439 | 0.17944 | 13.4994
0.147 | -0.072897 | 2.7551 | -14.4682 | 14.7673 | -0.058879 | -0.18131 | -0.093458 | 0.21308 | 14.4682 | 0.18131 | 13.6668
0.1488 | -0.06729 | 2.7822 | -14.2037 | 14.5112 | -0.058879 | -0.18411 | -0.094393 | 0.21682 | 14.2037 | 0.18411 | 13.8341
0.1506 | -0.051402 | 2.7617 | -13.8879 | 14.1953 | -0.059813 | -0.18598 | -0.095327 | 0.21869 | 13.8879 | 0.18598 | 14.0015
0.1524 | -0.066355 | 2.686 | -13.6804 | 13.9785 | -0.064486 | -0.18692 | -0.096262 | 0.22056 | 13.6804 | 0.18692 | 14.1688
0.1542 | -0.10748 | 2.6224 | -13.5645 | 13.8561 | -0.068224 | -0.18598 | -0.097196 | 0.2215 | 13.5645 | 0.18598 | 14.3362
0.156 | -0.15701 | 2.5364 | -13.6009 | 13.8813 | -0.072897 | -0.18505 | -0.098131 | 0.22243 | 13.6009 | 0.18505 | 14.5035
0.1578 | -0.20654 | 2.4523 | -13.7075 | 13.9757 | -0.076636 | -0.18411 | -0.099065 | 0.22336 | 13.7075 | 0.18411 | 14.6709
0.1595 | -0.25327 | 2.3897 | -13.9234 | 14.1813 | -0.080374 | -0.18318 | -0.1 | 0.2243 | 13.9234 | 0.18318 | 14.8289
0.1613 | -0.29252 | 2.3514 | -14.2056 | 14.4561 | -0.083178 | -0.18411 | -0.10093 | 0.22617 | 14.2056 | 0.18411 | 14.9963
0.1631 | -0.31682 | 2.3636 | -14.5533 | 14.8047 | -0.085047 | -0.18598 | -0.10093 | 0.22897 | 14.5533 | 0.18598 | 15.1636
0.1649 | -0.35607 | 2.414 | -14.929 | 15.186 | -0.084112 | -0.18598 | -0.10187 | 0.22897 | 14.929 | 0.18598 | 15.331
0.1667 | -0.38411 | 2.4869 | -15.2364 | 15.5 | -0.083178 | -0.18692 | -0.1028 | 0.22991 | 15.2364 | 0.18692 | 15.4983
0.1685 | -0.40374 | 2.572 | -15.5019 | 15.7748 | -0.081308 | -0.18879 | -0.10374 | 0.23178 | 15.5019 | 0.18879 | 15.6657
0.1703 | -0.40748 | 2.6832 | -15.7252 | 16.015 | -0.079439 | -0.19159 | -0.10467 | 0.23271 | 15.7252 | 0.19159 | 15.833
0.1721 | -0.41215 | 2.7879 | -15.8533 | 16.1561 | -0.076636 | -0.19439 | -0.10561 | 0.23458 | 15.8533 | 0.19439 | 16.0004
0.1739 | -0.40467 | 2.9 | -15.8888 | 16.2121 | -0.073832 | -0.19626 | -0.10561 | 0.23645 | 15.8888 | 0.19626 | 16.1677
0.1757 | -0.39159 | 2.9953 | -15.8645 | 16.2047 | -0.071963 | -0.19907 | -0.10654 | 0.23738 | 15.8645 | 0.19907 | 16.3351
0.1775 | -0.36822 | 3.0832 | -15.7673 | 16.1252 | -0.070093 | -0.20093 | -0.10748 | 0.23925 | 15.7673 | 0.20093 | 16.5024
0.1793 | -0.34393 | 3.157 | -15.5953 | 15.971 | -0.068224 | -0.2028 | -0.10748 | 0.24112 | 15.5953 | 0.2028 | 16.6698
0.1811 | -0.31869 | 3.1916 | -15.3654 | 15.7523 | -0.06729 | -0.20467 | -0.10841 | 0.24206 | 15.3654 | 0.20467 | 16.8371
0.1829 | -0.28972 | 3.2037 | -15.1037 | 15.4953 | -0.06729 | -0.20561 | -0.10841 | 0.24299 | 15.1037 | 0.20561 | 17.0045
0.1846 | -0.26449 | 3.1841 | -14.8514 | 15.243 | -0.06729 | -0.20561 | -0.10841 | 0.24299 | 14.8514 | 0.20561 | 17.1625
0.1864 | -0.25701 | 3.1449 | -14.7318 | 15.1234 | -0.069159 | -0.20374 | -0.10841 | 0.24206 | 14.7318 | 0.20374 | 17.3299
0.1882 | -0.26355 | 3.0514 | -14.672 | 15.0505 | -0.072897 | -0.20093 | -0.10841 | 0.24112 | 14.672 | 0.20093 | 17.4972
0.19 | -0.28411 | 2.9523 | -14.7336 | 15.0953 | -0.07757 | -0.19813 | -0.10841 | 0.24019 | 14.7336 | 0.19813 | 17.6646
0.1918 | -0.31121 | 2.8551 | -14.872 | 15.2121 | -0.081308 | -0.19626 | -0.10841 | 0.23925 | 14.872 | 0.19626 | 17.8319
0.1936 | -0.33645 | 2.7757 | -15.128 | 15.4533 | -0.085047 | -0.19346 | -0.10841 | 0.23832 | 15.128 | 0.19346 | 17.9993
0.1954 | -0.35607 | 2.7542 | -15.4318 | 15.7458 | -0.08785 | -0.19159 | -0.10841 | 0.23832 | 15.4318 | 0.19159 | 18.1666
0.1972 | -0.35514 | 2.8131 | -15.843 | 16.1598 | -0.086916 | -0.19252 | -0.10841 | 0.23925 | 15.843 | 0.19252 | 18.334
0.199 | -0.34393 | 2.8972 | -16.229 | 16.5533 | -0.085981 | -0.19439 | -0.10748 | 0.23925 | 16.229 | 0.19439 | 18.5013
0.2008 | -0.32804 | 2.9738 | -16.4785 | 16.814 | -0.084112 | -0.19626 | -0.10748 | 0.24019 | 16.4785 | 0.19626 | 18.6687
0.2026 | -0.32523 | 3.0009 | -16.7159 | 17.0589 | -0.083178 | -0.19533 | -0.10654 | 0.23832 | 16.7159 | 0.19533 | 18.836
0.2044 | -0.31963 | 3.0402 | -16.8907 | 17.243 | -0.081308 | -0.19533 | -0.10561 | 0.23738 | 16.8907 | 0.19533 | 19.0033
0.2062 | -0.3 | 3.0963 | -16.9841 | 17.3458 | -0.078505 | -0.19626 | -0.10467 | 0.23645 | 16.9841 | 0.19626 | 19.1707
0.208 | -0.28692 | 3.0916 | -17.0346 | 17.3935 | -0.076636 | -0.19346 | -0.10374 | 0.23364 | 17.0346 | 0.19346 | 19.338
0.2097 | -0.27477 | 3.0804 | -17.0822 | 17.4364 | -0.074766 | -0.19159 | -0.10187 | 0.23084 | 17.0822 | 0.19159 | 19.4961
0.2115 | -0.30748 | 3.0738 | -17.1551 | 17.5056 | -0.071963 | -0.18598 | -0.10093 | 0.2243 | 17.1551 | 0.18598 | 19.6634
0.2133 | -0.34953 | 3.0187 | -17.3897 | 17.7262 | -0.070093 | -0.17944 | -0.099065 | 0.21776 | 17.3897 | 0.17944 | 19.8308
0.2151 | -0.39626 | 2.914 | -17.7411 | 18.057 | -0.070093 | -0.17009 | -0.097196 | 0.20935 | 17.7411 | 0.17009 | 19.9981
0.2169 | -0.45421 | 2.8 | -18.2009 | 18.4916 | -0.069159 | -0.15981 | -0.096262 | 0.20093 | 18.2009 | 0.15981 | 20.1655
0.2187 | -0.51028 | 2.6935 | -18.8 | 19.0645 | -0.069159 | -0.15047 | -0.094393 | 0.19252 | 18.8 | 0.15047 | 20.3328
0.2205 | -0.58598 | 2.5776 | -19.5299 | 19.7766 | -0.06729 | -0.14019 | -0.093458 | 0.18318 | 19.5299 | 0.14019 | 20.5002
0.2223 | -0.67383 | 2.4514 | -20.315 | 20.5449 | -0.065421 | -0.12804 | -0.093458 | 0.17383 | 20.315 | 0.12804 | 20.6675
0.2241 | -0.76542 | 2.3065 | -21.2505 | 21.4626 | -0.062617 | -0.11402 | -0.092523 | 0.16355 | 21.2505 | 0.11402 | 20.8349
0.2259 | -0.87477 | 2.1495 | -22.3449 | 22.5374 | -0.059813 | -0.098131 | -0.091589 | 0.15234 | 22.3449 | 0.098131 | 21.0022
0.2277 | -1.0037 | 1.9841 | -23.6075 | 23.7832 | -0.054206 | -0.080374 | -0.091589 | 0.14019 | 23.6075 | 0.080374 | 21.1696
0.2295 | -1.1486 | 1.8103 | -25.0411 | 25.2047 | -0.048598 | -0.060748 | -0.091589 | 0.12991 | 25.0411 | 0.060748 | 21.3369
0.2313 | -1.2963 | 1.6206 | -26.6336 | 26.7869 | -0.042056 | -0.040187 | -0.093458 | 0.12336 | 26.6336 | 0.040187 | 21.5043
0.2331 | -1.4551 | 1.4028 | -28.3607 | 28.5093 | -0.034579 | -0.016822 | -0.094393 | 0.1215 | 28.3607 | 0.016822 | 21.6716
0.2348 | -1.6533 | 1.1449 | -30.3271 | 30.4701 | -0.026168 | 0.0093458 | -0.097196 | 0.1243 | 30.3271 | -0.0093458 | 21.8297
0.2366 | -1.8673 | 0.86542 | -32.4963 | 32.6308 | -0.017757 | 0.037383 | -0.1 | 0.12991 | 32.4963 | -0.037383 | 21.997
0.2384 | -2.0869 | 0.57103 | -34.8262 | 34.957 | -0.0084112 | 0.066355 | -0.1028 | 0.14112 | 34.8262 | -0.066355 | 22.1644
0.2402 | -2.3206 | 0.24953 | -37.4131 | 37.5458 | 0.0018692 | 0.099065 | -0.10748 | 0.15981 | 37.4131 | -0.099065 | 22.3317
0.242 | -2.5607 | -0.085981 | -40.1738 | 40.3103 | 0.01215 | 0.13271 | -0.11308 | 0.18411 | 40.1738 | -0.13271 | 22.4991
0.2438 | -2.8009 | -0.45047 | -43.2103 | 43.3533 | 0.024299 | 0.17009 | -0.11869 | 0.21589 | 43.2103 | -0.17009 | 22.6664
0.2456 | -3.0421 | -0.83925 | -46.4794 | 46.6336 | 0.037383 | 0.20841 | -0.12617 | 0.2514 | 46.4794 | -0.20841 | 22.8338
0.2474 | -3.2766 | -1.2467 | -50.0393 | 50.2075 | 0.051402 | 0.24953 | -0.13458 | 0.29159 | 50.0393 | -0.24953 | 23.0011
0.2492 | -3.5047 | -1.6607 | -53.8626 | 54.0439 | 0.066355 | 0.29159 | -0.14393 | 0.33551 | 53.8626 | -0.29159 | 23.1685
0.251 | -3.7308 | -2.0832 | -58.0037 | 58.2009 | 0.083178 | 0.33738 | -0.15421 | 0.38224 | 58.0037 | -0.33738 | 23.3358
0.2528 | -3.9533 | -2.5206 | -62.3841 | 62.5963 | 0.10093 | 0.38411 | -0.16542 | 0.43271 | 62.3841 | -0.38411 | 23.5032
0.2546 | -4.1551 | -2.9692 | -67.3234 | 67.5514 | 0.12056 | 0.43364 | -0.1785 | 0.48692 | 67.3234 | -0.43364 | 23.6705
0.2564 | -4.3243 | -3.3944 | -72.2654 | 72.5075 | 0.14112 | 0.48131 | -0.19159 | 0.53925 | 72.2654 | -0.48131 | 23.8379
0.2581 | -4.4813 | -3.8692 | -77.372 | 77.6318 | 0.16075 | 0.53271 | -0.20654 | 0.59439 | 77.372 | -0.53271 | 23.9959
0.2599 | -4.5953 | -4.3505 | -82.4533 | 82.7299 | 0.18037 | 0.58224 | -0.22056 | 0.64953 | 82.4533 | -0.58224 | 24.1633
0.2617 | -4.657 | -4.8701 | -87.7103 | 88.0019 | 0.2 | 0.63271 | -0.23551 | 0.70561 | 87.7103 | -0.63271 | 24.3306
0.2635 | -4.7019 | -5.3944 | -92.7692 | 93.0766 | 0.21869 | 0.68131 | -0.25047 | 0.75888 | 92.7692 | -0.68131 | 24.498
0.2653 | -4.7112 | -5.9206 | -97.8411 | 98.1636 | 0.23738 | 0.72897 | -0.26542 | 0.81215 | 97.8411 | -0.72897 | 24.6653
0.2671 | -4.7019 | -6.3738 | -103.0047 | 103.3374 | 0.25794 | 0.77757 | -0.28037 | 0.86729 | 103.0047 | -0.77757 | 24.8327
0.2689 | -4.672 | -6.8364 | -108.0738 | 108.4178 | 0.27757 | 0.8215 | -0.29533 | 0.91682 | 108.0738 | -0.8215 | 25
0.2707 | -4.6075 | -7.2467 | -112.7907 | 113.1439 | 0.29533 | 0.85794 | -0.30841 | 0.95888 | 112.7907 | -0.85794 | 25.1673
0.2725 | -4.5234 | -7.628 | -117.2748 | 117.6346 | 0.31215 | 0.88879 | -0.3215 | 0.99626 | 117.2748 | -0.88879 | 25.3347
0.2743 | -4.3981 | -7.985 | -121.7346 | 122.1 | 0.32804 | 0.91495 | -0.33364 | 1.028 | 121.7346 | -0.91495 | 25.502
0.2761 | -4.2776 | -8.3103 | -126.0477 | 126.4168 | 0.34393 | 0.94019 | -0.34579 | 1.0598 | 126.0477 | -0.94019 | 25.6694
0.2779 | -4.1542 | -8.6112 | -130.1439 | 130.5168 | 0.35888 | 0.96636 | -0.35701 | 1.0916 | 130.1439 | -0.96636 | 25.8367
0.2797 | -4.0196 | -8.8729 | -133.9243 | 134.3 | 0.37383 | 0.98972 | -0.36729 | 1.1206 | 133.9243 | -0.98972 | 26.0041
0.2815 | -3.8636 | -9.0972 | -137.4168 | 137.7935 | 0.38598 | 1.0075 | -0.3757 | 1.143 | 137.4168 | -1.0075 | 26.1714
0.2832 | -3.6916 | -9.2944 | -140.6766 | 141.0542 | 0.3972 | 1.0215 | -0.38411 | 1.1626 | 140.6766 | -1.0215 | 26.3295
0.285 | -3.5168 | -9.4589 | -143.6869 | 144.0645 | 0.40748 | 1.0318 | -0.39159 | 1.1776 | 143.6869 | -1.0318 | 26.4968
0.2868 | -3.3402 | -9.5664 | -146.2841 | 146.6589 | 0.41589 | 1.0364 | -0.39813 | 1.1869 | 146.2841 | -1.0364 | 26.6642
0.2886 | -3.1626 | -9.6589 | -148.6813 | 149.0542 | 0.42336 | 1.0402 | -0.40374 | 1.1944 | 148.6813 | -1.0402 | 26.8315
0.2904 | -2.986 | -9.7449 | -150.8738 | 151.2439 | 0.42991 | 1.0421 | -0.40935 | 1.2009 | 150.8738 | -1.0421 | 26.9989
0.2922 | -2.8178 | -9.7822 | -152.6271 | 152.9925 | 0.43551 | 1.0421 | -0.41215 | 1.2047 | 152.6271 | -1.0421 | 27.1662
0.294 | -2.6533 | -9.7542 | -154.1729 | 154.5318 | 0.44019 | 1.0364 | -0.41589 | 1.2019 | 154.1729 | -1.0364 | 27.3336
0.2958 | -2.4897 | -9.7168 | -155.6075 | 155.9598 | 0.44486 | 1.0318 | -0.41776 | 1.2009 | 155.6075 | -1.0318 | 27.5009
0.2976 | -2.3262 | -9.6093 | -156.9075 | 157.2477 | 0.44953 | 1.0234 | -0.41963 | 1.1963 | 156.9075 | -1.0234 | 27.6683
0.2994 | -2.1729 | -9.4692 | -158.1654 | 158.4925 | 0.45421 | 1.015 | -0.42056 | 1.1916 | 158.1654 | -1.015 | 27.8356
0.3012 | -1.9869 | -9.2664 | -159.2785 | 159.5907 | 0.45701 | 0.99626 | -0.4215 | 1.1785 | 159.2785 | -0.99626 | 28.003
0.303 | -1.8093 | -9.0514 | -160.2196 | 160.5178 | 0.45888 | 0.9785 | -0.42243 | 1.1645 | 160.2196 | -0.9785 | 28.1703
0.3048 | -1.6402 | -8.8318 | -160.9617 | 161.2477 | 0.45888 | 0.95888 | -0.42243 | 1.1486 | 160.9617 | -0.95888 | 28.3377
0.3066 | -1.4748 | -8.5981 | -161.5421 | 161.8159 | 0.45794 | 0.93832 | -0.4215 | 1.1318 | 161.5421 | -0.93832 | 28.505
0.3083 | -1.3075 | -8.3075 | -162.0037 | 162.2664 | 0.45607 | 0.91215 | -0.42056 | 1.1103 | 162.0037 | -0.91215 | 28.6631
0.3101 | -1.1514 | -8.0327 | -162.3243 | 162.5729 | 0.45514 | 0.89065 | -0.41869 | 1.0925 | 162.3243 | -0.89065 | 28.8304
0.3119 | -1.0028 | -7.7364 | -162.6065 | 162.8449 | 0.4514 | 0.85701 | -0.41776 | 1.0664 | 162.6065 | -0.85701 | 28.9978
0.3137 | -0.86636 | -7.4495 | -162.785 | 163.0131 | 0.44953 | 0.83271 | -0.41589 | 1.0477 | 162.785 | -0.83271 | 29.1651
0.3155 | -0.7215 | -7.1336 | -162.9449 | 163.1636 | 0.44673 | 0.8028 | -0.41402 | 1.0299 | 162.9449 | -0.8028 | 29.3325
0.3173 | -0.58505 | -6.7953 | -163.0991 | 163.3131 | 0.44299 | 0.76916 | -0.41215 | 1.015 | 163.0991 | -0.76916 | 29.4998
0.3191 | -0.45701 | -6.4505 | -163.2206 | 163.4327 | 0.43925 | 0.74019 | -0.41028 | 1.0056 | 163.2206 | -0.74019 | 29.6672
0.3209 | -0.32617 | -6.1308 | -163.3056 | 163.514 | 0.43738 | 0.71589 | -0.40748 | 1.0009 | 163.3056 | -0.71589 | 29.8345
0.3227 | -0.18037 | -5.7972 | -163.3243 | 163.5234 | 0.43458 | 0.69159 | -0.40467 | 0.98879 | 163.3243 | -0.69159 | 30.0019
0.3245 | -0.049533 | -5.4645 | -163.315 | 163.5075 | 0.43271 | 0.66729 | -0.40093 | 0.98131 | 163.315 | -0.66729 | 30.1692
0.3263 | 0.065421 | -5.1299 | -163.2364 | 163.4262 | 0.42897 | 0.64206 | -0.39813 | 0.97103 | 163.2364 | -0.64206 | 30.3366
0.3281 | 0.16262 | -4.8224 | -163.1019 | 163.2897 | 0.42523 | 0.61776 | -0.39533 | 0.95888 | 163.1019 | -0.61776 | 30.5039
0.3299 | 0.2785 | -4.457 | -162.8935 | 163.0729 | 0.42056 | 0.59159 | -0.39159 | 0.94112 | 162.8935 | -0.59159 | 30.6713
0.3316 | 0.37664 | -4.1159 | -162.6467 | 162.8196 | 0.41682 | 0.56636 | -0.38785 | 0.92336 | 162.6467 | -0.56636 | 30.8293
0.3334 | 0.51028 | -3.7598 | -162.2458 | 162.4121 | 0.41215 | 0.54019 | -0.38318 | 0.90654 | 162.2458 | -0.54019 | 30.9967
0.3352 | 0.62056 | -3.3682 | -161.8692 | 162.029 | 0.40654 | 0.51308 | -0.37944 | 0.88692 | 161.8692 | -0.51308 | 31.164
0.337 | 0.73271 | -3.0075 | -161.4692 | 161.6299 | 0.40093 | 0.48598 | -0.37477 | 0.87477 | 161.4692 | -0.48598 | 31.3313
0.3388 | 0.83738 | -2.6411 | -161.0318 | 161.1944 | 0.39533 | 0.45888 | -0.37103 | 0.86262 | 161.0318 | -0.45888 | 31.4987
0.3406 | 0.92897 | -2.3075 | -160.5589 | 160.7252 | 0.39065 | 0.43551 | -0.36636 | 0.85234 | 160.5589 | -0.43551 | 31.666
0.3424 | 1.0075 | -2.0178 | -160.0626 | 160.2336 | 0.38598 | 0.41495 | -0.36168 | 0.84579 | 160.0626 | -0.41495 | 31.8334
0.3442 | 1.1065 | -1.6701 | -159.4561 | 159.6215 | 0.38131 | 0.39252 | -0.35701 | 0.83178 | 159.4561 | -0.39252 | 32.0007
0.346 | 1.2019 | -1.2505 | -158.8991 | 159.0626 | 0.37383 | 0.36542 | -0.35234 | 0.81121 | 158.8991 | -0.36542 | 32.1681
0.3478 | 1.2991 | -0.86916 | -158.3794 | 158.5458 | 0.36636 | 0.34112 | -0.34766 | 0.79533 | 158.3794 | -0.34112 | 32.3354
0.3496 | 1.3738 | -0.57196 | -157.5953 | 157.7607 | 0.36168 | 0.3243 | -0.34206 | 0.78692 | 157.5953 | -0.3243 | 32.5028
0.3514 | 1.4393 | -0.21121 | -156.6645 | 156.828 | 0.35514 | 0.30374 | -0.33645 | 0.7729 | 156.6645 | -0.30374 | 32.6701
0.3532 | 1.5421 | 0.18972 | -155.8981 | 156.0598 | 0.34673 | 0.28131 | -0.33178 | 0.75514 | 155.8981 | -0.28131 | 32.8375
0.355 | 1.6486 | 0.59159 | -155.2617 | 155.4234 | 0.34019 | 0.25794 | -0.32617 | 0.73925 | 155.2617 | -0.25794 | 33.0048
0.3567 | 1.7383 | 0.95327 | -154.7084 | 154.8794 | 0.33178 | 0.23738 | -0.3215 | 0.72804 | 154.7084 | -0.23738 | 33.1629
0.3585 | 1.814 | 1.2991 | -154.2617 | 154.4458 | 0.3243 | 0.21869 | -0.31682 | 0.7215 | 154.2617 | -0.21869 | 33.3302
0.3603 | 1.8944 | 1.6393 | -153.7196 | 153.9121 | 0.31869 | 0.19813 | -0.31215 | 0.71495 | 153.7196 | -0.19813 | 33.4976
0.3621 | 1.9804 | 1.9645 | -152.9785 | 153.1776 | 0.31215 | 0.17664 | -0.30654 | 0.71121 | 152.9785 | -0.17664 | 33.6649
0.3639 | 2.0645 | 2.2822 | -152.2093 | 152.4168 | 0.30561 | 0.15514 | -0.30093 | 0.70841 | 152.2093 | -0.15514 | 33.8323
0.3657 | 2.1505 | 2.5869 | -151.3925 | 151.6084 | 0.29907 | 0.13458 | -0.29533 | 0.70467 | 151.3925 | -0.13458 | 33.9996
0.3675 | 2.286 | 2.914 | -150.2682 | 150.4785 | 0.29159 | 0.11215 | -0.28785 | 0.68879 | 150.2682 | -0.11215 | 34.167
0.3693 | 2.429 | 3.2626 | -149.0505 | 149.2514 | 0.28411 | 0.086916 | -0.28037 | 0.66822 | 149.0505 | -0.086916 | 34.3343
0.3711 | 2.5766 | 3.6215 | -147.8953 | 148.086 | 0.27383 | 0.057944 | -0.27196 | 0.64299 | 147.8953 | -0.057944 | 34.5017
0.3729 | 2.714 | 3.9991 | -146.8178 | 147.0009 | 0.26355 | 0.028037 | -0.26449 | 0.61495 | 146.8178 | -0.028037 | 34.669
0.3747 | 2.8252 | 4.3486 | -145.9579 | 146.1411 | 0.25421 | 0.0028037 | -0.25794 | 0.59533 | 145.9579 | -0.0028037 | 34.8364
0.3765 | 2.8991 | 4.5645 | -145.3084 | 145.5019 | 0.2486 | -0.014019 | -0.25327 | 0.59626 | 145.3084 | 0.014019 | 35.0037
0.3783 | 2.9832 | 4.7308 | -144.6458 | 144.8505 | 0.24206 | -0.029907 | -0.2486 | 0.59907 | 144.6458 | 0.029907 | 35.1711
0.3801 | 3.1056 | 4.9748 | -143.8869 | 144.0963 | 0.23551 | -0.051402 | -0.24299 | 0.58879 | 143.8869 | 0.051402 | 35.3384
0.3818 | 3.2009 | 5.1103 | -143.2832 | 143.4981 | 0.23271 | -0.065421 | -0.23832 | 0.58224 | 143.2832 | 0.065421 | 35.4965
0.3836 | 3.271 | 5.2187 | -142.7047 | 142.9224 | 0.23178 | -0.07757 | -0.23458 | 0.57757 | 142.7047 | 0.07757 | 35.6638
0.3854 | 3.3327 | 5.3636 | -141.9374 | 142.1589 | 0.22617 | -0.092523 | -0.22991 | 0.56822 | 141.9374 | 0.092523 | 35.8312
0.3872 | 3.3953 | 5.5093 | -141.1692 | 141.3953 | 0.2215 | -0.10748 | -0.22617 | 0.55888 | 141.1692 | 0.10748 | 35.9985
0.389 | 3.4654 | 5.586 | -140.2841 | 140.5178 | 0.21776 | -0.11682 | -0.22243 | 0.55794 | 140.2841 | 0.11682 | 36.1659
0.3908 | 3.5393 | 5.6766 | -139.3645 | 139.6037 | 0.21308 | -0.1271 | -0.21869 | 0.55607 | 139.3645 | 0.1271 | 36.3332
0.3926 | 3.614 | 5.8178 | -138.4028 | 138.6467 | 0.20841 | -0.14019 | -0.21402 | 0.54673 | 138.4028 | 0.14019 | 36.5006
0.3944 | 3.6879 | 5.9215 | -137.4206 | 137.6692 | 0.20467 | -0.1514 | -0.20935 | 0.54112 | 137.4206 | 0.1514 | 36.6679
0.3962 | 3.7617 | 6.028 | -136.4065 | 136.6607 | 0.20093 | -0.16262 | -0.20467 | 0.53832 | 136.4065 | 0.16262 | 36.8353
0.398 | 3.8224 | 6.1215 | -135.3393 | 135.6 | 0.19813 | -0.17196 | -0.2 | 0.53458 | 135.3393 | 0.17196 | 37.0026
0.3998 | 3.8673 | 6.129 | -134.1308 | 134.3963 | 0.19346 | -0.17196 | -0.19626 | 0.52897 | 134.1308 | 0.17196 | 37.17
0.4016 | 3.9187 | 6.1308 | -132.8757 | 133.1449 | 0.18879 | -0.16822 | -0.19159 | 0.5215 | 132.8757 | 0.16822 | 37.3373
0.4034 | 3.9879 | 6.1206 | -131.6523 | 131.9262 | 0.18318 | -0.16542 | -0.18972 | 0.51589 | 131.6523 | 0.16542 | 37.5046
0.4051 | 4.0766 | 6.0981 | -130.4617 | 130.7393 | 0.17757 | -0.16355 | -0.18972 | 0.51121 | 130.4617 | 0.16355 | 37.6627
0.4069 | 4.1598 | 6.0748 | -129.2953 | 129.5785 | 0.17196 | -0.16168 | -0.18972 | 0.50748 | 129.2953 | 0.16168 | 37.83
0.4087 | 4.229 | 6.0449 | -128.1916 | 128.4794 | 0.16542 | -0.16262 | -0.19065 | 0.50561 | 128.1916 | 0.16262 | 37.9974
0.4105 | 4.3131 | 6.0121 | -127.014 | 127.3075 | 0.15794 | -0.16168 | -0.19252 | 0.50187 | 127.014 | 0.16168 | 38.1647
0.4123 | 4.3579 | 5.9832 | -125.9131 | 126.2065 | 0.15047 | -0.16449 | -0.19533 | 0.49813 | 125.9131 | 0.16449 | 38.3321
0.4141 | 4.4037 | 5.9879 | -124.8589 | 125.1542 | 0.14206 | -0.17103 | -0.1972 | 0.49346 | 124.8589 | 0.17103 | 38.4994
0.4159 | 4.4505 | 6.0084 | -123.8944 | 124.1897 | 0.13458 | -0.18037 | -0.1972 | 0.48879 | 123.8944 | 0.18037 | 38.6668
0.4177 | 4.4916 | 6.0262 | -122.9888 | 123.2869 | 0.12804 | -0.19065 | -0.19626 | 0.48785 | 122.9888 | 0.19065 | 38.8341
0.4195 | 4.5383 | 6.0047 | -122.1196 | 122.4196 | 0.1215 | -0.1972 | -0.19626 | 0.48879 | 122.1196 | 0.1972 | 39.0015
0.4213 | 4.5935 | 5.9336 | -121.2579 | 121.5589 | 0.11402 | -0.19907 | -0.19813 | 0.48972 | 121.2579 | 0.19907 | 39.1688
0.4231 | 4.629 | 5.8393 | -120.4037 | 120.7056 | 0.10841 | -0.20374 | -0.19907 | 0.49346 | 120.4037 | 0.20374 | 39.3362
0.4249 | 4.7047 | 5.7935 | -119.5243 | 119.8308 | 0.10187 | -0.20748 | -0.20187 | 0.49813 | 119.5243 | 0.20748 | 39.5035
0.4267 | 4.7645 | 5.7822 | -118.6271 | 118.9374 | 0.094393 | -0.21308 | -0.20374 | 0.5 | 118.6271 | 0.21308 | 39.6709
0.4285 | 4.7925 | 5.7822 | -117.7393 | 118.0495 | 0.08972 | -0.22056 | -0.20467 | 0.49813 | 117.7393 | 0.22056 | 39.8382
0.4302 | 4.8093 | 5.7804 | -116.9084 | 117.2178 | 0.085047 | -0.22897 | -0.20374 | 0.49439 | 116.9084 | 0.22897 | 39.9963
0.432 | 4.8103 | 5.7561 | -116.1355 | 116.443 | 0.080374 | -0.23364 | -0.20374 | 0.49065 | 116.1355 | 0.23364 | 40.1636
0.4338 | 4.8065 | 5.7243 | -115.3832 | 115.6888 | 0.075701 | -0.23645 | -0.2028 | 0.48505 | 115.3832 | 0.23645 | 40.331
0.4356 | 4.8308 | 5.8047 | -114.6056 | 114.9121 | 0.071963 | -0.25327 | -0.20187 | 0.47196 | 114.6056 | 0.25327 | 40.4983
0.4374 | 4.9187 | 6.0439 | -113.8561 | 114.1692 | 0.063551 | -0.27477 | -0.1972 | 0.44953 | 113.8561 | 0.27477 | 40.6657
0.4392 | 4.9486 | 6.1832 | -113.143 | 113.4617 | 0.056075 | -0.28785 | -0.19346 | 0.43364 | 113.143 | 0.28785 | 40.833
0.441 | 4.9505 | 6.215 | -112.2393 | 112.5626 | 0.052336 | -0.28879 | -0.19346 | 0.42897 | 112.2393 | 0.28879 | 41.0004
0.4428 | 4.9206 | 6.1832 | -111.4131 | 111.7393 | 0.050467 | -0.28879 | -0.19533 | 0.42804 | 111.4131 | 0.28879 | 41.1677
0.4446 | 4.8682 | 6.1383 | -110.6804 | 111.0056 | 0.047664 | -0.28879 | -0.19533 | 0.42617 | 110.6804 | 0.28879 | 41.3351
0.4464 | 4.8215 | 6.0972 | -109.9449 | 110.2692 | 0.04486 | -0.28972 | -0.19533 | 0.4243 | 109.9449 | 0.28972 | 41.5024
0.4482 | 4.7308 | 5.9785 | -109.3383 | 109.6579 | 0.046729 | -0.28879 | -0.19533 | 0.42336 | 109.3383 | 0.28879 | 41.6698
0.45 | 4.6832 | 5.8607 | -108.7589 | 109.0729 | 0.048598 | -0.28785 | -0.19626 | 0.42523 | 108.7589 | 0.28785 | 41.8371
0.4518 | 4.6355 | 5.7065 | -108.343 | 108.6523 | 0.053271 | -0.28879 | -0.19626 | 0.42617 | 108.343 | 0.28879 | 42.0045
0.4536 | 4.5692 | 5.5159 | -107.9794 | 108.2822 | 0.058879 | -0.28692 | -0.19626 | 0.4271 | 107.9794 | 0.28692 | 42.1718
0.4553 | 4.5 | 5.3047 | -107.6841 | 107.9794 | 0.065421 | -0.28318 | -0.1972 | 0.4271 | 107.6841 | 0.28318 | 42.3299
0.4571 | 4.4692 | 5.1682 | -107.3103 | 107.5991 | 0.06729 | -0.28224 | -0.1972 | 0.4271 | 107.3103 | 0.28224 | 42.4972
0.4589 | 4.4093 | 4.9486 | -107.2215 | 107.4972 | 0.073832 | -0.2785 | -0.19813 | 0.42523 | 107.2215 | 0.2785 | 42.6646
0.4607 | 4.3467 | 4.7121 | -107.1561 | 107.4196 | 0.080374 | -0.27477 | -0.19813 | 0.4243 | 107.1561 | 0.27477 | 42.8319
0.4625 | 4.2841 | 4.4794 | -107.0879 | 107.3383 | 0.086916 | -0.27009 | -0.19813 | 0.42243 | 107.0879 | 0.27009 | 42.9993
0.4643 | 4.2056 | 4.286 | -106.9841 | 107.2252 | 0.094393 | -0.26636 | -0.19907 | 0.4215 | 106.9841 | 0.26636 | 43.1666
0.4661 | 4.128 | 4.0925 | -106.8804 | 107.1121 | 0.10187 | -0.26262 | -0.2 | 0.4215 | 106.8804 | 0.26262 | 43.334
0.4679 | 4.1009 | 3.9701 | -106.5355 | 106.7626 | 0.10841 | -0.26449 | -0.20093 | 0.4243 | 106.5355 | 0.26449 | 43.5013
0.4697 | 4.0822 | 3.8589 | -106.1533 | 106.3776 | 0.11402 | -0.26729 | -0.20093 | 0.4271 | 106.1533 | 0.26729 | 43.6687
0.4715 | 4.0729 | 3.7336 | -105.7355 | 105.9561 | 0.11963 | -0.26916 | -0.20187 | 0.43084 | 105.7355 | 0.26916 | 43.836
0.4733 | 4.0645 | 3.6065 | -105.3121 | 105.529 | 0.1243 | -0.27009 | -0.2028 | 0.43364 | 105.3121 | 0.27009 | 44.0033
0.4751 | 3.9935 | 3.4972 | -104.8925 | 105.1047 | 0.12897 | -0.26822 | -0.2028 | 0.43364 | 104.8925 | 0.26822 | 44.1707
0.4769 | 3.9131 | 3.4084 | -104.4327 | 104.6411 | 0.13271 | -0.26449 | -0.20374 | 0.43271 | 104.4327 | 0.26449 | 44.338
0.4786 | 3.8439 | 3.3336 | -103.8804 | 104.0869 | 0.13645 | -0.26168 | -0.20467 | 0.43364 | 103.8804 | 0.26168 | 44.4961
0.4804 | 3.7804 | 3.2673 | -103.2617 | 103.4664 | 0.13925 | -0.25794 | -0.20561 | 0.43271 | 103.2617 | 0.25794 | 44.6634
0.4822 | 3.7299 | 3.2159 | -102.515 | 102.7178 | 0.13925 | -0.25327 | -0.20654 | 0.42991 | 102.515 | 0.25327 | 44.8308
0.484 | 3.7075 | 3.1766 | -101.6346 | 101.8402 | 0.14019 | -0.24953 | -0.20841 | 0.4271 | 101.6346 | 0.24953 | 44.9981
0.4858 | 3.7383 | 3.1336 | -100.671 | 100.8804 | 0.13832 | -0.24766 | -0.21028 | 0.42523 | 100.671 | 0.24766 | 45.1655
0.4876 | 3.7869 | 3.0561 | -99.5318 | 99.7421 | 0.13271 | -0.24673 | -0.21495 | 0.4243 | 99.5318 | 0.24673 | 45.3328
0.4894 | 3.8215 | 2.9822 | -98.5879 | 98.7981 | 0.12804 | -0.24486 | -0.21963 | 0.42523 | 98.5879 | 0.24486 | 45.5002
0.4912 | 3.8112 | 2.9131 | -97.7748 | 97.9822 | 0.12523 | -0.24112 | -0.22243 | 0.42617 | 97.7748 | 0.24112 | 45.6675
0.493 | 3.7579 | 2.8402 | -97.057 | 97.2579 | 0.12243 | -0.23551 | -0.22336 | 0.42523 | 97.057 | 0.23551 | 45.8349
0.4948 | 3.7009 | 2.8252 | -96.4533 | 96.6505 | 0.12243 | -0.23084 | -0.2243 | 0.4215 | 96.4533 | 0.23084 | 46.0022
0.4966 | 3.6402 | 2.7822 | -95.9271 | 96.1178 | 0.12243 | -0.22804 | -0.22336 | 0.41963 | 95.9271 | 0.22804 | 46.1696
0.4984 | 3.6065 | 2.6551 | -95.5738 | 95.7533 | 0.1271 | -0.2243 | -0.22243 | 0.41495 | 95.5738 | 0.2243 | 46.3369
0.5002 | 3.5542 | 2.4514 | -95.2215 | 95.3841 | 0.13364 | -0.22056 | -0.2215 | 0.40935 | 95.2215 | 0.22056 | 46.5043
0.502 | 3.5056 | 2.2832 | -94.7841 | 94.9346 | 0.13925 | -0.21682 | -0.22056 | 0.40561 | 94.7841 | 0.21682 | 46.6716
0.5037 | 3.4355 | 2.1888 | -94.257 | 94.3972 | 0.14206 | -0.21308 | -0.21963 | 0.40093 | 94.257 | 0.21308 | 46.8297
0.5055 | 3.3318 | 2.2336 | -93.5533 | 93.686 | 0.14206 | -0.20841 | -0.21869 | 0.39626 | 93.5533 | 0.20841 | 46.997
0.5073 | 3.2318 | 2.3523 | -92.7822 | 92.9084 | 0.13925 | -0.20748 | -0.21776 | 0.39346 | 92.7822 | 0.20748 | 47.1644
0.5091 | 3.1355 | 2.343 | -92.2495 | 92.3701 | 0.13925 | -0.20374 | -0.21776 | 0.38972 | 92.2495 | 0.20374 | 47.3317
0.5109 | 3.0411 | 2.3187 | -91.815 | 91.9308 | 0.13925 | -0.2 | -0.21776 | 0.38692 | 91.815 | 0.2 | 47.4991
0.5127 | 2.9477 | 2.2869 | -91.4579 | 91.5701 | 0.13925 | -0.19626 | -0.21776 | 0.38411 | 91.4579 | 0.19626 | 47.6664
0.5145 | 2.8607 | 2.229 | -91.1196 | 91.2271 | 0.14019 | -0.19252 | -0.21869 | 0.38131 | 91.1196 | 0.19252 | 47.8338
0.5163 | 2.772 | 2.1664 | -90.815 | 90.9178 | 0.14112 | -0.18972 | -0.21963 | 0.37944 | 90.815 | 0.18972 | 48.0011
0.5181 | 2.6748 | 2.085 | -90.6243 | 90.7215 | 0.14206 | -0.18879 | -0.2215 | 0.38037 | 90.6243 | 0.18879 | 48.1685
0.5199 | 2.5916 | 2.0673 | -90.4822 | 90.5757 | 0.14206 | -0.18972 | -0.22336 | 0.38224 | 90.4822 | 0.18972 | 48.3358
0.5217 | 2.5318 | 1.9626 | -90.3907 | 90.4785 | 0.14579 | -0.19159 | -0.22617 | 0.38598 | 90.3907 | 0.19159 | 48.5032
0.5235 | 2.4785 | 1.8607 | -90.3346 | 90.4187 | 0.14953 | -0.19533 | -0.22991 | 0.39065 | 90.3346 | 0.19533 | 48.6705
0.5253 | 2.4374 | 1.7636 | -90.3336 | 90.415 | 0.15421 | -0.2028 | -0.23271 | 0.39813 | 90.3336 | 0.2028 | 48.8379
0.5271 | 2.3935 | 1.6748 | -90.3664 | 90.4458 | 0.15888 | -0.21028 | -0.23645 | 0.40561 | 90.3664 | 0.21028 | 49.0052
0.5288 | 2.3411 | 1.614 | -90.5 | 90.5766 | 0.16355 | -0.21776 | -0.23925 | 0.41402 | 90.5 | 0.21776 | 49.1633
0.5306 | 2.2841 | 1.5551 | -90.6925 | 90.7664 | 0.16636 | -0.22056 | -0.24299 | 0.41869 | 90.6925 | 0.22056 | 49.3306
0.5324 | 2.2393 | 1.4879 | -90.9355 | 91.0056 | 0.16916 | -0.22243 | -0.24766 | 0.42243 | 90.9355 | 0.22243 | 49.498
0.5342 | 2.2131 | 1.4131 | -91.1551 | 91.2243 | 0.17383 | -0.2271 | -0.2514 | 0.42617 | 91.1551 | 0.2271 | 49.6653
0.536 | 2.1514 | 1.3598 | -91.3056 | 91.372 | 0.17664 | -0.22897 | -0.25514 | 0.42991 | 91.3056 | 0.22897 | 49.8327
0.5378 | 2.1411 | 1.3421 | -91.5729 | 91.6383 | 0.17757 | -0.22991 | -0.25888 | 0.43458 | 91.5729 | 0.22991 | 50
0.5396 | 2.1019 | 1.3009 | -91.8271 | 91.8916 | 0.18037 | -0.23271 | -0.26262 | 0.43925 | 91.8271 | 0.23271 | 50.1673
0.5414 | 2.1 | 1.2879 | -92.1084 | 92.1748 | 0.18037 | -0.23364 | -0.26729 | 0.44393 | 92.1084 | 0.23364 | 50.3347
0.5432 | 2.1028 | 1.2701 | -92.3486 | 92.414 | 0.18131 | -0.23364 | -0.27196 | 0.4486 | 92.3486 | 0.23364 | 50.502
0.545 | 2.086 | 1.2271 | -92.485 | 92.5505 | 0.18411 | -0.23458 | -0.27664 | 0.45327 | 92.485 | 0.23458 | 50.6694
0.5468 | 2.0486 | 1.1813 | -92.6308 | 92.6963 | 0.18972 | -0.23645 | -0.28037 | 0.45701 | 92.6308 | 0.23645 | 50.8367
0.5486 | 2.0159 | 1.1542 | -92.8486 | 92.914 | 0.19346 | -0.23551 | -0.28411 | 0.46075 | 92.8486 | 0.23551 | 51.0041
0.5504 | 1.986 | 1.129 | -92.8393 | 92.9047 | 0.1972 | -0.24112 | -0.28692 | 0.46449 | 92.8393 | 0.24112 | 51.1714
0.5521 | 1.9383 | 1.1402 | -92.8215 | 92.886 | 0.20093 | -0.24579 | -0.28972 | 0.46916 | 92.8215 | 0.24579 | 51.3295
0.5539 | 1.9355 | 1.1308 | -92.9879 | 93.0514 | 0.21121 | -0.24112 | -0.29346 | 0.47477 | 92.9879 | 0.24112 | 51.4968
0.5557 | 1.9636 | 1.1299 | -93.1093 | 93.1729 | 0.22056 | -0.24112 | -0.29626 | 0.48224 | 93.1093 | 0.24112 | 51.6642
0.5575 | 1.9215 | 1.1364 | -93.1804 | 93.243 | 0.2243 | -0.24299 | -0.29907 | 0.48598 | 93.1804 | 0.24299 | 51.8315
0.5593 | 1.8953 | 1.157 | -93.0916 | 93.1523 | 0.2271 | -0.24579 | -0.30093 | 0.48972 | 93.0916 | 0.24579 | 51.9989
0.5611 | 1.8654 | 1.1794 | -92.9991 | 93.0561 | 0.22804 | -0.24953 | -0.30187 | 0.49346 | 92.9991 | 0.24953 | 52.1662
0.5629 | 1.8346 | 1.2103 | -92.9486 | 93.0037 | 0.22897 | -0.2514 | -0.30374 | 0.49626 | 92.9486 | 0.2514 | 52.3336
0.5647 | 1.8037 | 1.2458 | -92.9224 | 92.9729 | 0.22897 | -0.25421 | -0.30467 | 0.49907 | 92.9224 | 0.25421 | 52.5009
0.5665 | 1.7785 | 1.2112 | -92.9187 | 92.9654 | 0.22804 | -0.25327 | -0.30561 | 0.50093 | 92.9187 | 0.25327 | 52.6683
0.5683 | 1.7346 | 1.1355 | -93.0692 | 93.1131 | 0.2271 | -0.2514 | -0.30748 | 0.50187 | 93.0692 | 0.2514 | 52.8356
0.5701 | 1.671 | 1.0944 | -93.385 | 93.4271 | 0.2271 | -0.2514 | -0.31028 | 0.5028 | 93.385 | 0.2514 | 53.003
0.5719 | 1.6178 | 1.0972 | -93.7645 | 93.8037 | 0.22617 | -0.25327 | -0.31215 | 0.50467 | 93.7645 | 0.25327 | 53.1703
0.5737 | 1.6112 | 1.1439 | -94.1159 | 94.1533 | 0.22617 | -0.25794 | -0.31495 | 0.50935 | 94.1159 | 0.25794 | 53.3377
0.5755 | 1.5972 | 1.1019 | -94.343 | 94.3785 | 0.2271 | -0.26075 | -0.31589 | 0.51215 | 94.343 | 0.26075 | 53.505
0.5772 | 1.5673 | 1.0486 | -94.4907 | 94.5243 | 0.22804 | -0.26168 | -0.31776 | 0.51402 | 94.4907 | 0.26168 | 53.6631
0.579 | 1.5318 | 1.014 | -94.6794 | 94.7121 | 0.22897 | -0.26168 | -0.32056 | 0.51589 | 94.6794 | 0.26168 | 53.8304
0.5808 | 1.4757 | 1.0037 | -94.9888 | 95.0206 | 0.22991 | -0.26262 | -0.32243 | 0.51869 | 94.9888 | 0.26262 | 53.9978
0.5826 | 1.4084 | 0.9972 | -95.4019 | 95.4336 | 0.23084 | -0.26355 | -0.32617 | 0.52243 | 95.4019 | 0.26355 | 54.1651
0.5844 | 1.3477 | 1.0224 | -95.915 | 95.9458 | 0.23271 | -0.26636 | -0.33084 | 0.5271 | 95.915 | 0.26636 | 54.3325
0.5862 | 1.3009 | 1.0449 | -96.4916 | 96.5215 | 0.23645 | -0.27009 | -0.33458 | 0.53364 | 96.4916 | 0.27009 | 54.4998
0.588 | 1.2869 | 1.0252 | -96.8991 | 96.9299 | 0.24206 | -0.2757 | -0.33645 | 0.54206 | 96.8991 | 0.2757 | 54.6672
0.5898 | 1.2579 | 0.97944 | -97.257 | 97.2879 | 0.24766 | -0.27944 | -0.33832 | 0.5486 | 97.257 | 0.27944 | 54.8345
0.5916 | 1.228 | 0.95888 | -97.686 | 97.7168 | 0.2514 | -0.28037 | -0.34112 | 0.55327 | 97.686 | 0.28037 | 55.0019
0.5934 | 1.2271 | 0.95514 | -98.057 | 98.0879 | 0.25607 | -0.28318 | -0.34299 | 0.55794 | 98.057 | 0.28318 | 55.1692
0.5952 | 1.272 | 0.96822 | -98.3009 | 98.3327 | 0.26075 | -0.28972 | -0.34486 | 0.56636 | 98.3009 | 0.28972 | 55.3366
0.597 | 1.2654 | 0.97009 | -98.7075 | 98.7383 | 0.26542 | -0.29252 | -0.34766 | 0.57196 | 98.7075 | 0.29252 | 55.5039
0.5988 | 1.2393 | 0.99346 | -99.2364 | 99.2682 | 0.27103 | -0.29439 | -0.3514 | 0.5785 | 99.2364 | 0.29439 | 55.6713
0.6006 | 1.1626 | 0.99907 | -99.8589 | 99.8888 | 0.27664 | -0.29252 | -0.35514 | 0.58318 | 99.8589 | 0.29252 | 55.8386
0.6023 | 1.0794 | 0.99626 | -100.4682 | 100.4981 | 0.28224 | -0.28879 | -0.35794 | 0.58598 | 100.4682 | 0.28879 | 55.9967
0.6041 | 1.028 | 1.0168 | -101.2561 | 101.286 | 0.28972 | -0.28785 | -0.36262 | 0.59252 | 101.2561 | 0.28785 | 56.164
0.6059 | 1.028 | 1.0589 | -102.0523 | 102.0832 | 0.2972 | -0.29252 | -0.36729 | 0.60187 | 102.0523 | 0.29252 | 56.3313
0.6077 | 1.0355 | 1.1187 | -102.8327 | 102.8645 | 0.30654 | -0.29533 | -0.37103 | 0.61215 | 102.8327 | 0.29533 | 56.4987
0.6095 | 1.0262 | 1.172 | -103.5822 | 103.615 | 0.31589 | -0.29626 | -0.37477 | 0.62056 | 103.5822 | 0.29626 | 56.666
0.6113 | 0.99626 | 1.1832 | -104.228 | 104.2617 | 0.3215 | -0.29346 | -0.37757 | 0.6243 | 104.228 | 0.29346 | 56.8334
0.6131 | 0.98692 | 1.1953 | -104.8065 | 104.8411 | 0.32617 | -0.29346 | -0.38131 | 0.62897 | 104.8065 | 0.29346 | 57.0007
0.6149 | 0.95888 | 1.2 | -105.3542 | 105.3879 | 0.32897 | -0.29346 | -0.38505 | 0.63271 | 105.3542 | 0.29346 | 57.1681
0.6167 | 0.93458 | 1.2271 | -105.8822 | 105.9159 | 0.32991 | -0.29252 | -0.38972 | 0.63738 | 105.8822 | 0.29252 | 57.3354
0.6185 | 0.91308 | 1.2673 | -106.3981 | 106.4318 | 0.32991 | -0.29252 | -0.39626 | 0.64112 | 106.3981 | 0.29252 | 57.5028
0.6203 | 0.9486 | 1.3121 | -106.7776 | 106.8121 | 0.33364 | -0.29346 | -0.4 | 0.64673 | 106.7776 | 0.29346 | 57.6701
0.6221 | 0.94393 | 1.343 | -107.1785 | 107.2131 | 0.33551 | -0.29065 | -0.40374 | 0.6486 | 107.1785 | 0.29065 | 57.8375
0.6239 | 0.90748 | 1.3645 | -107.7056 | 107.7421 | 0.33738 | -0.28879 | -0.40841 | 0.65047 | 107.7056 | 0.28879 | 58.0048
0.6256 | 0.89159 | 1.4028 | -108.3374 | 108.3757 | 0.34019 | -0.28879 | -0.41121 | 0.65514 | 108.3374 | 0.28879 | 58.1629
0.6274 | 0.87477 | 1.4879 | -109.1626 | 109.2028 | 0.3486 | -0.28879 | -0.41308 | 0.66168 | 109.1626 | 0.28879 | 58.3302
0.6292 | 0.92523 | 1.6131 | -109.9869 | 110.0327 | 0.35981 | -0.29159 | -0.41776 | 0.6729 | 109.9869 | 0.29159 | 58.4976
0.631 | 0.99159 | 1.7486 | -110.8112 | 110.8636 | 0.37103 | -0.29439 | -0.42336 | 0.68505 | 110.8112 | 0.29439 | 58.6649
0.6328 | 1.0112 | 1.8495 | -111.6953 | 111.7514 | 0.38318 | -0.29533 | -0.42804 | 0.69626 | 111.6953 | 0.29533 | 58.8323
0.6346 | 1.0103 | 1.9364 | -112.6019 | 112.6607 | 0.39533 | -0.29439 | -0.43271 | 0.70748 | 112.6019 | 0.29439 | 58.9996
0.6364 | 1.0477 | 2.0551 | -113.3355 | 113.3991 | 0.40561 | -0.29439 | -0.43832 | 0.71776 | 113.3355 | 0.29439 | 59.167
0.6382 | 1.0692 | 2.1533 | -114.0804 | 114.1486 | 0.41589 | -0.29159 | -0.44393 | 0.7271 | 114.0804 | 0.29159 | 59.3343
0.64 | 1.0963 | 2.215 | -114.6327 | 114.7028 | 0.42056 | -0.28972 | -0.4486 | 0.73178 | 114.6327 | 0.28972 | 59.5017
0.6418 | 1.1327 | 2.2664 | -115.1449 | 115.2196 | 0.42523 | -0.28879 | -0.45327 | 0.73645 | 115.1449 | 0.28879 | 59.669
0.6436 | 1.1785 | 2.3308 | -115.6037 | 115.6813 | 0.42991 | -0.28692 | -0.45701 | 0.74112 | 115.6037 | 0.28692 | 59.8364
0.6454 | 1.2243 | 2.3953 | -116.0234 | 116.1037 | 0.43458 | -0.28505 | -0.46168 | 0.74486 | 116.0234 | 0.28505 | 60.0037
0.6472 | 1.2757 | 2.4533 | -116.3907 | 116.4738 | 0.43832 | -0.28411 | -0.46449 | 0.74766 | 116.3907 | 0.28411 | 60.1711
0.649 | 1.3318 | 2.5084 | -116.6692 | 116.7551 | 0.44299 | -0.28505 | -0.46729 | 0.7514 | 116.6692 | 0.28505 | 60.3384
0.6507 | 1.3794 | 2.5673 | -116.928 | 117.0178 | 0.44953 | -0.28505 | -0.47009 | 0.75421 | 116.928 | 0.28505 | 60.4965
0.6525 | 1.4234 | 2.6009 | -117.1327 | 117.2243 | 0.45234 | -0.28598 | -0.47196 | 0.75701 | 117.1327 | 0.28598 | 60.6638
0.6543 | 1.4355 | 2.6056 | -117.3103 | 117.4028 | 0.45421 | -0.28692 | -0.47383 | 0.75888 | 117.3103 | 0.28692 | 60.8312
0.6561 | 1.472 | 2.6505 | -117.4832 | 117.5766 | 0.45514 | -0.28785 | -0.4757 | 0.76075 | 117.4832 | 0.28785 | 60.9985
0.6579 | 1.5299 | 2.7523 | -117.6299 | 117.7271 | 0.45701 | -0.28692 | -0.4785 | 0.76262 | 117.6299 | 0.28692 | 61.1659
0.6597 | 1.5729 | 2.8477 | -117.785 | 117.885 | 0.45794 | -0.28598 | -0.48224 | 0.76449 | 117.785 | 0.28598 | 61.3332
0.6615 | 1.5495 | 2.9019 | -118.0075 | 118.1084 | 0.45701 | -0.28131 | -0.48505 | 0.76262 | 118.0075 | 0.28131 | 61.5006
0.6633 | 1.4963 | 2.9402 | -118.328 | 118.429 | 0.45701 | -0.27383 | -0.48785 | 0.76168 | 118.328 | 0.27383 | 61.6679
0.6651 | 1.4598 | 2.9458 | -118.6738 | 118.7748 | 0.45794 | -0.27103 | -0.49065 | 0.76355 | 118.6738 | 0.27103 | 61.8353
0.6669 | 1.4206 | 2.9449 | -119.1112 | 119.2131 | 0.45888 | -0.27103 | -0.49346 | 0.76636 | 119.1112 | 0.27103 | 62.0026
0.6687 | 1.4121 | 2.9327 | -119.5673 | 119.6701 | 0.46355 | -0.2729 | -0.4972 | 0.7729 | 119.5673 | 0.2729 | 62.17
0.6705 | 1.4168 | 2.9299 | -119.9617 | 120.0654 | 0.46636 | -0.2729 | -0.50093 | 0.77757 | 119.9617 | 0.2729 | 62.3373
0.6723 | 1.4224 | 2.9383 | -120.3084 | 120.414 | 0.46916 | -0.27477 | -0.50467 | 0.78318 | 120.3084 | 0.27477 | 62.5046
0.6741 | 1.4617 | 2.8963 | -120.7037 | 120.8121 | 0.47383 | -0.27757 | -0.50748 | 0.78972 | 120.7037 | 0.27757 | 62.672
0.6758 | 1.5075 | 2.8832 | -121.0159 | 121.1252 | 0.47477 | -0.2785 | -0.51028 | 0.79252 | 121.0159 | 0.2785 | 62.83
0.6776 | 1.5505 | 2.8794 | -121.2822 | 121.3935 | 0.4757 | -0.2785 | -0.51308 | 0.79533 | 121.2822 | 0.2785 | 62.9974
0.6794 | 1.5925 | 2.871 | -121.5383 | 121.6495 | 0.47757 | -0.27944 | -0.51589 | 0.79907 | 121.5383 | 0.27944 | 63.1647
0.6812 | 1.6364 | 2.8841 | -121.7467 | 121.8589 | 0.47944 | -0.28318 | -0.51869 | 0.8028 | 121.7467 | 0.28318 | 63.3321
0.683 | 1.6495 | 2.8477 | -121.8991 | 122.0093 | 0.47944 | -0.28411 | -0.5215 | 0.80561 | 121.8991 | 0.28411 | 63.4994
0.6848 | 1.6121 | 2.771 | -122.0841 | 122.1888 | 0.4785 | -0.28411 | -0.52243 | 0.80654 | 122.0841 | 0.28411 | 63.6668
0.6866 | 1.5748 | 2.6822 | -122.3028 | 122.4019 | 0.47664 | -0.28318 | -0.5243 | 0.80654 | 122.3028 | 0.28318 | 63.8341
0.6884 | 1.5486 | 2.5972 | -122.5364 | 122.6299 | 0.47664 | -0.28411 | -0.52617 | 0.80841 | 122.5364 | 0.28411 | 64.0015
0.6902 | 1.5542 | 2.5477 | -122.7598 | 122.8505 | 0.47757 | -0.28785 | -0.52804 | 0.81215 | 122.7598 | 0.28785 | 64.1688
0.692 | 1.6065 | 2.5617 | -122.9196 | 123.0112 | 0.47944 | -0.29252 | -0.52991 | 0.81682 | 122.9196 | 0.29252 | 64.3362
0.6938 | 1.6579 | 2.5804 | -123.0112 | 123.1037 | 0.47944 | -0.29439 | -0.53364 | 0.82056 | 123.0112 | 0.29439 | 64.5035
0.6956 | 1.7178 | 2.5981 | -123.0626 | 123.157 | 0.4785 | -0.2972 | -0.53645 | 0.82243 | 123.0626 | 0.2972 | 64.6709
0.6974 | 1.7822 | 2.6159 | -123.1047 | 123.2009 | 0.47664 | -0.29907 | -0.53925 | 0.8243 | 123.1047 | 0.29907 | 64.8382
0.6992 | 1.8533 | 2.6449 | -123.1224 | 123.2215 | 0.4757 | -0.30093 | -0.54206 | 0.8271 | 123.1224 | 0.30093 | 65.0056
0.7009 | 1.9252 | 2.6748 | -123.1355 | 123.2374 | 0.47383 | -0.30374 | -0.54486 | 0.82897 | 123.1355 | 0.30374 | 65.1636
0.7027 | 2.0084 | 2.6925 | -123.1243 | 123.228 | 0.47103 | -0.30561 | -0.54766 | 0.83084 | 123.1243 | 0.30561 | 65.331
0.7045 | 2.0617 | 2.6991 | -123.0888 | 123.1944 | 0.46916 | -0.30654 | -0.54953 | 0.83178 | 123.0888 | 0.30654 | 65.4983
0.7063 | 2.0935 | 2.6963 | -123.0243 | 123.1308 | 0.46822 | -0.30654 | -0.55047 | 0.83084 | 123.0243 | 0.30654 | 65.6657
0.7081 | 2.1196 | 2.6785 | -122.9598 | 123.0664 | 0.46449 | -0.30561 | -0.55047 | 0.82991 | 122.9598 | 0.30561 | 65.833
0.7099 | 2.1458 | 2.6617 | -122.8963 | 123.0028 | 0.46168 | -0.30467 | -0.5514 | 0.82804 | 122.8963 | 0.30467 | 66.0004
0.7117 | 2.1224 | 2.628 | -122.9056 | 123.0112 | 0.45701 | -0.30093 | -0.5514 | 0.8243 | 122.9056 | 0.30093 | 66.1677
0.7135 | 2.1028 | 2.5822 | -122.986 | 123.0907 | 0.45234 | -0.29907 | -0.5514 | 0.82056 | 122.986 | 0.29907 | 66.3351
0.7153 | 2.071 | 2.5243 | -123.1065 | 123.2075 | 0.44673 | -0.29626 | -0.5514 | 0.81589 | 123.1065 | 0.29626 | 66.5024
0.7171 | 2.0336 | 2.4673 | -123.2579 | 123.357 | 0.44206 | -0.29439 | -0.5514 | 0.81215 | 123.2579 | 0.29439 | 66.6698
0.7189 | 1.9673 | 2.3879 | -123.4121 | 123.5075 | 0.43738 | -0.29065 | -0.5514 | 0.80748 | 123.4121 | 0.29065 | 66.8371
0.7207 | 1.9159 | 2.328 | -123.5692 | 123.6607 | 0.43364 | -0.28879 | -0.55047 | 0.80374 | 123.5692 | 0.28879 | 67.0045
0.7225 | 1.8841 | 2.2963 | -123.728 | 123.8196 | 0.42991 | -0.28879 | -0.5486 | 0.80093 | 123.728 | 0.28879 | 67.1718
0.7242 | 1.8449 | 2.2748 | -123.8346 | 123.9243 | 0.4271 | -0.28785 | -0.5486 | 0.8 | 123.8346 | 0.28785 | 67.3299
0.726 | 1.7991 | 2.272 | -123.9393 | 124.029 | 0.4271 | -0.28598 | -0.5486 | 0.79907 | 123.9393 | 0.28598 | 67.4972
0.7278 | 1.7561 | 2.272 | -124.0364 | 124.1262 | 0.42617 | -0.28411 | -0.5486 | 0.79813 | 124.0364 | 0.28411 | 67.6646
0.7296 | 1.7336 | 2.285 | -124.0794 | 124.1701 | 0.42523 | -0.28411 | -0.5486 | 0.79813 | 124.0794 | 0.28411 | 67.8319
0.7314 | 1.729 | 2.2991 | -124.0813 | 124.172 | 0.42523 | -0.28505 | -0.54953 | 0.79813 | 124.0813 | 0.28505 | 67.9993
0.7332 | 1.7355 | 2.315 | -124.0523 | 124.1449 | 0.42336 | -0.28598 | -0.55047 | 0.79907 | 124.0523 | 0.28598 | 68.1666
0.735 | 1.7421 | 2.3318 | -124.0234 | 124.1159 | 0.4215 | -0.28692 | -0.55047 | 0.79813 | 124.0234 | 0.28692 | 68.334
0.7368 | 1.7486 | 2.3486 | -123.9935 | 124.0879 | 0.41869 | -0.28692 | -0.5514 | 0.79813 | 123.9935 | 0.28692 | 68.5013
0.7386 | 1.7495 | 2.3645 | -124.0421 | 124.1364 | 0.41589 | -0.28785 | -0.55234 | 0.79813 | 124.0421 | 0.28785 | 68.6687
0.7404 | 1.7383 | 2.3636 | -124.1364 | 124.229 | 0.41402 | -0.29159 | -0.55421 | 0.79907 | 124.1364 | 0.29159 | 68.836
0.7422 | 1.7093 | 2.3523 | -124.2766 | 124.3682 | 0.41215 | -0.29439 | -0.55514 | 0.80093 | 124.2766 | 0.29439 | 69.0033
0.744 | 1.6794 | 2.3542 | -124.4841 | 124.5748 | 0.41121 | -0.29346 | -0.55514 | 0.8 | 124.4841 | 0.29346 | 69.1707
0.7458 | 1.6598 | 2.3589 | -124.7336 | 124.8252 | 0.40935 | -0.29252 | -0.55514 | 0.79813 | 124.7336 | 0.29252 | 69.338
0.7476 | 1.7 | 2.3804 | -124.8346 | 124.928 | 0.40841 | -0.29533 | -0.55701 | 0.8 | 124.8346 | 0.29533 | 69.5054
0.7493 | 1.7505 | 2.4131 | -124.8916 | 124.986 | 0.40935 | -0.29907 | -0.55794 | 0.8028 | 124.8916 | 0.29907 | 69.6634
0.7511 | 1.7925 | 2.4589 | -124.9486 | 125.0439 | 0.40841 | -0.3028 | -0.55888 | 0.80467 | 124.9486 | 0.3028 | 69.8308
0.7529 | 1.8346 | 2.5178 | -124.9486 | 125.0449 | 0.40935 | -0.30561 | -0.55981 | 0.80654 | 124.9486 | 0.30561 | 69.9981
0.7547 | 1.8766 | 2.585 | -124.9093 | 125.0084 | 0.41028 | -0.30748 | -0.56075 | 0.80935 | 124.9093 | 0.30748 | 70.1655
0.7565 | 1.9271 | 2.6402 | -124.8271 | 124.928 | 0.41121 | -0.30935 | -0.56262 | 0.81121 | 124.8271 | 0.30935 | 70.3328
0.7583 | 1.9897 | 2.6785 | -124.6897 | 124.7925 | 0.41121 | -0.31121 | -0.56262 | 0.81308 | 124.6897 | 0.31121 | 70.5002
0.7601 | 2.0103 | 2.7075 | -124.5411 | 124.6449 | 0.41308 | -0.31028 | -0.56355 | 0.81402 | 124.5411 | 0.31028 | 70.6675
0.7619 | 2.0047 | 2.7308 | -124.3869 | 124.4888 | 0.41402 | -0.30935 | -0.56355 | 0.81495 | 124.3869 | 0.30935 | 70.8349
0.7637 | 2 | 2.7533 | -124.2327 | 124.3336 | 0.41589 | -0.30748 | -0.56262 | 0.81495 | 124.2327 | 0.30748 | 71.0022
0.7655 | 1.9467 | 2.7505 | -124.114 | 124.2112 | 0.41682 | -0.3028 | -0.56355 | 0.81308 | 124.114 | 0.3028 | 71.1696
0.7673 | 1.8925 | 2.7467 | -123.9963 | 124.0907 | 0.41869 | -0.29907 | -0.56355 | 0.81215 | 123.9963 | 0.29907 | 71.3369
0.7691 | 1.8336 | 2.7402 | -123.8822 | 123.9729 | 0.42056 | -0.29346 | -0.56355 | 0.81028 | 123.8822 | 0.29346 | 71.5043
0.7709 | 1.772 | 2.7308 | -123.772 | 123.8598 | 0.42243 | -0.28785 | -0.56262 | 0.80841 | 123.772 | 0.28785 | 71.6716
0.7727 | 1.7093 | 2.7215 | -123.6617 | 123.7458 | 0.4243 | -0.28224 | -0.56262 | 0.80654 | 123.6617 | 0.28224 | 71.839
0.7744 | 1.6617 | 2.7243 | -123.529 | 123.6112 | 0.42523 | -0.27664 | -0.56262 | 0.80374 | 123.529 | 0.27664 | 71.997
0.7762 | 1.6336 | 2.7449 | -123.3645 | 123.4458 | 0.4243 | -0.27103 | -0.56262 | 0.8 | 123.3645 | 0.27103 | 72.1644
0.778 | 1.6449 | 2.7607 | -123.0813 | 123.1626 | 0.4243 | -0.26636 | -0.56262 | 0.79813 | 123.0813 | 0.26636 | 72.3317
0.7798 | 1.6757 | 2.7935 | -122.7187 | 122.7991 | 0.42336 | -0.26168 | -0.56262 | 0.79439 | 122.7187 | 0.26168 | 72.4991
0.7816 | 1.7393 | 2.8252 | -122.3224 | 122.4047 | 0.42243 | -0.25701 | -0.56355 | 0.79159 | 122.3224 | 0.25701 | 72.6664
0.7834 | 1.7925 | 2.8439 | -121.8617 | 121.9458 | 0.41869 | -0.25234 | -0.56355 | 0.78692 | 121.8617 | 0.25234 | 72.8338
0.7852 | 1.8402 | 2.872 | -121.3056 | 121.3925 | 0.41308 | -0.24579 | -0.56262 | 0.7785 | 121.3056 | 0.24579 | 73.0011
0.787 | 1.8701 | 2.8879 | -120.7766 | 120.8645 | 0.40654 | -0.23738 | -0.56262 | 0.77103 | 120.7766 | 0.23738 | 73.1685
0.7888 | 1.8636 | 2.8794 | -120.3028 | 120.3907 | 0.40187 | -0.22617 | -0.56262 | 0.76262 | 120.3028 | 0.22617 | 73.3358
0.7906 | 1.843 | 2.9187 | -119.8121 | 119.9009 | 0.39159 | -0.21308 | -0.55981 | 0.74953 | 119.8121 | 0.21308 | 73.5032
0.7924 | 1.8252 | 2.9393 | -119.2364 | 119.3252 | 0.37944 | -0.19813 | -0.55607 | 0.73551 | 119.2364 | 0.19813 | 73.6705
0.7942 | 1.8103 | 2.9402 | -118.6346 | 118.7224 | 0.36822 | -0.18318 | -0.55327 | 0.72056 | 118.6346 | 0.18318 | 73.8379
0.796 | 1.7935 | 2.9374 | -118.2271 | 118.314 | 0.36262 | -0.1729 | -0.55327 | 0.71402 | 118.2271 | 0.1729 | 74.0052
0.7977 | 1.8131 | 2.9486 | -117.8897 | 117.9766 | 0.35981 | -0.16542 | -0.55327 | 0.71121 | 117.8897 | 0.16542 | 74.1633
0.7995 | 1.8393 | 2.9636 | -117.6402 | 117.7271 | 0.35794 | -0.15981 | -0.55327 | 0.70935 | 117.6402 | 0.15981 | 74.3306
0.8013 | 1.8364 | 2.9654 | -117.4411 | 117.5252 | 0.35514 | -0.15327 | -0.55421 | 0.70561 | 117.4411 | 0.15327 | 74.498
0.8031 | 1.8336 | 2.9673 | -117.2411 | 117.3243 | 0.35234 | -0.14673 | -0.55514 | 0.7028 | 117.2411 | 0.14673 | 74.6653
0.8049 | 1.8206 | 2.9514 | -117.0972 | 117.1785 | 0.34766 | -0.13925 | -0.55607 | 0.69907 | 117.0972 | 0.13925 | 74.8327
0.8067 | 1.7551 | 2.8813 | -117.0159 | 117.0925 | 0.34299 | -0.12897 | -0.55701 | 0.69346 | 117.0159 | 0.12897 | 75
0.8085 | 1.6607 | 2.8224 | -116.8972 | 116.971 | 0.33832 | -0.11869 | -0.55794 | 0.68879 | 116.8972 | 0.11869 | 75.1673
0.8103 | 1.6084 | 2.7701 | -116.8159 | 116.886 | 0.33458 | -0.11121 | -0.55981 | 0.68598 | 116.8159 | 0.11121 | 75.3347
0.8121 | 1.5925 | 2.7626 | -116.6523 | 116.7215 | 0.33084 | -0.10467 | -0.56075 | 0.68318 | 116.6523 | 0.10467 | 75.502
0.8139 | 1.5935 | 2.7645 | -116.4402 | 116.5084 | 0.32897 | -0.10093 | -0.56075 | 0.68131 | 116.4402 | 0.10093 | 75.6694
0.8157 | 1.5757 | 2.7486 | -116.1701 | 116.2364 | 0.32617 | -0.095327 | -0.56075 | 0.67757 | 116.1701 | 0.095327 | 75.8367
0.8175 | 1.5738 | 2.7561 | -115.8206 | 115.886 | 0.3215 | -0.088785 | -0.56075 | 0.67383 | 115.8206 | 0.088785 | 76.0041
0.8193 | 1.5748 | 2.7495 | -115.4028 | 115.4664 | 0.31402 | -0.081308 | -0.56168 | 0.66822 | 115.4028 | 0.081308 | 76.1714
0.8211 | 1.5159 | 2.6346 | -114.8879 | 114.9467 | 0.30093 | -0.068224 | -0.56075 | 0.66075 | 114.8879 | 0.068224 | 76.3388
0.8228 | 1.5093 | 2.5729 | -114.5019 | 114.557 | 0.29533 | -0.059813 | -0.56075 | 0.65794 | 114.5019 | 0.059813 | 76.4968
0.8246 | 1.5009 | 2.5262 | -114.1785 | 114.2318 | 0.29252 | -0.053271 | -0.55981 | 0.65607 | 114.1785 | 0.053271 | 76.6642
0.8264 | 1.4607 | 2.4477 | -113.8841 | 113.9346 | 0.28879 | -0.04486 | -0.55888 | 0.65327 | 113.8841 | 0.04486 | 76.8315
0.8282 | 1.3925 | 2.3953 | -113.6393 | 113.6869 | 0.28692 | -0.037383 | -0.55794 | 0.6514 | 113.6393 | 0.037383 | 76.9989
0.83 | 1.3944 | 2.3505 | -113.4645 | 113.5121 | 0.28785 | -0.034579 | -0.55794 | 0.6514 | 113.4645 | 0.034579 | 77.1662
0.8318 | 1.3981 | 2.2355 | -113.2393 | 113.2841 | 0.28692 | -0.030841 | -0.55701 | 0.65047 | 113.2393 | 0.030841 | 77.3336
0.8336 | 1.3748 | 2.071 | -112.9748 | 113.015 | 0.28224 | -0.027103 | -0.55514 | 0.64579 | 112.9748 | 0.027103 | 77.5009
0.8354 | 1.3299 | 1.9673 | -112.7944 | 112.8318 | 0.27944 | -0.021495 | -0.55514 | 0.64206 | 112.7944 | 0.021495 | 77.6683
0.8372 | 1.2832 | 1.8523 | -112.6617 | 112.6963 | 0.27664 | -0.018692 | -0.55421 | 0.63832 | 112.6617 | 0.018692 | 77.8356
0.839 | 1.2196 | 1.7692 | -112.557 | 112.5879 | 0.27944 | -0.013084 | -0.55234 | 0.63645 | 112.557 | 0.013084 | 78.003
0.8408 | 1.1664 | 1.6664 | -112.4729 | 112.5019 | 0.27757 | -0.0084112 | -0.5514 | 0.63271 | 112.4729 | 0.0084112 | 78.1703
0.8426 | 1.1121 | 1.5542 | -112.4972 | 112.5243 | 0.2757 | -0.0056075 | -0.55047 | 0.62991 | 112.4972 | 0.0056075 | 78.3377
0.8444 | 1.0888 | 1.4477 | -112.5963 | 112.6224 | 0.27383 | -0.0037383 | -0.54953 | 0.62617 | 112.5963 | 0.0037383 | 78.505
0.8462 | 1.0813 | 1.3449 | -112.6748 | 112.7019 | 0.2729 | -0.0018692 | -0.54766 | 0.6243 | 112.6748 | 0.0018692 | 78.6724
0.8479 | 1.1047 | 1.2523 | -112.685 | 112.7121 | 0.2729 | 0.0028037 | -0.54486 | 0.62336 | 112.685 | -0.0028037 | 78.8304
0.8497 | 1.1421 | 1.2019 | -112.6467 | 112.6757 | 0.27383 | 0.0046729 | -0.54299 | 0.62336 | 112.6467 | -0.0046729 | 78.9978
0.8515 | 1.1617 | 1.0991 | -112.6084 | 112.6383 | 0.27477 | 0.0046729 | -0.54112 | 0.62056 | 112.6084 | -0.0046729 | 79.1651
0.8533 | 1.1701 | 1.0168 | -112.5486 | 112.5794 | 0.27477 | 0.0046729 | -0.53832 | 0.61869 | 112.5486 | -0.0046729 | 79.3325
0.8551 | 1.1645 | 0.96075 | -112.4626 | 112.4925 | 0.27664 | 0.0037383 | -0.53738 | 0.61682 | 112.4626 | -0.0037383 | 79.4998
0.8569 | 1.1804 | 0.92243 | -112.3561 | 112.386 | 0.27757 | 0.0028037 | -0.53645 | 0.61589 | 112.3561 | -0.0028037 | 79.6672
0.8587 | 1.2196 | 0.90467 | -112.2262 | 112.2561 | 0.2785 | 0.0018692 | -0.53645 | 0.61589 | 112.2262 | -0.0018692 | 79.8345
0.8605 | 1.2421 | 0.87009 | -112.0523 | 112.0822 | 0.28131 | 0.00093458 | -0.53551 | 0.61589 | 112.0523 | -0.00093458 | 80.0019
0.8623 | 1.2523 | 0.82617 | -111.8551 | 111.8841 | 0.28505 | 0.00093458 | -0.53458 | 0.61589 | 111.8551 | -0.00093458 | 80.1692
0.8641 | 1.3037 | 0.81121 | -111.6243 | 111.6542 | 0.28692 | -0.0037383 | -0.53458 | 0.61495 | 111.6243 | 0.0037383 | 80.3366
0.8659 | 1.371 | 0.82991 | -111.3271 | 111.357 | 0.28785 | -0.0084112 | -0.53364 | 0.61402 | 111.3271 | 0.0084112 | 80.5039
0.8677 | 1.3935 | 0.81682 | -111.0421 | 111.072 | 0.28972 | -0.011215 | -0.53364 | 0.61402 | 111.0421 | 0.011215 | 80.6713
0.8695 | 1.415 | 0.80467 | -110.7598 | 110.7897 | 0.29252 | -0.014019 | -0.53271 | 0.61402 | 110.7598 | 0.014019 | 80.8386
0.8712 | 1.4262 | 0.81682 | -110.5785 | 110.6075 | 0.29252 | -0.017757 | -0.53271 | 0.61402 | 110.5785 | 0.017757 | 80.9967
0.873 | 1.3991 | 0.81682 | -110.4439 | 110.472 | 0.29252 | -0.017757 | -0.53271 | 0.61402 | 110.4439 | 0.017757 | 81.164
0.8748 | 1.3346 | 0.80374 | -110.3664 | 110.3925 | 0.29252 | -0.013084 | -0.53271 | 0.61402 | 110.3664 | 0.013084 | 81.3313
0.8766 | 1.2776 | 0.8 | -110.3234 | 110.3467 | 0.29252 | -0.01028 | -0.53271 | 0.61308 | 110.3234 | 0.01028 | 81.4987
0.8784 | 1.2626 | 0.88131 | -110.4458 | 110.4673 | 0.29439 | -0.013084 | -0.53271 | 0.61589 | 110.4458 | 0.013084 | 81.666
0.8802 | 1.2131 | 0.96636 | -110.6308 | 110.6514 | 0.3 | -0.016822 | -0.53364 | 0.6215 | 110.6308 | 0.016822 | 81.8334
0.882 | 1.1626 | 1.0243 | -110.7308 | 110.7514 | 0.30654 | -0.016822 | -0.53458 | 0.6271 | 110.7308 | 0.016822 | 82.0007
0.8838 | 1.1168 | 1.072 | -110.7869 | 110.8065 | 0.31215 | -0.015888 | -0.53551 | 0.63178 | 110.7869 | 0.015888 | 82.1681
0.8856 | 1.1495 | 1.1084 | -110.6794 | 110.7 | 0.31589 | -0.016822 | -0.53551 | 0.63364 | 110.6794 | 0.016822 | 82.3354
0.8874 | 1.2047 | 1.1523 | -110.4916 | 110.5121 | 0.31869 | -0.017757 | -0.53645 | 0.63645 | 110.4916 | 0.017757 | 82.5028
0.8892 | 1.2505 | 1.1991 | -110.2645 | 110.286 | 0.3215 | -0.018692 | -0.53645 | 0.63832 | 110.2645 | 0.018692 | 82.6701
0.891 | 1.3103 | 1.2645 | -109.9888 | 110.0103 | 0.32617 | -0.020561 | -0.53738 | 0.64393 | 109.9888 | 0.020561 | 82.8375
0.8928 | 1.3486 | 1.3159 | -109.7355 | 109.7579 | 0.32991 | -0.020561 | -0.53925 | 0.64673 | 109.7355 | 0.020561 | 83.0048
0.8946 | 1.3234 | 1.3308 | -109.5514 | 109.5738 | 0.33084 | -0.016822 | -0.53925 | 0.64766 | 109.5514 | 0.016822 | 83.1722
0.8963 | 1.2551 | 1.3206 | -109.4206 | 109.4421 | 0.32991 | -0.01028 | -0.53925 | 0.64766 | 109.4206 | 0.01028 | 83.3302
0.8981 | 1.2159 | 1.3308 | -109.3654 | 109.386 | 0.32991 | -0.0056075 | -0.54019 | 0.64766 | 109.3654 | 0.0056075 | 83.4976
0.8999 | 1.1935 | 1.3449 | -109.3935 | 109.414 | 0.32991 | -0.0037383 | -0.54112 | 0.6486 | 109.3935 | 0.0037383 | 83.6649
0.9017 | 1.171 | 1.3794 | -109.5112 | 109.5318 | 0.32897 | 0 | -0.54393 | 0.6486 | 109.5112 | 0 | 83.8323
0.9035 | 1.1598 | 1.4187 | -109.6056 | 109.6262 | 0.3271 | 0.0037383 | -0.54579 | 0.6486 | 109.6056 | -0.0037383 | 83.9996
0.9053 | 1.1766 | 1.4551 | -109.5701 | 109.5916 | 0.32617 | 0.0056075 | -0.54673 | 0.64766 | 109.5701 | -0.0056075 | 84.167
0.9071 | 1.2084 | 1.5 | -109.4607 | 109.4841 | 0.32617 | 0.0065421 | -0.54766 | 0.64766 | 109.4607 | -0.0065421 | 84.3343
0.9089 | 1.1673 | 1.515 | -109.3327 | 109.3561 | 0.32523 | 0.01215 | -0.54766 | 0.64766 | 109.3327 | -0.01215 | 84.5017
0.9107 | 1.1224 | 1.529 | -109.2047 | 109.2271 | 0.32336 | 0.017757 | -0.54766 | 0.64766 | 109.2047 | -0.017757 | 84.669
0.9125 | 1.1411 | 1.5841 | -109.0692 | 109.0925 | 0.32243 | 0.020561 | -0.54766 | 0.64673 | 109.0692 | -0.020561 | 84.8364
0.9143 | 1.1701 | 1.614 | -108.8551 | 108.8794 | 0.32056 | 0.02243 | -0.5486 | 0.64579 | 108.8551 | -0.02243 | 85.0037
0.9161 | 1.1551 | 1.6243 | -108.543 | 108.5682 | 0.31402 | 0.028037 | -0.54953 | 0.64206 | 108.543 | -0.028037 | 85.1711
0.9179 | 1.1411 | 1.6299 | -108.2738 | 108.2991 | 0.30841 | 0.033645 | -0.55047 | 0.63925 | 108.2738 | -0.033645 | 85.3384
0.9197 | 1.114 | 1.6168 | -108.0897 | 108.114 | 0.30561 | 0.037383 | -0.55047 | 0.63832 | 108.0897 | -0.037383 | 85.5058
0.9214 | 1.072 | 1.6112 | -107.9421 | 107.9664 | 0.3028 | 0.042056 | -0.5514 | 0.63738 | 107.9421 | -0.042056 | 85.6638
0.9232 | 1.0495 | 1.6393 | -107.8561 | 107.8794 | 0.30187 | 0.045794 | -0.5514 | 0.63738 | 107.8561 | -0.045794 | 85.8312
0.925 | 1.0766 | 1.6944 | -107.686 | 107.7112 | 0.3028 | 0.046729 | -0.55234 | 0.63832 | 107.686 | -0.046729 | 85.9985
0.9268 | 1.1121 | 1.7467 | -107.4813 | 107.5084 | 0.30374 | 0.046729 | -0.55234 | 0.63925 | 107.4813 | -0.046729 | 86.1659
0.9286 | 1.1065 | 1.7654 | -107.2299 | 107.2579 | 0.3028 | 0.049533 | -0.55234 | 0.63925 | 107.2299 | -0.049533 | 86.3332
0.9304 | 1.086 | 1.7813 | -107.028 | 107.0561 | 0.30187 | 0.053271 | -0.55327 | 0.63925 | 107.028 | -0.053271 | 86.5006
0.9322 | 1.0626 | 1.7972 | -106.9056 | 106.9327 | 0.3 | 0.057009 | -0.55327 | 0.63925 | 106.9056 | -0.057009 | 86.6679
0.934 | 1.057 | 1.8056 | -106.7869 | 106.815 | 0.29907 | 0.059813 | -0.55327 | 0.63832 | 106.7869 | -0.059813 | 86.8353
0.9358 | 1.028 | 1.8112 | -106.729 | 106.757 | 0.29626 | 0.063551 | -0.55327 | 0.63738 | 106.729 | -0.063551 | 87.0026
0.9376 | 1.0206 | 1.8178 | -106.7561 | 106.7832 | 0.29439 | 0.064486 | -0.55327 | 0.63738 | 106.7561 | -0.064486 | 87.17
0.9394 | 1.043 | 1.8187 | -106.7477 | 106.7748 | 0.29346 | 0.064486 | -0.55327 | 0.63645 | 106.7477 | -0.064486 | 87.3373
0.9412 | 1.0561 | 1.8121 | -106.7206 | 106.7486 | 0.29252 | 0.065421 | -0.55327 | 0.63551 | 106.7206 | -0.065421 | 87.5046
0.943 | 1.0505 | 1.8159 | -106.7271 | 106.7551 | 0.29159 | 0.066355 | -0.55327 | 0.63551 | 106.7271 | -0.066355 | 87.672
0.9447 | 1.0364 | 1.8243 | -106.7551 | 106.7832 | 0.28972 | 0.068224 | -0.55327 | 0.63551 | 106.7551 | -0.068224 | 87.83
0.9465 | 1.0318 | 1.8561 | -106.8168 | 106.8449 | 0.28972 | 0.068224 | -0.55421 | 0.63551 | 106.8168 | -0.068224 | 87.9974
0.9483 | 1.0598 | 1.872 | -106.8477 | 106.8766 | 0.28972 | 0.066355 | -0.55514 | 0.63645 | 106.8477 | -0.066355 | 88.1647
0.9501 | 1.0748 | 1.8766 | -106.8645 | 106.8935 | 0.28879 | 0.066355 | -0.55514 | 0.63551 | 106.8645 | -0.066355 | 88.3321
0.9519 | 1.085 | 1.8701 | -106.8449 | 106.8738 | 0.28692 | 0.066355 | -0.55607 | 0.63551 | 106.8449 | -0.066355 | 88.4994
0.9537 | 1.1047 | 1.8804 | -106.7991 | 106.828 | 0.28598 | 0.066355 | -0.55607 | 0.63551 | 106.7991 | -0.066355 | 88.6668
0.9555 | 1.1196 | 1.885 | -106.6935 | 106.7234 | 0.28411 | 0.066355 | -0.55701 | 0.63645 | 106.6935 | -0.066355 | 88.8341
0.9573 | 1.1944 | 1.8935 | -106.5626 | 106.5916 | 0.27944 | 0.066355 | -0.55981 | 0.63645 | 106.5626 | -0.066355 | 89.0015
0.9591 | 1.2701 | 1.8953 | -106.357 | 106.386 | 0.27757 | 0.064486 | -0.56168 | 0.63645 | 106.357 | -0.064486 | 89.1688
0.9609 | 1.3159 | 1.8813 | -106.1299 | 106.1598 | 0.2757 | 0.063551 | -0.56262 | 0.63551 | 106.1299 | -0.063551 | 89.3362
0.9627 | 1.3196 | 1.8645 | -106.0449 | 106.0748 | 0.27477 | 0.064486 | -0.56262 | 0.63551 | 106.0449 | -0.064486 | 89.5035
0.9645 | 1.3206 | 1.8505 | -106.0224 | 106.0523 | 0.27383 | 0.065421 | -0.56262 | 0.63458 | 106.0224 | -0.065421 | 89.6709
0.9663 | 1.3215 | 1.8383 | -106.0439 | 106.0748 | 0.2729 | 0.066355 | -0.56262 | 0.63364 | 106.0439 | -0.066355 | 89.8382
0.9681 | 1.3243 | 1.8271 | -106.0701 | 106.1 | 0.27103 | 0.066355 | -0.56262 | 0.63364 | 106.0701 | -0.066355 | 90.0056
0.9698 | 1.2925 | 1.7925 | -106.0682 | 106.0972 | 0.26916 | 0.068224 | -0.56262 | 0.63364 | 106.0682 | -0.068224 | 90.1636
0.9716 | 1.2701 | 1.7813 | -106.0533 | 106.0822 | 0.26822 | 0.070093 | -0.56262 | 0.63364 | 106.0533 | -0.070093 | 90.331
0.9734 | 1.2561 | 1.7804 | -106.0336 | 106.0626 | 0.26822 | 0.071028 | -0.56262 | 0.63364 | 106.0336 | -0.071028 | 90.4983
0.9752 | 1.2523 | 1.7794 | -106.0449 | 106.0748 | 0.26729 | 0.071963 | -0.56262 | 0.63271 | 106.0449 | -0.071963 | 90.6657
0.977 | 1.2626 | 1.7822 | -106.0654 | 106.0953 | 0.26636 | 0.071028 | -0.56355 | 0.63271 | 106.0654 | -0.071028 | 90.833
0.9788 | 1.2822 | 1.7916 | -106.0495 | 106.0794 | 0.26542 | 0.070093 | -0.56355 | 0.63271 | 106.0495 | -0.070093 | 91.0004
0.9806 | 1.3075 | 1.814 | -106.086 | 106.1178 | 0.26542 | 0.069159 | -0.56355 | 0.63271 | 106.086 | -0.069159 | 91.1677
0.9824 | 1.3206 | 1.8215 | -106.1262 | 106.1579 | 0.26449 | 0.069159 | -0.56449 | 0.63271 | 106.1262 | -0.069159 | 91.3351
0.9842 | 1.2869 | 1.8037 | -106.1963 | 106.228 | 0.26355 | 0.071028 | -0.56449 | 0.63271 | 106.1963 | -0.071028 | 91.5024
0.986 | 1.2346 | 1.7832 | -106.328 | 106.3589 | 0.26262 | 0.074766 | -0.56449 | 0.63271 | 106.328 | -0.074766 | 91.6698
0.9878 | 1.1813 | 1.7645 | -106.5449 | 106.5757 | 0.26168 | 0.07757 | -0.56542 | 0.63364 | 106.5449 | -0.07757 | 91.8371
0.9896 | 1.1458 | 1.7523 | -106.7879 | 106.8196 | 0.26168 | 0.079439 | -0.56636 | 0.63458 | 106.7879 | -0.079439 | 92.0045
0.9914 | 1.1514 | 1.7654 | -106.943 | 106.9748 | 0.26168 | 0.079439 | -0.56636 | 0.63551 | 106.943 | -0.079439 | 92.1718
0.9932 | 1.2 | 1.7897 | -107.0037 | 107.0364 | 0.26262 | 0.076636 | -0.56729 | 0.63551 | 107.0037 | -0.076636 | 92.3392
0.9949 | 1.2654 | 1.8093 | -106.9897 | 107.0243 | 0.26168 | 0.072897 | -0.56822 | 0.63551 | 106.9897 | -0.072897 | 92.4972
0.9967 | 1.2991 | 1.815 | -106.9346 | 106.9701 | 0.26075 | 0.071028 | -0.56822 | 0.63551 | 106.9346 | -0.071028 | 92.6646
0.9985 | 1.2804 | 1.8234 | -106.8757 | 106.9103 | 0.25888 | 0.071963 | -0.56916 | 0.63551 | 106.8757 | -0.071963 | 92.8319
1.0003 | 1.2636 | 1.8327 | -106.814 | 106.8486 | 0.25701 | 0.072897 | -0.56916 | 0.63551 | 106.814 | -0.072897 | 92.9993
1.0021 | 1.2561 | 1.8439 | -106.7449 | 106.7794 | 0.25421 | 0.072897 | -0.57009 | 0.63458 | 106.7449 | -0.072897 | 93.1666
1.0039 | 1.2495 | 1.8607 | -106.7112 | 106.7467 | 0.25234 | 0.072897 | -0.57103 | 0.63458 | 106.7112 | -0.072897 | 93.334
1.0057 | 1.2458 | 1.8841 | -106.7178 | 106.7542 | 0.2514 | 0.072897 | -0.57103 | 0.63458 | 106.7178 | -0.072897 | 93.5013
1.0075 | 1.2421 | 1.9075 | -106.7252 | 106.7617 | 0.24953 | 0.072897 | -0.57196 | 0.63458 | 106.7252 | -0.072897 | 93.6687
1.0093 | 1.2393 | 1.915 | -106.7645 | 106.8009 | 0.2486 | 0.072897 | -0.57196 | 0.63458 | 106.7645 | -0.072897 | 93.836
1.0111 | 1.2364 | 1.9215 | -106.8047 | 106.8411 | 0.2486 | 0.072897 | -0.5729 | 0.63458 | 106.8047 | -0.072897 | 94.0033
1.0129 | 1.2299 | 1.9299 | -106.8841 | 106.9206 | 0.24766 | 0.072897 | -0.5729 | 0.63458 | 106.8841 | -0.072897 | 94.1707
1.0147 | 1.2243 | 1.9449 | -106.9963 | 107.0336 | 0.24766 | 0.071963 | -0.57383 | 0.63551 | 106.9963 | -0.071963 | 94.338
1.0165 | 1.2243 | 1.9664 | -107.1168 | 107.1542 | 0.2486 | 0.071028 | -0.57383 | 0.63551 | 107.1168 | -0.071028 | 94.5054
1.0182 | 1.2692 | 1.9879 | -107.2364 | 107.2757 | 0.24953 | 0.06729 | -0.57477 | 0.63551 | 107.2364 | -0.06729 | 94.6634
1.02 | 1.2888 | 1.9944 | -107.3187 | 107.3579 | 0.25047 | 0.066355 | -0.57477 | 0.63645 | 107.3187 | -0.066355 | 94.8308
1.0218 | 1.2757 | 1.9729 | -107.4028 | 107.4411 | 0.25047 | 0.066355 | -0.5757 | 0.63738 | 107.4028 | -0.066355 | 94.9981
1.0236 | 1.2458 | 1.9654 | -107.5336 | 107.571 | 0.25047 | 0.06729 | -0.57664 | 0.63738 | 107.5336 | -0.06729 | 95.1655
1.0254 | 1.2477 | 1.9991 | -107.586 | 107.6243 | 0.24953 | 0.06729 | -0.57664 | 0.63832 | 107.586 | -0.06729 | 95.3328
1.0272 | 1.2514 | 2.015 | -107.6421 | 107.6813 | 0.24953 | 0.066355 | -0.57757 | 0.63832 | 107.6421 | -0.066355 | 95.5002
1.029 | 1.2907 | 2.0131 | -107.6757 | 107.715 | 0.24953 | 0.064486 | -0.57757 | 0.63832 | 107.6757 | -0.064486 | 95.6675
1.0308 | 1.3252 | 2.0196 | -107.672 | 107.7121 | 0.24953 | 0.061682 | -0.5785 | 0.63832 | 107.672 | -0.061682 | 95.8349
1.0326 | 1.3654 | 2.0449 | -107.6252 | 107.6664 | 0.25047 | 0.059813 | -0.5785 | 0.63925 | 107.6252 | -0.059813 | 96.0022
1.0344 | 1.4103 | 2.0551 | -107.5187 | 107.5617 | 0.25047 | 0.057944 | -0.5785 | 0.63925 | 107.5187 | -0.057944 | 96.1696
1.0362 | 1.4542 | 2.057 | -107.3766 | 107.4196 | 0.24953 | 0.056075 | -0.57944 | 0.63925 | 107.3766 | -0.056075 | 96.3369
1.038 | 1.4477 | 2.0374 | -107.1477 | 107.1907 | 0.24953 | 0.057944 | -0.5785 | 0.63925 | 107.1477 | -0.057944 | 96.5043
1.0398 | 1.4318 | 2.0084 | -106.8393 | 106.8822 | 0.25047 | 0.059813 | -0.57757 | 0.63832 | 106.8393 | -0.059813 | 96.6716
1.0416 | 1.3888 | 2.0075 | -106.7393 | 106.7813 | 0.2486 | 0.062617 | -0.57757 | 0.63832 | 106.7393 | -0.062617 | 96.839
1.0433 | 1.385 | 2.0252 | -106.6411 | 106.6841 | 0.24766 | 0.063551 | -0.5785 | 0.63832 | 106.6411 | -0.063551 | 96.997
1.0451 | 1.3794 | 2.0402 | -106.5916 | 106.6346 | 0.24673 | 0.064486 | -0.57944 | 0.63832 | 106.5916 | -0.064486 | 97.1644
1.0469 | 1.4224 | 2.0701 | -106.5598 | 106.6047 | 0.24766 | 0.060748 | -0.57944 | 0.63925 | 106.5598 | -0.060748 | 97.3317
1.0487 | 1.415 | 2.0654 | -106.5019 | 106.5467 | 0.24673 | 0.061682 | -0.58037 | 0.63832 | 106.5019 | -0.061682 | 97.4991
1.0505 | 1.3598 | 2.0533 | -106.4505 | 106.4944 | 0.24673 | 0.065421 | -0.57944 | 0.63925 | 106.4505 | -0.065421 | 97.6664
1.0523 | 1.3514 | 2.0514 | -106.4841 | 106.528 | 0.24673 | 0.064486 | -0.58037 | 0.63925 | 106.4841 | -0.064486 | 97.8338
1.0541 | 1.3346 | 2.043 | -106.5065 | 106.5495 | 0.24579 | 0.064486 | -0.58037 | 0.63925 | 106.5065 | -0.064486 | 98.0011
1.0559 | 1.314 | 2.057 | -106.5093 | 106.5533 | 0.24579 | 0.066355 | -0.58037 | 0.63925 | 106.5093 | -0.066355 | 98.1685
1.0577 | 1.2953 | 2.0729 | -106.5196 | 106.5636 | 0.24486 | 0.06729 | -0.58037 | 0.63925 | 106.5196 | -0.06729 | 98.3358
1.0595 | 1.2841 | 2.0776 | -106.515 | 106.5598 | 0.24486 | 0.06729 | -0.58037 | 0.63925 | 106.515 | -0.06729 | 98.5032
1.0613 | 1.2991 | 2.0972 | -106.5477 | 106.5925 | 0.24486 | 0.065421 | -0.58131 | 0.64019 | 106.5477 | -0.065421 | 98.6705
1.0631 | 1.3336 | 2.1364 | -106.5935 | 106.6393 | 0.24579 | 0.062617 | -0.58224 | 0.64019 | 106.5935 | -0.062617 | 98.8379
1.0649 | 1.3421 | 2.1477 | -106.5963 | 106.6439 | 0.24579 | 0.062617 | -0.58318 | 0.64112 | 106.5963 | -0.062617 | 99.0052
1.0667 | 1.3458 | 2.157 | -106.5907 | 106.6393 | 0.24579 | 0.062617 | -0.58318 | 0.64206 | 106.5907 | -0.062617 | 99.1726
1.0684 | 1.3477 | 2.1645 | -106.5785 | 106.6271 | 0.24766 | 0.063551 | -0.58318 | 0.64299 | 106.5785 | -0.063551 | 99.3306
1.0702 | 1.3617 | 2.1645 | -106.5673 | 106.6159 | 0.24766 | 0.062617 | -0.58411 | 0.64299 | 106.5673 | -0.062617 | 99.498
1.072 | 1.3411 | 2.1308 | -106.5308 | 106.5785 | 0.24766 | 0.064486 | -0.58505 | 0.64393 | 106.5308 | -0.064486 | 99.6653
1.0738 | 1.3 | 2.1131 | -106.528 | 106.5748 | 0.24766 | 0.066355 | -0.58505 | 0.64393 | 106.528 | -0.066355 | 99.8327
1.0756 | 1.2916 | 2.1252 | -106.528 | 106.5738 | 0.24766 | 0.066355 | -0.58505 | 0.64486 | 106.528 | -0.066355 | 100
| Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My
Minimum | -4.7112 | -9.7822 | -163.3243 | 11.9533 | -0.08785 | -0.31121 | -0.58505 | 0.048598 | 11.8598 | -1.0421
Maximum | 4.9505 | 6.215 | -11.8598 | 163.5234 | 0.47944 | 1.0421 | -0.027103 | 1.2047 | 163.3243 | 0.31121

## Diagr_BW
Loads in %BW / %BWm
Contact forces and moments, measured telemetrically in knee implants of 8 subjects.
Averages of one individual subject.
Fx  lateral, Fy  anterior, -Fz  distal, Fres = Resultant Force
Mx around x, My around y, Mz around z, Mres = Resultant Moment
z-axis = axis of implant stem!
From Bergmann et al, PlosOne (2014) "Standardized Loads Acting in Knee Implants"
Subject K9L
Subject K9L

## Dat_AVER75
Data File: | k9l_2246_863_6_i1.akf
Curve Alignment Program: | 1.4.11
------------------------MEASURING INFOS------------------------
Diagram Title #1 | Forces and Moments at Knee Joint
Diagram Title #2 | Measured Relative to Implant
Comment #1 | Knee Joint; Gaitanalysis; Stand up; Seat Height: 45 cm;
Comment #2 | K9L, 14,5 Months PO *
BodyWeight [N]: | 1070
Patient(en): | K9L | [7 empty] | AVER75 loads in N (forces) and Nm (moments)
Implant Type: | Knee Joint
Varus Angle Leg [°] | 7 | [7 empty] | Subject | K9L
Axes Directions X, Y, Z: | Lateral, Anterior, Superior
Axes Origin: | Tibial Plateau at Stem Axis | [7 empty] | The individual peak value of Fres was adapted to the average peak of Fres from all subjects!
Transformation X Y Z (°): | 6.5 | -3.5 | -2 | [5 empty] | The AVER75 forces Fres of all subjects therefore have the same peak values.
Displacement X, Y, Z (mm): | 0 | 0 | 0 | [5 empty] | All 6 load components were then adapted by the same Multiplication Factor.
Activity Code: | KneeActivities | 1P_1_15_8
[9 empty] | Multiplication Factor = |  | Max(Fres, AVER75 of 8 Subjects) / |  |  | Max(Fres, Individual Subject)
[9 empty] | Max(Fres, AVER75 of 8 Subjects) = |  |  | 2060.01 | N
[9 empty] | Max(Fres, Individual Subject) = |  |  | 1749.7 | N
---------------------------DATA--------------------------------
Number of Data Sets = | 601
Number of analog channels = | 0
Time-Offset(sec) = | 0
Time | Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My | Time
[s] | [N] | [N] | [N] | [N] | [Nm] | [Nm] | [Nm] | [Nm] | [N] | (Nm] | [%Cycle]
0 | -1.3422 | -9.1245 | -178.4039 | 178.7336 | 0.12951 | -0.42385 | -0.34143 | 0.61222 | 178.4039 | 0.42385 | 0
0.0018 | -1.4364 | -8.6064 | -179.699 | 179.9816 | 0.14128 | -0.42385 | -0.34143 | 0.61222 | 179.699 | 0.42385 | 0.16735
0.0036 | -1.8131 | -8.1708 | -182.0301 | 182.2892 | 0.15306 | -0.41207 | -0.35321 | 0.61222 | 182.0301 | 0.41207 | 0.3347
0.0054 | -2.1546 | -7.6175 | -183.8197 | 184.0434 | 0.1766 | -0.4003 | -0.35321 | 0.61222 | 183.8197 | 0.4003 | 0.50205
0.0072 | -2.3547 | -6.7227 | -183.9139 | 184.1023 | 0.20015 | -0.41207 | -0.35321 | 0.63577 | 183.9139 | 0.41207 | 0.66939
0.009 | -2.0957 | -6.6874 | -181.9006 | 182.089 | 0.20015 | -0.41207 | -0.35321 | 0.63577 | 181.9006 | 0.41207 | 0.83674
0.0108 | -1.9426 | -6.7109 | -180.5231 | 180.6997 | 0.18838 | -0.41207 | -0.36498 | 0.63577 | 180.5231 | 0.41207 | 1.0041
0.0125 | -2.2958 | -5.9692 | -181.8889 | 182.0537 | 0.20015 | -0.4003 | -0.36498 | 0.63577 | 181.8889 | 0.4003 | 1.1621
0.0143 | -2.5078 | -5.145 | -182.2892 | 182.4304 | 0.2237 | -0.4003 | -0.37675 | 0.64754 | 182.2892 | 0.4003 | 1.3295
0.0161 | -2.4842 | -4.2856 | -181.0412 | 181.1707 | 0.24724 | -0.4003 | -0.37675 | 0.65932 | 181.0412 | 0.4003 | 1.4968
0.0179 | -2.2605 | -3.6262 | -178.6394 | 178.7453 | 0.25902 | -0.41207 | -0.37675 | 0.68286 | 178.6394 | 0.41207 | 1.6642
0.0197 | -1.8484 | -3.273 | -175.5665 | 175.6607 | 0.25902 | -0.41207 | -0.38853 | 0.68286 | 175.5665 | 0.41207 | 1.8315
0.0215 | -1.2598 | -2.7903 | -171.2574 | 171.3398 | 0.25902 | -0.43562 | -0.38853 | 0.70641 | 171.2574 | 0.43562 | 1.9989
0.0233 | -0.50626 | -3.0022 | -166.8541 | 166.9365 | 0.2237 | -0.45917 | -0.4003 | 0.71818 | 166.8541 | 0.45917 | 2.1662
0.0251 | 0.2237 | -3.6262 | -163.2161 | 163.3103 | 0.1766 | -0.48271 | -0.4003 | 0.72996 | 163.2161 | 0.48271 | 2.3336
0.0269 | 0.78882 | -4.5681 | -160.8496 | 160.9791 | 0.11774 | -0.49449 | -0.41207 | 0.72996 | 160.8496 | 0.49449 | 2.5009
0.0287 | 1.1067 | -5.1332 | -160.2845 | 160.4375 | 0.070641 | -0.49449 | -0.41207 | 0.71818 | 160.2845 | 0.49449 | 2.6683
0.0305 | 1.3775 | -5.5453 | -161.2735 | 161.4383 | 0.023547 | -0.50626 | -0.42385 | 0.72996 | 161.2735 | 0.50626 | 2.8356
0.0323 | 1.4364 | -5.3452 | -163.3574 | 163.534 | 0.011774 | -0.51803 | -0.43562 | 0.72996 | 163.3574 | 0.51803 | 3.003
0.0341 | 1.3775 | -4.7683 | -166.7364 | 166.9012 | 0.011774 | -0.51803 | -0.43562 | 0.7535 | 166.7364 | 0.51803 | 3.1703
0.0359 | 0.9772 | -3.4732 | -170.7982 | 170.9277 | 0.035321 | -0.52981 | -0.44739 | 0.76528 | 170.7982 | 0.52981 | 3.3377
0.0376 | 0.54158 | -2.0604 | -174.3185 | 174.4127 | 0.070641 | -0.52981 | -0.44739 | 0.78882 | 174.3185 | 0.52981 | 3.4957
0.0394 | 0.27079 | -0.2237 | -176.1316 | 176.2258 | 0.10596 | -0.55335 | -0.45917 | 0.81237 | 176.1316 | 0.55335 | 3.6631
0.0412 | 0.070641 | 1.6483 | -177.2737 | 177.3678 | 0.14128 | -0.56513 | -0.45917 | 0.83592 | 177.2737 | 0.56513 | 3.8304
0.043 | 0.047094 | 3.273 | -176.6732 | 176.8027 | 0.16483 | -0.58868 | -0.45917 | 0.85947 | 176.6732 | 0.58868 | 3.9978
0.0448 | 0.14128 | 4.4857 | -174.9896 | 175.1426 | 0.1766 | -0.624 | -0.45917 | 0.88301 | 174.9896 | 0.624 | 4.1651
0.0466 | 0.31788 | 5.7337 | -172.3641 | 172.5642 | 0.18838 | -0.65932 | -0.45917 | 0.90656 | 172.3641 | 0.65932 | 4.3325
0.0484 | 0.61222 | 6.4519 | -168.832 | 169.0793 | 0.1766 | -0.68286 | -0.45917 | 0.91833 | 168.832 | 0.68286 | 4.4998
0.0502 | 1.1185 | 6.8404 | -164.3581 | 164.6642 | 0.14128 | -0.70641 | -0.44739 | 0.93011 | 164.3581 | 0.70641 | 4.6672
0.052 | 1.6247 | 6.9817 | -160.4846 | 160.8261 | 0.10596 | -0.74173 | -0.44739 | 0.94188 | 160.4846 | 0.74173 | 4.8345
0.0538 | 2.0957 | 6.5225 | -157.7179 | 158.0593 | 0.047094 | -0.76528 | -0.45917 | 0.96543 | 157.7179 | 0.76528 | 5.0019
0.0556 | 2.5313 | 5.9456 | -155.8341 | 156.1755 | -0.011774 | -0.8006 | -0.45917 | 0.9772 | 155.8341 | 0.8006 | 5.1692
0.0574 | 2.8021 | 5.4865 | -155.9989 | 156.3521 | -0.070641 | -0.82415 | -0.47094 | 1.0125 | 155.9989 | 0.82415 | 5.3366
0.0592 | 2.9081 | 5.2628 | -157.0232 | 157.3882 | -0.11774 | -0.83592 | -0.47094 | 1.0243 | 157.0232 | 0.83592 | 5.5039
0.061 | 2.9551 | 5.5218 | -159.1307 | 159.5192 | -0.15306 | -0.85947 | -0.48271 | 1.0596 | 159.1307 | 0.85947 | 5.6713
0.0627 | 2.7668 | 6.6285 | -162.286 | 162.7216 | -0.15306 | -0.87124 | -0.49449 | 1.0832 | 162.286 | 0.87124 | 5.8293
0.0645 | 2.5431 | 8.4298 | -165.4766 | 166.0535 | -0.12951 | -0.89479 | -0.50626 | 1.1185 | 165.4766 | 0.89479 | 5.9967
0.0663 | 2.3194 | 10.2429 | -167.0896 | 167.7842 | -0.11774 | -0.93011 | -0.51803 | 1.142 | 167.0896 | 0.93011 | 6.164
0.0681 | 2.1428 | 11.9266 | -167.8431 | 168.6201 | -0.094188 | -0.95365 | -0.51803 | 1.1774 | 167.8431 | 0.95365 | 6.3313
0.0699 | 2.1781 | 13.5278 | -166.7481 | 167.6429 | -0.070641 | -0.98897 | -0.52981 | 1.2127 | 166.7481 | 0.98897 | 6.4987
0.0717 | 2.3076 | 15.023 | -165.0881 | 166.1124 | -0.058868 | -1.0243 | -0.54158 | 1.248 | 165.0881 | 1.0243 | 6.666
0.0735 | 2.5666 | 15.9766 | -162.4273 | 163.5575 | -0.070641 | -1.0714 | -0.54158 | 1.2833 | 162.4273 | 1.0714 | 6.8334
0.0753 | 2.9905 | 16.789 | -158.8128 | 160.0137 | -0.070641 | -1.1067 | -0.55335 | 1.3186 | 158.8128 | 1.1067 | 7.0007
0.0771 | 3.4967 | 16.7066 | -154.7038 | 155.9165 | -0.11774 | -1.142 | -0.56513 | 1.3422 | 154.7038 | 1.142 | 7.1681
0.0789 | 3.9912 | 16.53 | -150.8068 | 152.043 | -0.15306 | -1.1774 | -0.5769 | 1.354 | 150.8068 | 1.1774 | 7.3354
0.0807 | 4.2502 | 16.1297 | -149.5588 | 150.7715 | -0.2237 | -1.2127 | -0.5769 | 1.3893 | 149.5588 | 1.2127 | 7.5028
0.0825 | 4.368 | 15.6117 | -149.4058 | 150.5831 | -0.28256 | -1.2362 | -0.58868 | 1.4246 | 149.4058 | 1.2362 | 7.6701
0.0843 | 4.4504 | 15.7883 | -150.6538 | 151.7958 | -0.32966 | -1.2833 | -0.60045 | 1.4717 | 150.6538 | 1.2833 | 7.8375
0.086 | 4.4386 | 15.7412 | -153.4676 | 154.5979 | -0.36498 | -1.3069 | -0.61222 | 1.507 | 153.4676 | 1.3069 | 7.9955
0.0878 | 4.2738 | 16.3063 | -157.4353 | 158.6126 | -0.38853 | -1.3186 | -0.624 | 1.5423 | 157.4353 | 1.3186 | 8.1629
0.0896 | 4.0148 | 17.6249 | -161.7797 | 163.0748 | -0.36498 | -1.3422 | -0.63577 | 1.5659 | 161.7797 | 1.3422 | 8.3302
0.0914 | 3.8146 | 19.1673 | -165.2058 | 166.6657 | -0.34143 | -1.3657 | -0.64754 | 1.5776 | 165.2058 | 1.3657 | 8.4976
0.0932 | 3.8382 | 20.8156 | -166.1595 | 167.8313 | -0.32966 | -1.4128 | -0.65932 | 1.6247 | 166.1595 | 1.4128 | 8.6649
0.095 | 3.8735 | 22.4168 | -165.9829 | 167.8666 | -0.31788 | -1.4599 | -0.67109 | 1.6718 | 165.9829 | 1.4599 | 8.8323
0.0968 | 3.9912 | 23.7707 | -165.0057 | 167.1013 | -0.31788 | -1.507 | -0.69464 | 1.7189 | 165.0057 | 1.507 | 8.9996
0.0986 | 4.1325 | 24.7126 | -163.1101 | 165.3706 | -0.32966 | -1.5541 | -0.70641 | 1.766 | 163.1101 | 1.5541 | 9.167
0.1004 | 4.3562 | 25.2071 | -160.3904 | 162.7451 | -0.32966 | -1.5894 | -0.71818 | 1.7896 | 160.3904 | 1.5894 | 9.3343
0.1022 | 4.5328 | 24.9951 | -157.4588 | 159.8253 | -0.36498 | -1.6012 | -0.72996 | 1.8249 | 157.4588 | 1.6012 | 9.5017
0.104 | 4.6152 | 24.5713 | -156.1284 | 158.4478 | -0.41207 | -1.6247 | -0.74173 | 1.8602 | 156.1284 | 1.6247 | 9.669
0.1058 | 4.6741 | 23.8767 | -155.6575 | 157.8945 | -0.48271 | -1.6483 | -0.7535 | 1.9073 | 155.6575 | 1.6483 | 9.8364
0.1076 | 4.6152 | 23.4881 | -156.8819 | 159.06 | -0.52981 | -1.6718 | -0.77705 | 1.9426 | 156.8819 | 1.6718 | 10.0037
0.1094 | 4.4033 | 23.1703 | -159.6487 | 161.7562 | -0.5769 | -1.6836 | -0.78882 | 1.9779 | 159.6487 | 1.6836 | 10.1711
0.1111 | 4.1678 | 23.2409 | -163.5928 | 165.7238 | -0.61222 | -1.7072 | -0.81237 | 2.0133 | 163.5928 | 1.7072 | 10.3291
0.1129 | 3.8853 | 23.4528 | -168.7025 | 170.8218 | -0.63577 | -1.7307 | -0.82415 | 2.0486 | 168.7025 | 1.7307 | 10.4965
0.1147 | 3.5321 | 23.9355 | -174.1301 | 176.3082 | -0.624 | -1.7543 | -0.83592 | 2.0721 | 174.1301 | 1.7543 | 10.6638
0.1165 | 3.2613 | 25.1011 | -177.827 | 180.1111 | -0.61222 | -1.7896 | -0.85947 | 2.1075 | 177.827 | 1.7896 | 10.8312
0.1183 | 3.0611 | 26.5493 | -180.0051 | 182.4658 | -0.58868 | -1.8249 | -0.87124 | 2.1428 | 180.0051 | 1.8249 | 10.9985
0.1201 | 3.014 | 28.0916 | -180.8175 | 183.5018 | -0.56513 | -1.872 | -0.89479 | 2.1781 | 180.8175 | 1.872 | 11.1659
0.1219 | 3.0258 | 29.3749 | -180.5114 | 183.3841 | -0.55335 | -1.9191 | -0.91833 | 2.2252 | 180.5114 | 1.9191 | 11.3332
0.1237 | 3.0493 | 30.2815 | -179.334 | 182.3834 | -0.55335 | -1.9544 | -0.94188 | 2.2605 | 179.334 | 1.9544 | 11.5006
0.1255 | 3.0964 | 30.6111 | -177.0146 | 180.1699 | -0.5769 | -1.9897 | -0.96543 | 2.3076 | 177.0146 | 1.9897 | 11.6679
0.1273 | 3.0258 | 30.4463 | -174.7423 | 177.9094 | -0.60045 | -2.0133 | -0.9772 | 2.3429 | 174.7423 | 2.0133 | 11.8353
0.1291 | 2.8374 | 29.9753 | -173.2118 | 176.32 | -0.64754 | -2.025 | -1.0007 | 2.3665 | 173.2118 | 2.025 | 12.0026
0.1309 | 2.5078 | 29.1394 | -173.306 | 176.2258 | -0.70641 | -2.025 | -1.0125 | 2.4018 | 173.306 | 2.025 | 12.17
0.1327 | 2.0839 | 28.6096 | -174.4716 | 177.2501 | -0.7535 | -2.0368 | -1.0361 | 2.4253 | 174.4716 | 2.0368 | 12.3373
0.1345 | 1.5659 | 28.4801 | -176.7674 | 179.4517 | -0.77705 | -2.0486 | -1.0478 | 2.4489 | 176.7674 | 2.0486 | 12.5046
0.1362 | 1.0125 | 28.7156 | -180.1464 | 182.7954 | -0.8006 | -2.0604 | -1.0714 | 2.4842 | 180.1464 | 2.0604 | 12.6627
0.138 | 0.45917 | 29.6575 | -182.8896 | 185.6564 | -0.8006 | -2.0957 | -1.0949 | 2.5195 | 182.8896 | 2.0957 | 12.83
0.1398 | -0.070641 | 30.7524 | -185.1148 | 188.0346 | -0.78882 | -2.131 | -1.1067 | 2.5431 | 185.1148 | 2.131 | 12.9974
0.1416 | -0.44739 | 31.9886 | -186.1626 | 189.3062 | -0.77705 | -2.1663 | -1.1185 | 2.5784 | 186.1626 | 2.1663 | 13.1647
0.1434 | -0.624 | 33.1306 | -185.7741 | 189.1413 | -0.7535 | -2.2134 | -1.142 | 2.6137 | 185.7741 | 2.2134 | 13.3321
0.1452 | -0.7535 | 34.0725 | -184.479 | 188.07 | -0.74173 | -2.2605 | -1.1538 | 2.6608 | 184.479 | 2.2605 | 13.4994
0.147 | -0.91833 | 34.7083 | -182.2656 | 186.0331 | -0.74173 | -2.2841 | -1.1774 | 2.6844 | 182.2656 | 2.2841 | 13.6668
0.1488 | -0.84769 | 35.0497 | -178.9337 | 182.8072 | -0.74173 | -2.3194 | -1.1891 | 2.7315 | 178.9337 | 2.3194 | 13.8341
0.1506 | -0.64754 | 34.7907 | -174.9543 | 178.8278 | -0.7535 | -2.3429 | -1.2009 | 2.755 | 174.9543 | 2.3429 | 14.0015
0.1524 | -0.83592 | 33.8371 | -172.3406 | 176.0963 | -0.81237 | -2.3547 | -1.2127 | 2.7785 | 172.3406 | 2.3547 | 14.1688
0.1542 | -1.354 | 33.0365 | -170.8806 | 174.554 | -0.85947 | -2.3429 | -1.2244 | 2.7903 | 170.8806 | 2.3429 | 14.3362
0.156 | -1.9779 | 31.9533 | -171.3398 | 174.8719 | -0.91833 | -2.3312 | -1.2362 | 2.8021 | 171.3398 | 2.3312 | 14.5035
0.1578 | -2.6019 | 30.8937 | -172.682 | 176.061 | -0.96543 | -2.3194 | -1.248 | 2.8139 | 172.682 | 2.3194 | 14.6709
0.1595 | -3.1906 | 30.1048 | -175.4017 | 178.6512 | -1.0125 | -2.3076 | -1.2598 | 2.8256 | 175.4017 | 2.3076 | 14.8289
0.1613 | -3.6851 | 29.6221 | -178.9573 | 182.1126 | -1.0478 | -2.3194 | -1.2715 | 2.8492 | 178.9573 | 2.3194 | 14.9963
0.1631 | -3.9912 | 29.7752 | -183.337 | 186.5041 | -1.0714 | -2.3429 | -1.2715 | 2.8845 | 183.337 | 2.3429 | 15.1636
0.1649 | -4.4857 | 30.411 | -188.07 | 191.3077 | -1.0596 | -2.3429 | -1.2833 | 2.8845 | 188.07 | 2.3429 | 15.331
0.1667 | -4.8389 | 31.3293 | -191.9434 | 195.2636 | -1.0478 | -2.3547 | -1.2951 | 2.8963 | 191.9434 | 2.3547 | 15.4983
0.1685 | -5.0862 | 32.4007 | -195.2871 | 198.725 | -1.0243 | -2.3782 | -1.3069 | 2.9198 | 195.2871 | 2.3782 | 15.6657
0.1703 | -5.1332 | 33.8017 | -198.101 | 201.7508 | -1.0007 | -2.4136 | -1.3186 | 2.9316 | 198.101 | 2.4136 | 15.833
0.1721 | -5.1921 | 35.1204 | -199.7139 | 203.5286 | -0.96543 | -2.4489 | -1.3304 | 2.9551 | 199.7139 | 2.4489 | 16.0004
0.1739 | -5.0979 | 36.5332 | -200.1613 | 204.235 | -0.93011 | -2.4724 | -1.3304 | 2.9787 | 200.1613 | 2.4724 | 16.1677
0.1757 | -4.9331 | 37.7341 | -199.8552 | 204.1408 | -0.90656 | -2.5078 | -1.3422 | 2.9905 | 199.8552 | 2.5078 | 16.3351
0.1775 | -4.6388 | 38.8408 | -198.6308 | 203.14 | -0.88301 | -2.5313 | -1.354 | 3.014 | 198.6308 | 2.5313 | 16.5024
0.1793 | -4.3326 | 39.7709 | -196.4645 | 201.1974 | -0.85947 | -2.5549 | -1.354 | 3.0376 | 196.4645 | 2.5549 | 16.6698
0.1811 | -4.0148 | 40.2065 | -193.5682 | 198.4424 | -0.84769 | -2.5784 | -1.3657 | 3.0493 | 193.5682 | 2.5784 | 16.8371
0.1829 | -3.6498 | 40.3596 | -190.2716 | 195.2047 | -0.84769 | -2.5902 | -1.3657 | 3.0611 | 190.2716 | 2.5902 | 17.0045
0.1846 | -3.3319 | 40.1123 | -187.0928 | 192.0259 | -0.84769 | -2.5902 | -1.3657 | 3.0611 | 187.0928 | 2.5902 | 17.1625
0.1864 | -3.2377 | 39.6178 | -185.5857 | 190.5188 | -0.87124 | -2.5666 | -1.3657 | 3.0493 | 185.5857 | 2.5666 | 17.3299
0.1882 | -3.3201 | 38.4405 | -184.8322 | 189.6005 | -0.91833 | -2.5313 | -1.3657 | 3.0376 | 184.8322 | 2.5313 | 17.4972
0.19 | -3.5791 | 37.1925 | -185.6093 | 190.1656 | -0.9772 | -2.496 | -1.3657 | 3.0258 | 185.6093 | 2.496 | 17.6646
0.1918 | -3.9206 | 35.9681 | -187.3518 | 191.6373 | -1.0243 | -2.4724 | -1.3657 | 3.014 | 187.3518 | 2.4724 | 17.8319
0.1936 | -4.2385 | 34.9673 | -190.5777 | 194.6749 | -1.0714 | -2.4371 | -1.3657 | 3.0022 | 190.5777 | 2.4371 | 17.9993
0.1954 | -4.4857 | 34.6965 | -194.4041 | 198.36 | -1.1067 | -2.4136 | -1.3657 | 3.0022 | 194.4041 | 2.4136 | 18.1666
0.1972 | -4.4739 | 35.4382 | -199.5844 | 203.5757 | -1.0949 | -2.4253 | -1.3657 | 3.014 | 199.5844 | 2.4253 | 18.334
0.199 | -4.3326 | 36.4979 | -204.4469 | 208.5323 | -1.0832 | -2.4489 | -1.354 | 3.014 | 204.4469 | 2.4489 | 18.5013
0.2008 | -4.1325 | 37.4633 | -207.5904 | 211.8171 | -1.0596 | -2.4724 | -1.354 | 3.0258 | 207.5904 | 2.4724 | 18.6687
0.2026 | -4.0972 | 37.8047 | -210.5809 | 214.9018 | -1.0478 | -2.4607 | -1.3422 | 3.0022 | 210.5809 | 2.4607 | 18.836
0.2044 | -4.0265 | 38.2992 | -212.7825 | 217.2211 | -1.0243 | -2.4607 | -1.3304 | 2.9905 | 212.7825 | 2.4607 | 19.0033
0.2062 | -3.7793 | 39.0056 | -213.9599 | 218.5162 | -0.98897 | -2.4724 | -1.3186 | 2.9787 | 213.9599 | 2.4724 | 19.1707
0.208 | -3.6145 | 38.9468 | -214.5957 | 219.1167 | -0.96543 | -2.4371 | -1.3069 | 2.9434 | 214.5957 | 2.4371 | 19.338
0.2097 | -3.4614 | 38.8055 | -215.1961 | 219.6583 | -0.94188 | -2.4136 | -1.2833 | 2.9081 | 215.1961 | 2.4136 | 19.4961
0.2115 | -3.8735 | 38.7231 | -216.1144 | 220.5295 | -0.90656 | -2.3429 | -1.2715 | 2.8256 | 216.1144 | 2.3429 | 19.6634
0.2133 | -4.4033 | 38.0284 | -219.0696 | 223.3081 | -0.88301 | -2.2605 | -1.248 | 2.7432 | 219.0696 | 2.2605 | 19.8308
0.2151 | -4.992 | 36.7098 | -223.4964 | 227.4759 | -0.88301 | -2.1428 | -1.2244 | 2.6373 | 223.4964 | 2.1428 | 19.9981
0.2169 | -5.7219 | 35.2734 | -229.289 | 232.9506 | -0.87124 | -2.0133 | -1.2127 | 2.5313 | 229.289 | 2.0133 | 20.1655
0.2187 | -6.4283 | 33.9312 | -236.8358 | 240.1677 | -0.87124 | -1.8955 | -1.1891 | 2.4253 | 236.8358 | 1.8955 | 20.3328
0.2205 | -7.382 | 32.4713 | -246.0309 | 249.1391 | -0.84769 | -1.766 | -1.1774 | 2.3076 | 246.0309 | 1.766 | 20.5002
0.2223 | -8.4887 | 30.8819 | -255.9207 | 258.8169 | -0.82415 | -1.613 | -1.1774 | 2.1899 | 255.9207 | 1.613 | 20.6675
0.2241 | -9.6425 | 29.057 | -267.7059 | 270.3785 | -0.78882 | -1.4364 | -1.1656 | 2.0604 | 267.7059 | 1.4364 | 20.8349
0.2259 | -11.02 | 27.0791 | -281.4927 | 283.918 | -0.7535 | -1.2362 | -1.1538 | 1.9191 | 281.4927 | 1.2362 | 21.0022
0.2277 | -12.6447 | 24.9951 | -297.3987 | 299.6121 | -0.68286 | -1.0125 | -1.1538 | 1.766 | 297.3987 | 1.0125 | 21.1696
0.2295 | -14.4696 | 22.8053 | -315.4593 | 317.5196 | -0.61222 | -0.76528 | -1.1538 | 1.6365 | 315.4593 | 0.76528 | 21.3369
0.2313 | -16.3299 | 20.4153 | -335.5213 | 337.4522 | -0.52981 | -0.50626 | -1.1774 | 1.5541 | 335.5213 | 0.50626 | 21.5043
0.2331 | -18.3313 | 17.672 | -357.2788 | 359.1507 | -0.43562 | -0.21192 | -1.1891 | 1.5306 | 357.2788 | 0.21192 | 21.6716
0.2348 | -20.8273 | 14.4225 | -382.0502 | 383.8516 | -0.32966 | 0.11774 | -1.2244 | 1.5659 | 382.0502 | -0.11774 | 21.8297
0.2366 | -23.5235 | 10.9023 | -409.3765 | 411.0719 | -0.2237 | 0.47094 | -1.2598 | 1.6365 | 409.3765 | -0.47094 | 21.997
0.2384 | -26.2902 | 7.1936 | -438.7279 | 440.3761 | -0.10596 | 0.83592 | -1.2951 | 1.7778 | 438.7279 | -0.83592 | 22.1644
0.2402 | -29.2336 | 3.1435 | -471.3169 | 472.9888 | 0.023547 | 1.248 | -1.354 | 2.0133 | 471.3169 | -1.248 | 22.3317
0.242 | -32.2594 | -1.0832 | -506.0958 | 507.8148 | 0.15306 | 1.6718 | -1.4246 | 2.3194 | 506.0958 | -1.6718 | 22.4991
0.2438 | -35.2852 | -5.6748 | -544.348 | 546.1493 | 0.30611 | 2.1428 | -1.4952 | 2.7197 | 544.348 | -2.1428 | 22.6664
0.2456 | -38.3228 | -10.5726 | -585.5317 | 587.4743 | 0.47094 | 2.6255 | -1.5894 | 3.1671 | 585.5317 | -2.6255 | 22.8338
0.2474 | -41.2779 | -15.7059 | -630.377 | 632.4962 | 0.64754 | 3.1435 | -1.6954 | 3.6733 | 630.377 | -3.1435 | 23.0011
0.2492 | -44.1506 | -20.9215 | -678.5424 | 680.8264 | 0.83592 | 3.6733 | -1.8131 | 4.2267 | 678.5424 | -3.6733 | 23.1685
0.251 | -46.9998 | -26.2431 | -730.7108 | 733.195 | 1.0478 | 4.2502 | -1.9426 | 4.8154 | 730.7108 | -4.2502 | 23.3358
0.2528 | -49.8019 | -31.7531 | -785.8932 | 788.5658 | 1.2715 | 4.8389 | -2.0839 | 5.4511 | 785.8932 | -4.8389 | 23.5032
0.2546 | -52.345 | -37.4044 | -848.1161 | 850.9889 | 1.5188 | 5.4629 | -2.2487 | 6.134 | 848.1161 | -5.4629 | 23.6705
0.2564 | -54.476 | -42.7614 | -910.3744 | 913.4238 | 1.7778 | 6.0634 | -2.4136 | 6.7933 | 910.3744 | -6.0634 | 23.8379
0.2581 | -56.454 | -48.7423 | -974.7049 | 977.9779 | 2.025 | 6.7109 | -2.6019 | 7.4879 | 974.7049 | -6.7109 | 23.9959
0.2599 | -57.8903 | -54.8057 | -1038.7174 | 1042.2024 | 2.2723 | 7.3349 | -2.7785 | 8.1826 | 1038.7174 | -7.3349 | 24.1633
0.2617 | -58.6674 | -61.3517 | -1104.9434 | 1108.6167 | 2.5195 | 7.9707 | -2.9669 | 8.889 | 1104.9434 | -7.9707 | 24.3306
0.2635 | -59.2325 | -67.9567 | -1168.6733 | 1172.5468 | 2.755 | 8.5829 | -3.1553 | 9.5601 | 1168.6733 | -8.5829 | 24.498
0.2653 | -59.3502 | -74.5851 | -1232.5681 | 1236.63 | 2.9905 | 9.1833 | -3.3437 | 10.2312 | 1232.5681 | -9.1833 | 24.6653
0.2671 | -59.2325 | -80.2953 | -1297.6167 | 1301.8081 | 3.2495 | 9.7956 | -3.5321 | 10.9258 | 1297.6167 | -9.7956 | 24.8327
0.2689 | -58.8557 | -86.1232 | -1361.4762 | 1365.8089 | 3.4967 | 10.3489 | -3.7204 | 11.5498 | 1361.4762 | -10.3489 | 25
0.2707 | -58.0434 | -91.2918 | -1420.8971 | 1425.3475 | 3.7204 | 10.8081 | -3.8853 | 12.0796 | 1420.8971 | -10.8081 | 25.1673
0.2725 | -56.9838 | -96.0953 | -1477.3864 | 1481.9192 | 3.9324 | 11.1966 | -4.0501 | 12.5506 | 1477.3864 | -11.1966 | 25.3347
0.2743 | -55.4061 | -100.5928 | -1533.5695 | 1538.173 | 4.1325 | 11.5263 | -4.2031 | 12.9509 | 1533.5695 | -11.5263 | 25.502
0.2761 | -53.8873 | -104.69 | -1587.9043 | 1592.5548 | 4.3326 | 11.8441 | -4.3562 | 13.3512 | 1587.9043 | -11.8441 | 25.6694
0.2779 | -52.3332 | -108.4811 | -1639.5075 | 1644.2052 | 4.521 | 12.1738 | -4.4975 | 13.7515 | 1639.5075 | -12.1738 | 25.8367
0.2797 | -50.6378 | -111.7776 | -1687.1314 | 1691.8643 | 4.7094 | 12.4681 | -4.627 | 14.1164 | 1687.1314 | -12.4681 | 26.0041
0.2815 | -48.6717 | -114.6033 | -1731.1289 | 1735.8737 | 4.8625 | 12.6918 | -4.7329 | 14.399 | 1731.1289 | -12.6918 | 26.1714
0.2832 | -46.5053 | -117.0875 | -1772.1949 | 1776.9514 | 5.0037 | 12.8684 | -4.8389 | 14.6462 | 1772.1949 | -12.8684 | 26.3295
0.285 | -44.3037 | -119.1596 | -1810.1174 | 1814.8739 | 5.1332 | 12.9979 | -4.9331 | 14.8346 | 1810.1174 | -12.9979 | 26.4968
0.2868 | -42.0785 | -120.5136 | -1842.8359 | 1847.5571 | 5.2392 | 13.0568 | -5.0155 | 14.9524 | 1842.8359 | -13.0568 | 26.6642
0.2886 | -39.8415 | -121.6792 | -1873.035 | 1877.7326 | 5.3334 | 13.1039 | -5.0862 | 15.0465 | 1873.035 | -13.1039 | 26.8315
0.2904 | -37.6163 | -122.7623 | -1900.6556 | 1905.3179 | 5.4158 | 13.1275 | -5.1568 | 15.129 | 1900.6556 | -13.1275 | 26.9989
0.2922 | -35.4971 | -123.2333 | -1922.7427 | 1927.3462 | 5.4865 | 13.1275 | -5.1921 | 15.176 | 1922.7427 | -13.1275 | 27.1662
0.294 | -33.425 | -122.8801 | -1942.2161 | 1946.7371 | 5.5453 | 13.0568 | -5.2392 | 15.1407 | 1942.2161 | -13.0568 | 27.3336
0.2958 | -31.3646 | -122.4091 | -1960.2884 | 1964.727 | 5.6042 | 12.9979 | -5.2628 | 15.129 | 1960.2884 | -12.9979 | 27.5009
0.2976 | -29.3043 | -121.0552 | -1976.6654 | 1980.9509 | 5.6631 | 12.892 | -5.2863 | 15.0701 | 1976.6654 | -12.892 | 27.6683
0.2994 | -27.3734 | -119.2891 | -1992.5125 | 1996.6332 | 5.7219 | 12.786 | -5.2981 | 15.0112 | 1992.5125 | -12.786 | 27.8356
0.3012 | -25.0305 | -116.7343 | -2006.5347 | 2010.4671 | 5.7572 | 12.5506 | -5.3099 | 14.8464 | 2006.5347 | -12.5506 | 28.003
0.303 | -22.7935 | -114.0264 | -2018.3907 | 2022.1464 | 5.7808 | 12.3269 | -5.3216 | 14.6698 | 2018.3907 | -12.3269 | 28.1703
0.3048 | -20.6625 | -111.2596 | -2027.7388 | 2031.3415 | 5.7808 | 12.0796 | -5.3216 | 14.4696 | 2027.7388 | -12.0796 | 28.3377
0.3066 | -18.5786 | -108.3162 | -2035.0502 | 2038.4998 | 5.769 | 11.8206 | -5.3099 | 14.2577 | 2035.0502 | -11.8206 | 28.505
0.3083 | -16.4711 | -104.6547 | -2040.8663 | 2044.1746 | 5.7455 | 11.4909 | -5.2981 | 13.9869 | 2040.8663 | -11.4909 | 28.6631
0.3101 | -14.505 | -101.1933 | -2044.9046 | 2048.0363 | 5.7337 | 11.2201 | -5.2745 | 13.7632 | 2044.9046 | -11.2201 | 28.8304
0.3119 | -12.633 | -97.4611 | -2048.4602 | 2051.4624 | 5.6866 | 10.7963 | -5.2628 | 13.4336 | 2048.4602 | -10.7963 | 28.9978
0.3137 | -10.914 | -93.8466 | -2050.7089 | 2053.5817 | 5.6631 | 10.4902 | -5.2392 | 13.1981 | 2050.7089 | -10.4902 | 29.1651
0.3155 | -9.0891 | -89.8672 | -2052.7222 | 2055.4772 | 5.6277 | 10.1134 | -5.2157 | 12.9744 | 2052.7222 | -10.1134 | 29.3325
0.3173 | -7.3702 | -85.6051 | -2054.6648 | 2057.361 | 5.5806 | 9.6896 | -5.1921 | 12.786 | 2054.6648 | -9.6896 | 29.4998
0.3191 | -5.7572 | -81.2607 | -2056.1954 | 2058.868 | 5.5335 | 9.3246 | -5.1686 | 12.6683 | 2056.1954 | -9.3246 | 29.6672
0.3209 | -4.109 | -77.2342 | -2057.2668 | 2059.8923 | 5.51 | 9.0185 | -5.1332 | 12.6094 | 2057.2668 | -9.0185 | 29.8345
0.3227 | -2.2723 | -73.031 | -2057.5022 | 2060.01 | 5.4747 | 8.7124 | -5.0979 | 12.4564 | 2057.5022 | -8.7124 | 30.0019
0.3245 | -0.624 | -68.8397 | -2057.3845 | 2059.8099 | 5.4511 | 8.4063 | -5.0508 | 12.3622 | 2057.3845 | -8.4063 | 30.1692
0.3263 | 0.82415 | -64.6248 | -2056.3955 | 2058.7856 | 5.404 | 8.0884 | -5.0155 | 12.2327 | 2056.3955 | -8.0884 | 30.3366
0.3281 | 2.0486 | -60.7513 | -2054.7001 | 2057.0666 | 5.3569 | 7.7823 | -4.9802 | 12.0796 | 2054.7001 | -7.7823 | 30.5039
0.3299 | 3.5085 | -56.1478 | -2052.0747 | 2054.3352 | 5.2981 | 7.4526 | -4.9331 | 11.8559 | 2052.0747 | -7.4526 | 30.6713
0.3316 | 4.7447 | -51.8505 | -2048.9665 | 2051.1446 | 5.251 | 7.1347 | -4.886 | 11.6322 | 2048.9665 | -7.1347 | 30.8293
0.3334 | 6.4283 | -47.3648 | -2043.9156 | 2046.0113 | 5.1921 | 6.8051 | -4.8271 | 11.4203 | 2043.9156 | -6.8051 | 30.9967
0.3352 | 7.8176 | -42.4317 | -2039.1709 | 2041.1842 | 5.1215 | 6.4637 | -4.78 | 11.1731 | 2039.1709 | -6.4637 | 31.164
0.337 | 9.2304 | -37.8871 | -2034.1318 | 2036.1569 | 5.0508 | 6.1222 | -4.7212 | 11.02 | 2034.1318 | -6.1222 | 31.3313
0.3388 | 10.5491 | -33.2719 | -2028.6218 | 2030.6704 | 4.9802 | 5.7808 | -4.6741 | 10.8669 | 2028.6218 | -5.7808 | 31.4987
0.3406 | 11.7029 | -29.0688 | -2022.6644 | 2024.7601 | 4.9213 | 5.4865 | -4.6152 | 10.7374 | 2022.6644 | -5.4865 | 31.666
0.3424 | 12.6918 | -25.419 | -2016.4127 | 2018.5673 | 4.8625 | 5.2274 | -4.5563 | 10.655 | 2016.4127 | -5.2274 | 31.8334
0.3442 | 13.9398 | -21.0393 | -2008.7717 | 2010.8556 | 4.8036 | 4.9449 | -4.4975 | 10.4784 | 2008.7717 | -4.9449 | 32.0007
0.346 | 15.1407 | -15.7529 | -2001.7547 | 2003.8151 | 4.7094 | 4.6034 | -4.4386 | 10.2194 | 2001.7547 | -4.6034 | 32.1681
0.3478 | 16.3652 | -10.9494 | -1995.2086 | 1997.3043 | 4.6152 | 4.2973 | -4.3797 | 10.0193 | 1995.2086 | -4.2973 | 32.3354
0.3496 | 17.3071 | -7.2054 | -1985.3307 | 1987.4146 | 4.5563 | 4.0854 | -4.3091 | 9.9133 | 1985.3307 | -4.0854 | 32.5028
0.3514 | 18.1312 | -2.6608 | -1973.6043 | 1975.6646 | 4.4739 | 3.8264 | -4.2385 | 9.7367 | 1973.6043 | -3.8264 | 32.6701
0.3532 | 19.4263 | 2.39 | -1963.95 | 1965.9868 | 4.368 | 3.5438 | -4.1796 | 9.513 | 1963.95 | -3.5438 | 32.8375
0.355 | 20.7685 | 7.4526 | -1955.9322 | 1957.969 | 4.2856 | 3.2495 | -4.109 | 9.3128 | 1955.9322 | -3.2495 | 33.0048
0.3567 | 21.8987 | 12.009 | -1948.9623 | 1951.1169 | 4.1796 | 2.9905 | -4.0501 | 9.1716 | 1948.9623 | -2.9905 | 33.1629
0.3585 | 22.8524 | 16.3652 | -1943.3346 | 1945.654 | 4.0854 | 2.755 | -3.9912 | 9.0891 | 1943.3346 | -2.755 | 33.3302
0.3603 | 23.8649 | 20.6507 | -1936.5059 | 1938.9313 | 4.0148 | 2.496 | -3.9324 | 9.0067 | 1936.5059 | -2.496 | 33.4976
0.3621 | 24.9481 | 24.7479 | -1927.1696 | 1929.6773 | 3.9324 | 2.2252 | -3.8617 | 8.9596 | 1927.1696 | -2.2252 | 33.6649
0.3639 | 26.0077 | 28.7509 | -1917.48 | 1920.0937 | 3.8499 | 1.9544 | -3.7911 | 8.9243 | 1917.48 | -1.9544 | 33.8323
0.3657 | 27.0908 | 32.5891 | -1907.1899 | 1909.9096 | 3.7675 | 1.6954 | -3.7204 | 8.8772 | 1907.1899 | -1.6954 | 33.9996
0.3675 | 28.798 | 36.7098 | -1893.0264 | 1895.6754 | 3.6733 | 1.4128 | -3.6262 | 8.6771 | 1893.0264 | -1.4128 | 34.167
0.3693 | 30.5993 | 41.1013 | -1877.6855 | 1880.2168 | 3.5791 | 1.0949 | -3.5321 | 8.4181 | 1877.6855 | -1.0949 | 34.3343
0.3711 | 32.4596 | 45.6223 | -1863.1335 | 1865.5353 | 3.4496 | 0.72996 | -3.4261 | 8.1002 | 1863.1335 | -0.72996 | 34.5017
0.3729 | 34.1903 | 50.3788 | -1849.5586 | 1851.8662 | 3.3201 | 0.35321 | -3.3319 | 7.747 | 1849.5586 | -0.35321 | 34.669
0.3747 | 35.5913 | 54.7821 | -1838.727 | 1841.0346 | 3.2024 | 0.035321 | -3.2495 | 7.4997 | 1838.727 | -0.035321 | 34.8364
0.3765 | 36.5214 | 57.5018 | -1830.5444 | 1832.9815 | 3.1318 | -0.1766 | -3.1906 | 7.5115 | 1830.5444 | 0.1766 | 35.0037
0.3783 | 37.581 | 59.5975 | -1822.197 | 1824.7754 | 3.0493 | -0.37675 | -3.1318 | 7.5468 | 1822.197 | 0.37675 | 35.1711
0.3801 | 39.1234 | 62.6704 | -1812.6369 | 1815.2742 | 2.9669 | -0.64754 | -3.0611 | 7.4173 | 1812.6369 | 0.64754 | 35.3384
0.3818 | 40.3243 | 64.3775 | -1805.0312 | 1807.7391 | 2.9316 | -0.82415 | -3.0022 | 7.3349 | 1805.0312 | 0.82415 | 35.4965
0.3836 | 41.2073 | 65.7432 | -1797.7434 | 1800.4867 | 2.9198 | -0.9772 | -2.9551 | 7.276 | 1797.7434 | 0.9772 | 35.6638
0.3854 | 41.9843 | 67.5681 | -1788.0774 | 1790.8677 | 2.8492 | -1.1656 | -2.8963 | 7.1583 | 1788.0774 | 1.1656 | 35.8312
0.3872 | 42.7731 | 69.4048 | -1778.3996 | 1781.2487 | 2.7903 | -1.354 | -2.8492 | 7.0406 | 1778.3996 | 1.354 | 35.9985
0.389 | 43.6562 | 70.3702 | -1767.25 | 1770.1934 | 2.7432 | -1.4717 | -2.8021 | 7.0288 | 1767.25 | 1.4717 | 36.1659
0.3908 | 44.5863 | 71.5123 | -1755.6649 | 1758.6789 | 2.6844 | -1.6012 | -2.755 | 7.0052 | 1755.6649 | 1.6012 | 36.3332
0.3926 | 45.5281 | 73.2901 | -1743.55 | 1746.6229 | 2.6255 | -1.766 | -2.6961 | 6.8875 | 1743.55 | 1.766 | 36.5006
0.3944 | 46.4582 | 74.5969 | -1731.176 | 1734.3078 | 2.5784 | -1.9073 | -2.6373 | 6.8169 | 1731.176 | 1.9073 | 36.6679
0.3962 | 47.3884 | 75.9391 | -1718.4018 | 1721.6042 | 2.5313 | -2.0486 | -2.5784 | 6.7815 | 1718.4018 | 2.0486 | 36.8353
0.398 | 48.1536 | 77.1165 | -1704.9564 | 1708.2412 | 2.496 | -2.1663 | -2.5195 | 6.7344 | 1704.9564 | 2.1663 | 37.0026
0.3998 | 48.7188 | 77.2106 | -1689.7333 | 1693.077 | 2.4371 | -2.1663 | -2.4724 | 6.6638 | 1689.7333 | 2.1663 | 37.17
0.4016 | 49.3663 | 77.2342 | -1673.9215 | 1677.3123 | 2.3782 | -2.1192 | -2.4136 | 6.5696 | 1673.9215 | 2.1192 | 37.3373
0.4034 | 50.2375 | 77.1047 | -1658.51 | 1661.9596 | 2.3076 | -2.0839 | -2.39 | 6.499 | 1658.51 | 2.0839 | 37.5046
0.4051 | 51.356 | 76.8221 | -1643.5105 | 1647.0073 | 2.237 | -2.0604 | -2.39 | 6.4401 | 1643.5105 | 2.0604 | 37.6627
0.4069 | 52.4039 | 76.5278 | -1628.8172 | 1632.3846 | 2.1663 | -2.0368 | -2.39 | 6.393 | 1628.8172 | 2.0368 | 37.83
0.4087 | 53.2751 | 76.151 | -1614.9127 | 1618.5389 | 2.0839 | -2.0486 | -2.4018 | 6.3695 | 1614.9127 | 2.0486 | 37.9974
0.4105 | 54.3347 | 75.739 | -1600.0781 | 1603.7749 | 1.9897 | -2.0368 | -2.4253 | 6.3224 | 1600.0781 | 2.0368 | 38.1647
0.4123 | 54.8998 | 75.374 | -1586.2089 | 1589.9058 | 1.8955 | -2.0721 | -2.4607 | 6.2753 | 1586.2089 | 2.0721 | 38.3321
0.4141 | 55.4768 | 75.4328 | -1572.9284 | 1576.6488 | 1.7896 | -2.1546 | -2.4842 | 6.2164 | 1572.9284 | 2.1546 | 38.4994
0.4159 | 56.0654 | 75.6919 | -1560.7781 | 1564.4985 | 1.6954 | -2.2723 | -2.4842 | 6.1575 | 1560.7781 | 2.2723 | 38.6668
0.4177 | 56.5835 | 75.9156 | -1549.3696 | 1553.1253 | 1.613 | -2.4018 | -2.4724 | 6.1458 | 1549.3696 | 2.4018 | 38.8341
0.4195 | 57.1721 | 75.6448 | -1538.4202 | 1542.1995 | 1.5306 | -2.4842 | -2.4724 | 6.1575 | 1538.4202 | 2.4842 | 39.0015
0.4213 | 57.8668 | 74.75 | -1527.5651 | 1531.3561 | 1.4364 | -2.5078 | -2.496 | 6.1693 | 1527.5651 | 2.5078 | 39.1688
0.4231 | 58.3142 | 73.5609 | -1516.8041 | 1520.6069 | 1.3657 | -2.5666 | -2.5078 | 6.2164 | 1516.8041 | 2.5666 | 39.3362
0.4249 | 59.2678 | 72.984 | -1505.7252 | 1509.5869 | 1.2833 | -2.6137 | -2.5431 | 6.2753 | 1505.7252 | 2.6137 | 39.5035
0.4267 | 60.0213 | 72.8427 | -1494.4226 | 1498.3314 | 1.1891 | -2.6844 | -2.5666 | 6.2988 | 1494.4226 | 2.6844 | 39.6709
0.4285 | 60.3745 | 72.8427 | -1483.2378 | 1487.1466 | 1.1303 | -2.7785 | -2.5784 | 6.2753 | 1483.2378 | 2.7785 | 39.8382
0.4302 | 60.5865 | 72.8191 | -1472.7712 | 1476.6682 | 1.0714 | -2.8845 | -2.5666 | 6.2282 | 1472.7712 | 2.8845 | 39.9963
0.432 | 60.5982 | 72.513 | -1463.0345 | 1466.908 | 1.0125 | -2.9434 | -2.5666 | 6.1811 | 1463.0345 | 2.9434 | 40.1636
0.4338 | 60.5511 | 72.1127 | -1453.5568 | 1457.4067 | 0.95365 | -2.9787 | -2.5549 | 6.1104 | 1453.5568 | 2.9787 | 40.331
0.4356 | 60.8572 | 73.1252 | -1443.7613 | 1447.623 | 0.90656 | -3.1906 | -2.5431 | 5.9456 | 1443.7613 | 3.1906 | 40.4983
0.4374 | 61.964 | 76.1393 | -1434.3189 | 1438.263 | 0.8006 | -3.4614 | -2.4842 | 5.6631 | 1434.3189 | 3.4614 | 40.6657
0.4392 | 62.3407 | 77.8935 | -1425.3357 | 1429.3505 | 0.70641 | -3.6262 | -2.4371 | 5.4629 | 1425.3357 | 3.6262 | 40.833
0.441 | 62.3643 | 78.2938 | -1413.9507 | 1418.0244 | 0.65932 | -3.638 | -2.4371 | 5.404 | 1413.9507 | 3.638 | 41.0004
0.4428 | 61.9875 | 77.8935 | -1403.543 | 1407.6519 | 0.63577 | -3.638 | -2.4607 | 5.3923 | 1403.543 | 3.638 | 41.1677
0.4446 | 61.3282 | 77.3284 | -1394.3125 | 1398.4097 | 0.60045 | -3.638 | -2.4607 | 5.3687 | 1394.3125 | 3.638 | 41.3351
0.4464 | 60.7395 | 76.8103 | -1385.0468 | 1389.1322 | 0.56513 | -3.6498 | -2.4607 | 5.3452 | 1385.0468 | 3.6498 | 41.5024
0.4482 | 59.5975 | 75.3151 | -1377.4058 | 1381.4323 | 0.58868 | -3.638 | -2.4607 | 5.3334 | 1377.4058 | 3.638 | 41.6698
0.45 | 58.997 | 73.8316 | -1370.1062 | 1374.0621 | 0.61222 | -3.6262 | -2.4724 | 5.3569 | 1370.1062 | 3.6262 | 41.8371
0.4518 | 58.3966 | 71.889 | -1364.867 | 1368.764 | 0.67109 | -3.638 | -2.4724 | 5.3687 | 1364.867 | 3.638 | 42.0045
0.4536 | 57.5607 | 69.4872 | -1360.2871 | 1364.1017 | 0.74173 | -3.6145 | -2.4724 | 5.3805 | 1360.2871 | 3.6145 | 42.1718
0.4553 | 56.6894 | 66.8264 | -1356.5667 | 1360.2871 | 0.82415 | -3.5674 | -2.4842 | 5.3805 | 1356.5667 | 3.5674 | 42.3299
0.4571 | 56.3009 | 65.1075 | -1351.8573 | 1355.4953 | 0.84769 | -3.5556 | -2.4842 | 5.3805 | 1351.8573 | 3.5556 | 42.4972
0.4589 | 55.5474 | 62.3407 | -1350.7388 | 1354.212 | 0.93011 | -3.5085 | -2.496 | 5.3569 | 1350.7388 | 3.5085 | 42.6646
0.4607 | 54.7586 | 59.362 | -1349.9147 | 1353.2348 | 1.0125 | -3.4614 | -2.496 | 5.3452 | 1349.9147 | 3.4614 | 42.8319
0.4625 | 53.9697 | 56.4304 | -1349.0552 | 1352.2105 | 1.0949 | -3.4025 | -2.496 | 5.3216 | 1349.0552 | 3.4025 | 42.9993
0.4643 | 52.9808 | 53.9933 | -1347.7483 | 1350.7859 | 1.1891 | -3.3554 | -2.5078 | 5.3099 | 1347.7483 | 3.3554 | 43.1666
0.4661 | 52.0036 | 51.5562 | -1346.4415 | 1349.3613 | 1.2833 | -3.3084 | -2.5195 | 5.3099 | 1346.4415 | 3.3084 | 43.334
0.4679 | 51.6621 | 50.0138 | -1342.097 | 1344.958 | 1.3657 | -3.3319 | -2.5313 | 5.3452 | 1342.097 | 3.3319 | 43.5013
0.4697 | 51.4267 | 48.6128 | -1337.2817 | 1340.1073 | 1.4364 | -3.3672 | -2.5313 | 5.3805 | 1337.2817 | 3.3672 | 43.6687
0.4715 | 51.3089 | 47.0351 | -1332.0189 | 1334.7975 | 1.507 | -3.3908 | -2.5431 | 5.4276 | 1332.0189 | 3.3908 | 43.836
0.4733 | 51.203 | 45.434 | -1326.6855 | 1329.417 | 1.5659 | -3.4025 | -2.5549 | 5.4629 | 1326.6855 | 3.4025 | 44.0033
0.4751 | 50.3082 | 44.0565 | -1321.3992 | 1324.0718 | 1.6247 | -3.379 | -2.5549 | 5.4629 | 1321.3992 | 3.379 | 44.1707
0.4769 | 49.2957 | 42.938 | -1315.6067 | 1318.2322 | 1.6718 | -3.3319 | -2.5666 | 5.4511 | 1315.6067 | 3.3319 | 44.338
0.4786 | 48.4244 | 41.9961 | -1308.6485 | 1311.2505 | 1.7189 | -3.2966 | -2.5784 | 5.4629 | 1308.6485 | 3.2966 | 44.4961
0.4804 | 47.6238 | 41.1602 | -1300.8545 | 1303.4329 | 1.7543 | -3.2495 | -2.5902 | 5.4511 | 1300.8545 | 3.2495 | 44.6634
0.4822 | 46.9881 | 40.5126 | -1291.4474 | 1294.0023 | 1.7543 | -3.1906 | -2.6019 | 5.4158 | 1291.4474 | 3.1906 | 44.8308
0.484 | 46.7055 | 40.0181 | -1280.3568 | 1282.947 | 1.766 | -3.1435 | -2.6255 | 5.3805 | 1280.3568 | 3.1435 | 44.9981
0.4858 | 47.094 | 39.4766 | -1268.2183 | 1270.8556 | 1.7425 | -3.12 | -2.649 | 5.3569 | 1268.2183 | 3.12 | 45.1655
0.4876 | 47.7062 | 38.4994 | -1253.8664 | 1256.5154 | 1.6718 | -3.1082 | -2.7079 | 5.3452 | 1253.8664 | 3.1082 | 45.3328
0.4894 | 48.1419 | 37.5693 | -1241.9752 | 1244.6242 | 1.613 | -3.0847 | -2.7668 | 5.3569 | 1241.9752 | 3.0847 | 45.5002
0.4912 | 48.0123 | 36.698 | -1231.7322 | 1234.3459 | 1.5776 | -3.0376 | -2.8021 | 5.3687 | 1231.7322 | 3.0376 | 45.6675
0.493 | 47.3413 | 35.7797 | -1222.6902 | 1225.2215 | 1.5423 | -2.9669 | -2.8139 | 5.3569 | 1222.6902 | 2.9669 | 45.8349
0.4948 | 46.6231 | 35.5913 | -1215.0845 | 1217.5687 | 1.5423 | -2.9081 | -2.8256 | 5.3099 | 1215.0845 | 2.9081 | 46.0022
0.4966 | 45.8578 | 35.0497 | -1208.456 | 1210.8578 | 1.5423 | -2.8727 | -2.8139 | 5.2863 | 1208.456 | 2.8727 | 46.1696
0.4984 | 45.434 | 33.4485 | -1204.0056 | 1206.2661 | 1.6012 | -2.8256 | -2.8021 | 5.2274 | 1204.0056 | 2.8256 | 46.3369
0.5002 | 44.7746 | 30.8819 | -1199.567 | 1201.6156 | 1.6836 | -2.7785 | -2.7903 | 5.1568 | 1199.567 | 2.7785 | 46.5043
0.502 | 44.1624 | 28.7627 | -1194.057 | 1195.9525 | 1.7543 | -2.7315 | -2.7785 | 5.1097 | 1194.057 | 2.7315 | 46.6716
0.5037 | 43.2794 | 27.5735 | -1187.4167 | 1189.1828 | 1.7896 | -2.6844 | -2.7668 | 5.0508 | 1187.4167 | 2.6844 | 46.8297
0.5055 | 41.9725 | 28.1387 | -1178.5513 | 1180.2231 | 1.7896 | -2.6255 | -2.755 | 4.992 | 1178.5513 | 2.6255 | 46.997
0.5073 | 40.7128 | 29.6339 | -1168.8382 | 1170.4276 | 1.7543 | -2.6137 | -2.7432 | 4.9566 | 1168.8382 | 2.6137 | 47.1644
0.5091 | 39.5001 | 29.5162 | -1162.1273 | 1163.646 | 1.7543 | -2.5666 | -2.7432 | 4.9096 | 1162.1273 | 2.5666 | 47.3317
0.5109 | 38.311 | 29.2101 | -1156.6526 | 1158.1125 | 1.7543 | -2.5195 | -2.7432 | 4.8742 | 1156.6526 | 2.5195 | 47.4991
0.5127 | 37.1336 | 28.8098 | -1152.1551 | 1153.5679 | 1.7543 | -2.4724 | -2.7432 | 4.8389 | 1152.1551 | 2.4724 | 47.6664
0.5145 | 36.0387 | 28.0798 | -1147.8931 | 1149.247 | 1.766 | -2.4253 | -2.755 | 4.8036 | 1147.8931 | 2.4253 | 47.8338
0.5163 | 34.9202 | 27.291 | -1144.0549 | 1145.35 | 1.7778 | -2.39 | -2.7668 | 4.78 | 1144.0549 | 2.39 | 48.0011
0.5181 | 33.6958 | 26.2667 | -1141.6531 | 1142.8776 | 1.7896 | -2.3782 | -2.7903 | 4.7918 | 1141.6531 | 2.3782 | 48.1685
0.5199 | 32.6479 | 26.043 | -1139.8636 | 1141.0409 | 1.7896 | -2.39 | -2.8139 | 4.8154 | 1139.8636 | 2.39 | 48.3358
0.5217 | 31.8944 | 24.7244 | -1138.7098 | 1139.8165 | 1.8367 | -2.4136 | -2.8492 | 4.8625 | 1138.7098 | 2.4136 | 48.5032
0.5235 | 31.2233 | 23.441 | -1138.0034 | 1139.063 | 1.8838 | -2.4607 | -2.8963 | 4.9213 | 1138.0034 | 2.4607 | 48.6705
0.5253 | 30.7053 | 22.2166 | -1137.9916 | 1139.0159 | 1.9426 | -2.5549 | -2.9316 | 5.0155 | 1137.9916 | 2.5549 | 48.8379
0.5271 | 30.1519 | 21.0981 | -1138.4037 | 1139.4044 | 2.0015 | -2.649 | -2.9787 | 5.1097 | 1138.4037 | 2.649 | 49.0052
0.5288 | 29.4926 | 20.3328 | -1140.0873 | 1141.0527 | 2.0604 | -2.7432 | -3.014 | 5.2157 | 1140.0873 | 2.7432 | 49.1633
0.5306 | 28.7744 | 19.5911 | -1142.5126 | 1143.4427 | 2.0957 | -2.7785 | -3.0611 | 5.2745 | 1142.5126 | 2.7785 | 49.3306
0.5324 | 28.2093 | 18.7434 | -1145.5737 | 1146.4567 | 2.131 | -2.8021 | -3.12 | 5.3216 | 1145.5737 | 2.8021 | 49.498
0.5342 | 27.8797 | 17.8015 | -1148.3405 | 1149.2117 | 2.1899 | -2.861 | -3.1671 | 5.3687 | 1148.3405 | 2.861 | 49.6653
0.536 | 27.1026 | 17.1304 | -1150.236 | 1151.0719 | 2.2252 | -2.8845 | -3.2142 | 5.4158 | 1150.236 | 2.8845 | 49.8327
0.5378 | 26.9731 | 16.9068 | -1153.6032 | 1154.4274 | 2.237 | -2.8963 | -3.2613 | 5.4747 | 1153.6032 | 2.8963 | 50
0.5396 | 26.4786 | 16.3887 | -1156.8056 | 1157.618 | 2.2723 | -2.9316 | -3.3084 | 5.5335 | 1156.8056 | 2.9316 | 50.1673
0.5414 | 26.4551 | 16.2239 | -1160.3495 | 1161.1854 | 2.2723 | -2.9434 | -3.3672 | 5.5924 | 1160.3495 | 2.9434 | 50.3347
0.5432 | 26.4904 | 16.0002 | -1163.3753 | 1164.1994 | 2.2841 | -2.9434 | -3.4261 | 5.6513 | 1163.3753 | 2.9434 | 50.502
0.545 | 26.2785 | 15.4586 | -1165.0942 | 1165.9183 | 2.3194 | -2.9551 | -3.485 | 5.7101 | 1165.0942 | 2.9551 | 50.6694
0.5468 | 25.8075 | 14.8817 | -1166.9309 | 1167.755 | 2.39 | -2.9787 | -3.5321 | 5.7572 | 1166.9309 | 2.9787 | 50.8367
0.5486 | 25.3954 | 14.5403 | -1169.6741 | 1170.4982 | 2.4371 | -2.9669 | -3.5791 | 5.8043 | 1169.6741 | 2.9669 | 51.0041
0.5504 | 25.0187 | 14.2224 | -1169.5563 | 1170.3805 | 2.4842 | -3.0376 | -3.6145 | 5.8514 | 1169.5563 | 3.0376 | 51.1714
0.5521 | 24.4182 | 14.3637 | -1169.3326 | 1170.145 | 2.5313 | -3.0964 | -3.6498 | 5.9103 | 1169.3326 | 3.0964 | 51.3295
0.5539 | 24.3829 | 14.2459 | -1171.4283 | 1172.2289 | 2.6608 | -3.0376 | -3.6969 | 5.9809 | 1171.4283 | 3.0376 | 51.4968
0.5557 | 24.7361 | 14.2342 | -1172.9589 | 1173.7595 | 2.7785 | -3.0376 | -3.7322 | 6.0751 | 1172.9589 | 3.0376 | 51.6642
0.5575 | 24.2063 | 14.3166 | -1173.8537 | 1174.6425 | 2.8256 | -3.0611 | -3.7675 | 6.1222 | 1173.8537 | 3.0611 | 51.8315
0.5593 | 23.8767 | 14.5756 | -1172.7352 | 1173.5005 | 2.861 | -3.0964 | -3.7911 | 6.1693 | 1172.7352 | 3.0964 | 51.9989
0.5611 | 23.4999 | 14.8582 | -1171.5696 | 1172.2878 | 2.8727 | -3.1435 | -3.8028 | 6.2164 | 1171.5696 | 3.1435 | 52.1662
0.5629 | 23.1114 | 15.2467 | -1170.9338 | 1171.6285 | 2.8845 | -3.1671 | -3.8264 | 6.2517 | 1170.9338 | 3.1671 | 52.3336
0.5647 | 22.7229 | 15.6941 | -1170.6042 | 1171.24 | 2.8845 | -3.2024 | -3.8382 | 6.2871 | 1170.6042 | 3.2024 | 52.5009
0.5665 | 22.405 | 15.2585 | -1170.5571 | 1171.1458 | 2.8727 | -3.1906 | -3.8499 | 6.3106 | 1170.5571 | 3.1906 | 52.6683
0.5683 | 21.8516 | 14.3048 | -1172.4526 | 1173.006 | 2.861 | -3.1671 | -3.8735 | 6.3224 | 1172.4526 | 3.1671 | 52.8356
0.5701 | 21.051 | 13.7868 | -1176.4321 | 1176.9619 | 2.861 | -3.1671 | -3.9088 | 6.3341 | 1176.4321 | 3.1671 | 53.003
0.5719 | 20.3799 | 13.8221 | -1181.2121 | 1181.7066 | 2.8492 | -3.1906 | -3.9324 | 6.3577 | 1181.2121 | 3.1906 | 53.1703
0.5737 | 20.2975 | 14.4108 | -1185.6389 | 1186.1099 | 2.8492 | -3.2495 | -3.9677 | 6.4166 | 1185.6389 | 3.2495 | 53.3377
0.5755 | 20.1209 | 13.881 | -1188.4999 | 1188.9473 | 2.861 | -3.2848 | -3.9794 | 6.4519 | 1188.4999 | 3.2848 | 53.505
0.5772 | 19.7442 | 13.2099 | -1190.3601 | 1190.784 | 2.8727 | -3.2966 | -4.003 | 6.4754 | 1190.3601 | 3.2966 | 53.6631
0.579 | 19.2968 | 12.7743 | -1192.7384 | 1193.1504 | 2.8845 | -3.2966 | -4.0383 | 6.499 | 1192.7384 | 3.2966 | 53.8304
0.5808 | 18.5904 | 12.6447 | -1196.6354 | 1197.0357 | 2.8963 | -3.3084 | -4.0619 | 6.5343 | 1196.6354 | 3.3084 | 53.9978
0.5826 | 17.7427 | 12.5623 | -1201.8393 | 1202.2396 | 2.9081 | -3.3201 | -4.109 | 6.5814 | 1201.8393 | 3.3201 | 54.1651
0.5844 | 16.9774 | 12.8802 | -1208.3029 | 1208.6915 | 2.9316 | -3.3554 | -4.1678 | 6.6403 | 1208.3029 | 3.3554 | 54.3325
0.5862 | 16.3887 | 13.1628 | -1215.5672 | 1215.9439 | 2.9787 | -3.4025 | -4.2149 | 6.7227 | 1215.5672 | 3.4025 | 54.4998
0.588 | 16.2121 | 12.9155 | -1220.7004 | 1221.089 | 3.0493 | -3.4732 | -4.2385 | 6.8286 | 1220.7004 | 3.4732 | 54.6672
0.5898 | 15.8471 | 12.3386 | -1225.2097 | 1225.5982 | 3.12 | -3.5203 | -4.262 | 6.911 | 1225.2097 | 3.5203 | 54.8345
0.5916 | 15.4704 | 12.0796 | -1230.6137 | 1231.0023 | 3.1671 | -3.5321 | -4.2973 | 6.9699 | 1230.6137 | 3.5321 | 55.0019
0.5934 | 15.4586 | 12.0325 | -1235.2878 | 1235.6763 | 3.2259 | -3.5674 | -4.3209 | 7.0288 | 1235.2878 | 3.5674 | 55.1692
0.5952 | 16.0237 | 12.1974 | -1238.3607 | 1238.761 | 3.2848 | -3.6498 | -4.3444 | 7.1347 | 1238.3607 | 3.6498 | 55.3366
0.597 | 15.9413 | 12.2209 | -1243.4822 | 1243.8707 | 3.3437 | -3.6851 | -4.3797 | 7.2054 | 1243.4822 | 3.6851 | 55.5039
0.5988 | 15.6117 | 12.5152 | -1250.146 | 1250.5463 | 3.4143 | -3.7087 | -4.4268 | 7.2878 | 1250.146 | 3.7087 | 55.6713
0.6006 | 14.6462 | 12.5859 | -1257.9871 | 1258.3639 | 3.485 | -3.6851 | -4.4739 | 7.3467 | 1257.9871 | 3.6851 | 55.8386
0.6023 | 13.5984 | 12.5506 | -1265.6635 | 1266.0402 | 3.5556 | -3.638 | -4.5093 | 7.382 | 1265.6635 | 3.638 | 55.9967
0.6041 | 12.9509 | 12.8096 | -1275.5885 | 1275.9653 | 3.6498 | -3.6262 | -4.5681 | 7.4644 | 1275.5885 | 3.6262 | 56.164
0.6059 | 12.9509 | 13.3394 | -1285.6195 | 1286.0081 | 3.744 | -3.6851 | -4.627 | 7.5821 | 1285.6195 | 3.6851 | 56.3313
0.6077 | 13.045 | 14.0929 | -1295.4504 | 1295.8507 | 3.8617 | -3.7204 | -4.6741 | 7.7116 | 1295.4504 | 3.7204 | 56.4987
0.6095 | 12.9273 | 14.764 | -1304.8928 | 1305.3048 | 3.9794 | -3.7322 | -4.7212 | 7.8176 | 1304.8928 | 3.7322 | 56.666
0.6113 | 12.5506 | 14.9053 | -1313.0283 | 1313.4521 | 4.0501 | -3.6969 | -4.7565 | 7.8647 | 1313.0283 | 3.6969 | 56.8334
0.6131 | 12.4328 | 15.0583 | -1320.3161 | 1320.7517 | 4.109 | -3.6969 | -4.8036 | 7.9236 | 1320.3161 | 3.6969 | 57.0007
0.6149 | 12.0796 | 15.1172 | -1327.2153 | 1327.6392 | 4.1443 | -3.6969 | -4.8507 | 7.9707 | 1327.2153 | 3.6969 | 57.1681
0.6167 | 11.7735 | 15.4586 | -1333.8674 | 1334.2912 | 4.156 | -3.6851 | -4.9096 | 8.0295 | 1333.8674 | 3.6851 | 57.3354
0.6185 | 11.5027 | 15.9649 | -1340.3663 | 1340.7902 | 4.156 | -3.6851 | -4.992 | 8.0766 | 1340.3663 | 3.6851 | 57.5028
0.6203 | 11.9501 | 16.53 | -1345.1464 | 1345.582 | 4.2031 | -3.6969 | -5.0391 | 8.1473 | 1345.1464 | 3.6969 | 57.6701
0.6221 | 11.8912 | 16.9185 | -1350.1972 | 1350.6328 | 4.2267 | -3.6616 | -5.0862 | 8.1708 | 1350.1972 | 3.6616 | 57.8375
0.6239 | 11.4321 | 17.1893 | -1356.8375 | 1357.2966 | 4.2502 | -3.638 | -5.145 | 8.1944 | 1356.8375 | 3.638 | 58.0048
0.6256 | 11.2319 | 17.672 | -1364.7964 | 1365.2791 | 4.2856 | -3.638 | -5.1803 | 8.2532 | 1364.7964 | 3.638 | 58.1629
0.6274 | 11.02 | 18.7434 | -1375.1924 | 1375.6986 | 4.3915 | -3.638 | -5.2039 | 8.3356 | 1375.1924 | 3.638 | 58.3302
0.6292 | 11.6558 | 20.3211 | -1385.5766 | 1386.1535 | 4.5328 | -3.6733 | -5.2628 | 8.4769 | 1385.5766 | 3.6733 | 58.4976
0.631 | 12.4917 | 22.0282 | -1395.9608 | 1396.6201 | 4.6741 | -3.7087 | -5.3334 | 8.63 | 1395.9608 | 3.7087 | 58.6649
0.6328 | 12.7389 | 23.2998 | -1407.0986 | 1407.805 | 4.8271 | -3.7204 | -5.3923 | 8.7713 | 1407.0986 | 3.7204 | 58.8323
0.6346 | 12.7272 | 24.3947 | -1418.5189 | 1419.2606 | 4.9802 | -3.7087 | -5.4511 | 8.9125 | 1418.5189 | 3.7087 | 58.9996
0.6364 | 13.1981 | 25.8899 | -1427.7611 | 1428.5617 | 5.1097 | -3.7087 | -5.5218 | 9.0421 | 1427.7611 | 3.7087 | 59.167
0.6382 | 13.4689 | 27.1262 | -1437.1445 | 1438.004 | 5.2392 | -3.6733 | -5.5924 | 9.1598 | 1437.1445 | 3.6733 | 59.3343
0.64 | 13.8103 | 27.9032 | -1444.1027 | 1444.9857 | 5.2981 | -3.6498 | -5.6513 | 9.2187 | 1444.1027 | 3.6498 | 59.5017
0.6418 | 14.2695 | 28.5507 | -1450.5546 | 1451.4964 | 5.3569 | -3.638 | -5.7101 | 9.2775 | 1450.5546 | 3.638 | 59.669
0.6436 | 14.8464 | 29.3631 | -1456.3354 | 1457.3126 | 5.4158 | -3.6145 | -5.7572 | 9.3364 | 1456.3354 | 3.6145 | 59.8364
0.6454 | 15.4233 | 30.1755 | -1461.6217 | 1462.6342 | 5.4747 | -3.5909 | -5.8161 | 9.3835 | 1461.6217 | 3.5909 | 60.0037
0.6472 | 16.0708 | 30.9054 | -1466.2486 | 1467.2965 | 5.5218 | -3.5791 | -5.8514 | 9.4188 | 1466.2486 | 3.5791 | 60.1711
0.649 | 16.7772 | 31.6001 | -1469.7571 | 1470.8403 | 5.5806 | -3.5909 | -5.8868 | 9.4659 | 1469.7571 | 3.5909 | 60.3384
0.6507 | 17.3777 | 32.3418 | -1473.0184 | 1474.1487 | 5.6631 | -3.5909 | -5.9221 | 9.5012 | 1473.0184 | 3.5909 | 60.4965
0.6525 | 17.931 | 32.7657 | -1475.5968 | 1476.7506 | 5.6984 | -3.6027 | -5.9456 | 9.5365 | 1475.5968 | 3.6027 | 60.6638
0.6543 | 18.0841 | 32.8245 | -1477.8338 | 1478.9993 | 5.7219 | -3.6145 | -5.9692 | 9.5601 | 1477.8338 | 3.6145 | 60.8312
0.6561 | 18.5433 | 33.3897 | -1480.0119 | 1481.1892 | 5.7337 | -3.6262 | -5.9927 | 9.5836 | 1480.0119 | 3.6262 | 60.9985
0.6579 | 19.2732 | 34.673 | -1481.8603 | 1483.0848 | 5.7572 | -3.6145 | -6.028 | 9.6072 | 1481.8603 | 3.6145 | 61.1659
0.6597 | 19.8148 | 35.8739 | -1483.8147 | 1485.0745 | 5.769 | -3.6027 | -6.0751 | 9.6307 | 1483.8147 | 3.6027 | 61.3332
0.6615 | 19.5205 | 36.5567 | -1486.6168 | 1487.8883 | 5.7572 | -3.5438 | -6.1104 | 9.6072 | 1486.6168 | 3.5438 | 61.5006
0.6633 | 18.8494 | 37.0394 | -1490.6551 | 1491.9267 | 5.7572 | -3.4496 | -6.1458 | 9.5954 | 1490.6551 | 3.4496 | 61.6679
0.6651 | 18.3902 | 37.1101 | -1495.0113 | 1496.2829 | 5.769 | -3.4143 | -6.1811 | 9.619 | 1495.0113 | 3.4143 | 61.8353
0.6669 | 17.8957 | 37.0983 | -1500.5213 | 1501.8046 | 5.7808 | -3.4143 | -6.2164 | 9.6543 | 1500.5213 | 3.4143 | 62.0026
0.6687 | 17.7898 | 36.9453 | -1506.2668 | 1507.5619 | 5.8397 | -3.4379 | -6.2635 | 9.7367 | 1506.2668 | 3.4379 | 62.17
0.6705 | 17.8486 | 36.9099 | -1511.2352 | 1512.5421 | 5.875 | -3.4379 | -6.3106 | 9.7956 | 1511.2352 | 3.4379 | 62.3373
0.6723 | 17.9193 | 37.0159 | -1515.6032 | 1516.9336 | 5.9103 | -3.4614 | -6.3577 | 9.8662 | 1515.6032 | 3.4614 | 62.5046
0.6741 | 18.4138 | 36.4861 | -1520.5834 | 1521.9491 | 5.9692 | -3.4967 | -6.393 | 9.9486 | 1520.5834 | 3.4967 | 62.672
0.6758 | 18.9907 | 36.3213 | -1524.5157 | 1525.8932 | 5.9809 | -3.5085 | -6.4283 | 9.9839 | 1524.5157 | 3.5085 | 62.83
0.6776 | 19.5322 | 36.2742 | -1527.8712 | 1529.2722 | 5.9927 | -3.5085 | -6.4637 | 10.0193 | 1527.8712 | 3.5085 | 62.9974
0.6794 | 20.0621 | 36.1682 | -1531.0971 | 1532.4982 | 6.0163 | -3.5203 | -6.499 | 10.0663 | 1531.0971 | 3.5203 | 63.1647
0.6812 | 20.6154 | 36.333 | -1533.7226 | 1535.1354 | 6.0398 | -3.5674 | -6.5343 | 10.1134 | 1533.7226 | 3.5674 | 63.3321
0.683 | 20.7802 | 35.8739 | -1535.6417 | 1537.031 | 6.0398 | -3.5791 | -6.5696 | 10.1488 | 1535.6417 | 3.5791 | 63.4994
0.6848 | 20.3093 | 34.9084 | -1537.9728 | 1539.2915 | 6.028 | -3.5791 | -6.5814 | 10.1605 | 1537.9728 | 3.5791 | 63.6668
0.6866 | 19.8384 | 33.79 | -1540.7278 | 1541.9758 | 6.0045 | -3.5674 | -6.6049 | 10.1605 | 1540.7278 | 3.5674 | 63.8341
0.6884 | 19.5087 | 32.7186 | -1543.6712 | 1544.8486 | 6.0045 | -3.5791 | -6.6285 | 10.1841 | 1543.6712 | 3.5791 | 64.0015
0.6902 | 19.5793 | 32.0946 | -1546.4851 | 1547.6271 | 6.0163 | -3.6262 | -6.652 | 10.2312 | 1546.4851 | 3.6262 | 64.1688
0.692 | 20.2387 | 32.2712 | -1548.4983 | 1549.6521 | 6.0398 | -3.6851 | -6.6756 | 10.29 | 1548.4983 | 3.6851 | 64.3362
0.6938 | 20.8862 | 32.5066 | -1549.6521 | 1550.8177 | 6.0398 | -3.7087 | -6.7227 | 10.3371 | 1549.6521 | 3.7087 | 64.5035
0.6956 | 21.6397 | 32.7303 | -1550.2997 | 1551.4888 | 6.028 | -3.744 | -6.758 | 10.3607 | 1550.2997 | 3.744 | 64.6709
0.6974 | 22.4521 | 32.954 | -1550.8295 | 1552.0422 | 6.0045 | -3.7675 | -6.7933 | 10.3842 | 1550.8295 | 3.7675 | 64.8382
0.6992 | 23.3469 | 33.319 | -1551.0532 | 1552.3012 | 5.9927 | -3.7911 | -6.8286 | 10.4196 | 1551.0532 | 3.7911 | 65.0056
0.7009 | 24.2534 | 33.6958 | -1551.218 | 1552.5013 | 5.9692 | -3.8264 | -6.864 | 10.4431 | 1551.218 | 3.8264 | 65.1636
0.7027 | 25.3013 | 33.9195 | -1551.0767 | 1552.3836 | 5.9338 | -3.8499 | -6.8993 | 10.4666 | 1551.0767 | 3.8499 | 65.331
0.7045 | 25.9723 | 34.0019 | -1550.6293 | 1551.9598 | 5.9103 | -3.8617 | -6.9228 | 10.4784 | 1550.6293 | 3.8617 | 65.4983
0.7063 | 26.3726 | 33.9666 | -1549.817 | 1551.1592 | 5.8985 | -3.8617 | -6.9346 | 10.4666 | 1549.817 | 3.8617 | 65.6657
0.7081 | 26.7023 | 33.7429 | -1549.0046 | 1550.3468 | 5.8514 | -3.8499 | -6.9346 | 10.4549 | 1549.0046 | 3.8499 | 65.833
0.7099 | 27.032 | 33.5309 | -1548.204 | 1549.5462 | 5.8161 | -3.8382 | -6.9464 | 10.4313 | 1548.204 | 3.8382 | 66.0004
0.7117 | 26.7376 | 33.1071 | -1548.3217 | 1549.6521 | 5.7572 | -3.7911 | -6.9464 | 10.3842 | 1548.3217 | 3.7911 | 66.1677
0.7135 | 26.4904 | 32.5302 | -1549.3343 | 1550.6529 | 5.6984 | -3.7675 | -6.9464 | 10.3371 | 1549.3343 | 3.7675 | 66.3351
0.7153 | 26.0901 | 31.8002 | -1550.853 | 1552.1246 | 5.6277 | -3.7322 | -6.9464 | 10.2783 | 1550.853 | 3.7322 | 66.5024
0.7171 | 25.6191 | 31.0821 | -1552.7604 | 1554.0083 | 5.5689 | -3.7087 | -6.9464 | 10.2312 | 1552.7604 | 3.7087 | 66.6698
0.7189 | 24.7832 | 30.0813 | -1554.703 | 1555.9039 | 5.51 | -3.6616 | -6.9464 | 10.1723 | 1554.703 | 3.6616 | 66.8371
0.7207 | 24.1357 | 29.3278 | -1556.6809 | 1557.8347 | 5.4629 | -3.638 | -6.9346 | 10.1252 | 1556.6809 | 3.638 | 67.0045
0.7225 | 23.7354 | 28.9275 | -1558.6824 | 1559.8362 | 5.4158 | -3.638 | -6.911 | 10.0899 | 1558.6824 | 3.638 | 67.1718
0.7242 | 23.2409 | 28.6567 | -1560.0246 | 1561.1549 | 5.3805 | -3.6262 | -6.911 | 10.0781 | 1560.0246 | 3.6262 | 67.3299
0.726 | 22.664 | 28.6214 | -1561.3432 | 1562.4735 | 5.3805 | -3.6027 | -6.911 | 10.0663 | 1561.3432 | 3.6027 | 67.4972
0.7278 | 22.1224 | 28.6214 | -1562.5677 | 1563.6979 | 5.3687 | -3.5791 | -6.911 | 10.0546 | 1562.5677 | 3.5791 | 67.6646
0.7296 | 21.8398 | 28.7862 | -1563.1093 | 1564.2513 | 5.3569 | -3.5791 | -6.911 | 10.0546 | 1563.1093 | 3.5791 | 67.8319
0.7314 | 21.781 | 28.9628 | -1563.1328 | 1564.2748 | 5.3569 | -3.5909 | -6.9228 | 10.0546 | 1563.1328 | 3.5909 | 67.9993
0.7332 | 21.8634 | 29.163 | -1562.7678 | 1563.9334 | 5.3334 | -3.6027 | -6.9346 | 10.0663 | 1562.7678 | 3.6027 | 68.1666
0.735 | 21.9458 | 29.3749 | -1562.4029 | 1563.5684 | 5.3099 | -3.6145 | -6.9346 | 10.0546 | 1562.4029 | 3.6145 | 68.334
0.7368 | 22.0282 | 29.5868 | -1562.0261 | 1563.2152 | 5.2745 | -3.6145 | -6.9464 | 10.0546 | 1562.0261 | 3.6145 | 68.5013
0.7386 | 22.04 | 29.787 | -1562.6383 | 1563.8274 | 5.2392 | -3.6262 | -6.9581 | 10.0546 | 1562.6383 | 3.6262 | 68.6687
0.7404 | 21.8987 | 29.7752 | -1563.8274 | 1564.993 | 5.2157 | -3.6733 | -6.9817 | 10.0663 | 1563.8274 | 3.6733 | 68.836
0.7422 | 21.5337 | 29.6339 | -1565.5935 | 1566.7473 | 5.1921 | -3.7087 | -6.9935 | 10.0899 | 1565.5935 | 3.7087 | 69.0033
0.744 | 21.157 | 29.6575 | -1568.2072 | 1569.3492 | 5.1803 | -3.6969 | -6.9935 | 10.0781 | 1568.2072 | 3.6969 | 69.1707
0.7458 | 20.9097 | 29.7163 | -1571.3507 | 1572.5045 | 5.1568 | -3.6851 | -6.9935 | 10.0546 | 1571.3507 | 3.6851 | 69.338
0.7476 | 21.416 | 29.9871 | -1572.6223 | 1573.7996 | 5.145 | -3.7204 | -7.017 | 10.0781 | 1572.6223 | 3.7204 | 69.5054
0.7493 | 22.0518 | 30.3992 | -1573.3404 | 1574.5296 | 5.1568 | -3.7675 | -7.0288 | 10.1134 | 1573.3404 | 3.7675 | 69.6634
0.7511 | 22.5816 | 30.9761 | -1574.0586 | 1575.2595 | 5.145 | -3.8146 | -7.0406 | 10.137 | 1574.0586 | 3.8146 | 69.8308
0.7529 | 23.1114 | 31.7178 | -1574.0586 | 1575.2713 | 5.1568 | -3.8499 | -7.0523 | 10.1605 | 1574.0586 | 3.8499 | 69.9981
0.7547 | 23.6412 | 32.5655 | -1573.5641 | 1574.8121 | 5.1686 | -3.8735 | -7.0641 | 10.1959 | 1573.5641 | 3.8735 | 70.1655
0.7565 | 24.277 | 33.2601 | -1572.5281 | 1573.7996 | 5.1803 | -3.897 | -7.0876 | 10.2194 | 1572.5281 | 3.897 | 70.3328
0.7583 | 25.0658 | 33.7429 | -1570.7974 | 1572.0924 | 5.1803 | -3.9206 | -7.0876 | 10.2429 | 1570.7974 | 3.9206 | 70.5002
0.7601 | 25.3248 | 34.1078 | -1568.9254 | 1570.2322 | 5.2039 | -3.9088 | -7.0994 | 10.2547 | 1568.9254 | 3.9088 | 70.6675
0.7619 | 25.2542 | 34.4022 | -1566.9827 | 1568.2661 | 5.2157 | -3.897 | -7.0994 | 10.2665 | 1566.9827 | 3.897 | 70.8349
0.7637 | 25.1953 | 34.6847 | -1565.0401 | 1566.3117 | 5.2392 | -3.8735 | -7.0876 | 10.2665 | 1565.0401 | 3.8735 | 71.0022
0.7655 | 24.5242 | 34.6494 | -1563.5449 | 1564.7693 | 5.251 | -3.8146 | -7.0994 | 10.2429 | 1563.5449 | 3.8146 | 71.1696
0.7673 | 23.8413 | 34.6023 | -1562.0614 | 1563.2505 | 5.2745 | -3.7675 | -7.0994 | 10.2312 | 1562.0614 | 3.7675 | 71.3369
0.7691 | 23.0996 | 34.5199 | -1560.6251 | 1561.7671 | 5.2981 | -3.6969 | -7.0994 | 10.2076 | 1560.6251 | 3.6969 | 71.5043
0.7709 | 22.3226 | 34.4022 | -1559.2358 | 1560.3425 | 5.3216 | -3.6262 | -7.0876 | 10.1841 | 1559.2358 | 3.6262 | 71.6716
0.7727 | 21.5337 | 34.2844 | -1557.8465 | 1558.9061 | 5.3452 | -3.5556 | -7.0876 | 10.1605 | 1557.8465 | 3.5556 | 71.839
0.7744 | 20.9333 | 34.3198 | -1556.1747 | 1557.2107 | 5.3569 | -3.485 | -7.0876 | 10.1252 | 1556.1747 | 3.485 | 71.997
0.7762 | 20.5801 | 34.5788 | -1554.1025 | 1555.1268 | 5.3452 | -3.4143 | -7.0876 | 10.0781 | 1554.1025 | 3.4143 | 72.1644
0.778 | 20.7214 | 34.7789 | -1550.5352 | 1551.5595 | 5.3452 | -3.3554 | -7.0876 | 10.0546 | 1550.5352 | 3.3554 | 72.3317
0.7798 | 21.1099 | 35.191 | -1545.967 | 1546.9796 | 5.3334 | -3.2966 | -7.0876 | 10.0075 | 1545.967 | 3.2966 | 72.4991
0.7816 | 21.9105 | 35.5913 | -1540.9751 | 1542.0111 | 5.3216 | -3.2377 | -7.0994 | 9.9722 | 1540.9751 | 3.2377 | 72.6664
0.7834 | 22.5816 | 35.8268 | -1535.1707 | 1536.2304 | 5.2745 | -3.1788 | -7.0994 | 9.9133 | 1535.1707 | 3.1788 | 72.8338
0.7852 | 23.182 | 36.18 | -1528.1655 | 1529.2604 | 5.2039 | -3.0964 | -7.0876 | 9.8073 | 1528.1655 | 3.0964 | 73.0011
0.787 | 23.5588 | 36.3801 | -1521.5017 | 1522.6084 | 5.1215 | -2.9905 | -7.0876 | 9.7131 | 1521.5017 | 2.9905 | 73.1685
0.7888 | 23.4764 | 36.2742 | -1515.5325 | 1516.6392 | 5.0626 | -2.8492 | -7.0876 | 9.6072 | 1515.5325 | 2.8492 | 73.3358
0.7906 | 23.2173 | 36.7687 | -1509.3514 | 1510.4699 | 4.9331 | -2.6844 | -7.0523 | 9.4424 | 1509.3514 | 2.6844 | 73.5032
0.7924 | 22.9937 | 37.0277 | -1502.099 | 1503.2174 | 4.78 | -2.496 | -7.0052 | 9.2657 | 1502.099 | 2.496 | 73.6705
0.7942 | 22.8053 | 37.0394 | -1494.5168 | 1495.6235 | 4.6388 | -2.3076 | -6.9699 | 9.0774 | 1494.5168 | 2.3076 | 73.8379
0.796 | 22.5934 | 37.0041 | -1489.3836 | 1490.4785 | 4.5681 | -2.1781 | -6.9699 | 8.995 | 1489.3836 | 2.1781 | 74.0052
0.7977 | 22.8406 | 37.1454 | -1485.1333 | 1486.2283 | 4.5328 | -2.0839 | -6.9699 | 8.9596 | 1485.1333 | 2.0839 | 74.1633
0.7995 | 23.1703 | 37.3338 | -1481.9898 | 1483.0848 | 4.5093 | -2.0133 | -6.9699 | 8.9361 | 1481.9898 | 2.0133 | 74.3306
0.8013 | 23.1349 | 37.3573 | -1479.4821 | 1480.5417 | 4.4739 | -1.9309 | -6.9817 | 8.889 | 1479.4821 | 1.9309 | 74.498
0.8031 | 23.0996 | 37.3809 | -1476.9625 | 1478.0104 | 4.4386 | -1.8484 | -6.9935 | 8.8537 | 1476.9625 | 1.8484 | 74.6653
0.8049 | 22.9348 | 37.1807 | -1475.1494 | 1476.1737 | 4.3797 | -1.7543 | -7.0052 | 8.8066 | 1475.1494 | 1.7543 | 74.8327
0.8067 | 22.1106 | 36.2977 | -1474.1251 | 1475.0905 | 4.3209 | -1.6247 | -7.017 | 8.7359 | 1474.1251 | 1.6247 | 75
0.8085 | 20.9215 | 35.556 | -1472.6299 | 1473.56 | 4.262 | -1.4952 | -7.0288 | 8.6771 | 1472.6299 | 1.4952 | 75.1673
0.8103 | 20.2622 | 34.8967 | -1471.6056 | 1472.4886 | 4.2149 | -1.401 | -7.0523 | 8.6418 | 1471.6056 | 1.401 | 75.3347
0.8121 | 20.0621 | 34.8025 | -1469.5452 | 1470.4165 | 4.1678 | -1.3186 | -7.0641 | 8.6064 | 1469.5452 | 1.3186 | 75.502
0.8139 | 20.0738 | 34.826 | -1466.8726 | 1467.7321 | 4.1443 | -1.2715 | -7.0641 | 8.5829 | 1466.8726 | 1.2715 | 75.6694
0.8157 | 19.8501 | 34.6259 | -1463.4701 | 1464.306 | 4.109 | -1.2009 | -7.0641 | 8.5358 | 1463.4701 | 1.2009 | 75.8367
0.8175 | 19.8266 | 34.7201 | -1459.0668 | 1459.891 | 4.0501 | -1.1185 | -7.0641 | 8.4887 | 1459.0668 | 1.1185 | 76.0041
0.8193 | 19.8384 | 34.6376 | -1453.8041 | 1454.6046 | 3.9559 | -1.0243 | -7.0759 | 8.4181 | 1453.8041 | 1.0243 | 76.1714
0.8211 | 19.0966 | 33.1895 | -1447.3169 | 1448.0586 | 3.7911 | -0.85947 | -7.0641 | 8.3239 | 1447.3169 | 0.85947 | 76.3388
0.8228 | 19.0142 | 32.4125 | -1442.4544 | 1443.149 | 3.7204 | -0.7535 | -7.0641 | 8.2885 | 1442.4544 | 0.7535 | 76.4968
0.8246 | 18.9082 | 31.8238 | -1438.3808 | 1439.0519 | 3.6851 | -0.67109 | -7.0523 | 8.265 | 1438.3808 | 0.67109 | 76.6642
0.8264 | 18.402 | 30.8348 | -1434.6721 | 1435.3079 | 3.638 | -0.56513 | -7.0406 | 8.2297 | 1434.6721 | 0.56513 | 76.8315
0.8282 | 17.5425 | 30.1755 | -1431.5874 | 1432.1879 | 3.6145 | -0.47094 | -7.0288 | 8.2061 | 1431.5874 | 0.47094 | 76.9989
0.83 | 17.5661 | 29.6104 | -1429.3858 | 1429.9863 | 3.6262 | -0.43562 | -7.0288 | 8.2061 | 1429.3858 | 0.43562 | 77.1662
0.8318 | 17.6132 | 28.1622 | -1426.5484 | 1427.1135 | 3.6145 | -0.38853 | -7.017 | 8.1944 | 1426.5484 | 0.38853 | 77.3336
0.8336 | 17.3188 | 26.0901 | -1423.2165 | 1423.7227 | 3.5556 | -0.34143 | -6.9935 | 8.1355 | 1423.2165 | 0.34143 | 77.5009
0.8354 | 16.7537 | 24.7832 | -1420.9442 | 1421.4151 | 3.5203 | -0.27079 | -6.9935 | 8.0884 | 1420.9442 | 0.27079 | 77.6683
0.8372 | 16.165 | 23.3351 | -1419.2724 | 1419.708 | 3.485 | -0.23547 | -6.9817 | 8.0413 | 1419.2724 | 0.23547 | 77.8356
0.839 | 15.3644 | 22.2872 | -1417.9537 | 1418.3423 | 3.5203 | -0.16483 | -6.9581 | 8.0178 | 1417.9537 | 0.16483 | 78.003
0.8408 | 14.6933 | 20.9922 | -1416.8941 | 1417.2591 | 3.4967 | -0.10596 | -6.9464 | 7.9707 | 1416.8941 | 0.10596 | 78.1703
0.8426 | 14.0105 | 19.5793 | -1417.2002 | 1417.5417 | 3.4732 | -0.070641 | -6.9346 | 7.9353 | 1417.2002 | 0.070641 | 78.3377
0.8444 | 13.7161 | 18.2372 | -1418.4482 | 1418.7779 | 3.4496 | -0.047094 | -6.9228 | 7.8882 | 1418.4482 | 0.047094 | 78.505
0.8462 | 13.6219 | 16.9421 | -1419.4372 | 1419.7786 | 3.4379 | -0.023547 | -6.8993 | 7.8647 | 1419.4372 | 0.023547 | 78.6724
0.8479 | 13.9163 | 15.7765 | -1419.5667 | 1419.9081 | 3.4379 | 0.035321 | -6.864 | 7.8529 | 1419.5667 | -0.035321 | 78.8304
0.8497 | 14.3872 | 15.1407 | -1419.084 | 1419.449 | 3.4496 | 0.058868 | -6.8404 | 7.8529 | 1419.084 | -0.058868 | 78.9978
0.8515 | 14.6345 | 13.8456 | -1418.6013 | 1418.978 | 3.4614 | 0.058868 | -6.8169 | 7.8176 | 1418.6013 | -0.058868 | 79.1651
0.8533 | 14.7404 | 12.8096 | -1417.8478 | 1418.2363 | 3.4614 | 0.058868 | -6.7815 | 7.7941 | 1417.8478 | -0.058868 | 79.3325
0.8551 | 14.6698 | 12.1032 | -1416.7646 | 1417.1414 | 3.485 | 0.047094 | -6.7698 | 7.7705 | 1416.7646 | -0.047094 | 79.4998
0.8569 | 14.8699 | 11.6204 | -1415.4224 | 1415.7992 | 3.4967 | 0.035321 | -6.758 | 7.7587 | 1415.4224 | -0.035321 | 79.6672
0.8587 | 15.3644 | 11.3968 | -1413.7859 | 1414.1627 | 3.5085 | 0.023547 | -6.758 | 7.7587 | 1413.7859 | -0.023547 | 79.8345
0.8605 | 15.647 | 10.9611 | -1411.596 | 1411.9728 | 3.5438 | 0.011774 | -6.7462 | 7.7587 | 1411.596 | -0.011774 | 80.0019
0.8623 | 15.7765 | 10.4078 | -1409.1118 | 1409.4768 | 3.5909 | 0.011774 | -6.7344 | 7.7587 | 1409.1118 | -0.011774 | 80.1692
0.8641 | 16.424 | 10.2194 | -1406.2038 | 1406.5805 | 3.6145 | -0.047094 | -6.7344 | 7.747 | 1406.2038 | 0.047094 | 80.3366
0.8659 | 17.2717 | 10.4549 | -1402.4598 | 1402.8366 | 3.6262 | -0.10596 | -6.7227 | 7.7352 | 1402.4598 | 0.10596 | 80.5039
0.8677 | 17.5543 | 10.29 | -1398.8689 | 1399.2456 | 3.6498 | -0.14128 | -6.7227 | 7.7352 | 1398.8689 | 0.14128 | 80.6713
0.8695 | 17.8251 | 10.137 | -1395.3133 | 1395.69 | 3.6851 | -0.1766 | -6.7109 | 7.7352 | 1395.3133 | 0.1766 | 80.8386
0.8712 | 17.9664 | 10.29 | -1393.0292 | 1393.3942 | 3.6851 | -0.2237 | -6.7109 | 7.7352 | 1393.0292 | 0.2237 | 80.9967
0.873 | 17.6249 | 10.29 | -1391.3338 | 1391.687 | 3.6851 | -0.2237 | -6.7109 | 7.7352 | 1391.3338 | 0.2237 | 81.164
0.8748 | 16.8126 | 10.1252 | -1390.3566 | 1390.6863 | 3.6851 | -0.16483 | -6.7109 | 7.7352 | 1390.3566 | 0.16483 | 81.3313
0.8766 | 16.0944 | 10.0781 | -1389.8151 | 1390.1094 | 3.6851 | -0.12951 | -6.7109 | 7.7234 | 1389.8151 | 0.12951 | 81.4987
0.8784 | 15.906 | 11.1024 | -1391.3574 | 1391.6282 | 3.7087 | -0.16483 | -6.7109 | 7.7587 | 1391.3574 | 0.16483 | 81.666
0.8802 | 15.282 | 12.1738 | -1393.6885 | 1393.9476 | 3.7793 | -0.21192 | -6.7227 | 7.8294 | 1393.6885 | 0.21192 | 81.8334
0.882 | 14.6462 | 12.9038 | -1394.9483 | 1395.2073 | 3.8617 | -0.21192 | -6.7344 | 7.9 | 1394.9483 | 0.21192 | 82.0007
0.8838 | 14.0693 | 13.5042 | -1395.6547 | 1395.902 | 3.9324 | -0.20015 | -6.7462 | 7.9589 | 1395.6547 | 0.20015 | 82.1681
0.8856 | 14.4814 | 13.9634 | -1394.3008 | 1394.5598 | 3.9794 | -0.21192 | -6.7462 | 7.9824 | 1394.3008 | 0.21192 | 82.3354
0.8874 | 15.176 | 14.5167 | -1391.9343 | 1392.1933 | 4.0148 | -0.2237 | -6.758 | 8.0178 | 1391.9343 | 0.2237 | 82.5028
0.8892 | 15.7529 | 15.1054 | -1389.0733 | 1389.3441 | 4.0501 | -0.23547 | -6.758 | 8.0413 | 1389.0733 | 0.23547 | 82.6701
0.891 | 16.5065 | 15.9296 | -1385.6001 | 1385.8709 | 4.109 | -0.25902 | -6.7698 | 8.1119 | 1385.6001 | 0.25902 | 82.8375
0.8928 | 16.9892 | 16.5771 | -1382.4095 | 1382.6921 | 4.156 | -0.25902 | -6.7933 | 8.1473 | 1382.4095 | 0.25902 | 83.0048
0.8946 | 16.6713 | 16.7655 | -1380.0901 | 1380.3727 | 4.1678 | -0.21192 | -6.7933 | 8.159 | 1380.0901 | 0.21192 | 83.1722
0.8963 | 15.8118 | 16.636 | -1378.4419 | 1378.7126 | 4.156 | -0.12951 | -6.7933 | 8.159 | 1378.4419 | 0.12951 | 83.3302
0.8981 | 15.3173 | 16.7655 | -1377.7472 | 1378.0062 | 4.156 | -0.070641 | -6.8051 | 8.159 | 1377.7472 | 0.070641 | 83.4976
0.8999 | 15.0348 | 16.9421 | -1378.1004 | 1378.3594 | 4.156 | -0.047094 | -6.8169 | 8.1708 | 1378.1004 | 0.047094 | 83.6649
0.9017 | 14.7522 | 17.3777 | -1379.5839 | 1379.8429 | 4.1443 | 0 | -6.8522 | 8.1708 | 1379.5839 | 0 | 83.8323
0.9035 | 14.6109 | 17.8722 | -1380.773 | 1381.032 | 4.1207 | 0.047094 | -6.8757 | 8.1708 | 1380.773 | -0.047094 | 83.9996
0.9053 | 14.8228 | 18.3313 | -1380.3256 | 1380.5964 | 4.109 | 0.070641 | -6.8875 | 8.159 | 1380.3256 | -0.070641 | 84.167
0.9071 | 15.2231 | 18.8965 | -1378.9481 | 1379.2425 | 4.109 | 0.082415 | -6.8993 | 8.159 | 1378.9481 | -0.082415 | 84.3343
0.9089 | 14.7051 | 19.0849 | -1377.3351 | 1377.6295 | 4.0972 | 0.15306 | -6.8993 | 8.159 | 1377.3351 | -0.15306 | 84.5017
0.9107 | 14.14 | 19.2615 | -1375.7222 | 1376.0047 | 4.0736 | 0.2237 | -6.8993 | 8.159 | 1375.7222 | -0.2237 | 84.669
0.9125 | 14.3754 | 19.9561 | -1374.015 | 1374.3094 | 4.0619 | 0.25902 | -6.8993 | 8.1473 | 1374.015 | -0.25902 | 84.8364
0.9143 | 14.7404 | 20.3328 | -1371.3189 | 1371.625 | 4.0383 | 0.28256 | -6.911 | 8.1355 | 1371.3189 | -0.28256 | 85.0037
0.9161 | 14.5521 | 20.4624 | -1367.3865 | 1367.7044 | 3.9559 | 0.35321 | -6.9228 | 8.0884 | 1367.3865 | -0.35321 | 85.1711
0.9179 | 14.3754 | 20.533 | -1363.9958 | 1364.3136 | 3.8853 | 0.42385 | -6.9346 | 8.0531 | 1363.9958 | -0.42385 | 85.3384
0.9197 | 14.034 | 20.3682 | -1361.6764 | 1361.9825 | 3.8499 | 0.47094 | -6.9346 | 8.0413 | 1361.6764 | -0.47094 | 85.5058
0.9214 | 13.5042 | 20.2975 | -1359.8162 | 1360.1223 | 3.8146 | 0.52981 | -6.9464 | 8.0295 | 1359.8162 | -0.52981 | 85.6638
0.9232 | 13.2216 | 20.6507 | -1358.733 | 1359.0273 | 3.8028 | 0.5769 | -6.9464 | 8.0295 | 1358.733 | -0.5769 | 85.8312
0.925 | 13.5631 | 21.3454 | -1356.5902 | 1356.9081 | 3.8146 | 0.58868 | -6.9581 | 8.0413 | 1356.5902 | -0.58868 | 85.9985
0.9268 | 14.0105 | 22.0047 | -1354.0118 | 1354.3533 | 3.8264 | 0.58868 | -6.9581 | 8.0531 | 1354.0118 | -0.58868 | 86.1659
0.9286 | 13.9398 | 22.2401 | -1350.8448 | 1351.198 | 3.8146 | 0.624 | -6.9581 | 8.0531 | 1350.8448 | -0.624 | 86.3332
0.9304 | 13.6808 | 22.4403 | -1348.3017 | 1348.6549 | 3.8028 | 0.67109 | -6.9699 | 8.0531 | 1348.3017 | -0.67109 | 86.5006
0.9322 | 13.3865 | 22.6404 | -1346.7594 | 1347.1008 | 3.7793 | 0.71818 | -6.9699 | 8.0531 | 1346.7594 | -0.71818 | 86.6679
0.934 | 13.3158 | 22.7464 | -1345.2641 | 1345.6173 | 3.7675 | 0.7535 | -6.9699 | 8.0413 | 1345.2641 | -0.7535 | 86.8353
0.9358 | 12.9509 | 22.8171 | -1344.5342 | 1344.8874 | 3.7322 | 0.8006 | -6.9699 | 8.0295 | 1344.5342 | -0.8006 | 87.0026
0.9376 | 12.8567 | 22.8995 | -1344.8756 | 1345.217 | 3.7087 | 0.81237 | -6.9699 | 8.0295 | 1344.8756 | -0.81237 | 87.17
0.9394 | 13.1392 | 22.9112 | -1344.7696 | 1345.1111 | 3.6969 | 0.81237 | -6.9699 | 8.0178 | 1344.7696 | -0.81237 | 87.3373
0.9412 | 13.3041 | 22.8288 | -1344.4282 | 1344.7814 | 3.6851 | 0.82415 | -6.9699 | 8.006 | 1344.4282 | -0.82415 | 87.5046
0.943 | 13.2334 | 22.8759 | -1344.5106 | 1344.8638 | 3.6733 | 0.83592 | -6.9699 | 8.006 | 1344.5106 | -0.83592 | 87.672
0.9447 | 13.0568 | 22.9819 | -1344.8638 | 1345.217 | 3.6498 | 0.85947 | -6.9699 | 8.006 | 1344.8638 | -0.85947 | 87.83
0.9465 | 12.9979 | 23.3822 | -1345.6409 | 1345.9941 | 3.6498 | 0.85947 | -6.9817 | 8.006 | 1345.6409 | -0.85947 | 87.9974
0.9483 | 13.3512 | 23.5823 | -1346.0294 | 1346.3944 | 3.6498 | 0.83592 | -6.9935 | 8.0178 | 1346.0294 | -0.83592 | 88.1647
0.9501 | 13.5395 | 23.6412 | -1346.2413 | 1346.6063 | 3.638 | 0.83592 | -6.9935 | 8.006 | 1346.2413 | -0.83592 | 88.3321
0.9519 | 13.669 | 23.5588 | -1345.9941 | 1346.3591 | 3.6145 | 0.83592 | -7.0052 | 8.006 | 1345.9941 | -0.83592 | 88.4994
0.9537 | 13.9163 | 23.6883 | -1345.4172 | 1345.7822 | 3.6027 | 0.83592 | -7.0052 | 8.006 | 1345.4172 | -0.83592 | 88.6668
0.9555 | 14.1047 | 23.7472 | -1344.0868 | 1344.4635 | 3.5791 | 0.83592 | -7.017 | 8.0178 | 1344.0868 | -0.83592 | 88.8341
0.9573 | 15.0465 | 23.8531 | -1342.4385 | 1342.8035 | 3.5203 | 0.83592 | -7.0523 | 8.0178 | 1342.4385 | -0.83592 | 89.0015
0.9591 | 16.0002 | 23.8767 | -1339.8483 | 1340.2133 | 3.4967 | 0.81237 | -7.0759 | 8.0178 | 1339.8483 | -0.81237 | 89.1688
0.9609 | 16.5771 | 23.7001 | -1336.9873 | 1337.3641 | 3.4732 | 0.8006 | -7.0876 | 8.006 | 1336.9873 | -0.8006 | 89.3362
0.9627 | 16.6242 | 23.4881 | -1335.916 | 1336.2927 | 3.4614 | 0.81237 | -7.0876 | 8.006 | 1335.916 | -0.81237 | 89.5035
0.9645 | 16.636 | 23.3115 | -1335.6334 | 1336.0101 | 3.4496 | 0.82415 | -7.0876 | 7.9942 | 1335.6334 | -0.82415 | 89.6709
0.9663 | 16.6477 | 23.1585 | -1335.9042 | 1336.2927 | 3.4379 | 0.83592 | -7.0876 | 7.9824 | 1335.9042 | -0.83592 | 89.8382
0.9681 | 16.6831 | 23.0172 | -1336.2338 | 1336.6106 | 3.4143 | 0.83592 | -7.0876 | 7.9824 | 1336.2338 | -0.83592 | 90.0056
0.9698 | 16.2828 | 22.5816 | -1336.2103 | 1336.5753 | 3.3908 | 0.85947 | -7.0876 | 7.9824 | 1336.2103 | -0.85947 | 90.1636
0.9716 | 16.0002 | 22.4403 | -1336.0219 | 1336.3869 | 3.379 | 0.88301 | -7.0876 | 7.9824 | 1336.0219 | -0.88301 | 90.331
0.9734 | 15.8236 | 22.4285 | -1335.7747 | 1336.1397 | 3.379 | 0.89479 | -7.0876 | 7.9824 | 1335.7747 | -0.89479 | 90.4983
0.9752 | 15.7765 | 22.4168 | -1335.916 | 1336.2927 | 3.3672 | 0.90656 | -7.0876 | 7.9707 | 1335.916 | -0.90656 | 90.6657
0.977 | 15.906 | 22.4521 | -1336.175 | 1336.5517 | 3.3554 | 0.89479 | -7.0994 | 7.9707 | 1336.175 | -0.89479 | 90.833
0.9788 | 16.1532 | 22.5698 | -1335.9748 | 1336.3516 | 3.3437 | 0.88301 | -7.0994 | 7.9707 | 1335.9748 | -0.88301 | 91.0004
0.9806 | 16.4711 | 22.8524 | -1336.434 | 1336.8343 | 3.3437 | 0.87124 | -7.0994 | 7.9707 | 1336.434 | -0.87124 | 91.1677
0.9824 | 16.636 | 22.9466 | -1336.9403 | 1337.3405 | 3.3319 | 0.87124 | -7.1112 | 7.9707 | 1336.9403 | -0.87124 | 91.3351
0.9842 | 16.2121 | 22.7229 | -1337.8233 | 1338.2236 | 3.3201 | 0.89479 | -7.1112 | 7.9707 | 1337.8233 | -0.89479 | 91.5024
0.986 | 15.5528 | 22.4638 | -1339.4833 | 1339.8719 | 3.3084 | 0.94188 | -7.1112 | 7.9707 | 1339.4833 | -0.94188 | 91.6698
0.9878 | 14.8817 | 22.2284 | -1342.2148 | 1342.6033 | 3.2966 | 0.9772 | -7.123 | 7.9824 | 1342.2148 | -0.9772 | 91.8371
0.9896 | 14.4343 | 22.0753 | -1345.2759 | 1345.6762 | 3.2966 | 1.0007 | -7.1347 | 7.9942 | 1345.2759 | -1.0007 | 92.0045
0.9914 | 14.505 | 22.2401 | -1347.2303 | 1347.6306 | 3.2966 | 1.0007 | -7.1347 | 8.006 | 1347.2303 | -1.0007 | 92.1718
0.9932 | 15.1172 | 22.5463 | -1347.9956 | 1348.4076 | 3.3084 | 0.96543 | -7.1465 | 8.006 | 1347.9956 | -0.96543 | 92.3392
0.9949 | 15.9413 | 22.7935 | -1347.819 | 1348.2546 | 3.2966 | 0.91833 | -7.1583 | 8.006 | 1347.819 | -0.91833 | 92.4972
0.9967 | 16.3652 | 22.8641 | -1347.1243 | 1347.5717 | 3.2848 | 0.89479 | -7.1583 | 8.006 | 1347.1243 | -0.89479 | 92.6646
0.9985 | 16.1297 | 22.9701 | -1346.3826 | 1346.8182 | 3.2613 | 0.90656 | -7.1701 | 8.006 | 1346.3826 | -0.90656 | 92.8319
1.0003 | 15.9178 | 23.0878 | -1345.6055 | 1346.0412 | 3.2377 | 0.91833 | -7.1701 | 8.006 | 1345.6055 | -0.91833 | 92.9993
1.0021 | 15.8236 | 23.2291 | -1344.7343 | 1345.1699 | 3.2024 | 0.91833 | -7.1818 | 7.9942 | 1344.7343 | -0.91833 | 93.1666
1.0039 | 15.7412 | 23.441 | -1344.3105 | 1344.7579 | 3.1788 | 0.91833 | -7.1936 | 7.9942 | 1344.3105 | -0.91833 | 93.334
1.0057 | 15.6941 | 23.7354 | -1344.3929 | 1344.852 | 3.1671 | 0.91833 | -7.1936 | 7.9942 | 1344.3929 | -0.91833 | 93.5013
1.0075 | 15.647 | 24.0297 | -1344.4871 | 1344.9462 | 3.1435 | 0.91833 | -7.2054 | 7.9942 | 1344.4871 | -0.91833 | 93.6687
1.0093 | 15.6117 | 24.1239 | -1344.9816 | 1345.4407 | 3.1318 | 0.91833 | -7.2054 | 7.9942 | 1344.9816 | -0.91833 | 93.836
1.0111 | 15.5763 | 24.2063 | -1345.4878 | 1345.947 | 3.1318 | 0.91833 | -7.2172 | 7.9942 | 1345.4878 | -0.91833 | 94.0033
1.0129 | 15.4939 | 24.3123 | -1346.4886 | 1346.9477 | 3.12 | 0.91833 | -7.2172 | 7.9942 | 1346.4886 | -0.91833 | 94.1707
1.0147 | 15.4233 | 24.5007 | -1347.9014 | 1348.3723 | 3.12 | 0.90656 | -7.2289 | 8.006 | 1347.9014 | -0.90656 | 94.338
1.0165 | 15.4233 | 24.7715 | -1349.4202 | 1349.8911 | 3.1318 | 0.89479 | -7.2289 | 8.006 | 1349.4202 | -0.89479 | 94.5054
1.0182 | 15.9884 | 25.0422 | -1350.9272 | 1351.4217 | 3.1435 | 0.84769 | -7.2407 | 8.006 | 1350.9272 | -0.84769 | 94.6634
1.02 | 16.2357 | 25.1247 | -1351.9632 | 1352.4577 | 3.1553 | 0.83592 | -7.2407 | 8.0178 | 1351.9632 | -0.83592 | 94.8308
1.0218 | 16.0708 | 24.8539 | -1353.0229 | 1353.5056 | 3.1553 | 0.83592 | -7.2525 | 8.0295 | 1353.0229 | -0.83592 | 94.9981
1.0236 | 15.6941 | 24.7597 | -1354.6711 | 1355.1421 | 3.1553 | 0.84769 | -7.2643 | 8.0295 | 1354.6711 | -0.84769 | 95.1655
1.0254 | 15.7176 | 25.1835 | -1355.3305 | 1355.8132 | 3.1435 | 0.84769 | -7.2643 | 8.0413 | 1355.3305 | -0.84769 | 95.3328
1.0272 | 15.7647 | 25.3837 | -1356.0369 | 1356.5314 | 3.1435 | 0.83592 | -7.276 | 8.0413 | 1356.0369 | -0.83592 | 95.5002
1.029 | 16.2592 | 25.3601 | -1356.4607 | 1356.9552 | 3.1435 | 0.81237 | -7.276 | 8.0413 | 1356.4607 | -0.81237 | 95.6675
1.0308 | 16.6948 | 25.4425 | -1356.4136 | 1356.9199 | 3.1435 | 0.77705 | -7.2878 | 8.0413 | 1356.4136 | -0.77705 | 95.8349
1.0326 | 17.2011 | 25.7604 | -1355.825 | 1356.343 | 3.1553 | 0.7535 | -7.2878 | 8.0531 | 1355.825 | -0.7535 | 96.0022
1.0344 | 17.7662 | 25.8899 | -1354.4828 | 1355.0244 | 3.1553 | 0.72996 | -7.2878 | 8.0531 | 1354.4828 | -0.72996 | 96.1696
1.0362 | 18.3196 | 25.9135 | -1352.6932 | 1353.2348 | 3.1435 | 0.70641 | -7.2996 | 8.0531 | 1352.6932 | -0.70641 | 96.3369
1.038 | 18.2372 | 25.6662 | -1349.8087 | 1350.3503 | 3.1435 | 0.72996 | -7.2878 | 8.0531 | 1349.8087 | -0.72996 | 96.5043
1.0398 | 18.037 | 25.3013 | -1345.9234 | 1346.465 | 3.1553 | 0.7535 | -7.276 | 8.0413 | 1345.9234 | -0.7535 | 96.6716
1.0416 | 17.4954 | 25.2895 | -1344.6637 | 1345.1935 | 3.1318 | 0.78882 | -7.276 | 8.0413 | 1344.6637 | -0.78882 | 96.839
1.0433 | 17.4483 | 25.5132 | -1343.4275 | 1343.969 | 3.12 | 0.8006 | -7.2878 | 8.0413 | 1343.4275 | -0.8006 | 96.997
1.0451 | 17.3777 | 25.7016 | -1342.8035 | 1343.345 | 3.1082 | 0.81237 | -7.2996 | 8.0413 | 1342.8035 | -0.81237 | 97.1644
1.0469 | 17.9193 | 26.0783 | -1342.4032 | 1342.9683 | 3.12 | 0.76528 | -7.2996 | 8.0531 | 1342.4032 | -0.76528 | 97.3317
1.0487 | 17.8251 | 26.0194 | -1341.6732 | 1342.2383 | 3.1082 | 0.77705 | -7.3113 | 8.0413 | 1341.6732 | -0.77705 | 97.4991
1.0505 | 17.1304 | 25.8664 | -1341.0257 | 1341.579 | 3.1082 | 0.82415 | -7.2996 | 8.0531 | 1341.0257 | -0.82415 | 97.6664
1.0523 | 17.0245 | 25.8428 | -1341.4495 | 1342.0029 | 3.1082 | 0.81237 | -7.3113 | 8.0531 | 1341.4495 | -0.81237 | 97.8338
1.0541 | 16.8126 | 25.7369 | -1341.7321 | 1342.2736 | 3.0964 | 0.81237 | -7.3113 | 8.0531 | 1341.7321 | -0.81237 | 98.0011
1.0559 | 16.5535 | 25.9135 | -1341.7674 | 1342.3207 | 3.0964 | 0.83592 | -7.3113 | 8.0531 | 1341.7674 | -0.83592 | 98.1685
1.0577 | 16.3181 | 26.1136 | -1341.8969 | 1342.4502 | 3.0847 | 0.84769 | -7.3113 | 8.0531 | 1341.8969 | -0.84769 | 98.3358
1.0595 | 16.1768 | 26.1725 | -1341.838 | 1342.4032 | 3.0847 | 0.84769 | -7.3113 | 8.0531 | 1341.838 | -0.84769 | 98.5032
1.0613 | 16.3652 | 26.4197 | -1342.2501 | 1342.8152 | 3.0847 | 0.82415 | -7.3231 | 8.0649 | 1342.2501 | -0.82415 | 98.6705
1.0631 | 16.8008 | 26.9142 | -1342.827 | 1343.4039 | 3.0964 | 0.78882 | -7.3349 | 8.0649 | 1342.827 | -0.78882 | 98.8379
1.0649 | 16.9068 | 27.0555 | -1342.8623 | 1343.4628 | 3.0964 | 0.78882 | -7.3467 | 8.0766 | 1342.8623 | -0.78882 | 99.0052
1.0667 | 16.9538 | 27.1732 | -1342.7917 | 1343.4039 | 3.0964 | 0.78882 | -7.3467 | 8.0884 | 1342.7917 | -0.78882 | 99.1726
1.0684 | 16.9774 | 27.2674 | -1342.6386 | 1343.2508 | 3.12 | 0.8006 | -7.3467 | 8.1002 | 1342.6386 | -0.8006 | 99.3306
1.0702 | 17.154 | 27.2674 | -1342.4973 | 1343.1096 | 3.12 | 0.78882 | -7.3584 | 8.1002 | 1342.4973 | -0.78882 | 99.498
1.072 | 16.895 | 26.8436 | -1342.0382 | 1342.6386 | 3.12 | 0.81237 | -7.3702 | 8.1119 | 1342.0382 | -0.81237 | 99.6653
1.0738 | 16.3769 | 26.6199 | -1342.0029 | 1342.5915 | 3.12 | 0.83592 | -7.3702 | 8.1119 | 1342.0029 | -0.83592 | 99.8327
1.0756 | 16.271 | 26.7729 | -1342.0029 | 1342.5798 | 3.12 | 0.83592 | -7.3702 | 8.1237 | 1342.0029 | -0.83592 | 100
| Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My
Minimum | -59.3502 | -123.2333 | -2057.5022 | 150.5831 | -1.1067 | -3.9206 | -7.3702 | 0.61222 | 149.4058 | -13.1275
Maximum | 62.3643 | 78.2938 | -149.4058 | 2060.01 | 6.0398 | 13.1275 | -0.34143 | 15.176 | 2057.5022 | 3.9206

## Diagr_AVER75
AVER75 = Loads at Average Level in Subject with 75 kg Body Weight.
Contact forces and moments, measured telemetrically in knee implants of 8 subjects.
Averages of one individual subject.
Fx  lateral, Fy  anterior, -Fz  distal, Fres = Resultant Force
Mx around x, My around y, Mz around z, Mres = Resultant Moment
z-axis = axis of implant stem!
From Bergmann et al, PlosOne (2014) "Standardized Loads Acting in Knee Implants"
Subject K9L
Subject K9L
G.Bergmann et al., PlosOne, 2014

## Dat_HIGH100
Data File: | k9l_2246_863_6_i1.akf
Curve Alignment Program: | 1.4.11
------------------------MEASURING INFOS------------------------
Diagram Title #1 | Forces and Moments at Knee Joint
Diagram Title #2 | Measured Relative to Implant
Comment #1 | Knee Joint; Gaitanalysis; Stand up; Seat Height: 45 cm;
Comment #2 | K9L, 14,5 Months PO *
BodyWeight [N]: | 1070
Patient(en): | K9L | [7 empty] | Subject | K9L
Implant Type: | Knee Joint
Varus Angle Leg [°] | 7 | [7 empty] | HIGH100 Loads in N (forces) and Nm (moments)
Axes Directions X, Y, Z: | Lateral, Anterior, Superior
Axes Origin: | Tibial Plateau at Stem Axis | [7 empty] | = AVER75 Loads * 1/Caver
Transformation X Y Z (°): | 6.5 | -3.5 | -2
Displacement X, Y, Z (mm): | 0 | 0 | 0 | [5 empty] | Caver = | 0.54
Activity Code: | KneeActivities | 1P_1_15_8
---------------------------DATA--------------------------------
Number of Data Sets = | 601
Number of analog channels = | 0
Time-Offset(sec) = | 0
Time | Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My | Time
[s] | [N] | [N] | [N] | [N] | [Nm] | [Nm] | [Nm] | [Nm] | [N] | (Nm] | [%Cycle]
0 | -2.4855 | -16.8972 | -330.3776 | 330.9881 | 0.23983 | -0.7849 | -0.63228 | 1.1337 | 330.3776 | 0.7849 | 0
0.0018 | -2.6599 | -15.9378 | -332.7759 | 333.2992 | 0.26163 | -0.7849 | -0.63228 | 1.1337 | 332.7759 | 0.7849 | 0.16735
0.0036 | -3.3576 | -15.1311 | -337.0929 | 337.5725 | 0.28344 | -0.7631 | -0.65408 | 1.1337 | 337.0929 | 0.7631 | 0.3347
0.0054 | -3.9899 | -14.1064 | -340.4069 | 340.8211 | 0.32704 | -0.74129 | -0.65408 | 1.1337 | 340.4069 | 0.74129 | 0.50205
0.0072 | -4.3606 | -12.4494 | -340.5813 | 340.9302 | 0.37065 | -0.7631 | -0.65408 | 1.1774 | 340.5813 | 0.7631 | 0.66939
0.009 | -3.8809 | -12.384 | -336.853 | 337.2019 | 0.37065 | -0.7631 | -0.65408 | 1.1774 | 336.853 | 0.7631 | 0.83674
0.0108 | -3.5975 | -12.4276 | -334.3021 | 334.6291 | 0.34884 | -0.7631 | -0.67589 | 1.1774 | 334.3021 | 0.7631 | 1.0041
0.0125 | -4.2515 | -11.054 | -336.8312 | 337.1365 | 0.37065 | -0.74129 | -0.67589 | 1.1774 | 336.8312 | 0.74129 | 1.1621
0.0143 | -4.644 | -9.5278 | -337.5725 | 337.8342 | 0.41425 | -0.74129 | -0.69769 | 1.1992 | 337.5725 | 0.74129 | 1.3295
0.0161 | -4.6004 | -7.9362 | -335.2614 | 335.5013 | 0.45786 | -0.74129 | -0.69769 | 1.221 | 335.2614 | 0.74129 | 1.4968
0.0179 | -4.1861 | -6.7153 | -330.8137 | 331.0099 | 0.47966 | -0.7631 | -0.69769 | 1.2646 | 330.8137 | 0.7631 | 1.6642
0.0197 | -3.423 | -6.0612 | -325.1231 | 325.2976 | 0.47966 | -0.7631 | -0.71949 | 1.2646 | 325.1231 | 0.7631 | 1.8315
0.0215 | -2.3329 | -5.1673 | -317.1433 | 317.2959 | 0.47966 | -0.8067 | -0.71949 | 1.3082 | 317.1433 | 0.8067 | 1.9989
0.0233 | -0.93752 | -5.5597 | -308.9891 | 309.1417 | 0.41425 | -0.85031 | -0.74129 | 1.33 | 308.9891 | 0.85031 | 2.1662
0.0251 | 0.41425 | -6.7153 | -302.252 | 302.4264 | 0.32704 | -0.89391 | -0.74129 | 1.3518 | 302.252 | 0.89391 | 2.3336
0.0269 | 1.4608 | -8.4595 | -297.8697 | 298.1095 | 0.21803 | -0.91572 | -0.7631 | 1.3518 | 297.8697 | 0.91572 | 2.5009
0.0287 | 2.0495 | -9.506 | -296.8231 | 297.1066 | 0.13082 | -0.91572 | -0.7631 | 1.33 | 296.8231 | 0.91572 | 2.6683
0.0305 | 2.5509 | -10.2691 | -298.6546 | 298.9598 | 0.043606 | -0.93752 | -0.7849 | 1.3518 | 298.6546 | 0.93752 | 2.8356
0.0323 | 2.6599 | -9.8985 | -302.5136 | 302.8407 | 0.021803 | -0.95932 | -0.8067 | 1.3518 | 302.5136 | 0.95932 | 3.003
0.0341 | 2.5509 | -8.8301 | -308.771 | 309.0763 | 0.021803 | -0.95932 | -0.8067 | 1.3954 | 308.771 | 0.95932 | 3.1703
0.0359 | 1.8096 | -6.4318 | -316.293 | 316.5328 | 0.065408 | -0.98113 | -0.82851 | 1.4172 | 316.293 | 0.98113 | 3.3377
0.0376 | 1.0029 | -3.8155 | -322.812 | 322.9865 | 0.13082 | -0.98113 | -0.82851 | 1.4608 | 322.812 | 0.98113 | 3.4957
0.0394 | 0.50146 | -0.41425 | -326.1697 | 326.3441 | 0.19623 | -1.0247 | -0.85031 | 1.5044 | 326.1697 | 1.0247 | 3.6631
0.0412 | 0.13082 | 3.0524 | -328.2845 | 328.459 | 0.26163 | -1.0465 | -0.85031 | 1.548 | 328.2845 | 1.0465 | 3.8304
0.043 | 0.087211 | 6.0612 | -327.1726 | 327.4124 | 0.30524 | -1.0901 | -0.85031 | 1.5916 | 327.1726 | 1.0901 | 3.9978
0.0448 | 0.26163 | 8.3069 | -324.0548 | 324.3382 | 0.32704 | -1.1555 | -0.85031 | 1.6352 | 324.0548 | 1.1555 | 4.1651
0.0466 | 0.58868 | 10.618 | -319.1928 | 319.5634 | 0.34884 | -1.221 | -0.85031 | 1.6788 | 319.1928 | 1.221 | 4.3325
0.0484 | 1.1337 | 11.9479 | -312.6519 | 313.1098 | 0.32704 | -1.2646 | -0.85031 | 1.7006 | 312.6519 | 1.2646 | 4.4998
0.0502 | 2.0713 | 12.6674 | -304.3669 | 304.9338 | 0.26163 | -1.3082 | -0.82851 | 1.7224 | 304.3669 | 1.3082 | 4.6672
0.052 | 3.0088 | 12.9291 | -297.1938 | 297.826 | 0.19623 | -1.3736 | -0.82851 | 1.7442 | 297.1938 | 1.3736 | 4.8345
0.0538 | 3.8809 | 12.0787 | -292.0701 | 292.7024 | 0.087211 | -1.4172 | -0.85031 | 1.7878 | 292.0701 | 1.4172 | 5.0019
0.0556 | 4.6876 | 11.0104 | -288.5817 | 289.2139 | -0.021803 | -1.4826 | -0.85031 | 1.8096 | 288.5817 | 1.4826 | 5.1692
0.0574 | 5.1891 | 10.1601 | -288.8869 | 289.541 | -0.13082 | -1.5262 | -0.87211 | 1.875 | 288.8869 | 1.5262 | 5.3366
0.0592 | 5.3853 | 9.7458 | -290.7837 | 291.4596 | -0.21803 | -1.548 | -0.87211 | 1.8968 | 290.7837 | 1.548 | 5.5039
0.061 | 5.4725 | 10.2255 | -294.6864 | 295.4059 | -0.28344 | -1.5916 | -0.89391 | 1.9623 | 294.6864 | 1.5916 | 5.6713
0.0627 | 5.1237 | 12.275 | -300.5296 | 301.3363 | -0.28344 | -1.6134 | -0.91572 | 2.0059 | 300.5296 | 1.6134 | 5.8293
0.0645 | 4.7094 | 15.6108 | -306.4381 | 307.5065 | -0.23983 | -1.657 | -0.93752 | 2.0713 | 306.4381 | 1.657 | 5.9967
0.0663 | 4.2951 | 18.9684 | -309.4251 | 310.7115 | -0.21803 | -1.7224 | -0.95932 | 2.1149 | 309.4251 | 1.7224 | 6.164
0.0681 | 3.9681 | 22.0862 | -310.8205 | 312.2595 | -0.17442 | -1.766 | -0.95932 | 2.1803 | 310.8205 | 1.766 | 6.3313
0.0699 | 4.0335 | 25.0514 | -308.7928 | 310.4499 | -0.13082 | -1.8314 | -0.98113 | 2.2457 | 308.7928 | 1.8314 | 6.4987
0.0717 | 4.2733 | 27.8204 | -305.7187 | 307.6155 | -0.10901 | -1.8968 | -1.0029 | 2.3111 | 305.7187 | 1.8968 | 6.666
0.0735 | 4.753 | 29.5864 | -300.7912 | 302.8843 | -0.13082 | -1.9841 | -1.0029 | 2.3765 | 300.7912 | 1.9841 | 6.8334
0.0753 | 5.5379 | 31.0908 | -294.0978 | 296.3217 | -0.13082 | -2.0495 | -1.0247 | 2.4419 | 294.0978 | 2.0495 | 7.0007
0.0771 | 6.4754 | 30.9382 | -286.4886 | 288.7343 | -0.21803 | -2.1149 | -1.0465 | 2.4855 | 286.4886 | 2.1149 | 7.1681
0.0789 | 7.3911 | 30.6111 | -279.2719 | 281.5612 | -0.28344 | -2.1803 | -1.0683 | 2.5073 | 279.2719 | 2.1803 | 7.3354
0.0807 | 7.8708 | 29.8698 | -276.9608 | 279.2065 | -0.41425 | -2.2457 | -1.0683 | 2.5727 | 276.9608 | 2.2457 | 7.5028
0.0825 | 8.0888 | 28.9105 | -276.6773 | 278.8576 | -0.52327 | -2.2893 | -1.0901 | 2.6381 | 276.6773 | 2.2893 | 7.6701
0.0843 | 8.2415 | 29.2375 | -278.9884 | 281.1033 | -0.61048 | -2.3765 | -1.1119 | 2.7253 | 278.9884 | 2.3765 | 7.8375
0.086 | 8.2197 | 29.1503 | -284.1993 | 286.2924 | -0.67589 | -2.4201 | -1.1337 | 2.7908 | 284.1993 | 2.4201 | 7.9955
0.0878 | 7.9144 | 30.1969 | -291.5468 | 293.7271 | -0.71949 | -2.4419 | -1.1555 | 2.8562 | 291.5468 | 2.4419 | 8.1629
0.0896 | 7.4347 | 32.6388 | -299.5921 | 301.9904 | -0.67589 | -2.4855 | -1.1774 | 2.8998 | 299.5921 | 2.4855 | 8.3302
0.0914 | 7.0641 | 35.4949 | -305.9367 | 308.6402 | -0.63228 | -2.5291 | -1.1992 | 2.9216 | 305.9367 | 2.5291 | 8.4976
0.0932 | 7.1077 | 38.5473 | -307.7027 | 310.7987 | -0.61048 | -2.6163 | -1.221 | 3.0088 | 307.7027 | 2.6163 | 8.6649
0.095 | 7.1731 | 41.5125 | -307.3757 | 310.8641 | -0.58868 | -2.7035 | -1.2428 | 3.096 | 307.3757 | 2.7035 | 8.8323
0.0968 | 7.3911 | 44.0198 | -305.566 | 309.4469 | -0.58868 | -2.7908 | -1.2864 | 3.1832 | 305.566 | 2.7908 | 8.9996
0.0986 | 7.6528 | 45.764 | -302.0558 | 306.2419 | -0.61048 | -2.878 | -1.3082 | 3.2704 | 302.0558 | 2.878 | 9.167
0.1004 | 8.067 | 46.6798 | -297.0193 | 301.3799 | -0.61048 | -2.9434 | -1.33 | 3.314 | 297.0193 | 2.9434 | 9.3343
0.1022 | 8.3941 | 46.2873 | -291.5904 | 295.9728 | -0.67589 | -2.9652 | -1.3518 | 3.3794 | 291.5904 | 2.9652 | 9.5017
0.104 | 8.5467 | 45.5024 | -289.1267 | 293.4219 | -0.7631 | -3.0088 | -1.3736 | 3.4448 | 289.1267 | 3.0088 | 9.669
0.1058 | 8.6557 | 44.216 | -288.2546 | 292.3972 | -0.89391 | -3.0524 | -1.3954 | 3.5321 | 288.2546 | 3.0524 | 9.8364
0.1076 | 8.5467 | 43.4966 | -290.5221 | 294.5556 | -0.98113 | -3.096 | -1.439 | 3.5975 | 290.5221 | 3.096 | 10.0037
0.1094 | 8.1542 | 42.9079 | -295.6458 | 299.5485 | -1.0683 | -3.1178 | -1.4608 | 3.6629 | 295.6458 | 3.1178 | 10.1711
0.1111 | 7.7182 | 43.0387 | -302.9497 | 306.896 | -1.1337 | -3.1614 | -1.5044 | 3.7283 | 302.9497 | 3.1614 | 10.3291
0.1129 | 7.1949 | 43.4311 | -312.4121 | 316.3366 | -1.1774 | -3.205 | -1.5262 | 3.7937 | 312.4121 | 3.205 | 10.4965
0.1147 | 6.5408 | 44.3251 | -322.4632 | 326.4967 | -1.1555 | -3.2486 | -1.548 | 3.8373 | 322.4632 | 3.2486 | 10.6638
0.1165 | 6.0394 | 46.4835 | -329.3093 | 333.539 | -1.1337 | -3.314 | -1.5916 | 3.9027 | 329.3093 | 3.314 | 10.8312
0.1183 | 5.6687 | 49.1653 | -333.3428 | 337.8996 | -1.0901 | -3.3794 | -1.6134 | 3.9681 | 333.3428 | 3.3794 | 10.9985
0.1201 | 5.5815 | 52.0214 | -334.8472 | 339.8182 | -1.0465 | -3.4666 | -1.657 | 4.0335 | 334.8472 | 3.4666 | 11.1659
0.1219 | 5.6033 | 54.3979 | -334.2803 | 339.6002 | -1.0247 | -3.5539 | -1.7006 | 4.1207 | 334.2803 | 3.5539 | 11.3332
0.1237 | 5.6469 | 56.0768 | -332.1 | 337.7469 | -1.0247 | -3.6193 | -1.7442 | 4.1861 | 332.1 | 3.6193 | 11.5006
0.1255 | 5.7341 | 56.6872 | -327.8049 | 333.648 | -1.0683 | -3.6847 | -1.7878 | 4.2733 | 327.8049 | 3.6847 | 11.6679
0.1273 | 5.6033 | 56.382 | -323.5969 | 329.4619 | -1.1119 | -3.7283 | -1.8096 | 4.3388 | 323.5969 | 3.7283 | 11.8353
0.1291 | 5.2545 | 55.5099 | -320.7626 | 326.5185 | -1.1992 | -3.7501 | -1.8532 | 4.3824 | 320.7626 | 3.7501 | 12.0026
0.1309 | 4.644 | 53.9619 | -320.937 | 326.3441 | -1.3082 | -3.7501 | -1.875 | 4.4478 | 320.937 | 3.7501 | 12.17
0.1327 | 3.8591 | 52.9808 | -323.0955 | 328.2409 | -1.3954 | -3.7719 | -1.9186 | 4.4914 | 323.0955 | 3.7719 | 12.3373
0.1345 | 2.8998 | 52.7409 | -327.347 | 332.3181 | -1.439 | -3.7937 | -1.9404 | 4.535 | 327.347 | 3.7937 | 12.5046
0.1362 | 1.875 | 53.177 | -333.6044 | 338.51 | -1.4826 | -3.8155 | -1.9841 | 4.6004 | 333.6044 | 3.8155 | 12.6627
0.138 | 0.85031 | 54.9212 | -338.6845 | 343.8081 | -1.4826 | -3.8809 | -2.0277 | 4.6658 | 338.6845 | 3.8809 | 12.83
0.1398 | -0.13082 | 56.9489 | -342.8052 | 348.2123 | -1.4608 | -3.9463 | -2.0495 | 4.7094 | 342.8052 | 3.9463 | 12.9974
0.1416 | -0.82851 | 59.2382 | -344.7456 | 350.567 | -1.439 | -4.0117 | -2.0713 | 4.7748 | 344.7456 | 4.0117 | 13.1647
0.1434 | -1.1555 | 61.353 | -344.0261 | 350.2617 | -1.3954 | -4.0989 | -2.1149 | 4.8402 | 344.0261 | 4.0989 | 13.3321
0.1452 | -1.3954 | 63.0973 | -341.6278 | 348.2777 | -1.3736 | -4.1861 | -2.1367 | 4.9274 | 341.6278 | 4.1861 | 13.4994
0.147 | -1.7006 | 64.2746 | -337.5289 | 344.5058 | -1.3736 | -4.2297 | -2.1803 | 4.971 | 337.5289 | 4.2297 | 13.6668
0.1488 | -1.5698 | 64.9069 | -331.3587 | 338.5318 | -1.3736 | -4.2951 | -2.2021 | 5.0582 | 331.3587 | 4.2951 | 13.8341
0.1506 | -1.1992 | 64.4272 | -323.9894 | 331.1625 | -1.3954 | -4.3388 | -2.2239 | 5.1019 | 323.9894 | 4.3388 | 14.0015
0.1524 | -1.548 | 62.6612 | -319.1492 | 326.1043 | -1.5044 | -4.3606 | -2.2457 | 5.1455 | 319.1492 | 4.3606 | 14.1688
0.1542 | -2.5073 | 61.1786 | -316.4456 | 323.2481 | -1.5916 | -4.3388 | -2.2675 | 5.1673 | 316.4456 | 4.3388 | 14.3362
0.156 | -3.6629 | 59.1728 | -317.2959 | 323.8368 | -1.7006 | -4.317 | -2.2893 | 5.1891 | 317.2959 | 4.317 | 14.5035
0.1578 | -4.8184 | 57.2105 | -319.7815 | 326.0389 | -1.7878 | -4.2951 | -2.3111 | 5.2109 | 319.7815 | 4.2951 | 14.6709
0.1595 | -5.9086 | 55.7497 | -324.8179 | 330.8355 | -1.875 | -4.2733 | -2.3329 | 5.2327 | 324.8179 | 4.2733 | 14.8289
0.1613 | -6.8243 | 54.8558 | -331.4023 | 337.2455 | -1.9404 | -4.2951 | -2.3547 | 5.2763 | 331.4023 | 4.2951 | 14.9963
0.1631 | -7.3911 | 55.1392 | -339.513 | 345.3779 | -1.9841 | -4.3388 | -2.3547 | 5.3417 | 339.513 | 4.3388 | 15.1636
0.1649 | -8.3069 | 56.3166 | -348.2777 | 354.2735 | -1.9623 | -4.3388 | -2.3765 | 5.3417 | 348.2777 | 4.3388 | 15.331
0.1667 | -8.9609 | 58.0172 | -355.4508 | 361.5992 | -1.9404 | -4.3606 | -2.3983 | 5.3635 | 355.4508 | 4.3606 | 15.4983
0.1685 | -9.4188 | 60.0013 | -361.6428 | 368.0092 | -1.8968 | -4.4042 | -2.4201 | 5.4071 | 361.6428 | 4.4042 | 15.6657
0.1703 | -9.506 | 62.5958 | -366.8537 | 373.6125 | -1.8532 | -4.4696 | -2.4419 | 5.4289 | 366.8537 | 4.4696 | 15.833
0.1721 | -9.615 | 65.0377 | -369.8406 | 376.9047 | -1.7878 | -4.535 | -2.4637 | 5.4725 | 369.8406 | 4.535 | 16.0004
0.1739 | -9.4406 | 67.654 | -370.6692 | 378.2129 | -1.7224 | -4.5786 | -2.4637 | 5.5161 | 370.6692 | 4.5786 | 16.1677
0.1757 | -9.1354 | 69.8779 | -370.1023 | 378.0385 | -1.6788 | -4.644 | -2.4855 | 5.5379 | 370.1023 | 4.644 | 16.3351
0.1775 | -8.5903 | 71.9274 | -367.8348 | 376.1853 | -1.6352 | -4.6876 | -2.5073 | 5.5815 | 367.8348 | 4.6876 | 16.5024
0.1793 | -8.0234 | 73.6498 | -363.8231 | 372.5878 | -1.5916 | -4.7312 | -2.5073 | 5.6251 | 363.8231 | 4.7312 | 16.6698
0.1811 | -7.4347 | 74.4565 | -358.4596 | 367.4859 | -1.5698 | -4.7748 | -2.5291 | 5.6469 | 358.4596 | 4.7748 | 16.8371
0.1829 | -6.7589 | 74.7399 | -352.3548 | 361.4902 | -1.5698 | -4.7966 | -2.5291 | 5.6687 | 352.3548 | 4.7966 | 17.0045
0.1846 | -6.1702 | 74.2821 | -346.4681 | 355.6034 | -1.5698 | -4.7966 | -2.5291 | 5.6687 | 346.4681 | 4.7966 | 17.1625
0.1864 | -5.9958 | 73.3664 | -343.6773 | 352.8127 | -1.6134 | -4.753 | -2.5291 | 5.6469 | 343.6773 | 4.753 | 17.3299
0.1882 | -6.1484 | 71.1861 | -342.2819 | 351.1121 | -1.7006 | -4.6876 | -2.5291 | 5.6251 | 342.2819 | 4.6876 | 17.4972
0.19 | -6.628 | 68.875 | -343.7209 | 352.1586 | -1.8096 | -4.6222 | -2.5291 | 5.6033 | 343.7209 | 4.6222 | 17.6646
0.1918 | -7.2603 | 66.6075 | -346.9477 | 354.8839 | -1.8968 | -4.5786 | -2.5291 | 5.5815 | 346.9477 | 4.5786 | 17.8319
0.1936 | -7.849 | 64.7543 | -352.9217 | 360.5091 | -1.9841 | -4.5132 | -2.5291 | 5.5597 | 352.9217 | 4.5132 | 17.9993
0.1954 | -8.3069 | 64.2528 | -360.0076 | 367.3333 | -2.0495 | -4.4696 | -2.5291 | 5.5597 | 360.0076 | 4.4696 | 18.1666
0.1972 | -8.2851 | 65.6264 | -369.6008 | 376.992 | -2.0277 | -4.4914 | -2.5291 | 5.5815 | 369.6008 | 4.4914 | 18.334
0.199 | -8.0234 | 67.5886 | -378.6054 | 386.1709 | -2.0059 | -4.535 | -2.5073 | 5.5815 | 378.6054 | 4.535 | 18.5013
0.2008 | -7.6528 | 69.3765 | -384.4267 | 392.2539 | -1.9623 | -4.5786 | -2.5073 | 5.6033 | 384.4267 | 4.5786 | 18.6687
0.2026 | -7.5874 | 70.0087 | -389.9646 | 397.9662 | -1.9404 | -4.5568 | -2.4855 | 5.5597 | 389.9646 | 4.5568 | 18.836
0.2044 | -7.4566 | 70.9245 | -394.0417 | 402.2614 | -1.8968 | -4.5568 | -2.4637 | 5.5379 | 394.0417 | 4.5568 | 19.0033
0.2062 | -6.9987 | 72.2326 | -396.222 | 404.6597 | -1.8314 | -4.5786 | -2.4419 | 5.5161 | 396.222 | 4.5786 | 19.1707
0.208 | -6.6935 | 72.1236 | -397.3994 | 405.7716 | -1.7878 | -4.5132 | -2.4201 | 5.4507 | 397.3994 | 4.5132 | 19.338
0.2097 | -6.41 | 71.862 | -398.5113 | 406.7746 | -1.7442 | -4.4696 | -2.3765 | 5.3853 | 398.5113 | 4.4696 | 19.4961
0.2115 | -7.1731 | 71.7094 | -400.2119 | 408.388 | -1.6788 | -4.3388 | -2.3547 | 5.2327 | 400.2119 | 4.3388 | 19.6634
0.2133 | -8.1542 | 70.423 | -405.6844 | 413.5334 | -1.6352 | -4.1861 | -2.3111 | 5.08 | 405.6844 | 4.1861 | 19.8308
0.2151 | -9.2444 | 67.9811 | -413.8823 | 421.2516 | -1.6352 | -3.9681 | -2.2675 | 4.8838 | 413.8823 | 3.9681 | 19.9981
0.2169 | -10.5962 | 65.3211 | -424.6092 | 431.3899 | -1.6134 | -3.7283 | -2.2457 | 4.6876 | 424.6092 | 3.7283 | 20.1655
0.2187 | -11.9043 | 62.8356 | -438.5848 | 444.755 | -1.6134 | -3.5102 | -2.2021 | 4.4914 | 438.5848 | 3.5102 | 20.3328
0.2205 | -13.6703 | 60.1321 | -455.6128 | 461.3687 | -1.5698 | -3.2704 | -2.1803 | 4.2733 | 455.6128 | 3.2704 | 20.5002
0.2223 | -15.7198 | 57.1887 | -473.9271 | 479.2906 | -1.5262 | -2.987 | -2.1803 | 4.0553 | 473.9271 | 2.987 | 20.6675
0.2241 | -17.8565 | 53.8093 | -495.7517 | 500.701 | -1.4608 | -2.6599 | -2.1585 | 3.8155 | 495.7517 | 2.6599 | 20.8349
0.2259 | -20.4074 | 50.1464 | -521.2828 | 525.7742 | -1.3954 | -2.2893 | -2.1367 | 3.5539 | 521.2828 | 2.2893 | 21.0022
0.2277 | -23.4162 | 46.2873 | -550.7384 | 554.8373 | -1.2646 | -1.875 | -2.1367 | 3.2704 | 550.7384 | 1.875 | 21.1696
0.2295 | -26.7956 | 42.232 | -584.1838 | 587.9993 | -1.1337 | -1.4172 | -2.1367 | 3.0306 | 584.1838 | 1.4172 | 21.3369
0.2313 | -30.2405 | 37.806 | -621.3358 | 624.9114 | -0.98113 | -0.93752 | -2.1803 | 2.878 | 621.3358 | 0.93752 | 21.5043
0.2331 | -33.9469 | 32.726 | -661.6273 | 665.094 | -0.8067 | -0.39245 | -2.2021 | 2.8344 | 661.6273 | 0.39245 | 21.6716
0.2348 | -38.5691 | 26.7084 | -707.5004 | 710.8362 | -0.61048 | 0.21803 | -2.2675 | 2.8998 | 707.5004 | -0.21803 | 21.8297
0.2366 | -43.562 | 20.1894 | -758.1046 | 761.2442 | -0.41425 | 0.87211 | -2.3329 | 3.0306 | 758.1046 | -0.87211 | 21.997
0.2384 | -48.6856 | 13.3215 | -812.459 | 815.5114 | -0.19623 | 1.548 | -2.3983 | 3.2922 | 812.459 | -1.548 | 22.1644
0.2402 | -54.1363 | 5.8213 | -872.8091 | 875.9051 | 0.043606 | 2.3111 | -2.5073 | 3.7283 | 872.8091 | -2.3111 | 22.3317
0.242 | -59.7396 | -2.0059 | -937.2145 | 940.3977 | 0.28344 | 3.096 | -2.6381 | 4.2951 | 937.2145 | -3.096 | 22.4991
0.2438 | -65.3429 | -10.5089 | -1008.0518 | 1011.3876 | 0.56687 | 3.9681 | -2.769 | 5.0364 | 1008.0518 | -3.9681 | 22.6664
0.2456 | -70.9681 | -19.5789 | -1084.3179 | 1087.9154 | 0.87211 | 4.862 | -2.9434 | 5.8649 | 1084.3179 | -4.862 | 22.8338
0.2474 | -76.4406 | -29.0849 | -1167.3647 | 1171.2892 | 1.1992 | 5.8213 | -3.1396 | 6.8025 | 1167.3647 | -5.8213 | 23.0011
0.2492 | -81.7604 | -38.7435 | -1256.5599 | 1260.7897 | 1.548 | 6.8025 | -3.3576 | 7.8272 | 1256.5599 | -6.8025 | 23.1685
0.251 | -87.0367 | -48.5984 | -1353.1681 | 1357.7685 | 1.9404 | 7.8708 | -3.5975 | 8.9173 | 1353.1681 | -7.8708 | 23.3358
0.2528 | -92.2258 | -58.8021 | -1455.3577 | 1460.3069 | 2.3547 | 8.9609 | -3.8591 | 10.0947 | 1455.3577 | -8.9609 | 23.5032
0.2546 | -96.9352 | -69.2674 | -1570.5854 | 1575.9053 | 2.8126 | 10.1165 | -4.1643 | 11.3593 | 1570.5854 | -10.1165 | 23.6705
0.2564 | -100.8815 | -79.1877 | -1685.8786 | 1691.5255 | 3.2922 | 11.2284 | -4.4696 | 12.5802 | 1685.8786 | -11.2284 | 23.8379
0.2581 | -104.5444 | -90.2635 | -1805.009 | 1811.0702 | 3.7501 | 12.4276 | -4.8184 | 13.8666 | 1805.009 | -12.4276 | 23.9959
0.2599 | -107.2043 | -101.492 | -1923.5507 | 1930.0044 | 4.2079 | 13.5831 | -5.1455 | 15.1529 | 1923.5507 | -13.5831 | 24.1633
0.2617 | -108.6433 | -113.6143 | -2046.1914 | 2052.9939 | 4.6658 | 14.7605 | -5.4943 | 16.4611 | 2046.1914 | -14.7605 | 24.3306
0.2635 | -109.6898 | -125.8457 | -2164.2099 | 2171.383 | 5.1019 | 15.8942 | -5.8431 | 17.7039 | 2164.2099 | -15.8942 | 24.498
0.2653 | -109.9078 | -138.1206 | -2282.5336 | 2290.0555 | 5.5379 | 17.0062 | -6.192 | 18.9466 | 2282.5336 | -17.0062 | 24.6653
0.2671 | -109.6898 | -148.695 | -2402.994 | 2410.7558 | 6.0176 | 18.1399 | -6.5408 | 20.233 | 2402.994 | -18.1399 | 24.8327
0.2689 | -108.9921 | -159.4874 | -2521.2523 | 2529.2757 | 6.4754 | 19.1646 | -6.8897 | 21.3885 | 2521.2523 | -19.1646 | 25
0.2707 | -107.4877 | -169.0588 | -2631.2909 | 2639.5324 | 6.8897 | 20.015 | -7.1949 | 22.3697 | 2631.2909 | -20.015 | 25.1673
0.2725 | -105.5255 | -177.9543 | -2735.9007 | 2744.2948 | 7.2821 | 20.7344 | -7.5002 | 23.2418 | 2735.9007 | -20.7344 | 25.3347
0.2743 | -102.6039 | -186.283 | -2839.9436 | 2848.4685 | 7.6528 | 21.3449 | -7.7836 | 23.9831 | 2839.9436 | -21.3449 | 25.502
0.2761 | -99.7913 | -193.8704 | -2940.5634 | 2949.1755 | 8.0234 | 21.9336 | -8.067 | 24.7244 | 2940.5634 | -21.9336 | 25.6694
0.2779 | -96.9134 | -200.8909 | -3036.1251 | 3044.8244 | 8.3723 | 22.5441 | -8.3287 | 25.4657 | 3036.1251 | -22.5441 | 25.8367
0.2797 | -93.7738 | -206.9956 | -3124.3173 | 3133.082 | 8.7211 | 23.0891 | -8.5685 | 26.1415 | 3124.3173 | -23.0891 | 26.0041
0.2815 | -90.1327 | -212.2283 | -3205.7943 | 3214.5809 | 9.0046 | 23.5034 | -8.7647 | 26.6648 | 3205.7943 | -23.5034 | 26.1714
0.2832 | -86.121 | -216.8287 | -3281.8424 | 3290.6508 | 9.2662 | 23.8304 | -8.9609 | 27.1227 | 3281.8424 | -23.8304 | 26.3295
0.285 | -82.0439 | -220.666 | -3352.0692 | 3360.8775 | 9.506 | 24.0703 | -9.1354 | 27.4715 | 3352.0692 | -24.0703 | 26.4968
0.2868 | -77.9232 | -223.1733 | -3412.6592 | 3421.4021 | 9.7022 | 24.1793 | -9.288 | 27.6895 | 3412.6592 | -24.1793 | 26.6642
0.2886 | -73.7806 | -225.3318 | -3468.5833 | 3477.2826 | 9.8767 | 24.2665 | -9.4188 | 27.864 | 3468.5833 | -24.2665 | 26.8315
0.2904 | -69.6599 | -227.3376 | -3519.7326 | 3528.3665 | 10.0293 | 24.3101 | -9.5496 | 28.0166 | 3519.7326 | -24.3101 | 26.9989
0.2922 | -65.7354 | -228.2098 | -3560.6347 | 3569.1595 | 10.1601 | 24.3101 | -9.615 | 28.1038 | 3560.6347 | -24.3101 | 27.1662
0.294 | -61.8981 | -227.5557 | -3596.6965 | 3605.0687 | 10.2691 | 24.1793 | -9.7022 | 28.0384 | 3596.6965 | -24.1793 | 27.3336
0.2958 | -58.0826 | -226.6836 | -3630.1637 | 3638.3834 | 10.3781 | 24.0703 | -9.7458 | 28.0166 | 3630.1637 | -24.0703 | 27.5009
0.2976 | -54.2671 | -224.1762 | -3660.4914 | 3668.4276 | 10.4871 | 23.874 | -9.7895 | 27.9076 | 3660.4914 | -23.874 | 27.6683
0.2994 | -50.6915 | -220.9058 | -3689.838 | 3697.4689 | 10.5962 | 23.6778 | -9.8113 | 27.7986 | 3689.838 | -23.6778 | 27.8356
0.3012 | -46.3527 | -216.1746 | -3715.8051 | 3723.0872 | 10.6616 | 23.2418 | -9.8331 | 27.4933 | 3715.8051 | -23.2418 | 28.003
0.303 | -42.2102 | -211.16 | -3737.7605 | 3744.7156 | 10.7052 | 22.8275 | -9.8549 | 27.1663 | 3737.7605 | -22.8275 | 28.1703
0.3048 | -38.2639 | -206.0363 | -3755.0719 | 3761.7436 | 10.7052 | 22.3697 | -9.8549 | 26.7956 | 3755.0719 | -22.3697 | 28.3377
0.3066 | -34.4048 | -200.5856 | -3768.6114 | 3774.9996 | 10.6834 | 21.89 | -9.8331 | 26.4032 | 3768.6114 | -21.89 | 28.505
0.3083 | -30.5021 | -193.805 | -3779.382 | 3785.5086 | 10.6398 | 21.2795 | -9.8113 | 25.9017 | 3779.382 | -21.2795 | 28.6631
0.3101 | -26.861 | -187.3949 | -3786.8604 | 3792.6599 | 10.618 | 20.7781 | -9.7676 | 25.4875 | 3786.8604 | -20.7781 | 28.8304
0.3119 | -23.3944 | -180.4835 | -3793.4448 | 3799.0045 | 10.5307 | 19.9932 | -9.7458 | 24.877 | 3793.4448 | -19.9932 | 28.9978
0.3137 | -20.2112 | -173.79 | -3797.6091 | 3802.929 | 10.4871 | 19.4263 | -9.7022 | 24.4409 | 3797.6091 | -19.4263 | 29.1651
0.3155 | -16.8318 | -166.4207 | -3801.3374 | 3806.4393 | 10.4217 | 18.7286 | -9.6586 | 24.0267 | 3801.3374 | -18.7286 | 29.3325
0.3173 | -13.6485 | -158.5281 | -3804.9349 | 3809.9277 | 10.3345 | 17.9437 | -9.615 | 23.6778 | 3804.9349 | -17.9437 | 29.4998
0.3191 | -10.6616 | -150.4828 | -3807.7692 | 3812.7185 | 10.2473 | 17.2678 | -9.5714 | 23.4598 | 3807.7692 | -17.2678 | 29.6672
0.3209 | -7.6092 | -143.0263 | -3809.7533 | 3814.6153 | 10.2037 | 16.7009 | -9.506 | 23.3508 | 3809.7533 | -16.7009 | 29.8345
0.3227 | -4.2079 | -135.2427 | -3810.1893 | 3814.8333 | 10.1383 | 16.1341 | -9.4406 | 23.0673 | 3810.1893 | -16.1341 | 30.0019
0.3245 | -1.1555 | -127.4809 | -3809.9713 | 3814.4627 | 10.0947 | 15.5672 | -9.3534 | 22.8929 | 3809.9713 | -15.5672 | 30.1692
0.3263 | 1.5262 | -119.6755 | -3808.1399 | 3812.5658 | 10.0075 | 14.9785 | -9.288 | 22.6531 | 3808.1399 | -14.9785 | 30.3366
0.3281 | 3.7937 | -112.5024 | -3805.0003 | 3809.3826 | 9.9203 | 14.4116 | -9.2226 | 22.3697 | 3805.0003 | -14.4116 | 30.5039
0.3299 | 6.4972 | -103.9775 | -3800.1383 | 3804.3244 | 9.8113 | 13.8012 | -9.1354 | 21.9554 | 3800.1383 | -13.8012 | 30.6713
0.3316 | 8.7865 | -96.0195 | -3794.3823 | 3798.4158 | 9.724 | 13.2125 | -9.0482 | 21.5412 | 3794.3823 | -13.2125 | 30.8293
0.3334 | 11.9043 | -87.7126 | -3785.0289 | 3788.9098 | 9.615 | 12.602 | -8.9391 | 21.1487 | 3785.0289 | -12.602 | 30.9967
0.3352 | 14.477 | -78.5772 | -3776.2424 | 3779.9707 | 9.4842 | 11.9697 | -8.8519 | 20.6908 | 3776.2424 | -11.9697 | 31.164
0.337 | 17.0934 | -70.1614 | -3766.9108 | 3770.6609 | 9.3534 | 11.3374 | -8.7429 | 20.4074 | 3766.9108 | -11.3374 | 31.3313
0.3388 | 19.5353 | -61.6147 | -3756.7071 | 3760.5008 | 9.2226 | 10.7052 | -8.6557 | 20.124 | 3756.7071 | -10.7052 | 31.4987
0.3406 | 21.672 | -53.8311 | -3745.6749 | 3749.5558 | 9.1136 | 10.1601 | -8.5467 | 19.8841 | 3745.6749 | -10.1601 | 31.666
0.3424 | 23.5034 | -47.0722 | -3734.0976 | 3738.0875 | 9.0046 | 9.6804 | -8.4377 | 19.7315 | 3734.0976 | -9.6804 | 31.8334
0.3442 | 25.8145 | -38.9616 | -3719.9476 | 3723.8067 | 8.8955 | 9.1572 | -8.3287 | 19.4045 | 3719.9476 | -9.1572 | 32.0007
0.346 | 28.0384 | -29.1721 | -3706.9532 | 3710.7686 | 8.7211 | 8.5249 | -8.2197 | 18.9248 | 3706.9532 | -8.5249 | 32.1681
0.3478 | 30.3059 | -20.2766 | -3694.8308 | 3698.7117 | 8.5467 | 7.958 | -8.1106 | 18.5542 | 3694.8308 | -7.958 | 32.3354
0.3496 | 32.0501 | -13.3433 | -3676.5383 | 3680.3974 | 8.4377 | 7.5656 | -7.9798 | 18.3579 | 3676.5383 | -7.5656 | 32.5028
0.3514 | 33.5763 | -4.9274 | -3654.8227 | 3658.6382 | 8.2851 | 7.0859 | -7.849 | 18.0309 | 3654.8227 | -7.0859 | 32.6701
0.3532 | 35.9746 | 4.426 | -3636.9444 | 3640.7163 | 8.0888 | 6.5626 | -7.74 | 17.6167 | 3636.9444 | -6.5626 | 32.8375
0.355 | 38.4601 | 13.8012 | -3622.0967 | 3625.8686 | 7.9362 | 6.0176 | -7.6092 | 17.246 | 3622.0967 | -6.0176 | 33.0048
0.3567 | 40.5532 | 22.2388 | -3609.1895 | 3613.1794 | 7.74 | 5.5379 | -7.5002 | 16.9844 | 3609.1895 | -5.5379 | 33.1629
0.3585 | 42.3192 | 30.3059 | -3598.7677 | 3603.0629 | 7.5656 | 5.1019 | -7.3911 | 16.8318 | 3598.7677 | -5.1019 | 33.3302
0.3603 | 44.1942 | 38.2421 | -3586.1221 | 3590.6135 | 7.4347 | 4.6222 | -7.2821 | 16.6791 | 3586.1221 | -4.6222 | 33.4976
0.3621 | 46.2001 | 45.8295 | -3568.8325 | 3573.4765 | 7.2821 | 4.1207 | -7.1513 | 16.5919 | 3568.8325 | -4.1207 | 33.6649
0.3639 | 48.1624 | 53.2424 | -3550.8888 | 3555.729 | 7.1295 | 3.6193 | -7.0205 | 16.5265 | 3550.8888 | -3.6193 | 33.8323
0.3657 | 50.1682 | 60.3501 | -3531.8332 | 3536.8696 | 6.9769 | 3.1396 | -6.8897 | 16.4393 | 3531.8332 | -3.1396 | 33.9996
0.3675 | 53.3296 | 67.9811 | -3505.6044 | 3510.5101 | 6.8025 | 2.6163 | -6.7153 | 16.0687 | 3505.6044 | -2.6163 | 34.167
0.3693 | 56.6654 | 76.1135 | -3477.1954 | 3481.883 | 6.628 | 2.0277 | -6.5408 | 15.589 | 3477.1954 | -2.0277 | 34.3343
0.3711 | 60.1103 | 84.4858 | -3450.2472 | 3454.6949 | 6.3882 | 1.3518 | -6.3446 | 15.0003 | 3450.2472 | -1.3518 | 34.5017
0.3729 | 63.3153 | 93.2941 | -3425.1085 | 3429.3819 | 6.1484 | 0.65408 | -6.1702 | 14.3462 | 3425.1085 | -0.65408 | 34.669
0.3747 | 65.9098 | 101.4484 | -3405.05 | 3409.3233 | 5.9304 | 0.065408 | -6.0176 | 13.8884 | 3405.05 | -0.065408 | 34.8364
0.3765 | 67.6322 | 106.4848 | -3389.897 | 3394.4102 | 5.7995 | -0.32704 | -5.9086 | 13.9102 | 3389.897 | 0.32704 | 35.0037
0.3783 | 69.5945 | 110.3657 | -3374.4389 | 3379.2137 | 5.6469 | -0.69769 | -5.7995 | 13.9756 | 3374.4389 | 0.69769 | 35.1711
0.3801 | 72.4507 | 116.0562 | -3356.735 | 3361.6188 | 5.4943 | -1.1992 | -5.6687 | 13.7358 | 3356.735 | 1.1992 | 35.3384
0.3818 | 74.6745 | 119.2176 | -3342.6504 | 3347.6651 | 5.4289 | -1.5262 | -5.5597 | 13.5831 | 3342.6504 | 1.5262 | 35.4965
0.3836 | 76.3097 | 121.7468 | -3329.1545 | 3334.2345 | 5.4071 | -1.8096 | -5.4725 | 13.4741 | 3329.1545 | 1.8096 | 35.6638
0.3854 | 77.7487 | 125.1262 | -3311.2544 | 3316.4217 | 5.2763 | -2.1585 | -5.3635 | 13.2561 | 3311.2544 | 2.1585 | 35.8312
0.3872 | 79.2095 | 128.5274 | -3293.3325 | 3298.6088 | 5.1673 | -2.5073 | -5.2763 | 13.0381 | 3293.3325 | 2.5073 | 35.9985
0.389 | 80.8447 | 130.3152 | -3272.6853 | 3278.136 | 5.08 | -2.7253 | -5.1891 | 13.0163 | 3272.6853 | 2.7253 | 36.1659
0.3908 | 82.5671 | 132.4301 | -3251.2313 | 3256.8128 | 4.971 | -2.9652 | -5.1019 | 12.9727 | 3251.2313 | 2.9652 | 36.3332
0.3926 | 84.3114 | 135.7223 | -3228.7963 | 3234.4868 | 4.862 | -3.2704 | -4.9928 | 12.7546 | 3228.7963 | 3.2704 | 36.5006
0.3944 | 86.0338 | 138.1424 | -3205.8815 | 3211.6811 | 4.7748 | -3.5321 | -4.8838 | 12.6238 | 3205.8815 | 3.5321 | 36.6679
0.3962 | 87.7562 | 140.628 | -3182.2255 | 3188.1559 | 4.6876 | -3.7937 | -4.7748 | 12.5584 | 3182.2255 | 3.7937 | 36.8353
0.398 | 89.1734 | 142.8082 | -3157.3267 | 3163.4097 | 4.6222 | -4.0117 | -4.6658 | 12.4712 | 3157.3267 | 4.0117 | 37.0026
0.3998 | 90.2199 | 142.9827 | -3129.1357 | 3135.3277 | 4.5132 | -4.0117 | -4.5786 | 12.3404 | 3129.1357 | 4.0117 | 37.17
0.4016 | 91.4191 | 143.0263 | -3099.8546 | 3106.1338 | 4.4042 | -3.9245 | -4.4696 | 12.166 | 3099.8546 | 3.9245 | 37.3373
0.4034 | 93.0325 | 142.7864 | -3071.3148 | 3077.703 | 4.2733 | -3.8591 | -4.426 | 12.0351 | 3071.3148 | 3.8591 | 37.5046
0.4051 | 95.1037 | 142.2632 | -3043.538 | 3050.0134 | 4.1425 | -3.8155 | -4.426 | 11.9261 | 3043.538 | 3.8155 | 37.6627
0.4069 | 97.0442 | 141.7181 | -3016.3281 | 3022.9344 | 4.0117 | -3.7719 | -4.426 | 11.8389 | 3016.3281 | 3.7719 | 37.83
0.4087 | 98.6576 | 141.0204 | -2990.579 | 2997.2943 | 3.8591 | -3.7937 | -4.4478 | 11.7953 | 2990.579 | 3.7937 | 37.9974
0.4105 | 100.6199 | 140.2573 | -2963.1075 | 2969.9536 | 3.6847 | -3.7719 | -4.4914 | 11.7081 | 2963.1075 | 3.7719 | 38.1647
0.4123 | 101.6664 | 139.5814 | -2937.4238 | 2944.2699 | 3.5102 | -3.8373 | -4.5568 | 11.6209 | 2937.4238 | 3.8373 | 38.3321
0.4141 | 102.7347 | 139.6904 | -2912.8303 | 2919.72 | 3.314 | -3.9899 | -4.6004 | 11.5119 | 2912.8303 | 3.9899 | 38.4994
0.4159 | 103.8249 | 140.1701 | -2890.3298 | 2897.2195 | 3.1396 | -4.2079 | -4.6004 | 11.4029 | 2890.3298 | 4.2079 | 38.6668
0.4177 | 104.7842 | 140.5844 | -2869.2029 | 2876.158 | 2.987 | -4.4478 | -4.5786 | 11.3811 | 2869.2029 | 4.4478 | 38.8341
0.4195 | 105.8743 | 140.0829 | -2848.9263 | 2855.925 | 2.8344 | -4.6004 | -4.5786 | 11.4029 | 2848.9263 | 4.6004 | 39.0015
0.4213 | 107.1607 | 138.4259 | -2828.8242 | 2835.8447 | 2.6599 | -4.644 | -4.6222 | 11.4247 | 2828.8242 | 4.644 | 39.1688
0.4231 | 107.9892 | 136.2238 | -2808.8964 | 2815.9387 | 2.5291 | -4.753 | -4.644 | 11.5119 | 2808.8964 | 4.753 | 39.3362
0.4249 | 109.7552 | 135.1555 | -2788.38 | 2795.5313 | 2.3765 | -4.8402 | -4.7094 | 11.6209 | 2788.38 | 4.8402 | 39.5035
0.4267 | 111.1506 | 134.8938 | -2767.4493 | 2774.6879 | 2.2021 | -4.971 | -4.753 | 11.6645 | 2767.4493 | 4.971 | 39.6709
0.4285 | 111.8047 | 134.8938 | -2746.7367 | 2753.9752 | 2.0931 | -5.1455 | -4.7748 | 11.6209 | 2746.7367 | 5.1455 | 39.8382
0.4302 | 112.1971 | 134.8502 | -2727.354 | 2734.5707 | 1.9841 | -5.3417 | -4.753 | 11.5337 | 2727.354 | 5.3417 | 39.9963
0.432 | 112.2189 | 134.2834 | -2709.3231 | 2716.4962 | 1.875 | -5.4507 | -4.753 | 11.4465 | 2709.3231 | 5.4507 | 40.1636
0.4338 | 112.1317 | 133.5421 | -2691.7719 | 2698.9014 | 1.766 | -5.5161 | -4.7312 | 11.3156 | 2691.7719 | 5.5161 | 40.331
0.4356 | 112.6986 | 135.4171 | -2673.6319 | 2680.7833 | 1.6788 | -5.9086 | -4.7094 | 11.0104 | 2673.6319 | 5.9086 | 40.4983
0.4374 | 114.7481 | 140.9986 | -2656.1461 | 2663.45 | 1.4826 | -6.41 | -4.6004 | 10.4871 | 2656.1461 | 6.41 | 40.6657
0.4392 | 115.4457 | 144.2472 | -2639.5106 | 2646.9453 | 1.3082 | -6.7153 | -4.5132 | 10.1165 | 2639.5106 | 6.7153 | 40.833
0.441 | 115.4894 | 144.9885 | -2618.4273 | 2625.9711 | 1.221 | -6.7371 | -4.5132 | 10.0075 | 2618.4273 | 6.7371 | 41.0004
0.4428 | 114.7917 | 144.2472 | -2599.1536 | 2606.7628 | 1.1774 | -6.7371 | -4.5568 | 9.9857 | 2599.1536 | 6.7371 | 41.1677
0.4446 | 113.5707 | 143.2007 | -2582.0603 | 2589.6476 | 1.1119 | -6.7371 | -4.5568 | 9.9421 | 2582.0603 | 6.7371 | 41.3351
0.4464 | 112.4806 | 142.2414 | -2564.9015 | 2572.467 | 1.0465 | -6.7589 | -4.5568 | 9.8985 | 2564.9015 | 6.7589 | 41.5024
0.4482 | 110.3657 | 139.4724 | -2550.7515 | 2558.208 | 1.0901 | -6.7371 | -4.5568 | 9.8767 | 2550.7515 | 6.7371 | 41.6698
0.45 | 109.2538 | 136.7253 | -2537.2337 | 2544.5595 | 1.1337 | -6.7153 | -4.5786 | 9.9203 | 2537.2337 | 6.7153 | 41.8371
0.4518 | 108.1418 | 133.1278 | -2527.5315 | 2534.7482 | 1.2428 | -6.7371 | -4.5786 | 9.9421 | 2527.5315 | 6.7371 | 42.0045
0.4536 | 106.5938 | 128.68 | -2519.0502 | 2526.1143 | 1.3736 | -6.6935 | -4.5786 | 9.9639 | 2519.0502 | 6.6935 | 42.1718
0.4553 | 104.9804 | 123.7526 | -2512.1605 | 2519.0502 | 1.5262 | -6.6062 | -4.6004 | 9.9639 | 2512.1605 | 6.6062 | 42.3299
0.4571 | 104.2609 | 120.5694 | -2503.4394 | 2510.1765 | 1.5698 | -6.5844 | -4.6004 | 9.9639 | 2503.4394 | 6.5844 | 42.4972
0.4589 | 102.8655 | 115.4457 | -2501.3681 | 2507.8 | 1.7224 | -6.4972 | -4.6222 | 9.9203 | 2501.3681 | 6.4972 | 42.6646
0.4607 | 101.4048 | 109.9296 | -2499.8419 | 2505.9903 | 1.875 | -6.41 | -4.6222 | 9.8985 | 2499.8419 | 6.41 | 42.8319
0.4625 | 99.944 | 104.5007 | -2498.2503 | 2504.0935 | 2.0277 | -6.301 | -4.6222 | 9.8549 | 2498.2503 | 6.301 | 42.9993
0.4643 | 98.1125 | 99.9876 | -2495.8302 | 2501.4554 | 2.2021 | -6.2138 | -4.644 | 9.8331 | 2495.8302 | 6.2138 | 43.1666
0.4661 | 96.3029 | 95.4744 | -2493.4101 | 2498.8172 | 2.3765 | -6.1266 | -4.6658 | 9.8331 | 2493.4101 | 6.1266 | 43.334
0.4679 | 95.6706 | 92.6182 | -2485.3649 | 2490.663 | 2.5291 | -6.1702 | -4.6876 | 9.8985 | 2485.3649 | 6.1702 | 43.5013
0.4697 | 95.2346 | 90.0237 | -2476.4476 | 2481.6802 | 2.6599 | -6.2356 | -4.6876 | 9.9639 | 2476.4476 | 6.2356 | 43.6687
0.4715 | 95.0165 | 87.1021 | -2466.7017 | 2471.8472 | 2.7908 | -6.2792 | -4.7094 | 10.0511 | 2466.7017 | 6.2792 | 43.836
0.4733 | 94.8203 | 84.1369 | -2456.8251 | 2461.8833 | 2.8998 | -6.301 | -4.7312 | 10.1165 | 2456.8251 | 6.301 | 44.0033
0.4751 | 93.1633 | 81.586 | -2447.0356 | 2451.9848 | 3.0088 | -6.2574 | -4.7312 | 10.1165 | 2447.0356 | 6.2574 | 44.1707
0.4769 | 91.2883 | 79.5148 | -2436.3086 | 2441.1707 | 3.096 | -6.1702 | -4.753 | 10.0947 | 2436.3086 | 6.1702 | 44.338
0.4786 | 89.6749 | 77.7705 | -2423.4232 | 2428.2416 | 3.1832 | -6.1048 | -4.7748 | 10.1165 | 2423.4232 | 6.1048 | 44.4961
0.4804 | 88.1923 | 76.2225 | -2408.9897 | 2413.7646 | 3.2486 | -6.0176 | -4.7966 | 10.0947 | 2408.9897 | 6.0176 | 44.6634
0.4822 | 87.0149 | 75.0234 | -2391.5693 | 2396.3005 | 3.2486 | -5.9086 | -4.8184 | 10.0293 | 2391.5693 | 5.9086 | 44.8308
0.484 | 86.4916 | 74.1077 | -2371.0311 | 2375.8277 | 3.2704 | -5.8213 | -4.862 | 9.9639 | 2371.0311 | 5.8213 | 44.9981
0.4858 | 87.2111 | 73.1047 | -2348.5524 | 2353.4362 | 3.2268 | -5.7777 | -4.9056 | 9.9203 | 2348.5524 | 5.7777 | 45.1655
0.4876 | 88.3449 | 71.2951 | -2321.9748 | 2326.8805 | 3.096 | -5.7559 | -5.0146 | 9.8985 | 2321.9748 | 5.7559 | 45.3328
0.4894 | 89.1516 | 69.5727 | -2299.954 | 2304.8596 | 2.987 | -5.7123 | -5.1237 | 9.9203 | 2299.954 | 5.7123 | 45.5002
0.4912 | 88.9118 | 67.9593 | -2280.9856 | 2285.8258 | 2.9216 | -5.6251 | -5.1891 | 9.9421 | 2280.9856 | 5.6251 | 45.6675
0.493 | 87.669 | 66.2587 | -2264.241 | 2268.9286 | 2.8562 | -5.4943 | -5.2109 | 9.9203 | 2264.241 | 5.4943 | 45.8349
0.4948 | 86.339 | 65.9098 | -2250.1565 | 2254.7568 | 2.8562 | -5.3853 | -5.2327 | 9.8331 | 2250.1565 | 5.3853 | 46.0022
0.4966 | 84.9218 | 64.9069 | -2237.8815 | 2242.3293 | 2.8562 | -5.3199 | -5.2109 | 9.7895 | 2237.8815 | 5.3199 | 46.1696
0.4984 | 84.1369 | 61.9417 | -2229.64 | 2233.8262 | 2.9652 | -5.2327 | -5.1891 | 9.6804 | 2229.64 | 5.2327 | 46.3369
0.5002 | 82.916 | 57.1887 | -2221.4204 | 2225.2141 | 3.1178 | -5.1455 | -5.1673 | 9.5496 | 2221.4204 | 5.1455 | 46.5043
0.502 | 81.7822 | 53.2642 | -2211.2167 | 2214.7269 | 3.2486 | -5.0582 | -5.1455 | 9.4624 | 2211.2167 | 5.0582 | 46.6716
0.5037 | 80.147 | 51.0621 | -2198.9199 | 2202.1903 | 3.314 | -4.971 | -5.1237 | 9.3534 | 2198.9199 | 4.971 | 46.8297
0.5055 | 77.7269 | 52.1087 | -2182.5024 | 2185.5984 | 3.314 | -4.862 | -5.1019 | 9.2444 | 2182.5024 | 4.862 | 46.997
0.5073 | 75.394 | 54.8776 | -2164.5151 | 2167.4585 | 3.2486 | -4.8402 | -5.08 | 9.179 | 2164.5151 | 4.8402 | 47.1644
0.5091 | 73.1483 | 54.6596 | -2152.0875 | 2154.9001 | 3.2486 | -4.753 | -5.08 | 9.0918 | 2152.0875 | 4.753 | 47.3317
0.5109 | 70.9463 | 54.0927 | -2141.9492 | 2144.6528 | 3.2486 | -4.6658 | -5.08 | 9.0264 | 2141.9492 | 4.6658 | 47.4991
0.5127 | 68.766 | 53.3514 | -2133.6206 | 2136.2369 | 3.2486 | -4.5786 | -5.08 | 8.9609 | 2133.6206 | 4.5786 | 47.6664
0.5145 | 66.7383 | 51.9996 | -2125.728 | 2128.2353 | 3.2704 | -4.4914 | -5.1019 | 8.8955 | 2125.728 | 4.4914 | 47.8338
0.5163 | 64.6671 | 50.5389 | -2118.6202 | 2121.0186 | 3.2922 | -4.426 | -5.1237 | 8.8519 | 2118.6202 | 4.426 | 48.0011
0.5181 | 62.3996 | 48.642 | -2114.1725 | 2116.44 | 3.314 | -4.4042 | -5.1673 | 8.8737 | 2114.1725 | 4.4042 | 48.1685
0.5199 | 60.4591 | 48.2278 | -2110.8585 | 2113.0387 | 3.314 | -4.426 | -5.2109 | 8.9173 | 2110.8585 | 4.426 | 48.3358
0.5217 | 59.0637 | 45.7858 | -2108.7218 | 2110.7712 | 3.4012 | -4.4696 | -5.2763 | 9.0046 | 2108.7218 | 4.4696 | 48.5032
0.5235 | 57.821 | 43.4093 | -2107.4136 | 2109.3759 | 3.4884 | -4.5568 | -5.3635 | 9.1136 | 2107.4136 | 4.5568 | 48.6705
0.5253 | 56.8617 | 41.1419 | -2107.3918 | 2109.2887 | 3.5975 | -4.7312 | -5.4289 | 9.288 | 2107.3918 | 4.7312 | 48.8379
0.5271 | 55.8369 | 39.0706 | -2108.1549 | 2110.0081 | 3.7065 | -4.9056 | -5.5161 | 9.4624 | 2108.1549 | 4.9056 | 49.0052
0.5288 | 54.616 | 37.6534 | -2111.2727 | 2113.0605 | 3.8155 | -5.08 | -5.5815 | 9.6586 | 2111.2727 | 5.08 | 49.1633
0.5306 | 53.286 | 36.2798 | -2115.7641 | 2117.4865 | 3.8809 | -5.1455 | -5.6687 | 9.7676 | 2115.7641 | 5.1455 | 49.3306
0.5324 | 52.2395 | 34.71 | -2121.4328 | 2123.068 | 3.9463 | -5.1891 | -5.7777 | 9.8549 | 2121.4328 | 5.1891 | 49.498
0.5342 | 51.629 | 32.9658 | -2126.5565 | 2128.1699 | 4.0553 | -5.2981 | -5.8649 | 9.9421 | 2126.5565 | 5.2981 | 49.6653
0.536 | 50.19 | 31.7231 | -2130.0667 | 2131.6147 | 4.1207 | -5.3417 | -5.9522 | 10.0293 | 2130.0667 | 5.3417 | 49.8327
0.5378 | 49.9502 | 31.3088 | -2136.3023 | 2137.8285 | 4.1425 | -5.3635 | -6.0394 | 10.1383 | 2136.3023 | 5.3635 | 50
0.5396 | 49.0345 | 30.3495 | -2142.2327 | 2143.7371 | 4.2079 | -5.4289 | -6.1266 | 10.2473 | 2142.2327 | 5.4289 | 50.1673
0.5414 | 48.9909 | 30.0442 | -2148.7953 | 2150.3433 | 4.2079 | -5.4507 | -6.2356 | 10.3563 | 2148.7953 | 5.4507 | 50.3347
0.5432 | 49.0563 | 29.63 | -2154.3986 | 2155.9248 | 4.2297 | -5.4507 | -6.3446 | 10.4653 | 2154.3986 | 5.4507 | 50.502
0.545 | 48.6638 | 28.6271 | -2157.5818 | 2159.108 | 4.2951 | -5.4725 | -6.4536 | 10.5744 | 2157.5818 | 5.4725 | 50.6694
0.5468 | 47.7917 | 27.5587 | -2160.9831 | 2162.5093 | 4.426 | -5.5161 | -6.5408 | 10.6616 | 2160.9831 | 5.5161 | 50.8367
0.5486 | 47.0286 | 26.9264 | -2166.0631 | 2167.5893 | 4.5132 | -5.4943 | -6.628 | 10.7488 | 2166.0631 | 5.4943 | 51.0041
0.5504 | 46.3309 | 26.3378 | -2165.8451 | 2167.3713 | 4.6004 | -5.6251 | -6.6935 | 10.836 | 2165.8451 | 5.6251 | 51.1714
0.5521 | 45.219 | 26.5994 | -2165.4308 | 2166.9352 | 4.6876 | -5.7341 | -6.7589 | 10.945 | 2165.4308 | 5.7341 | 51.3295
0.5539 | 45.1536 | 26.3814 | -2169.3117 | 2170.7943 | 4.9274 | -5.6251 | -6.8461 | 11.0758 | 2169.3117 | 5.6251 | 51.4968
0.5557 | 45.8077 | 26.3596 | -2172.1461 | 2173.6287 | 5.1455 | -5.6251 | -6.9115 | 11.2502 | 2172.1461 | 5.6251 | 51.6642
0.5575 | 44.8265 | 26.5122 | -2173.8031 | 2175.2639 | 5.2327 | -5.6687 | -6.9769 | 11.3374 | 2173.8031 | 5.6687 | 51.8315
0.5593 | 44.216 | 26.9918 | -2171.7318 | 2173.149 | 5.2981 | -5.7341 | -7.0205 | 11.4247 | 2171.7318 | 5.7341 | 51.9989
0.5611 | 43.5184 | 27.5151 | -2169.5734 | 2170.9033 | 5.3199 | -5.8213 | -7.0423 | 11.5119 | 2169.5734 | 5.8213 | 52.1662
0.5629 | 42.7989 | 28.2346 | -2168.396 | 2169.6824 | 5.3417 | -5.8649 | -7.0859 | 11.5773 | 2168.396 | 5.8649 | 52.3336
0.5647 | 42.0794 | 29.0631 | -2167.7855 | 2168.9629 | 5.3417 | -5.9304 | -7.1077 | 11.6427 | 2167.7855 | 5.9304 | 52.5009
0.5665 | 41.4907 | 28.2564 | -2167.6983 | 2168.7885 | 5.3199 | -5.9086 | -7.1295 | 11.6863 | 2167.6983 | 5.9086 | 52.6683
0.5683 | 40.466 | 26.4904 | -2171.2086 | 2172.2333 | 5.2981 | -5.8649 | -7.1731 | 11.7081 | 2171.2086 | 5.8649 | 52.8356
0.5701 | 38.9834 | 25.5311 | -2178.5779 | 2179.559 | 5.2981 | -5.8649 | -7.2385 | 11.7299 | 2178.5779 | 5.8649 | 53.003
0.5719 | 37.7406 | 25.5965 | -2187.4298 | 2188.3456 | 5.2763 | -5.9086 | -7.2821 | 11.7735 | 2187.4298 | 5.9086 | 53.1703
0.5737 | 37.588 | 26.6866 | -2195.6277 | 2196.4998 | 5.2763 | -6.0176 | -7.3475 | 11.8825 | 2195.6277 | 6.0176 | 53.3377
0.5755 | 37.261 | 25.7055 | -2200.9258 | 2201.7543 | 5.2981 | -6.083 | -7.3693 | 11.9479 | 2200.9258 | 6.083 | 53.505
0.5772 | 36.5633 | 24.4627 | -2204.3706 | 2205.1555 | 5.3199 | -6.1048 | -7.4129 | 11.9915 | 2204.3706 | 6.1048 | 53.6631
0.579 | 35.7348 | 23.656 | -2208.7748 | 2209.5379 | 5.3417 | -6.1048 | -7.4784 | 12.0351 | 2208.7748 | 6.1048 | 53.8304
0.5808 | 34.4266 | 23.4162 | -2215.9915 | 2216.7328 | 5.3635 | -6.1266 | -7.522 | 12.1005 | 2215.9915 | 6.1266 | 53.9978
0.5826 | 32.8568 | 23.2636 | -2225.6283 | 2226.3696 | 5.3853 | -6.1484 | -7.6092 | 12.1878 | 2225.6283 | 6.1484 | 54.1651
0.5844 | 31.4396 | 23.8522 | -2237.598 | 2238.3175 | 5.4289 | -6.2138 | -7.7182 | 12.2968 | 2237.598 | 6.2138 | 54.3325
0.5862 | 30.3495 | 24.3755 | -2251.0504 | 2251.7481 | 5.5161 | -6.301 | -7.8054 | 12.4494 | 2251.0504 | 6.301 | 54.4998
0.588 | 30.0224 | 23.9177 | -2260.5564 | 2261.2759 | 5.6469 | -6.4318 | -7.849 | 12.6456 | 2260.5564 | 6.4318 | 54.6672
0.5898 | 29.3465 | 22.8493 | -2268.9068 | 2269.6263 | 5.7777 | -6.519 | -7.8926 | 12.7982 | 2268.9068 | 6.519 | 54.8345
0.5916 | 28.6489 | 22.3697 | -2278.9143 | 2279.6338 | 5.8649 | -6.5408 | -7.958 | 12.9072 | 2278.9143 | 6.5408 | 55.0019
0.5934 | 28.6271 | 22.2824 | -2287.57 | 2288.2895 | 5.974 | -6.6062 | -8.0016 | 13.0163 | 2287.57 | 6.6062 | 55.1692
0.5952 | 29.6736 | 22.5877 | -2293.2606 | 2294.0019 | 6.083 | -6.7589 | -8.0452 | 13.2125 | 2293.2606 | 6.7589 | 55.3366
0.597 | 29.521 | 22.6313 | -2302.7448 | 2303.4643 | 6.192 | -6.8243 | -8.1106 | 13.3433 | 2302.7448 | 6.8243 | 55.5039
0.5988 | 28.9105 | 23.1764 | -2315.0851 | 2315.8264 | 6.3228 | -6.8679 | -8.1978 | 13.4959 | 2315.0851 | 6.8679 | 55.6713
0.6006 | 27.1227 | 23.3072 | -2329.6058 | 2330.3035 | 6.4536 | -6.8243 | -8.2851 | 13.6049 | 2329.6058 | 6.8243 | 55.8386
0.6023 | 25.1822 | 23.2418 | -2343.8212 | 2344.5189 | 6.5844 | -6.7371 | -8.3505 | 13.6703 | 2343.8212 | 6.7371 | 55.9967
0.6041 | 23.9831 | 23.7214 | -2362.201 | 2362.8987 | 6.7589 | -6.7153 | -8.4595 | 13.823 | 2362.201 | 6.7153 | 56.164
0.6059 | 23.9831 | 24.7026 | -2380.7769 | 2381.4964 | 6.9333 | -6.8243 | -8.5685 | 14.041 | 2380.7769 | 6.8243 | 56.3313
0.6077 | 24.1575 | 26.0979 | -2398.9823 | 2399.7236 | 7.1513 | -6.8897 | -8.6557 | 14.2808 | 2398.9823 | 6.8897 | 56.4987
0.6095 | 23.9395 | 27.3407 | -2416.4681 | 2417.2312 | 7.3693 | -6.9115 | -8.7429 | 14.477 | 2416.4681 | 6.9115 | 56.666
0.6113 | 23.2418 | 27.6023 | -2431.5338 | 2432.3187 | 7.5002 | -6.8461 | -8.8083 | 14.5643 | 2431.5338 | 6.8461 | 56.8334
0.6131 | 23.0237 | 27.8858 | -2445.0297 | 2445.8364 | 7.6092 | -6.8461 | -8.8955 | 14.6733 | 2445.0297 | 6.8461 | 57.0007
0.6149 | 22.3697 | 27.9948 | -2457.8062 | 2458.5911 | 7.6746 | -6.8461 | -8.9827 | 14.7605 | 2457.8062 | 6.8461 | 57.1681
0.6167 | 21.8028 | 28.6271 | -2470.1248 | 2470.9097 | 7.6964 | -6.8243 | -9.0918 | 14.8695 | 2470.1248 | 6.8243 | 57.3354
0.6185 | 21.3013 | 29.5646 | -2482.1599 | 2482.9448 | 7.6964 | -6.8243 | -9.2444 | 14.9567 | 2482.1599 | 6.8243 | 57.5028
0.6203 | 22.1298 | 30.6111 | -2491.0118 | 2491.8185 | 7.7836 | -6.8461 | -9.3316 | 15.0875 | 2491.0118 | 6.8461 | 57.6701
0.6221 | 22.0208 | 31.3306 | -2500.3652 | 2501.1719 | 7.8272 | -6.7807 | -9.4188 | 15.1311 | 2500.3652 | 6.7807 | 57.8375
0.6239 | 21.1705 | 31.8321 | -2512.662 | 2513.5123 | 7.8708 | -6.7371 | -9.5278 | 15.1747 | 2512.662 | 6.7371 | 58.0048
0.6256 | 20.7999 | 32.726 | -2527.4007 | 2528.2946 | 7.9362 | -6.7371 | -9.5932 | 15.2838 | 2527.4007 | 6.7371 | 58.1629
0.6274 | 20.4074 | 34.71 | -2546.6525 | 2547.59 | 8.1324 | -6.7371 | -9.6368 | 15.4364 | 2546.6525 | 6.7371 | 58.3302
0.6292 | 21.5848 | 37.6316 | -2565.8826 | 2566.9509 | 8.3941 | -6.8025 | -9.7458 | 15.698 | 2565.8826 | 6.8025 | 58.4976
0.631 | 23.1328 | 40.793 | -2585.1126 | 2586.3336 | 8.6557 | -6.8679 | -9.8767 | 15.9814 | 2585.1126 | 6.8679 | 58.6649
0.6328 | 23.5906 | 43.1477 | -2605.7381 | 2607.0462 | 8.9391 | -6.8897 | -9.9857 | 16.2431 | 2605.7381 | 6.8897 | 58.8323
0.6346 | 23.5688 | 45.1754 | -2626.8868 | 2628.2604 | 9.2226 | -6.8679 | -10.0947 | 16.5047 | 2626.8868 | 6.8679 | 58.9996
0.6364 | 24.4409 | 47.9443 | -2644.002 | 2645.4846 | 9.4624 | -6.8679 | -10.2255 | 16.7445 | 2644.002 | 6.8679 | 59.167
0.6382 | 24.9424 | 50.2336 | -2661.3788 | 2662.9704 | 9.7022 | -6.8025 | -10.3563 | 16.9626 | 2661.3788 | 6.8025 | 59.3343
0.64 | 25.5747 | 51.6726 | -2674.2642 | 2675.8994 | 9.8113 | -6.7589 | -10.4653 | 17.0716 | 2674.2642 | 6.7589 | 59.5017
0.6418 | 26.425 | 52.8718 | -2686.2122 | 2687.9564 | 9.9203 | -6.7371 | -10.5744 | 17.1806 | 2686.2122 | 6.7371 | 59.669
0.6436 | 27.4933 | 54.3761 | -2696.9173 | 2698.727 | 10.0293 | -6.6935 | -10.6616 | 17.2896 | 2696.9173 | 6.6935 | 59.8364
0.6454 | 28.5616 | 55.8805 | -2706.7068 | 2708.5818 | 10.1383 | -6.6498 | -10.7706 | 17.3768 | 2706.7068 | 6.6498 | 60.0037
0.6472 | 29.7608 | 57.2323 | -2715.2753 | 2717.2157 | 10.2255 | -6.628 | -10.836 | 17.4422 | 2715.2753 | 6.628 | 60.1711
0.649 | 31.069 | 58.5187 | -2721.7725 | 2723.7784 | 10.3345 | -6.6498 | -10.9014 | 17.5294 | 2721.7725 | 6.6498 | 60.3384
0.6507 | 32.1809 | 59.8923 | -2727.8119 | 2729.9049 | 10.4871 | -6.6498 | -10.9668 | 17.5948 | 2727.8119 | 6.6498 | 60.4965
0.6525 | 33.2056 | 60.6772 | -2732.5867 | 2734.7234 | 10.5525 | -6.6717 | -11.0104 | 17.6603 | 2732.5867 | 6.6717 | 60.6638
0.6543 | 33.4891 | 60.7862 | -2736.7292 | 2738.8877 | 10.5962 | -6.6935 | -11.054 | 17.7039 | 2736.7292 | 6.6935 | 60.8312
0.6561 | 34.3394 | 61.8327 | -2740.7627 | 2742.943 | 10.618 | -6.7153 | -11.0976 | 17.7475 | 2740.7627 | 6.7153 | 60.9985
0.6579 | 35.6912 | 64.2092 | -2744.1858 | 2746.4533 | 10.6616 | -6.6935 | -11.163 | 17.7911 | 2744.1858 | 6.6935 | 61.1659
0.6597 | 36.6941 | 66.4331 | -2747.805 | 2750.1379 | 10.6834 | -6.6717 | -11.2502 | 17.8347 | 2747.805 | 6.6717 | 61.3332
0.6615 | 36.149 | 67.6976 | -2752.9941 | 2755.3488 | 10.6616 | -6.5626 | -11.3156 | 17.7911 | 2752.9941 | 6.5626 | 61.5006
0.6633 | 34.9063 | 68.5916 | -2760.4724 | 2762.8271 | 10.6616 | -6.3882 | -11.3811 | 17.7693 | 2760.4724 | 6.3882 | 61.6679
0.6651 | 34.056 | 68.7224 | -2768.5395 | 2770.8942 | 10.6834 | -6.3228 | -11.4465 | 17.8129 | 2768.5395 | 6.3228 | 61.8353
0.6669 | 33.1402 | 68.7006 | -2778.7432 | 2781.1197 | 10.7052 | -6.3228 | -11.5119 | 17.8783 | 2778.7432 | 6.3228 | 62.0026
0.6687 | 32.944 | 68.4171 | -2789.3829 | 2791.7812 | 10.8142 | -6.3664 | -11.5991 | 18.0309 | 2789.3829 | 6.3664 | 62.17
0.6705 | 33.053 | 68.3517 | -2798.5837 | 2801.0038 | 10.8796 | -6.3664 | -11.6863 | 18.1399 | 2798.5837 | 6.3664 | 62.3373
0.6723 | 33.1838 | 68.548 | -2806.6725 | 2809.1363 | 10.945 | -6.41 | -11.7735 | 18.2707 | 2806.6725 | 6.41 | 62.5046
0.6741 | 34.0996 | 67.5668 | -2815.8951 | 2818.4242 | 11.054 | -6.4754 | -11.8389 | 18.4234 | 2815.8951 | 6.4754 | 62.672
0.6758 | 35.1679 | 67.2616 | -2823.1773 | 2825.7282 | 11.0758 | -6.4972 | -11.9043 | 18.4888 | 2823.1773 | 6.4972 | 62.83
0.6776 | 36.1708 | 67.1744 | -2829.391 | 2831.9856 | 11.0976 | -6.4972 | -11.9697 | 18.5542 | 2829.391 | 6.4972 | 62.9974
0.6794 | 37.1519 | 66.9782 | -2835.365 | 2837.9595 | 11.1412 | -6.519 | -12.0351 | 18.6414 | 2835.365 | 6.519 | 63.1647
0.6812 | 38.1767 | 67.2834 | -2840.227 | 2842.8434 | 11.1848 | -6.6062 | -12.1005 | 18.7286 | 2840.227 | 6.6062 | 63.3321
0.683 | 38.4819 | 66.4331 | -2843.7809 | 2846.3536 | 11.1848 | -6.628 | -12.166 | 18.794 | 2843.7809 | 6.628 | 63.4994
0.6848 | 37.6098 | 64.6453 | -2848.0978 | 2850.5397 | 11.163 | -6.628 | -12.1878 | 18.8158 | 2848.0978 | 6.628 | 63.6668
0.6866 | 36.7377 | 62.574 | -2853.1997 | 2855.5108 | 11.1194 | -6.6062 | -12.2314 | 18.8158 | 2853.1997 | 6.6062 | 63.8341
0.6884 | 36.1272 | 60.5899 | -2858.6504 | 2860.8307 | 11.1194 | -6.628 | -12.275 | 18.8594 | 2858.6504 | 6.628 | 64.0015
0.6902 | 36.258 | 59.4344 | -2863.8612 | 2865.9761 | 11.1412 | -6.7153 | -12.3186 | 18.9466 | 2863.8612 | 6.7153 | 64.1688
0.692 | 37.479 | 59.7614 | -2867.5895 | 2869.7262 | 11.1848 | -6.8243 | -12.3622 | 19.0556 | 2867.5895 | 6.8243 | 64.3362
0.6938 | 38.6781 | 60.1975 | -2869.7262 | 2871.8847 | 11.1848 | -6.8679 | -12.4494 | 19.1428 | 2869.7262 | 6.8679 | 64.5035
0.6956 | 40.0735 | 60.6117 | -2870.9254 | 2873.1274 | 11.163 | -6.9333 | -12.5148 | 19.1865 | 2870.9254 | 6.9333 | 64.6709
0.6974 | 41.5779 | 61.026 | -2871.9065 | 2874.1522 | 11.1194 | -6.9769 | -12.5802 | 19.2301 | 2871.9065 | 6.9769 | 64.8382
0.6992 | 43.2349 | 61.7019 | -2872.3207 | 2874.6318 | 11.0976 | -7.0205 | -12.6456 | 19.2955 | 2872.3207 | 7.0205 | 65.0056
0.7009 | 44.9137 | 62.3996 | -2872.626 | 2875.0025 | 11.054 | -7.0859 | -12.711 | 19.3391 | 2872.626 | 7.0859 | 65.1636
0.7027 | 46.8542 | 62.8138 | -2872.3643 | 2874.7844 | 10.9886 | -7.1295 | -12.7764 | 19.3827 | 2872.3643 | 7.1295 | 65.331
0.7045 | 48.0969 | 62.9664 | -2871.5358 | 2873.9995 | 10.945 | -7.1513 | -12.82 | 19.4045 | 2871.5358 | 7.1513 | 65.4983
0.7063 | 48.8382 | 62.901 | -2870.0314 | 2872.517 | 10.9232 | -7.1513 | -12.8418 | 19.3827 | 2870.0314 | 7.1513 | 65.6657
0.7081 | 49.4487 | 62.4868 | -2868.527 | 2871.0126 | 10.836 | -7.1295 | -12.8418 | 19.3609 | 2868.527 | 7.1295 | 65.833
0.7099 | 50.0592 | 62.0943 | -2867.0445 | 2869.53 | 10.7706 | -7.1077 | -12.8636 | 19.3173 | 2867.0445 | 7.1077 | 66.0004
0.7117 | 49.5141 | 61.3094 | -2867.2625 | 2869.7262 | 10.6616 | -7.0205 | -12.8636 | 19.2301 | 2867.2625 | 7.0205 | 66.1677
0.7135 | 49.0563 | 60.2411 | -2869.1375 | 2871.5794 | 10.5525 | -6.9769 | -12.8636 | 19.1428 | 2869.1375 | 6.9769 | 66.3351
0.7153 | 48.315 | 58.8893 | -2871.9501 | 2874.3048 | 10.4217 | -6.9115 | -12.8636 | 19.0338 | 2871.9501 | 6.9115 | 66.5024
0.7171 | 47.4429 | 57.5594 | -2875.4821 | 2877.7932 | 10.3127 | -6.8679 | -12.8636 | 18.9466 | 2875.4821 | 6.8679 | 66.6698
0.7189 | 45.8949 | 55.7061 | -2879.0796 | 2881.3035 | 10.2037 | -6.7807 | -12.8636 | 18.8376 | 2879.0796 | 6.7807 | 66.8371
0.7207 | 44.6957 | 54.3107 | -2882.7425 | 2884.8791 | 10.1165 | -6.7371 | -12.8418 | 18.7504 | 2882.7425 | 6.7371 | 67.0045
0.7225 | 43.9544 | 53.5694 | -2886.4489 | 2888.5856 | 10.0293 | -6.7371 | -12.7982 | 18.685 | 2886.4489 | 6.7371 | 67.1718
0.7242 | 43.0387 | 53.068 | -2888.9345 | 2891.0275 | 9.9639 | -6.7153 | -12.7982 | 18.6632 | 2888.9345 | 6.7153 | 67.3299
0.726 | 41.9704 | 53.0026 | -2891.3764 | 2893.4694 | 9.9639 | -6.6717 | -12.7982 | 18.6414 | 2891.3764 | 6.6717 | 67.4972
0.7278 | 40.9674 | 53.0026 | -2893.6439 | 2895.7369 | 9.9421 | -6.628 | -12.7982 | 18.6196 | 2893.6439 | 6.628 | 67.6646
0.7296 | 40.4442 | 53.3078 | -2894.6468 | 2896.7617 | 9.9203 | -6.628 | -12.7982 | 18.6196 | 2894.6468 | 6.628 | 67.8319
0.7314 | 40.3352 | 53.6349 | -2894.6904 | 2896.8053 | 9.9203 | -6.6498 | -12.82 | 18.6196 | 2894.6904 | 6.6498 | 67.9993
0.7332 | 40.4878 | 54.0055 | -2894.0145 | 2896.173 | 9.8767 | -6.6717 | -12.8418 | 18.6414 | 2894.0145 | 6.6717 | 68.1666
0.735 | 40.6404 | 54.3979 | -2893.3386 | 2895.4971 | 9.8331 | -6.6935 | -12.8418 | 18.6196 | 2893.3386 | 6.6935 | 68.334
0.7368 | 40.793 | 54.7904 | -2892.6409 | 2894.843 | 9.7676 | -6.6935 | -12.8636 | 18.6196 | 2892.6409 | 6.6935 | 68.5013
0.7386 | 40.8148 | 55.161 | -2893.7747 | 2895.9768 | 9.7022 | -6.7153 | -12.8854 | 18.6196 | 2893.7747 | 6.7153 | 68.6687
0.7404 | 40.5532 | 55.1392 | -2895.9768 | 2898.1352 | 9.6586 | -6.8025 | -12.9291 | 18.6414 | 2895.9768 | 6.8025 | 68.836
0.7422 | 39.8773 | 54.8776 | -2899.2472 | 2901.3838 | 9.615 | -6.8679 | -12.9509 | 18.685 | 2899.2472 | 6.8679 | 69.0033
0.744 | 39.1796 | 54.9212 | -2904.0874 | 2906.2023 | 9.5932 | -6.8461 | -12.9509 | 18.6632 | 2904.0874 | 6.8461 | 69.1707
0.7458 | 38.7217 | 55.0302 | -2909.9087 | 2912.0454 | 9.5496 | -6.8243 | -12.9509 | 18.6196 | 2909.9087 | 6.8243 | 69.338
0.7476 | 39.6593 | 55.5317 | -2912.2634 | 2914.4437 | 9.5278 | -6.8897 | -12.9945 | 18.6632 | 2912.2634 | 6.8897 | 69.5054
0.7493 | 40.8366 | 56.2948 | -2913.5934 | 2915.7955 | 9.5496 | -6.9769 | -13.0163 | 18.7286 | 2913.5934 | 6.9769 | 69.6634
0.7511 | 41.8177 | 57.3631 | -2914.9234 | 2917.1473 | 9.5278 | -7.0641 | -13.0381 | 18.7722 | 2914.9234 | 7.0641 | 69.8308
0.7529 | 42.7989 | 58.7367 | -2914.9234 | 2917.1691 | 9.5496 | -7.1295 | -13.0599 | 18.8158 | 2914.9234 | 7.1295 | 69.9981
0.7547 | 43.78 | 60.3065 | -2914.0077 | 2916.3188 | 9.5714 | -7.1731 | -13.0817 | 18.8812 | 2914.0077 | 7.1731 | 70.1655
0.7565 | 44.9573 | 61.5929 | -2912.089 | 2914.4437 | 9.5932 | -7.2167 | -13.1253 | 18.9248 | 2912.089 | 7.2167 | 70.3328
0.7583 | 46.4181 | 62.4868 | -2908.884 | 2911.2823 | 9.5932 | -7.2603 | -13.1253 | 18.9684 | 2908.884 | 7.2603 | 70.5002
0.7601 | 46.8978 | 63.1627 | -2905.4174 | 2907.8375 | 9.6368 | -7.2385 | -13.1471 | 18.9902 | 2905.4174 | 7.2385 | 70.6675
0.7619 | 46.767 | 63.7077 | -2901.8199 | 2904.1964 | 9.6586 | -7.2167 | -13.1471 | 19.012 | 2901.8199 | 7.2167 | 70.8349
0.7637 | 46.658 | 64.231 | -2898.2224 | 2900.5771 | 9.7022 | -7.1731 | -13.1253 | 19.012 | 2898.2224 | 7.1731 | 71.0022
0.7655 | 45.4152 | 64.1656 | -2895.4535 | 2897.721 | 9.724 | -7.0641 | -13.1471 | 18.9684 | 2895.4535 | 7.0641 | 71.1696
0.7673 | 44.1506 | 64.0784 | -2892.7063 | 2894.9084 | 9.7676 | -6.9769 | -13.1471 | 18.9466 | 2892.7063 | 6.9769 | 71.3369
0.7691 | 42.7771 | 63.9258 | -2890.0464 | 2892.1613 | 9.8113 | -6.8461 | -13.1471 | 18.903 | 2890.0464 | 6.8461 | 71.5043
0.7709 | 41.3381 | 63.7077 | -2887.4737 | 2889.5231 | 9.8549 | -6.7153 | -13.1253 | 18.8594 | 2887.4737 | 6.7153 | 71.6716
0.7727 | 39.8773 | 63.4897 | -2884.9009 | 2886.8632 | 9.8985 | -6.5844 | -13.1253 | 18.8158 | 2884.9009 | 6.5844 | 71.839
0.7744 | 38.7654 | 63.5551 | -2881.8049 | 2883.7236 | 9.9203 | -6.4536 | -13.1253 | 18.7504 | 2881.8049 | 6.4536 | 71.997
0.7762 | 38.1113 | 64.0348 | -2877.9677 | 2879.8645 | 9.8985 | -6.3228 | -13.1253 | 18.6632 | 2877.9677 | 6.3228 | 72.1644
0.778 | 38.3729 | 64.4054 | -2871.3614 | 2873.2582 | 9.8985 | -6.2138 | -13.1253 | 18.6196 | 2871.3614 | 6.2138 | 72.3317
0.7798 | 39.0924 | 65.1685 | -2862.9019 | 2864.777 | 9.8767 | -6.1048 | -13.1253 | 18.5324 | 2862.9019 | 6.1048 | 72.4991
0.7816 | 40.575 | 65.9098 | -2853.6575 | 2855.5762 | 9.8549 | -5.9958 | -13.1471 | 18.467 | 2853.6575 | 5.9958 | 72.6664
0.7834 | 41.8177 | 66.3459 | -2842.9088 | 2844.871 | 9.7676 | -5.8868 | -13.1471 | 18.3579 | 2842.9088 | 5.8868 | 72.8338
0.7852 | 42.9297 | 67 | -2829.9361 | 2831.9638 | 9.6368 | -5.7341 | -13.1253 | 18.1617 | 2829.9361 | 5.7341 | 73.0011
0.787 | 43.6274 | 67.3706 | -2817.5957 | 2819.6452 | 9.4842 | -5.5379 | -13.1253 | 17.9873 | 2817.5957 | 5.5379 | 73.1685
0.7888 | 43.4748 | 67.1744 | -2806.5417 | 2808.5912 | 9.3752 | -5.2763 | -13.1253 | 17.7911 | 2806.5417 | 5.2763 | 73.3358
0.7906 | 42.9951 | 68.0901 | -2795.0953 | 2797.1665 | 9.1354 | -4.971 | -13.0599 | 17.4858 | 2795.0953 | 4.971 | 73.5032
0.7924 | 42.5808 | 68.5698 | -2781.6647 | 2783.736 | 8.8519 | -4.6222 | -12.9727 | 17.1588 | 2781.6647 | 4.6222 | 73.6705
0.7942 | 42.232 | 68.5916 | -2767.6238 | 2769.6732 | 8.5903 | -4.2733 | -12.9072 | 16.8099 | 2767.6238 | 4.2733 | 73.8379
0.796 | 41.8395 | 68.5262 | -2758.1177 | 2760.1454 | 8.4595 | -4.0335 | -12.9072 | 16.6573 | 2758.1177 | 4.0335 | 74.0052
0.7977 | 42.2974 | 68.7878 | -2750.2469 | 2752.2746 | 8.3941 | -3.8591 | -12.9072 | 16.5919 | 2750.2469 | 3.8591 | 74.1633
0.7995 | 42.9079 | 69.1366 | -2744.4256 | 2746.4533 | 8.3505 | -3.7283 | -12.9072 | 16.5483 | 2744.4256 | 3.7283 | 74.3306
0.8013 | 42.8425 | 69.1802 | -2739.7816 | 2741.7438 | 8.2851 | -3.5757 | -12.9291 | 16.4611 | 2739.7816 | 3.5757 | 74.498
0.8031 | 42.7771 | 69.2238 | -2735.1158 | 2737.0563 | 8.2197 | -3.423 | -12.9509 | 16.3957 | 2735.1158 | 3.423 | 74.6653
0.8049 | 42.4718 | 68.8532 | -2731.7582 | 2733.655 | 8.1106 | -3.2486 | -12.9727 | 16.3085 | 2731.7582 | 3.2486 | 74.8327
0.8067 | 40.9456 | 67.218 | -2729.8613 | 2731.6492 | 8.0016 | -3.0088 | -12.9945 | 16.1777 | 2729.8613 | 3.0088 | 75
0.8085 | 38.7435 | 65.8444 | -2727.0924 | 2728.8148 | 7.8926 | -2.769 | -13.0163 | 16.0687 | 2727.0924 | 2.769 | 75.1673
0.8103 | 37.5226 | 64.6235 | -2725.1955 | 2726.8307 | 7.8054 | -2.5945 | -13.0599 | 16.0032 | 2725.1955 | 2.5945 | 75.3347
0.8121 | 37.1519 | 64.449 | -2721.38 | 2722.9935 | 7.7182 | -2.4419 | -13.0817 | 15.9378 | 2721.38 | 2.4419 | 75.502
0.8139 | 37.1737 | 64.4926 | -2716.4308 | 2718.0224 | 7.6746 | -2.3547 | -13.0817 | 15.8942 | 2716.4308 | 2.3547 | 75.6694
0.8157 | 36.7595 | 64.122 | -2710.1298 | 2711.6778 | 7.6092 | -2.2239 | -13.0817 | 15.807 | 2710.1298 | 2.2239 | 75.8367
0.8175 | 36.7159 | 64.2964 | -2701.9756 | 2703.5018 | 7.5002 | -2.0713 | -13.0817 | 15.7198 | 2701.9756 | 2.0713 | 76.0041
0.8193 | 36.7377 | 64.1438 | -2692.2297 | 2693.7123 | 7.3257 | -1.8968 | -13.1035 | 15.589 | 2692.2297 | 1.8968 | 76.1714
0.8211 | 35.3641 | 61.4621 | -2680.2164 | 2681.59 | 7.0205 | -1.5916 | -13.0817 | 15.4146 | 2680.2164 | 1.5916 | 76.3388
0.8228 | 35.2115 | 60.0231 | -2671.2118 | 2672.4982 | 6.8897 | -1.3954 | -13.0817 | 15.3492 | 2671.2118 | 1.3954 | 76.4968
0.8246 | 35.0153 | 58.9329 | -2663.6681 | 2664.9108 | 6.8243 | -1.2428 | -13.0599 | 15.3056 | 2663.6681 | 1.2428 | 76.6642
0.8264 | 34.0778 | 57.1015 | -2656.8002 | 2657.9775 | 6.7371 | -1.0465 | -13.0381 | 15.2401 | 2656.8002 | 1.0465 | 76.8315
0.8282 | 32.4862 | 55.8805 | -2651.0879 | 2652.1998 | 6.6935 | -0.87211 | -13.0163 | 15.1965 | 2651.0879 | 0.87211 | 76.9989
0.83 | 32.5298 | 54.834 | -2647.0107 | 2648.1227 | 6.7153 | -0.8067 | -13.0163 | 15.1965 | 2647.0107 | 0.8067 | 77.1662
0.8318 | 32.617 | 52.1523 | -2641.7563 | 2642.8028 | 6.6935 | -0.71949 | -12.9945 | 15.1747 | 2641.7563 | 0.71949 | 77.3336
0.8336 | 32.0719 | 48.315 | -2635.5861 | 2636.5236 | 6.5844 | -0.63228 | -12.9509 | 15.0657 | 2635.5861 | 0.63228 | 77.5009
0.8354 | 31.0254 | 45.8949 | -2631.3782 | 2632.2503 | 6.519 | -0.50146 | -12.9509 | 14.9785 | 2631.3782 | 0.50146 | 77.6683
0.8372 | 29.9352 | 43.2131 | -2628.2822 | 2629.0889 | 6.4536 | -0.43606 | -12.9291 | 14.8913 | 2628.2822 | 0.43606 | 77.8356
0.839 | 28.4526 | 41.2727 | -2625.8402 | 2626.5597 | 6.519 | -0.30524 | -12.8854 | 14.8477 | 2625.8402 | 0.30524 | 78.003
0.8408 | 27.2099 | 38.8744 | -2623.878 | 2624.5539 | 6.4754 | -0.19623 | -12.8636 | 14.7605 | 2623.878 | 0.19623 | 78.1703
0.8426 | 25.9453 | 36.258 | -2624.4449 | 2625.0771 | 6.4318 | -0.13082 | -12.8418 | 14.6951 | 2624.4449 | 0.13082 | 78.3377
0.8444 | 25.4002 | 33.7725 | -2626.756 | 2627.3664 | 6.3882 | -0.087211 | -12.82 | 14.6079 | 2626.756 | 0.087211 | 78.505
0.8462 | 25.2258 | 31.3742 | -2628.5874 | 2629.2197 | 6.3664 | -0.043606 | -12.7764 | 14.5643 | 2628.5874 | 0.043606 | 78.6724
0.8479 | 25.7709 | 29.2157 | -2628.8272 | 2629.4595 | 6.3664 | 0.065408 | -12.711 | 14.5425 | 2628.8272 | -0.065408 | 78.8304
0.8497 | 26.643 | 28.0384 | -2627.9333 | 2628.6092 | 6.3882 | 0.10901 | -12.6674 | 14.5425 | 2627.9333 | -0.10901 | 78.9978
0.8515 | 27.1009 | 25.6401 | -2627.0394 | 2627.7371 | 6.41 | 0.10901 | -12.6238 | 14.477 | 2627.0394 | -0.10901 | 79.1651
0.8533 | 27.2971 | 23.7214 | -2625.644 | 2626.3635 | 6.41 | 0.10901 | -12.5584 | 14.4334 | 2625.644 | -0.10901 | 79.3325
0.8551 | 27.1663 | 22.4133 | -2623.6382 | 2624.3359 | 6.4536 | 0.087211 | -12.5366 | 14.3898 | 2623.6382 | -0.087211 | 79.4998
0.8569 | 27.5369 | 21.5193 | -2621.1526 | 2621.8503 | 6.4754 | 0.065408 | -12.5148 | 14.368 | 2621.1526 | -0.065408 | 79.6672
0.8587 | 28.4526 | 21.1051 | -2618.1221 | 2618.8197 | 6.4972 | 0.043606 | -12.5148 | 14.368 | 2618.1221 | -0.043606 | 79.8345
0.8605 | 28.9759 | 20.2984 | -2614.0667 | 2614.7644 | 6.5626 | 0.021803 | -12.493 | 14.368 | 2614.0667 | -0.021803 | 80.0019
0.8623 | 29.2157 | 19.2737 | -2609.4664 | 2610.1422 | 6.6498 | 0.021803 | -12.4712 | 14.368 | 2609.4664 | -0.021803 | 80.1692
0.8641 | 30.4149 | 18.9248 | -2604.0811 | 2604.7788 | 6.6935 | -0.087211 | -12.4712 | 14.3462 | 2604.0811 | 0.087211 | 80.3366
0.8659 | 31.9847 | 19.3609 | -2597.1478 | 2597.8455 | 6.7153 | -0.19623 | -12.4494 | 14.3244 | 2597.1478 | 0.19623 | 80.5039
0.8677 | 32.508 | 19.0556 | -2590.4979 | 2591.1956 | 6.7589 | -0.26163 | -12.4494 | 14.3244 | 2590.4979 | 0.26163 | 80.6713
0.8695 | 33.0094 | 18.7722 | -2583.9135 | 2584.6112 | 6.8243 | -0.32704 | -12.4276 | 14.3244 | 2583.9135 | 0.32704 | 80.8386
0.8712 | 33.2711 | 19.0556 | -2579.6837 | 2580.3596 | 6.8243 | -0.41425 | -12.4276 | 14.3244 | 2579.6837 | 0.41425 | 80.9967
0.873 | 32.6388 | 19.0556 | -2576.5441 | 2577.1982 | 6.8243 | -0.41425 | -12.4276 | 14.3244 | 2576.5441 | 0.41425 | 81.164
0.8748 | 31.1344 | 18.7504 | -2574.7345 | 2575.345 | 6.8243 | -0.30524 | -12.4276 | 14.3244 | 2574.7345 | 0.30524 | 81.3313
0.8766 | 29.8044 | 18.6632 | -2573.7316 | 2574.2767 | 6.8243 | -0.23983 | -12.4276 | 14.3026 | 2573.7316 | 0.23983 | 81.4987
0.8784 | 29.4556 | 20.56 | -2576.5878 | 2577.0892 | 6.8679 | -0.30524 | -12.4276 | 14.368 | 2576.5878 | 0.30524 | 81.666
0.8802 | 28.3 | 22.5441 | -2580.9047 | 2581.3844 | 6.9987 | -0.39245 | -12.4494 | 14.4989 | 2580.9047 | 0.39245 | 81.8334
0.882 | 27.1227 | 23.8959 | -2583.2376 | 2583.7173 | 7.1513 | -0.39245 | -12.4712 | 14.6297 | 2583.2376 | 0.39245 | 82.0007
0.8838 | 26.0543 | 25.0078 | -2584.5458 | 2585.0036 | 7.2821 | -0.37065 | -12.493 | 14.7387 | 2584.5458 | 0.37065 | 82.1681
0.8856 | 26.8174 | 25.8581 | -2582.0384 | 2582.5181 | 7.3693 | -0.39245 | -12.493 | 14.7823 | 2582.0384 | 0.39245 | 82.3354
0.8874 | 28.1038 | 26.8828 | -2577.6561 | 2578.1357 | 7.4347 | -0.41425 | -12.5148 | 14.8477 | 2577.6561 | 0.41425 | 82.5028
0.8892 | 29.1721 | 27.973 | -2572.358 | 2572.8595 | 7.5002 | -0.43606 | -12.5148 | 14.8913 | 2572.358 | 0.43606 | 82.6701
0.891 | 30.5675 | 29.4992 | -2565.9262 | 2566.4277 | 7.6092 | -0.47966 | -12.5366 | 15.0221 | 2565.9262 | 0.47966 | 82.8375
0.8928 | 31.4614 | 30.6983 | -2560.0176 | 2560.5409 | 7.6964 | -0.47966 | -12.5802 | 15.0875 | 2560.0176 | 0.47966 | 83.0048
0.8946 | 30.8727 | 31.0472 | -2555.7225 | 2556.2458 | 7.7182 | -0.39245 | -12.5802 | 15.1093 | 2555.7225 | 0.39245 | 83.1722
0.8963 | 29.2811 | 30.8073 | -2552.6701 | 2553.1716 | 7.6964 | -0.23983 | -12.5802 | 15.1093 | 2552.6701 | 0.23983 | 83.3302
0.8981 | 28.3654 | 31.0472 | -2551.3837 | 2551.8634 | 7.6964 | -0.13082 | -12.602 | 15.1093 | 2551.3837 | 0.13082 | 83.4976
0.8999 | 27.8422 | 31.3742 | -2552.0378 | 2552.5175 | 7.6964 | -0.087211 | -12.6238 | 15.1311 | 2552.0378 | 0.087211 | 83.6649
0.9017 | 27.3189 | 32.1809 | -2554.785 | 2555.2646 | 7.6746 | 0 | -12.6892 | 15.1311 | 2554.785 | 0 | 83.8323
0.9035 | 27.0573 | 33.0966 | -2556.987 | 2557.4667 | 7.631 | 0.087211 | -12.7328 | 15.1311 | 2556.987 | -0.087211 | 83.9996
0.9053 | 27.4497 | 33.9469 | -2556.1585 | 2556.66 | 7.6092 | 0.13082 | -12.7546 | 15.1093 | 2556.1585 | -0.13082 | 84.167
0.9071 | 28.191 | 34.9935 | -2553.6076 | 2554.1527 | 7.6092 | 0.15262 | -12.7764 | 15.1093 | 2553.6076 | -0.15262 | 84.3343
0.9089 | 27.2317 | 35.3423 | -2550.6206 | 2551.1657 | 7.5874 | 0.28344 | -12.7764 | 15.1093 | 2550.6206 | -0.28344 | 84.5017
0.9107 | 26.1851 | 35.6694 | -2547.6337 | 2548.1569 | 7.5438 | 0.41425 | -12.7764 | 15.1093 | 2547.6337 | -0.41425 | 84.669
0.9125 | 26.6212 | 36.9557 | -2544.4722 | 2545.0173 | 7.522 | 0.47966 | -12.7764 | 15.0875 | 2544.4722 | -0.47966 | 84.8364
0.9143 | 27.2971 | 37.6534 | -2539.4794 | 2540.0463 | 7.4784 | 0.52327 | -12.7982 | 15.0657 | 2539.4794 | -0.52327 | 85.0037
0.9161 | 26.9482 | 37.8932 | -2532.1973 | 2532.786 | 7.3257 | 0.65408 | -12.82 | 14.9785 | 2532.1973 | -0.65408 | 85.1711
0.9179 | 26.6212 | 38.0241 | -2525.9181 | 2526.5068 | 7.1949 | 0.7849 | -12.8418 | 14.9131 | 2525.9181 | -0.7849 | 85.3384
0.9197 | 25.9889 | 37.7188 | -2521.6229 | 2522.1898 | 7.1295 | 0.87211 | -12.8418 | 14.8913 | 2521.6229 | -0.87211 | 85.5058
0.9214 | 25.0078 | 37.588 | -2518.1781 | 2518.745 | 7.0641 | 0.98113 | -12.8636 | 14.8695 | 2518.1781 | -0.98113 | 85.6638
0.9232 | 24.4845 | 38.2421 | -2516.1722 | 2516.7173 | 7.0423 | 1.0683 | -12.8636 | 14.8695 | 2516.1722 | -1.0683 | 85.8312
0.925 | 25.1168 | 39.5284 | -2512.2041 | 2512.7928 | 7.0641 | 1.0901 | -12.8854 | 14.8913 | 2512.2041 | -1.0901 | 85.9985
0.9268 | 25.9453 | 40.7494 | -2507.4293 | 2508.0616 | 7.0859 | 1.0901 | -12.8854 | 14.9131 | 2507.4293 | -1.0901 | 86.1659
0.9286 | 25.8145 | 41.1855 | -2501.5644 | 2502.2185 | 7.0641 | 1.1555 | -12.8854 | 14.9131 | 2501.5644 | -1.1555 | 86.3332
0.9304 | 25.3348 | 41.5561 | -2496.855 | 2497.509 | 7.0423 | 1.2428 | -12.9072 | 14.9131 | 2496.855 | -1.2428 | 86.5006
0.9322 | 24.7898 | 41.9268 | -2493.9988 | 2494.6311 | 6.9987 | 1.33 | -12.9072 | 14.9131 | 2493.9988 | -1.33 | 86.6679
0.934 | 24.659 | 42.123 | -2491.2298 | 2491.8839 | 6.9769 | 1.3954 | -12.9072 | 14.8913 | 2491.2298 | -1.3954 | 86.8353
0.9358 | 23.9831 | 42.2538 | -2489.8781 | 2490.5322 | 6.9115 | 1.4826 | -12.9072 | 14.8695 | 2489.8781 | -1.4826 | 87.0026
0.9376 | 23.8086 | 42.4064 | -2490.5104 | 2491.1426 | 6.8679 | 1.5044 | -12.9072 | 14.8695 | 2490.5104 | -1.5044 | 87.17
0.9394 | 24.3319 | 42.4282 | -2490.3141 | 2490.9464 | 6.8461 | 1.5044 | -12.9072 | 14.8477 | 2490.3141 | -1.5044 | 87.3373
0.9412 | 24.6371 | 42.2756 | -2489.6818 | 2490.3359 | 6.8243 | 1.5262 | -12.9072 | 14.8259 | 2489.6818 | -1.5262 | 87.5046
0.943 | 24.5063 | 42.3628 | -2489.8345 | 2490.4886 | 6.8025 | 1.548 | -12.9072 | 14.8259 | 2489.8345 | -1.548 | 87.672
0.9447 | 24.1793 | 42.559 | -2490.4886 | 2491.1426 | 6.7589 | 1.5916 | -12.9072 | 14.8259 | 2490.4886 | -1.5916 | 87.83
0.9465 | 24.0703 | 43.3003 | -2491.9275 | 2492.5816 | 6.7589 | 1.5916 | -12.9291 | 14.8259 | 2491.9275 | -1.5916 | 87.9974
0.9483 | 24.7244 | 43.671 | -2492.647 | 2493.3229 | 6.7589 | 1.548 | -12.9509 | 14.8477 | 2492.647 | -1.548 | 88.1647
0.9501 | 25.0732 | 43.78 | -2493.0395 | 2493.7154 | 6.7371 | 1.548 | -12.9509 | 14.8259 | 2493.0395 | -1.548 | 88.3321
0.9519 | 25.313 | 43.6274 | -2492.5816 | 2493.2575 | 6.6935 | 1.548 | -12.9727 | 14.8259 | 2492.5816 | -1.548 | 88.4994
0.9537 | 25.7709 | 43.8672 | -2491.5133 | 2492.1892 | 6.6717 | 1.548 | -12.9727 | 14.8259 | 2491.5133 | -1.548 | 88.6668
0.9555 | 26.1197 | 43.9762 | -2489.0496 | 2489.7473 | 6.628 | 1.548 | -12.9945 | 14.8477 | 2489.0496 | -1.548 | 88.8341
0.9573 | 27.864 | 44.1724 | -2485.9972 | 2486.6731 | 6.519 | 1.548 | -13.0599 | 14.8477 | 2485.9972 | -1.548 | 89.0015
0.9591 | 29.63 | 44.216 | -2481.2006 | 2481.8765 | 6.4754 | 1.5044 | -13.1035 | 14.8477 | 2481.2006 | -1.5044 | 89.1688
0.9609 | 30.6983 | 43.889 | -2475.9025 | 2476.6002 | 6.4318 | 1.4826 | -13.1253 | 14.8259 | 2475.9025 | -1.4826 | 89.3362
0.9627 | 30.7855 | 43.4966 | -2473.9184 | 2474.6161 | 6.41 | 1.5044 | -13.1253 | 14.8259 | 2473.9184 | -1.5044 | 89.5035
0.9645 | 30.8073 | 43.1695 | -2473.3952 | 2474.0929 | 6.3882 | 1.5262 | -13.1253 | 14.8041 | 2473.3952 | -1.5262 | 89.6709
0.9663 | 30.8291 | 42.8861 | -2473.8966 | 2474.6161 | 6.3664 | 1.548 | -13.1253 | 14.7823 | 2473.8966 | -1.548 | 89.8382
0.9681 | 30.8945 | 42.6244 | -2474.5071 | 2475.2048 | 6.3228 | 1.548 | -13.1253 | 14.7823 | 2474.5071 | -1.548 | 90.0056
0.9698 | 30.1533 | 41.8177 | -2474.4635 | 2475.1394 | 6.2792 | 1.5916 | -13.1253 | 14.7823 | 2474.4635 | -1.5916 | 90.1636
0.9716 | 29.63 | 41.5561 | -2474.1147 | 2474.7905 | 6.2574 | 1.6352 | -13.1253 | 14.7823 | 2474.1147 | -1.6352 | 90.331
0.9734 | 29.3029 | 41.5343 | -2473.6568 | 2474.3327 | 6.2574 | 1.657 | -13.1253 | 14.7823 | 2473.6568 | -1.657 | 90.4983
0.9752 | 29.2157 | 41.5125 | -2473.9184 | 2474.6161 | 6.2356 | 1.6788 | -13.1253 | 14.7605 | 2473.9184 | -1.6788 | 90.6657
0.977 | 29.4556 | 41.5779 | -2474.3981 | 2475.0958 | 6.2138 | 1.657 | -13.1471 | 14.7605 | 2474.3981 | -1.657 | 90.833
0.9788 | 29.9134 | 41.7959 | -2474.0274 | 2474.7251 | 6.192 | 1.6352 | -13.1471 | 14.7605 | 2474.0274 | -1.6352 | 91.0004
0.9806 | 30.5021 | 42.3192 | -2474.8778 | 2475.6191 | 6.192 | 1.6134 | -13.1471 | 14.7605 | 2474.8778 | -1.6134 | 91.1677
0.9824 | 30.8073 | 42.4936 | -2475.8153 | 2476.5566 | 6.1702 | 1.6134 | -13.1689 | 14.7605 | 2475.8153 | -1.6134 | 91.3351
0.9842 | 30.0224 | 42.0794 | -2477.4505 | 2478.1918 | 6.1484 | 1.657 | -13.1689 | 14.7605 | 2477.4505 | -1.657 | 91.5024
0.986 | 28.8015 | 41.5997 | -2480.5247 | 2481.2442 | 6.1266 | 1.7442 | -13.1689 | 14.7605 | 2480.5247 | -1.7442 | 91.6698
0.9878 | 27.5587 | 41.1637 | -2485.5829 | 2486.3024 | 6.1048 | 1.8096 | -13.1907 | 14.7823 | 2485.5829 | -1.8096 | 91.8371
0.9896 | 26.7302 | 40.8802 | -2491.2516 | 2491.9929 | 6.1048 | 1.8532 | -13.2125 | 14.8041 | 2491.2516 | -1.8532 | 92.0045
0.9914 | 26.861 | 41.1855 | -2494.8709 | 2495.6122 | 6.1048 | 1.8532 | -13.2125 | 14.8259 | 2494.8709 | -1.8532 | 92.1718
0.9932 | 27.9948 | 41.7523 | -2496.2881 | 2497.0512 | 6.1266 | 1.7878 | -13.2343 | 14.8259 | 2496.2881 | -1.7878 | 92.3392
0.9949 | 29.521 | 42.2102 | -2495.9611 | 2496.7678 | 6.1048 | 1.7006 | -13.2561 | 14.8259 | 2495.9611 | -1.7006 | 92.4972
0.9967 | 30.3059 | 42.341 | -2494.6747 | 2495.5032 | 6.083 | 1.657 | -13.2561 | 14.8259 | 2494.6747 | -1.657 | 92.6646
0.9985 | 29.8698 | 42.5372 | -2493.3011 | 2494.1078 | 6.0394 | 1.6788 | -13.2779 | 14.8259 | 2493.3011 | -1.6788 | 92.8319
1.0003 | 29.4774 | 42.7553 | -2491.8621 | 2492.6688 | 5.9958 | 1.7006 | -13.2779 | 14.8259 | 2491.8621 | -1.7006 | 92.9993
1.0021 | 29.3029 | 43.0169 | -2490.2487 | 2491.0554 | 5.9304 | 1.7006 | -13.2997 | 14.8041 | 2490.2487 | -1.7006 | 93.1666
1.0039 | 29.1503 | 43.4093 | -2489.4638 | 2490.2923 | 5.8868 | 1.7006 | -13.3215 | 14.8041 | 2489.4638 | -1.7006 | 93.334
1.0057 | 29.0631 | 43.9544 | -2489.6164 | 2490.4667 | 5.8649 | 1.7006 | -13.3215 | 14.8041 | 2489.6164 | -1.7006 | 93.5013
1.0075 | 28.9759 | 44.4995 | -2489.7909 | 2490.6412 | 5.8213 | 1.7006 | -13.3433 | 14.8041 | 2489.7909 | -1.7006 | 93.6687
1.0093 | 28.9105 | 44.6739 | -2490.7066 | 2491.5569 | 5.7995 | 1.7006 | -13.3433 | 14.8041 | 2490.7066 | -1.7006 | 93.836
1.0111 | 28.8451 | 44.8265 | -2491.6441 | 2492.4944 | 5.7995 | 1.7006 | -13.3651 | 14.8041 | 2491.6441 | -1.7006 | 94.0033
1.0129 | 28.6925 | 45.0228 | -2493.4973 | 2494.3476 | 5.7777 | 1.7006 | -13.3651 | 14.8041 | 2493.4973 | -1.7006 | 94.1707
1.0147 | 28.5616 | 45.3716 | -2496.1137 | 2496.9858 | 5.7777 | 1.6788 | -13.3869 | 14.8259 | 2496.1137 | -1.6788 | 94.338
1.0165 | 28.5616 | 45.8731 | -2498.9262 | 2499.7983 | 5.7995 | 1.657 | -13.3869 | 14.8259 | 2498.9262 | -1.657 | 94.5054
1.0182 | 29.6082 | 46.3745 | -2501.717 | 2502.6327 | 5.8213 | 1.5698 | -13.4087 | 14.8259 | 2501.717 | -1.5698 | 94.6634
1.02 | 30.066 | 46.5271 | -2503.6356 | 2504.5513 | 5.8431 | 1.548 | -13.4087 | 14.8477 | 2503.6356 | -1.548 | 94.8308
1.0218 | 29.7608 | 46.0257 | -2505.5979 | 2506.4918 | 5.8431 | 1.548 | -13.4305 | 14.8695 | 2505.5979 | -1.548 | 94.9981
1.0236 | 29.0631 | 45.8513 | -2508.6503 | 2509.5224 | 5.8431 | 1.5698 | -13.4523 | 14.8695 | 2508.6503 | -1.5698 | 95.1655
1.0254 | 29.1067 | 46.6362 | -2509.8712 | 2510.7651 | 5.8213 | 1.5698 | -13.4523 | 14.8913 | 2509.8712 | -1.5698 | 95.3328
1.0272 | 29.1939 | 47.0068 | -2511.1794 | 2512.0951 | 5.8213 | 1.548 | -13.4741 | 14.8913 | 2511.1794 | -1.548 | 95.5002
1.029 | 30.1096 | 46.9632 | -2511.9643 | 2512.88 | 5.8213 | 1.5044 | -13.4741 | 14.8913 | 2511.9643 | -1.5044 | 95.6675
1.0308 | 30.9163 | 47.1158 | -2511.8771 | 2512.8146 | 5.8213 | 1.439 | -13.4959 | 14.8913 | 2511.8771 | -1.439 | 95.8349
1.0326 | 31.8539 | 47.7045 | -2510.7869 | 2511.7463 | 5.8431 | 1.3954 | -13.4959 | 14.9131 | 2510.7869 | -1.3954 | 96.0022
1.0344 | 32.9004 | 47.9443 | -2508.3014 | 2509.3044 | 5.8431 | 1.3518 | -13.4959 | 14.9131 | 2508.3014 | -1.3518 | 96.1696
1.0362 | 33.9251 | 47.9879 | -2504.9874 | 2505.9903 | 5.8213 | 1.3082 | -13.5177 | 14.9131 | 2504.9874 | -1.3082 | 96.3369
1.038 | 33.7725 | 47.5301 | -2499.6457 | 2500.6487 | 5.8213 | 1.3518 | -13.4959 | 14.9131 | 2499.6457 | -1.3518 | 96.5043
1.0398 | 33.4019 | 46.8542 | -2492.4508 | 2493.4537 | 5.8431 | 1.3954 | -13.4741 | 14.8913 | 2492.4508 | -1.3954 | 96.6716
1.0416 | 32.3989 | 46.8324 | -2490.1179 | 2491.099 | 5.7995 | 1.4608 | -13.4741 | 14.8913 | 2490.1179 | -1.4608 | 96.839
1.0433 | 32.3117 | 47.2466 | -2487.8286 | 2488.8315 | 5.7777 | 1.4826 | -13.4959 | 14.8913 | 2487.8286 | -1.4826 | 96.997
1.0451 | 32.1809 | 47.5955 | -2486.6731 | 2487.676 | 5.7559 | 1.5044 | -13.5177 | 14.8913 | 2486.6731 | -1.5044 | 97.1644
1.0469 | 33.1838 | 48.2932 | -2485.9318 | 2486.9783 | 5.7777 | 1.4172 | -13.5177 | 14.9131 | 2485.9318 | -1.4172 | 97.3317
1.0487 | 33.0094 | 48.1842 | -2484.58 | 2485.6265 | 5.7559 | 1.439 | -13.5395 | 14.8913 | 2484.58 | -1.439 | 97.4991
1.0505 | 31.7231 | 47.9007 | -2483.3808 | 2484.4056 | 5.7559 | 1.5262 | -13.5177 | 14.9131 | 2483.3808 | -1.5262 | 97.6664
1.0523 | 31.5268 | 47.8571 | -2484.1657 | 2485.1905 | 5.7559 | 1.5044 | -13.5395 | 14.9131 | 2484.1657 | -1.5044 | 97.8338
1.0541 | 31.1344 | 47.6609 | -2484.689 | 2485.6919 | 5.7341 | 1.5044 | -13.5395 | 14.9131 | 2484.689 | -1.5044 | 98.0011
1.0559 | 30.6547 | 47.9879 | -2484.7544 | 2485.7792 | 5.7341 | 1.548 | -13.5395 | 14.9131 | 2484.7544 | -1.548 | 98.1685
1.0577 | 30.2187 | 48.3586 | -2484.9943 | 2486.019 | 5.7123 | 1.5698 | -13.5395 | 14.9131 | 2484.9943 | -1.5698 | 98.3358
1.0595 | 29.957 | 48.4676 | -2484.8852 | 2485.9318 | 5.7123 | 1.5698 | -13.5395 | 14.9131 | 2484.8852 | -1.5698 | 98.5032
1.0613 | 30.3059 | 48.9255 | -2485.6483 | 2486.6949 | 5.7123 | 1.5262 | -13.5613 | 14.9349 | 2485.6483 | -1.5262 | 98.6705
1.0631 | 31.1126 | 49.8412 | -2486.7167 | 2487.785 | 5.7341 | 1.4608 | -13.5831 | 14.9349 | 2486.7167 | -1.4608 | 98.8379
1.0649 | 31.3088 | 50.1028 | -2486.7821 | 2487.894 | 5.7341 | 1.4608 | -13.6049 | 14.9567 | 2486.7821 | -1.4608 | 99.0052
1.0667 | 31.396 | 50.3208 | -2486.6513 | 2487.785 | 5.7341 | 1.4608 | -13.6049 | 14.9785 | 2486.6513 | -1.4608 | 99.1726
1.0684 | 31.4396 | 50.4953 | -2486.3678 | 2487.5016 | 5.7777 | 1.4826 | -13.6049 | 15.0003 | 2486.3678 | -1.4826 | 99.3306
1.0702 | 31.7667 | 50.4953 | -2486.1062 | 2487.2399 | 5.7777 | 1.4608 | -13.6267 | 15.0003 | 2486.1062 | -1.4608 | 99.498
1.072 | 31.287 | 49.7104 | -2485.2559 | 2486.3678 | 5.7777 | 1.5044 | -13.6485 | 15.0221 | 2485.2559 | -1.5044 | 99.6653
1.0738 | 30.3277 | 49.2961 | -2485.1905 | 2486.2806 | 5.7777 | 1.548 | -13.6485 | 15.0221 | 2485.1905 | -1.548 | 99.8327
1.0756 | 30.1314 | 49.5795 | -2485.1905 | 2486.2588 | 5.7777 | 1.548 | -13.6485 | 15.0439 | 2485.1905 | -1.548 | 100
| Fx | Fy | Fz | Fres | Mx | My | Mz | Mres | -Fz | -My
Minimum | -109.9078 | -228.2098 | -3810.1893 | 278.8576 | -2.0495 | -7.2603 | -13.6485 | 1.1337 | 276.6773 | -24.3101
Maximum | 115.4894 | 144.9885 | -276.6773 | 3814.8333 | 11.1848 | 24.3101 | -0.63228 | 28.1038 | 3810.1893 | 7.2603

## Diagr_HIGH100
HIGH100 = Loads at High Level in Subject with 100kg Body Weight.
Contact forces and moments, measured telemetrically in knee implants of 8 subjects.
Averages of one individual subject.
Fx  lateral, Fy  anterior, -Fz  distal, Fres = Resultant Force
Mx around x, My around y, Mz around z, Mres = Resultant Moment
z-axis = axis of implant stem!
From Bergmann et al, PlosOne (2014) "Standardized Loads Acting in Knee Implants"
Subject K9L
Subject K9L
G.Bergmann et al., PlosOne, 2014
G.Bergmann et al., PlosOne, 2014